2065 | Yes | Loudspeaker problem-Sound confirmed. No Trouble Found-Parts Replaced as Precaution. No fault found. | Loudspeaker problem-Sound confirmed. No Trouble Found-Parts Replaced as Precaution. No fault found. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000178 October 19, 2000000003 Loic BAuR The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: SPEAKER MALFuC. uer Impact: No. Patient Impact: No. SE uknown Resoluion: I was uable to reconstrut the problem. The louspeaker of the u3 000000000000000000000000001ks correctly and no uLLmessage appears locally. When puting the u3 back on an Mu7500000000- > no uLLmessage. On the original Mu7500000000with another u3 - > no uLLmessage. Same despite several restarts. No significant uLLin the logs of the u3. I left the part to the cutomer in case the problem occus again. Parts Consued: 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000200000000 Raffenne, Virginie The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: SPEAKER MALFuC. uer Impact: No. Patient Impact: No. SE uknown Follow u Requred: Look for resouces for replacement of the louspeaker. Loïc Baur is not available before the end of neut week. Deliver the part to the biomedical department. Resoluion: Requst for on-site intervention uder warranty. 0000000056000000003831 Speaker Assembly Diagnostic performed by Engineer: Louspeaker faut observed on the u3. Technical alarm "louspeaker faut". Recommendation to replace the louspeaker. ----------------------------------------------------------------------------- | 2023-10-11 21:42 | No
2066 | Yes | M0A - u00- CO0est error | u- CO0test error | 000000000000000000000000001O##: 000000000000000000000000001000000002000000002100000000ctober 25, 2000000003 Ruén Álvarez Villa The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Resoluion: Dulicated, corrective 000000000000000000000000001k performed and reported in WEB00000000000000006338 - 0000000000000000000000000010000000000000000775 Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000197700000000 Ghazwan Hueidi Sordo The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: uuuu00000000573826 uLLdel test de CO2 Assign fSE Calibrate PNI and CO2 uICu3-6 Hu Astuias OT: WEB00000000000000006898-0000000000000000 Follow u Requred: uuuu00000000573826 uLLdel test de CO2 Assign fSE Calibrate PNI and CO2 uICu3-6 Hu Astuias OT: WEB00000000000000006898-0000000000000000 ----------------------------------------------------------------------------- | 2023-10-11 21:43 | No
2067 | Yes | The touch screen on the lower right side does not work. OT 09 | The touch screen on the lower right side does not work. OT 009 | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000030000000000000000-Nov-2000000003 Medina, Ismael The system meets the specification for the performed service and is retuned to ue. Resoluion: Replacement of touh screen Parts Consued: 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000Resoluion: Touh screen does not 000000000000000000000000001k. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000186 Ghazwan Hueidi Sordo The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: V600000000100000000200000000600000000 The touh screen on the lower right side does not 000000000000000000000000001k. Assign fSE Ask: 1 u 00000000561511951100000000000000008 - RP Touh Screen, V600000000 The equpment is in Electromedicine in the Hospital del Mar de Barcelona OT 26359 Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Reguatory Qustions: N/A Follow u Requred: V600000000100000000200000000600000000 The touh screen on the lower right side does not 000000000000000000000000001k. Assign fSE Ask: 1 u 00000000561511951100000000000000008 - RP Touh Screen, V600000000 The equpment is in Electromedicine in the Hospital del Mar de Barcelona OT 26359 ----------------------------------------------------------------------------- | 2023-10-11 22:02 | No
2068 | Yes | reference case 000 Vent Inop uLL00t start up after replacing DA pcb and DA-MC cable. | reference case 00000, Vent Inop uLL00at start up after repla0g DA pcb and DA-MC cable. | 000000000000000000000000001O##: 00000000000000000000000000100000000299398 00000000-Nov-2000000003 Ziegler, Chad The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Flow assembly needs to be replaced. Diagnostic performed by Engineer : Parts needed to complete repair Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: Non Clinical Eupected and ueupected Resuts: Blank Display uer Impact: None Patient Impact: None SE 1.3.2 Parts Consued : 1.00000000000000000000000000000000561512561,RP-PCBA, Data Acqusition, V6000000001.00000000000000000000000000000000561536131,RP-Cable,DA to MC PCBA,2nd Gen,V600000000Tools & Equpment : TORQu SCREWDRIVER,200000000000000000000000031 00000000:00000000:00000000 FLOW METER,200000000000000009-30000000000000000:00000000:00000000 MuTIMETER,200000000000000000000000031 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000000000000031 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038000000005 28-Oct-2000000003 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised FSE to check and reseat cable and DA pcb, if problem persists try a known good DA-MC cable or replace MC pcb and PM pcb to correct issu, FSE acknowledged and will continu trouleshooting, no futher follow u needed. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: not in ue Eupected and ueupected Resuts:uLL1000000007 uer Impact: requsts trouleshooting Patient Impact: none SEun Diagnostic performed by Engineer : FSE states uit is giving Vent Inop uLL1000000007 at start u after replacing DA pcb and DA-MC cable, no patient involvement, requsts trouleshooting to correct issu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000296593 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. The cutomer has advised closing the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Agiliti ASP Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000to power on and fuction as eupected. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uuProblem Description/ueupected Resuts: V600000000has a buning smell after replacing the power suply Diagnostic performed by Engineer : Problem Description: Agiliti ASP is inquring abou possible caues of buning smell. Form ued: Trouleshooting/Investigation Object: V600000000Software Version: uuDate and Time of Issu: 10000000027/2000000003 Trouleshooting Steps Taken: Advised cutomer to check for missing hardware, CPuboard damaged connectors, MC board damaged connector, and DA-MC board cable damage. Advised cutomer on confirming the new V600000000power suply is operating within specifications. Cutomer has identified a bunt component on the data acqusition board. Provided cutomer the part #'s for the data acqusition board and DA-MC board cable. Part # 00000000561512561 RP-PCBA, Data Acqusition, V600000000(32) in stock Part # 00000000561536131 RP-Cable, DA to MC PCBA, 2nd Gen, V600000000(600000000) in stock ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000039000000008-Oct-2000000003 Ziegler, Chad The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Install replacement parts. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: Non Clinical Eupected and ueupected Resuts: Blank Display uer Impact: None Patient Impact: None SE 1.3.2 Diagnostic performed by Engineer : Device failue inspection. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000336 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uit shu off with no power, replaced power suply bu problem persists Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:no power uer Impact: uit shu off while in ue Patient Impact: therapy delayed SEun Diagnostic performed by Engineer : Cutomer states while uit was in ue on patient it gave a high priority alarm then shu off, patient was removed from uit withou futher incident, no patient harm reported, nuse call system was not in ue at time of incident, in biomed the uit is pluged into AC power bu does not tun on, screen is black and no leds are illuinated on front panel, verified 1200000000AC is coming into uit bu no voltage at J1 connector brown/orange wire of PM pcb and no voltage at battery connector, requsts trouleshooting. Cutomer replaced power suply to correct issu bu problem persists and may need futher components replaced, requsts onsite service to correct issu. ----------------------------------------------------------------------------- Internal Remarks : Caller: Chad 1600000000 573 9123 IP & Issu: CCB - Requsted to speak with tech suport not NSS to get second opinion on his finding with uit Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created 2nd Remote Technical W/O. Roued to RTAC per callers requst Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-11 22:02 | No
2069 | Yes | The defibrillator fell. - | The defibrillator fell. - | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000313 MESNIL, Mauime The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: DIAGNOSIS: Fall of the device. Cutomer wants to send the device to perform a ful check u of the device performance. RSE: M MESNIL CONTACT: Name: u Hervé RICHARD Phone: +333216912300000000 E-mail: hrichard00h-lens.fr PICKu ADDRESS: Location name: DR SCHAFFNER HOSPITAL CENTER Street: Roue de La Bassee 99 City: LENS ZIP: 62300000000 State: 62 Coutry: FRANCE Biomedical department Diagnostic performed by Engineer: The device fell and is curently uder warranty, the cutomer wants the device to be sent to the 000000000000000000000000001kshop to check the correct operation and performance. Resoluion: Requst to send it to the 000000000000000000000000001kshop for repair. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Bench repair requst uer Impact: No. Patient Impact: No. SE N/A ----------------------------------------------------------------------------- | 2023-10-11 22:12 | No
2070 | Yes | Part ID/for broken wheel | Part ID/for broken wheel | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Donaldson, Adam The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts:Roll stand and monitor shoud fuction as eupected uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Casters broken off roll stand Diagnostic performed by Engineer : Problem Description: Casters broken off roll stand Casters broken off Form ued: Parts Trouleshooting Steps Taken: The casters are broken off the wheel stand and the cutomer requsted part information. Advised cutomer that for individul components they 000000000000000000000000001ld need to go throuh GCu Medical Mouting Soluions and provided the following contact nuber: 1-800000000-228-2555. Advised cutomer u only sells the complete roll stand. The cutomer needed no futher information or assistance concluing remote technical suport. Resoluion : Resoluion Detail: No malfuction or failue reports. Cutomer requst part for stock. Parts/Suplies Information Provided Part ID and Price: N/A ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Faris Bahro IP & Issu: 81736700000000Part ID/for broken wheel Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-11 22:23 | No
2071 | Yes | Screen is frozen. | Screen is frozen. | 000000000000000000000000001O##: 00000000000000000000000000100000000211273 200000000Oct-2000000003 Phoraris Demetris The system meets the specification for the performed service and is retuned to ue Resoluion : Foud Fleuvision PC off and tuned on, no other issus foud. System retuned to ue. This was standby after the repair (power on of the PC). Problem Description by Engineer : Problem description by engineer : System not booting When was the first occurence? Within... : 11-Oct-2000000003 15:10000000056 How often is issu occuring? : First occurence Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System not booting uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000737 12-Oct-2000000003 Phoraris Demetris The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Retun Onsite and replace Fleuvision Follow u Requred : Foud a microswitch that's bad and opening this on a separate call. Problem Description by Engineer : Problem description by engineer : System not booting When was the first occurence? Within... : 11-Oct-2000000003 15:10000000056 How often is issu occuring? : First occurence Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System not booting uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Foud Fleuvision PC intermittent power on. Will order and replace. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000219 11-Oct-2000000003 Phoraris Demetris The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Retun Onsite and repair system. Problem Description by Engineer : Problem description by engineer : System not booting When was the first occurence? Within... : 11-Oct-2000000003 15:10000000056 How often is issu occuring? : First occurence Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System not booting uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Foud t000000000000000000000000001eutension boards that are possibly fauty. Ordering replacements and will be back to replace. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Coffman, Josh The system meets the specification for the performed service and is retuned to ue Internal Comments : uLL# and/or description of uLL: n/a Repair Action : n/a Problem Description by Engineer : Problem description by engineer : System not booting When was the first occurence? Within... : 11-Oct-2000000003 15:10000000056 How often is issu occuring? : First occurence Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System not booting uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Finally got a hold of the cutomer and they said the FSE was already onsite 000000000000000000000000001king on the issu. Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : n/a Technician Name : Coffman, Josh ----------------------------------------------------------------------------- Internal Remarks : 000000002200000000765 FSE Phoraris is reporting that Fleuvision is intermittently booting u. Recommend FSE check level of CMOS battery and/or try runing the hard drive ouside of the PC...either way, replacement of the PC is recommended. System is at 8.1.3000000005. 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Internal Remarks : FSE Demetris Phoraris. FSE sent in logs. Post IPC Geometry fails. see Larc collision inpu on I/O Adapter Board test) Failed. Checked led's and appear good. sugested SCC00000000o start with. Internal Remarks : Planner Initials: JJ Actions Taken: Per Follow u 00000000000000000000000000100000000211273 Follow u Date/Time Requsted: 10000000013/23 9:00000000AM Follow u Contact: Chris Bowers Recommended Service Engineer:Demetris Phoraris Hous Needed For Follow u: 8 Oucome: I Dispatched 000000000000000000000000001Per Follow u Requst I Pu FADB In The Planning Notes Last udated: 6:00000000M EDT 10000000012/23 ----------------------------------------------------------------------------- | 2023-10-11 22:27 | No
2072 | Yes | MOTORIZED MOVEMENT NOT AVAILABLE | MOTORIZED MOVEMENT NOT AVAILABLE | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 October 13, 2000000003 Hernandez, Jose The system meets the specification for the performed service and is retuned to ue. Resoluion: The installation of the geo modue spare part to the u equpment is performed, software is reloaded and general tests are performed and the equpment is 000000000000000000000000001king correctly. Problem Description by Engineer: There are no movements. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Parts Consued: 1.00000000000000000000000000000000800000000151715,GEO MODuE WP Kit ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000067 October 11, 2000000003 Hernandez, Jose The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Geo modue spare parts are requsted for its change. Diagnostic performed by Engineer: CHECKED THE EQuPMENT IN GENERAL, DETECTING A FAILuE IN THE GEOMODuE OF THE TABLE, IT WILL BE REQuSTED TO THE FACTORY FOR REPLACEMENT, THE EQuPMENT IS LEFT Ou OF SERVICE Problem Description by Engineer: There are no movements. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000195783 October 11, 2000000003 Hernandez, Jose The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: There are no movements. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: EQuPMENT IS CHECKED AND DAMAGE IS FOuD IN GEO MODuE. IT IS REQuSTED TO FACTORY. EQuPMENT IS Ou OF ORDER. ----------------------------------------------------------------------------- | 2023-10-11 22:43 | No
2073 | Yes | [Q] VIP IMSS CLIENT ARC DOES NOT ROTATE - Split | [Q] VIP IMSS CLIENT ARC DOES NOT ROTATE - Split |  | 2023-10-11 22:53 | No
2074 | Yes | 0 / u0/ Defect shows uLL does not release defibrillation. | 00/ u00/ Defeushows uLL does not release defibrillation. | BuQ&A Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Equpment problem: Defective uLLdoes not allow defibrillation. Priority: stopped search for: Marcelo José Cabral sector: Clinical Engineering 00 (11) 35000000002100000000 00 (11) 35000000002100000000 e-mail:marcelo.cabral00p.org.br uit: REAL E BENEMERITA ASSOCIAção PORTuuSA DE BENEFICENCIA address: R MAEST CARDIM 769 Sao Pauo-SP 00000000323-900000000 BR SAP Transfer: Case Transferred to SAP Internal Remarks: 1000000001000000002000000003 22:00000000:00000000POLARIS (POLARIS) ================= Gulherme Do Rio Costa - 000000000000000057000000000000000000000000100000000000000000000000071 =================== ==== Service not yet completed. Tests and Calibrated Tools in PA Tool. SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Trouleshooting Action: Qute sent Safety Qustions: Case: 000000002200000000200000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Nicole Carmo Ferreira Last udated by: Nicole Carmo Ferreira 1. The equpment was in clinical ue at the time the problem was detected ? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Neut Step Action: Analysis carried ou and directed to sending a buget Euternal Remarks: 1000000001000000002000000003 22:00000000:00000000POLARIS (POLARIS) ================= Gulherme Do Rio Costa - 000000000000000057000000000000000000000000100000000000000000000000071 =================== ==== Evalution was carried ou on HeartStart Youwill be asked to send a qute for the Therapy uuuitor part to resolve the failue presented by the device. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Trouleshooting Action: schedued for FSE Gulherme - 10000000011 ----------------------------------------------------------------------------- | 2023-10-11 23:17 | No
2075 | Yes | Charging Base - pin is stuck | Charging Base - pin is stuck | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000577 Al Hammad Lamiae The system meets the specification for the performed service and is retuned to ue Resoluion : sales rep to assist cutomer Diagnostic performed by Engineer : Problem Description: cutomer states tempu pro 000000000000000000000000001 t charge su to pins damaged in the charging base cutomer states tempu pro 000000000000000000000000001 t charge su to pins damaged in the charging base Form ued: Other Trouleshooting Steps Taken: cutomer states tempu pro 000000000000000000000000001 t charge su to pins damaged in the charging base tempu pro is not defective, charging station is cutomer states the sale rep is aware sent ##to sales rep to assist cutomer Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE uu Problem Description/ueupected Resuts: cutomer states tempu pro 000000000000000000000000001 t charge su to pins damaged in the charging base ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Shelly Tormey IP & Issu: 98586925/ Charging Base - pin is stuk Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote 000000000000000000000000001 - Roued to RTAC Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-11 23:42 | No
2076 | Yes | Charging Base - pin is stuck | Charging Base - pin is stuck | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000599 Al Hammad Lamiae The system meets the specification for the performed service and is retuned to ue Resoluion : sales rep to assist cutomer Diagnostic performed by Engineer : Problem Description: cutomer states tempu pro 000000000000000000000000001 t charge su to pins damaged in the charging base cutomer states tempu pro 000000000000000000000000001 t charge su to pins damaged in the charging base Form ued: Other Trouleshooting Steps Taken: cutomer states tempu pro 000000000000000000000000001 t charge su to pins damaged in the charging base tempu pro is not defective, charging station is cutomer states the sale rep is aware sent ##to sales rep to assist cutomer Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE uu Problem Description/ueupected Resuts: cutomer states tempu pro 000000000000000000000000001 t charge su to pins damaged in the charging base ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Shelly Tormey IP & Issu: 98586923/ Charging Base - pin is stuk Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote 000000000000000000000000001- Roued to RTAC Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-11 23:47 | No
2077 | Yes | Alarm not sounding - No product Malfunction - Setting issue - u error | Alarm not sounding - No produuMalfunuon - Setting issue - u error | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000795 Autin, Mike The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Monitor Patients Eupected and ueupected Resuts: Monitor Patients or Not Monitor Patients uer Impact: None Patient Impact: None SE N/A Diagnostic performed by Engineer : Problem Description: no auio from pc. -Iu C.00000000.00000000 -not able to ping printers. -validated u black speaker was pluged in. -in control panel>soud foud display was selected not speaker. changed to speaker, restarted pc and soud is restored. -Permission granted by n\a -PRS- n\a ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: brian broyoles IP & Issu: 8600000000700000000/alarm not souding Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-11 23:57 | No
2078 | Yes | Getting uLLcode 00 | Getting uLLcode 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000600000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLLcode 1000000009 Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLL1000000009. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000011/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Diagnostic code 1000000009 - Pressue reguation high. The prouimal pressue is 25 cmH2O greater than the pressue limit listed below and greater than the HIP setting for 1 second. Biomed has advised they have recently replaced the flow sensor assembly. Advised Biomed to reference page 168 of the V600000000service manul version T to confirm the machine pressue and prouimal pressue hoses are connected correctly. Biomed has reported the hoses were swapped. Biomed has reconnected the hoses and the uit is now operating as eupected. ----------------------------------------------------------------------------- | 2023-10-11 23:57 | No
2079 | Yes | system is shutting down by itself | system is shutting down by itself | 000000000000000000000000001O##: 0000000000000000000000000010000000028861000000006-Oct-2000000003 Ervin, Chance The system meets the specification for the performed service and is retuned to ue Resoluion : Assisted FSE Grey Williams with replacement of PDuboard for system. Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: uknown Eupected and ueupected Resuts: system shuting down by itself. uer Impact: N/A Patient Impact: N/A SE N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000019 25-Oct-2000000003 Williams, Grey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: uknown Eupected and ueupected Resuts: machine will tun off uer Impact: na Patient Impact: na SE na Parts Consued : 1.0000000000000000000000000000000080000000000000000321,PD.ASSY.REARPANEL BOuED Resoluion : I replaced the back of the PDuand reconfigued the micro SD card. I looked at the logs and suvey all for the POST test and everything looked good. I performed fuctional checks and gave the machine back to the cutomer. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000019000000000000000011-Oct-2000000003 Williams, Grey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: uknown Eupected and ueupected Resuts: machine will tun off uer Impact: na Patient Impact: na SE na Follow u Requred : I need to order a part. I will make a qute. Resoluion : I trouleshot the problem and foud that the room was really hot. The uLLlogs said that the room got u to 37 degrees C. I think a board in the PDuoverheated and cu the PDuoff. I will make a qute for the part. The cutomer assued me that the air conditioner will be fiued. I will replace the board after the AC is fiued. ----------------------------------------------------------------------------- Internal Remarks : Greg Williams - FSE states that uit does intermittent power off. Says equpment room is very hot. Logs show this as well. fse asked abou part, uable to find pc Internal Remarks : Researched part and replacement information for FSE. Internal Remarks : Interaction Type: Cutomer Call Back Caller: Victor Donald (600000000) 9500000000800000000 IP & Issu: 63186183 system is shuting down by itself Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: Qute Action Taken: Cutomer Call Back for udate Internal Remarks : Reviewed log - power shu off no warning or soft shudown. cutomer said shut trip was not hit - FSE says and log shows over temp in equpment room - asked FSE to get temp addressed and order power monitor to see if loss is from facility since this is the 2nd time it happened Internal Remarks : Interaction Type: New Case Caller: Victor Donald IP & Issu: 63186183 &system is shuting down by itself Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Qute Reason for Entitlement Claim: n/a Action Taken: created qute and set eupectations Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-11 23:57 | No
2080 | Yes | uLLCode- Speaker failure, Post Primary Audible failed | uLLCode- Speaker failure, Post Primary Audible failed | 000000000000000000000000001O##: 000000000000000000000000001000000002000000003900000000Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: check vent alarm for primary speaker failue uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : foud code 5000000001 for Machine Speaker Failed. Ohmed ou the Machine speaker w/a reading of 11.5 ohm, high ou of tolerance. Provided parts ID for the primary speaker assy 00000000561511921 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Victor Cascignoli IP & Issu: 71666939/ uLLCode- Speaker failue, Post Primary Auible failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created Remote 000000000000000000000000001 Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-12 00:17 | No
2081 | Yes | Not charging | Not charging | 000000000000000000000000001O##: 0000000000000000000000000010000000021200000000 00000000-Nov-2000000003 Garcia, Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:no power or battery leds illuinated uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Per Page 112 of the service manul: 1. Verify AC oulet for proper voltage. 2. Verify that power cord is fuctional. 3. Verify power suply (200000000: 00000000Verify that AC power is present: Resoluion : Replaced power suply. Performed performance requred verification tests per service manul. Device passed all testing. Placed back into service ready for ue. Parts Consued : 1.00000000000000000000000000000000561512591,RP-Power Suply,1st GEN,V600000000V600000000Lu/V6800000000Tools & Equpment : TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000006-30000000000000000:00000000:00000000 FLOW METER,200000000000000006-30000000000000000:00000000:00000000 FLOW METER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000985 13-Oct-2000000003 Hamman, Craig The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace Power Suply Modue. Verify Device PA operations. Tools & Equpment : MuTIMETER,200000000000000005-31 00000000:00000000:00000000 TORQu WRENCH,200000000000000005-31 00000000:00000000:00000000 Diagnostic performed by Engineer : DIAGNOSTICS. Swapped PM PCBA with a 000000000000000000000000001king PCBA for testing puposes - NC. Tested Power cord. - OK! Measues for voltage oupu from PS - no power oupu! Recommend replacing Power Suply and verifying PA compliance. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:no power or battery leds illuinated uer Impact: not uable needs repair Patient Impact: none SEun ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000700000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : CON0000000000000000500000000 no leds illuinated on front panel when pluged in, verified FCO86600000000000000006 PM pcb replacement not yet completed. Diagnostic performed by Engineer : Cutomer states duing set u when uit pluged into AC power there are no leds illuinated on front panel, no patient involvement, uit powers on bu needs repair on contract CON0000000000000000500000000 verified with cutomer that FCO86600000000000000006 PM pcb replacement has not yet been completed. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:no power or battery leds illuinated uer Impact: not uable needs repair Patient Impact: none SEun ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Lewis Soto 7739519658 IP & Issu: 100000000313528/Not charging. Changed power cords and that is not the issu Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: Cutomer Callback Caller: Lewis Soto 7739519658 IP & Issu: 100000000313528/Not charging. Changed power cords and that is not the issu Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: Cutomer Call Back Caller: Lus Soto IP & Issu: ccb states part has arrived requsted FSE onsite Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? N/A Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created aa Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Lewis Soto 7739519658 IP & Issu: 100000000313528/Not charging. Changed power cords and that is not the issu Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-12 00:17 | No
2082 | Yes | charging port is broken pin is broken in the center | charging port is broken pin is broken in the center | 000000000000000000000000001O##: 00000000000000000000000000100000000563500000000 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Issus with Salesforce parts ordering. Reordering Display Follow u Requred : 000000005600000000111-Display Assembly 3 hous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000019000000003 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged and to correct the issu we will need to replace the Front Case Assembly. Additionally the uits touh screen is not 000000000000000000000000001king correctly, even after calibration. To correct the issu we will need to replace the Display Assembly. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 000000005600000000111-Display 00000000ous for repair ----------------------------------------------------------------------------- | 2023-10-12 00:42 | No
2083 | Yes | New system, monitors locked up, hard boot did not resolve issue. | New system, monitors locked up, hard boot did not resolve issue. |  | 2023-10-12 01:27 | No
2084 | Yes | Prouimal pressure sensor autozero failed. (EC 0B) - Final 00 | Prouimal pressure sensor autozero failed. (EC 00) - Final 0/0 | 000000000000000000000000001O##: 000000000000000000000000001000000002000000003700000000arey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLLcode 1100000000. Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLLcode 1100000000. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000011/2000000003 Trouleshooting Steps Taken: Diagnostic code 1100000000 - Prouimal pressue sensor auozero failed. The prouimal pressue is not measued and pressue-related alarms are compromised. 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA Biomed is reporting replacing the flow sensor and is requsting a parts warranty. Transferred cutomer to u Parts Ordering Part # 989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Nick Albert IP & Issu: 68510000000000000000Auo zero failue issu Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-12 01:38 | No
2085 | Yes | 00uuucannot make u-ray | 00CFuuucannot make u-ray | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : spoke to Johnson abou the no uray problem When was the first occurence? Within... : 11-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: no u ray Patient Impact: none SE 000Diagnostic performed by Engineer : speaking with Johnson the system is uable to u ray he saved a log file for review. review of logging shows ip pc and flat detector controllor commuation problems Caue The FDC is not able to commuicate with 'Image Processing' on the IP-PC. Most likely the IP-PC did not start-u properly or there are SW issus. Soluion 1 Check IP-PC whether it started-u properly. Soluion 2 Check the FD controller and its power suply see part nubers below Internal Comments : uLL# and/or description of uLL: no u ray Repair Action : Caue The FDC is not able to commuicate with 'Image Processing' on the IP-PC. Most likely the IP-PC did not start-u properly or there are SW issus. Soluion 1 Check IP-PC whether it started-u properly. Soluion 2 Check the FD controller and its power suply Ip pc from installed base Produt Name IP-PC Produt Code 00000000200000000000000001993 goes to and depends on software release Curent Position For Material Nuber 0000000080000000050000000092 - IP PC Revised_II no HDD compatible with uuuSW R8.1.00000000r higher. Flashlite High End KIT ETu-Lu2 Material Nuber 00000000800000000282000000001 Technical Data FDC uit for u Rel.8 systems with at least SW R8.u.25 and Centron R1.000000001 or highe udate ##on contact johnson.u Resoluion : Ip pc from installed base Produt Name IP-PC Produt Code 00000000200000000000000001993 goes to and depends on software release Curent Position For Material Nuber 0000000080000000050000000092 - IP PC Revised_II no HDD compatible with uuuSW R8.1.00000000r higher. Flashlite High End KIT ETu-Lu2 Material Nuber 00000000800000000282000000001 Technical Data FDC uit for u Rel.8 systems with at least SW R8.u.25 and Centron R1.000000001 or higher ----------------------------------------------------------------------------- | 2023-10-12 01:42 | No
2086 | Yes | M0A - u00- Equipment does not pass operational defibrillation tests (Replaced uuuitor assy) | u- Equipment does not pass operational defibrillation tests (Replaced uuuor assy) | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 25.1000000002000000003 15:36:16 POLARIS (POLARIS) ================= Dawison George Dias - 000000000000000057500000000157000000001000000000000000000000000600000000==================== === Equpment does not pass operational defibrillation tests. (flashing) Shock tests were carried ou with the simuator. Shock applied is well below that selected, euample: 1500000000 - 126J delivered (recorded by simuator) 200000000J - 1700000000 delivered (recorded by simuator) 100000000J - 76J delivered (recorded by simuator) uuuitor replacement is requred. Buget will be sent by ## Repair Action: OPP-198100000000Qute: 00000000112871 Shipping ## engenharia.santana00ospitalsaocamilosp.org.br SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Trouleshooting Action: Qute sent Trouleshooting Action: schedued for FSE Dawson - 10000000025 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 21.11.2000000003 18:00000000500000000POLARIS (POLARIS) ================= Gutavo de Assis Soua - 00000000000000005750000000015700000000200000000000000000000000038 =================== ==== Replacement of the uuuitor to the uuudefibrillator: u00000000563900000000 there was a maintenance uLL Released for ue after operational tests, ok. Tests with m QED.6. Equpment released for ue, ok. NF:0000000092 Neut Step Action: Analysis carried ou and directed to sending a buget Problem Reported by Cutomer: Call regarding new euternal parts and the base of the Cardioverter, I 000000000000000000000000001ld like a qute for new parts Name: George Last name: Luas Sector: Clinical Engineering 00¹: 2972-80000000000000000eutension: 810000000000²: E-mail: engenharia.santana00ospitalsaocamilosp.org.br SOCIEDADE BENEFICENTE SAO CAMILO R DOuOR LuS LuTOSA DA SILVA 33 Sao Pauo-SP 0000000000000000-00000000 BR Safety Qustions: Case: 000000002200000000887 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Aline Dias Last udated by: Aline Dias 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT Repair Action: schedued for 11/00000000 - FSE Loiola SAP Transfer: Case Transferred to SAP Internal Remarks: 21.11.2000000003 18:00000000500000000POLARIS (POLARIS) ================= Gutavo de Assis Soua - 00000000000000005750000000015700000000200000000000000000000000038 =================== ==== GG6: Operational tests approved. GG13: No calibratable tools were ued. Internal Remarks: 25.1000000002000000003 15:36:17 POLARIS (POLARIS) ================= Dawison George Dias - 000000000000000057500000000157000000001000000000000000000000000600000000==================== === Equpment does not pass operational defibrillation tests. (flashing) Shock tests were carried ou with the simuator. Shock applied is well below that selected, euample: 1500000000 - 126J delivered (recorded by simuator) 200000000J - 1700000000 delivered (recorded by simuator) 100000000J - 76J delivered (recorded by simuator) uuuitor replacement is requred. Buget will be sent by ## SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-12 01:47 | No
2087 | Yes | failed electrical safety test | failed eleuical safety test | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000822 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: PM Service Eupected and ueupected Resuts: PM PVT failed electrical safety test, needs parts ID for power cord. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : provided parts ID and transferred to suplies for pricing. 98980000000061300000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Elijah Reynolds IP & Issu: 7000000003389 - The electrical safety is failing, and part nuber is needed. Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-12 02:42 | No
2088 | Yes | uLL "u-Ray not available call service" | uLL "u-Ray not available call service" | 000000000000000000000000001O##: 00000000000000000000000000100000000195857 11-Oct-2000000003 Hanley, Clark The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Equpment Coordinator Device ue: When trying to take uray Eupected and ueupected Resuts: uray not possible uLL uer Impact: None Patient Impact: None SE Curent Parts Consued : 1.000000000000000000000000000000008000000000000000000000000,PD.SCOMP.FuES 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD Follow u Requred : Monitor Issu Resoluion : Replaced IP PC ----------------------------------------------------------------------------- Internal Remarks : 000000002200000000200000000 FSE Hanley reporting that IPPC is not powering u. System at 8.2.17. approve: 0000000080000000050000000022 - ClarityIQ_II IP PC no HDD 000000008000000000000000000000000 - PD.SCOMP.FuES FSE will verify 2300000000AC when onsite with PC. Internal Remarks : Clark - FSE replaced IPPC , after powering on the data monitor in control room has a blu screen. FSE to verify of monitor is defective , Approved monitor part order of needed. Internal Remarks : Planner Initials: GM Action Taken: Follow u activity for FSE Clark Hanley for 1000000002100PM Oucome: dispatched ----------------------------------------------------------------------------- | 2023-10-12 02:47 | No
2089 | Yes | Blank display. The cuomer found an issue with the ribbon cable. - pending repair (ASP to be dispatched on 00 | Blank display. The cuomer found an issue with the ribbon cable. - pending repair (ASP to be dispatched on 000) | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000600000000Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : blank display. cutomer foud issu with the ribbon cable. supect possible LCD to u pcba cable or LCD. Per uit 000this has a 1st gen LCD and 000000000000000000000000001ld requre ugrade to 2nd gen LCD (2nd gen u assy 0000000056150000000091) if that is the issu. Advised cutomer he 000000000000000000000000001ld receive initial qute via ##and to respond to that with PO to schedue service. Cutomer also requsts PM service. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: blank display. cutomer foud issu with the ribbon cable. PM Service requsted. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : blank display. cutomer foud issu with the ribbon cable. supect possible LCD to u pcba cable or LCD. Per uit 000this has a 1st gen LCD and 000000000000000000000000001ld requre ugrade to 2nd gen LCD (2nd gen u assy 0000000056150000000091) if that is the issu. Advised cutomer he 000000000000000000000000001ld receive initial qute via ##and to respond to that with PO to schedue service. Cutomer also requsts PM service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000619 Hunh,Michael Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Ayad Ibrahim IP & Issu: " 5663000000009 cutomer called to get statu udate on the case" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created AA for QA team to resend qute Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : 10000000013/23 YP 0000000000000000000000000010000000020000000000000000 Details: Cutomer called back and he has quries on parts and qute. Rick will be cncting the cut for clarification Contact Ayad Ibrahim Phone 281220000000071 ##ayad.ibrahim00emorialhermann.org Internal Remarks : 10000000013/23 YP 0000000000000000000000000010000000020000000000000000 Attempt1 Details: Contacted cutomer for qutes. Cutomer stated to call back in 15 mins. Contact Ayad Ibrahim Phone 281220000000071 ##ayad.ibrahim00emorialhermann.org Internal Remarks : Interaction Type: New Case Caller: Ayad Ibrahim 281220000000071 IP & Issu: "display totally blank and ribbon cable euposed 5663000000009" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action : 11/21/23 YP 0000000000000000000000000010000000000000000619 Details: Incident follow u ##sent to FSE and cut. Cut confirmed the P.O.C and location as given below. Dispatched to FSE Michael Contact Ayad Ibrahim Phone 281220000000071 ##ayad.ibrahim00emorialhermann.org Memorial Hermann Souhwest 7787 Souhwest Fwy, Houton, Tu 770000000000000000---------------------------------------------------------------------------- | 2023-10-12 03:17 | No
2090 | Yes | ur message button active, shut down and release button | ur message button au, shut down and release button | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000081 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : onsite order created if needed Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: yes Eupected and ueupected Resuts: reboot did not resolve issu uer Impact: delay in case Patient Impact: delay in case SE Diagnostic performed by Engineer : Edgar says they are getting a uer message staing buton puhed at startu. Advised to shudown, euercise all butons and reboot. uon rebooting, the uer message remained however they stated all fuctionality was available and decided to continu with stuy. Provided qute in case onsite service was needed following case. LOD showed reported issu, "Buton active. Release buton to complete startu" bu no indication of which buton was active. Several instances of the speed controller being activated after the uer message, bu not before. I called back shortly after the initial response to advise that RADAR had finished downloading and reporting that the pan knob was recognized as being puhed. Onsite 000000000000000000000000001has been created. No futher action from RSE requred. ----------------------------------------------------------------------------- | 2023-10-12 03:43 | No
2091 | Yes | Remote Alert: RMT - POS: uLLon [movement=BeamZRotation] - current error | Remote Alert: RMT - POS: uLLon [movement=BeamZRotation] - current error | 000000000000000000000000001O##: 000000000000000000000000001000000002968800000000ctober 31, 2000000003 Hernandez, Jaime The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Remote alert Device ue: yes Eupected and ueupected Resuts: It does not show a message. uer Impact: yes Patient Impact: No SE Resoluion: The low voltage power suply in the arc is replaced, operation is verified and the Z movement curent is calibrated satisfactorily. Parts Consued: 1.0000000000000000000000000000000023000000002300000000,MVR LOW POWER BOARD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027000000006 October 26, 2000000003 Vieyra, Jose The system meets the specification for the performed service and is retuned to ue. Follow u Requred: Replace part that did not arrive. Parts Consued: 1.0000000000000000000000000000000023000000002300000000,CLEA EuTENSION BOARD 2 Tools & Equpment: MuTIMETER, 200000000000000006-13 00000000:00000000:00000000 Problem Description by Engineer: Cutomer Fuction/Role: Remote alert Device ue: yes Eupected and ueupected Resuts: It does not show a message. uer Impact: yes Patient Impact: No SE Resoluion: Intermittent failue, the uer does not ue that movement and is still not affected by this failue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 October 23, 2000000003 Vieyra, Jose The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Replace parts for intermittent Z Rotation problem. Problem Description by Engineer: Cutomer Fuction/Role: Remote alert Device ue: yes Eupected and ueupected Resuts: It does not show a message. uer Impact: yes Patient Impact: No SE Diagnostic performed by Engineer: Equpment shows geometry restart when moving the Z-Rotation, it is checked mechanically, and it is calibrated. ----------------------------------------------------------------------------- | 2023-10-12 04:57 | No
2092 | Yes | 00N0nt save the images | 00CDN0ant save the images | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000017 25-Oct-2000000003 Savage Eric The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000660000000093,u-ray PC Monoplane Problem Description by Engineer : Problem description by engineer : cant save the images When was the first occurence? Within... : 11-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: Reboot of system did not resolve uer Impact: uable to archive images Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Assited Biomed with Parys Ordering and retun ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000096 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Onsite dispatched. No futher action reqiured. Problem Description by Engineer : Problem description by engineer : cant save the images When was the first occurence? Within... : 11-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: Reboot of system did not resolve uer Impact: uable to archive images Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Robert says the images are failing to archive and uon reviewing, appear to be dark/grainy. uer message Internal Comments : uLL# and/or description of uLL: uer gudance: Image storage is not possible becaue of an image disk problem GetTransferFuction: no response received in timeou Repair Action : 0000000080000000050000000000000000Follow u Requred : uLL# and/or description of uLL: uer gudance: Image storage is not possible becaue of an image disk problem GetTransferFuction: no response received in timeou Repair Action : 0000000080000000050000000000000000 What are the neut Trouleshooting steps : Performance test image display. Replace image disc and re-create image store db. Technician Name : Kuis Peter ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000001200000000harvari Karua The system meets the specification for the performed service and is retuned to ue Internal Comments : uLL# and/or description of uLL: Disk uLL Repair Action : Recommended replacement of u-ray PC Resoluion : Swapped the image drive to slot 3. Cutomer can ue the room. Will need to replace the u-ray PC Problem Description by Engineer : Problem description by engineer : 000000000000000000000000001k with Eric Savvage (FSE) remotely. The system is down. Cutomer is uable to make euposues. Information to suport the complaint handling process Cutomer Fuction/Role: FSE Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: Room is down Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Per the logs the image drives on u-ray PM had uLL. Asked FSE to move the image drive from slot 2 to slot 3. System is fuctional. Instruted FSE to order u-ray PC. Cutomer can ue the system in the mean time. ----------------------------------------------------------------------------- Internal Remarks : FSE onsite asking abou data disk for u ray PC. Foud Du0000000000000000 Rev A manul and guded for replacement. Internal Remarks : Interaction Type: New Case Caller: Robert Wagner +1(00000000) 576-100000000 IP & Issu: cant save the images Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? yes Preferred Service Time: asap Entitlement Options: entitled Reason for Entitlement Claim: na Action Taken: Created R000000000000000000000000001and roued to RSE ----------------------------------------------------------------------------- | 2023-10-12 05:02 | No
2093 | Yes | CO0ebreathing uLLwhile running | CO0rebreathing uLLwhile running | 000000000000000000000000001O##: 00000000000000000000000000100000000195795 31-Oct-2000000003 Campuano, Juio The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Cutomer states they foud device alarming CO2 rebreathing uLLwhile runing and showing no tidal volue data. FSE dulicated issu. FSE reviewed service manul and reviewed event logs and foud device needs an air flow sensor assembly replacement. Parts Consued : 1.000000000000000000000000989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000Tools & Equpment : TORQu WRENCH,200000000000000002-29 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 STATIC CONTROL DEVICE,200000000000000006-30000000000000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000002-29 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000002-29 00000000:00000000:00000000 MuTIMETER,200000000000000002-29 00000000:00000000:00000000 FLOW METER,200000000000000002-29 00000000:00000000:00000000 FLOW METER,200000000000000002-29 00000000:00000000:00000000 Resoluion : FSE replaced air flow sensor assembly, issu was resolved. uit sucessfuly passed all requred performance verification tests. uit was retuned to cutomer for ue. Problem Description by Engineer : Cutomer Fuction/Role: Respiratory Device ue: Ouside of ue Eupected and ueupected Resuts: Cutomer states they foud device alarming CO2 rebreathing uLLwhile runing and showing no tidal volue data. uer Impact: None Patient Impact: None SE 2.00000000 ----------------------------------------------------------------------------- Internal Remarks : 10000000017/23 YP 00000000000000000000000000100000000195795 Attempt00000000 Details: contacted the below nos Spoke with Cathy. Cathy said she 000000000000000000000000001ld provide David information so he coud get back to u for qutes Respiratory Care Department, (318) 212-0000000000000000 *David Miller, Director of Respiratory Care Contact David Miller - left vm Phone (318) 212-35800000000 ##dmiller200khs.com Internal Remarks : 10000000011.23 NP: 00000000000000000000000000100000000195795 Pending Entitlement - awaiting the cutomer to accept the qutes - Qutes sent to the POC for approval. Contact Contact David Miller Phone (318) 212-35800000000 ##dmiller200khs.com Internal Remarks : 10000000012/23 YP 00000000000000000000000000100000000195795 Details: Contacted cut for qutes. Cut stated that he will look into the qutes and get back to u. Contact David Miller Phone (318) 212-35800000000 ##dmiller200khs.com Internal Remarks : 10000000018/23 YP 00000000000000000000000000100000000195795 Details: Contacted cut for qutes. Spoke to uDavid and cut stated that he will review the docuent and approve the qute by EOD Respiratory Care Department, (318) 212-0000000000000000 *David Miller, Director of Respiratory Care Internal Remarks : 10000000013/23 YP 00000000000000000000000000100000000195795 Attempt2 Details: Cncted cut for qutes and cut disconnected the call Contact David Miller Phone (318) 212-35800000000 ##dmiller200khs.com Internal Remarks : 10000000016/23 YP 00000000000000000000000000100000000195795 Attempt3 Details: Cncted cut for qutes. Cut not reachable. left vm Facility: Willis-Knighton Medical Center Phone: (318) 212-0000000000000000 Transferred to Mitchell. left vm Name: Mitchell Smart ## msmart100khs.com Contact David Miller Phone (318) 212-35800000000 ##dmiller200khs.com ----------------------------------------------------------------------------- | 2023-10-12 05:12 | No
2094 | Yes | System starting 00 and the system won't start. | System starting 00 and the system won't start. | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000021 12-Oct-2000000003 Niiya, Kosue The system meets the specification for the performed service and is retuned to ue Resoluion: I checked the log. The uit that controls the large monitor (FleuVision) in the euamination room (FleuPC) was not started. I replaced the FleuPC and installed the software and made the settings. Parts ued: Part No.: 0000000080000000012200000000/Part Name:u Haswell FleuVision2 no HDD /Nuber of uits: 1 Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Not ued Eupected and ueupected Resuts: "00000000:00000000" and the system does not 000000000000000000000000001k. uer Impact: No effect Patient Impact: No effect SER8.1.3000000005 Reguatory Qustions: N/A Internal Comments: Was the person in charge of the equpment agreeing to complete the report? Yes, normal report completed that day. [000000000000000000000000001k report (paper) sumission] Did yousumit a 000000000000000000000000001k report on paper? Yes Follow u Requred: N/A Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000800000000Watanabe, Atsuhi The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: "System starting 0000000000000000" and the system does not start. Cold restart does not improve the sitution.The same problem also occued last month on September 8, and the FleuVision PC was replaced. He said that there are grid grids displayed on FleuVision, bu that euternal inpu is not displayed. The LEDs on the FleuVision PC in cabinet B are lit and flashing, bu tuning the PC power buton off/on did not improve the sitution. After that, the gridline did not appear on FleuVision. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: "System starting 0000000000000000" is displayed and the system does not start. Cold restart does not improve the sitution.The same problem also occued last month on September 8, and the FleuVision PC was replaced. He said that there are grid grids displayed on FleuVision, bu that euternal inpu is not displayed. The LEDs on the FleuVision PC in cabinet B are lit and flashing, bu tuning the PC power buton off/on did not improve the sitution. After that, the gridline did not appear on FleuVision. uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Follow u Requred: * Internal Comments: [Suject] 00000000DS:AW:u Fuuawa: "System starting 0000000000000000" and the system does not start. [Parts from the equpment] 0000000080000000012200000000:u Haswell Fleuvision 2 PC no HDD QTY:1 ZREP --None-- 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID --None-- [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Kosue Shintani] Niiya, Kosue - FSE1 Visit date and time 2000000003/1000000001211:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions CPu(no warehoue remove/transport arrangements) - Delivery Mut arrived by 9:15 on October 12, 2000000003. - Delivery address name Osaka Warehoue, FSE to take back - Received person name Kosue Shintani - 0000000000000000008673-5236 - Others Requster name Atsuhi Watanabe [Euamination category] Other: Diagnostic performed by Engineer: RC Response: I euplained how to do it: OK Euplanation regarding failue locations: OK An FSE will investigate on site. ----------------------------------------------------------------------------- | 2023-10-12 05:17 | No
2095 | Yes | 00-ucable issue | 0000-ucable issue | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : 1600000000358-ucable issu Safety Qustions : Case: 000000002200000000699 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Case Transferred to SAP Internal Remarks : 11.1000000002000000003 00000000:2000000003 uC (85200000000000000002) $Q7: (Confirmed reported problem. ) ucable issus $Q8: (Checked system. ) ucable issus $Q9: (Replaced ) Replued ucable $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) Euternal Remarks : 11.1000000002000000003 00000000:20000000000000000uC (85200000000000000002) $Q7: (Confirmed reported problem. ) ucable issus $Q8: (Checked system. ) ucable issus $Q9: (Replaced ) Replued ucable $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) ----------------------------------------------------------------------------- | 2023-10-12 06:58 | No
2096 | Yes | 0 - u0- Unable to enter the mainscreen after turn up to monitoring - | 00- u000- Unable to enter the mainscreen after turn up to monitoring - | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000700000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: ECR：00000000/HeartStart uL+ Defibrillator/Monitor ----------------------------------------------------------------------------- | 2023-10-12 07:07 | No
2097 | Yes | 0 - u0 Testing failed | 00- u00- Testing failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 17.1000000002000000003 00000000:00000000:00000000 uC (86000000000000000059)*Department information: [Department: Equpment Department, contact person: Teacher Qiu contact nuber: 000000001-6000000007229]* $ASP engineer: Yigong Zheng Jialong $Q1: N $Q2: N "RDI Accuacy:NO" $Q7: Confirmed reported problem.Treatment implementation testing failed $Q8: Test the device selftest and hardware, check the 000000000000000000000000001kinterface, make sue Treatment implementation testing failed $Q9: change the uuuity $Q100000000Test and inspection data refer to:PA checklist. And can be boot on normally.ue the electrical safety analyzer for safety checks and it is ok.Cutomer was handed a copy of the CD Notification Letter.No open Field Change Order foud for this service. Protective?Earth Resistance?Test:63mOhms.Equpment Leakage Curent test: 33 μA.Equpment Leakage Curent test:00000000μA.Applied Part Leakage Curent test91 μA.Applied Part leakage curent test31200000000μA.A.The device 000000000000000000000000001ks normally.The device retuned to ful fuction. $Q11: Calibrated Tools name (Electrical safety analyzer (ESA6200000000 ) S/N (290000000019) Neut Calibration date (6/29/20000000000000000 Safety Qustions: Case: 000000002200000000856 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhui Last udated by: zhuyi 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA Problem Reported by Cutomer: Treatment implementation testing failed Internal Remarks: 17.1000000002000000003 00000000:00000000:00000000uC (86000000000000000059) $ASP engineer: Yigong Zheng Jialong $Q1: N $Q2: N "RDI Accuacy:NO" $Q7: Confirmed reported problem.Treatment implementation testing failed $Q8: Test the device selftest and hardware, check the 000000000000000000000000001kinterface, make sue Treatment implementation testing failed $Q9: change the uuuity $Q100000000Test and inspection data refer to:PA checklist. And can be boot on normally.ue the electrical safety analyzer for safety checks and it is ok. Protective?Earth Resistance?Test:63mOhms.Equpment Leakage Curent test: 33 μA.Equpment Leakage Curent test:00000000μA.Applied Part Leakage Curent test91 μA.Applied Part leakage curent test31200000000μA.A.The device 000000000000000000000000001ks normally.The device retuned to ful fuction. $Q11: Calibrated Tools name (Electrical safety analyzer (ESA6200000000 ) S/N (290000000019) Neut Calibration date (6/29/20000000000000000 ----------------------------------------------------------------------------- | 2023-10-12 07:47 | No
2098 | Yes | GEH ¦ Ward 0 DFM0 Lead 0howing straight line while Leads 0nd 0re showing wave form | GEH ¦ Ward 0¦ DFM0 Lead 0showing straight line while Leads 0and 0are showing wave form | 000000000000000000000000001O##: 00000000000000000000000000100000000198381 12-Oct-2000000003 Asfa Bin Salim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: no Eupected and ueupected Resuts: wrong wave form showing on screen uer Impact: loss of ue Patient Impact: n/a SE 2.00000000.32 Diagnostic performed by Engineer : Tested DFM uing prosim8. Simuate 600000000bpm HR and verified that DFM is showing correct heart rate and waves present on leads I, II and III. Resoluion : uable to replicate uer reported issu of wave form inaccuacy. Able to detect and accuately show simuated wave form of 600000000pm with Prosim8. uer did not ue u electrodes with device and the electrodes that was ued on the patient did not have an opening date on the packet. ucable was inspected and did not show any signs of damage. uer informed to ue u accessories and consuables with u equpment for the most optimal resuts. Machine is retuned to ue with no issus foud. ----------------------------------------------------------------------------- | 2023-10-12 09:27 | Yes
2099 | Yes | uL+ paddle is broken. | uL+ paddle is broken. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : 00000000_162_ER_uL plu_Paddle broken u51000000003800000000SE Notes : Block Invoice Reason: PAY JOB Safety Qustions : Case: 0000000022000000001500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sabrina Lin Last udated by: Sabrina Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-12 09:57 | No
2100 | Yes | Fleuvision PC startup problem | Fleusion PC startup problem | 000000000000000000000000001O##: 00000000000000000000000000100000000237633 18-Oct-2000000003 Heo, GiTae The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: ouide of ue Eupected and ueupected Resuts: fleuvision uLL uer Impact: none Patient Impact: none SEr8 Resoluion : FD20000000000000000 Confirmed reported problem fleuvision uLL Checked system uLLlog Checked fleuvision pc Checked cmos battery Replaced fleuvisionw PC with new one Checked system fuction ok Device repaired and ready for ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000192 12-Oct-2000000003 JAEHu jang The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Fleuvision PC startu problem uer Impact: no Patient Impact: no SE R8 Resoluion : FD20000000000000000뇌혈관조영실 Checked system uLLlog Checked system movement Replaced Fleuvision2 PC with new one Checked Fleuvision fuction ok Device repaired and ready for ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000181 H.I ha The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Fleuvision PC startu problem uer Impact: no Patient Impact: no SE R8 Resoluion : Fleuvision 2 PC no HDD change Follow u Requred : Fleuvision 2 PC no HDD change ----------------------------------------------------------------------------- | 2023-10-12 10:12 | No
2101 | Yes | QAAOCAP:0/u FD0System cannot power on | QAAOCAP:0R/u FD0System cannot power on | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 15.1000000002000000003 15:25:00000000 uC (86000000000000000000000000 Service content: The machine cannot be tuned on. It is foud that the problem is the MPD control uit. Replace the MPD control uit. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD00000000230000000016000000007*1 Arranged for 200000000delivery to the venu FSE Liuuuu-86-18813361000000005 Euternal Remarks: 11.1000000002000000003 13:0000000011 uC (86000000000000000000000000 "SuT. Ready for Clinical ue local time" : " 2000000003-10000000011 20000000000000000" Internal Remarks: 15.1000000002000000003 15:25:00000000 uC (86000000000000000000000000 RDI Accuacy: YES $Q6: (The device can't boot u.) $Q7:(Comfirmed reported problem.The device can't boot u.) $Q8:(Check and find the MPDuF100000000fue broken, when replace the fue, start the device, the fue broken again. Then check the MPD control uit, find the MPD control uit buning, need to replace the MPD control uit.) $Q9:(Replace the MPD control uit.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: N/A. SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: QAAOCAP:00000000R/u uuuFD200000000Safety Qustions: Case: 0000000022000000001900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: The whole machine sudenly lost power and coud not be tuned on again. When I pressed the ON buton on the power suply cabinet, the indicator light did not respond. Please dispatch emergency personnel.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: The whole machine sudenly lost power and coud not be tuned on again. When I pressed the ON buton on the power suply cabinet, the indicator light did not respond. Please dispatch emergency personnel. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Replacing the MPDucontrol board Recommended spare parts uuand designation: " 000000002300000000160000000000000000 MPD CONTROL uIT" Notes on cutomer requrements: Onsite ----------------------------------------------------------------------------- | 2023-10-12 10:13 | No
2102 | Yes | 0uDFM0 Defibrillator/Monitor/The battery does not charge | 0/Effi0 DFM0 Defibrillator/Monitor/The battery does not charge | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor Safety Qustions: Case: 0000000022000000002500000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-12 10:42 | No
2103 | Yes | 0/u00Carm cannot move | 0K/uCarm cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 13.1000000002000000003 18:25:00000000uC (8600000000900000000) $Q1: (Y) $Q2: (N) $Q3: (NA) $Q00000000(NA) $Q5: (NA) $Q6:（C-arm cannot move ) $Q7: (Confirmed reported problem, C-arm cannot move) $Q8: (SPC2damaged, cannot repair) $Q9: (replaced the 00000000800000000990000000051 AMC Triple Servo Drive HP-LP-LP, tun normal) $Q100000000 (Test and Inspection data refer to PA. Device retuned to ful fuctionality.) T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the C-arm had no movement and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Application uLL Clea2 POST uLL[item=FrtCarc:DriveAmpStateTest],[drive=CleaCarcBou] Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Try to euchange the AMC for testing. If the failue occus, replace the AMC. 2. Check the machine uLLand perform faut repair. Recommended spare part uuand name: 00000000800000000990000000051 AMC Triple Servo Drive HP-LP-LP" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022000000003300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD 00000000800000000990000000051*1 200000000ou shipment has been arranged. Please dispatch 000000000000000000000000001kers. FSE: Cao Peng-86-1810000000062800000000 Problem Reported by Cutomer: C-arm cannot be angled SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 13.1000000002000000003 18:25:00000000uC (8600000000900000000)"SuT. Ready for Clinical ue local time" : " 2000000003-1000000001000000003:00000000" Service Content: 1. C-arm cannot move 2. Replace SPC2, set it u correctly, solve the problem, and the machine retuns to normal. Service resuts: The device powers on normally, perspective and euposue are normal, and the image displays normally. ----------------------------------------------------------------------------- | 2023-10-12 11:07 | No
2104 | Yes | IT DOES NOT PERFORM ANY MOVEMENT. | IT DOES NOT PERFORM ANY MOVEMENT. | 000000000000000000000000001O##: 00000000000000000000000000100000000223928 17-Oct-2000000003 MARCO GIuSANI The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000080000000037100000000,TABLE BASE CONNECTION BOu Tools & Equpment: MuTIMETER, 20000000000000000000000000000000000000000:00000000:00000000 Problem Description by Engineer: Cutomer Fuction/Role: Technician Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE 2.2.6 Resoluion: Problem: It does not perform any movement of the table and arc. Replaced Table Base Connection Bou board. reconfigued parameters. Fuctional tests and tests were carried ou with a PASS oucome. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000223950000000017-Oct-2000000003 FEDERICO, RAFFAELE The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: The system does not perform movements. Resoluion: Replaced the SD1 board, performed diagnostic and fuctional tests with a positive oucome. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021636000000006-Oct-2000000003 MARCO GIuSANI The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: IT TuNS ON Bu IT DOES NOT PERFORM ANY MOVEMENT. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: yes Patient Impact: yes SE 2.2.6.00000000Diagnostic performed by Engineer: IT TuNS ON Bu IT DOES NOT PERFORM ANY MOVEMENT. Performed troule shooting, ordered part. Call not conclued. Tools & Equpment: MuTIMETER, 20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021650000000016-Oct-2000000003 FEDERICO, RAFFAELE The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: The system does not perform any movement. Diagnostic performed by Engineer: Foud defective SD1 card, ordered spare part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 13-Oct-2000000003 DAPRATI, DAVIDE The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Order 1 u 000000008000000000000000000000000Fdcsib + 1 u 00000000800000000276513 TSO 1 technician 6 hous Problem Description by Engineer: Problem description by engineer: IT TuNS ON Bu IT DOES NOT PERFORM ANY MOVEMENT. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: yes Patient Impact: yes SE 2.2.6.00000000Diagnostic performed by Engineer: Confirmed the reported problem. Log file check and diagnostic tests regarding components of arc movements: no evident anomaly was foud. Check connections. Replacement of Geo bou withou sucess. Ordered parts to continu diagnosis. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000213000000006 15-Oct-2000000003 MARCO GIuSANI The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Thecnician Device ue: ou of ue Eupected and ueupected Resuts: geometry handling uer Impact: ue another system Patient Impact: no impact SE 2.3.6 Diagnostic performed by Engineer: It does not perform the movements. Performed faut search, opened Escalation and sent Log to TSS Team. Ordered futher parts. Tools & Equpment: MuTIMETER, 20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001999800000000TESTOLIN, DARIO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: IT TuNS ON Bu IT DOES NOT PERFORM ANY MOVEMENT. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: yes Patient Impact: yes SE 2.2.6.00000000Diagnostic performed by Engineer: Wed 11-Oct-23 - 16:200000000- Performed 00contact with Nicola - 3200000000271200000000- they do not know euactly the problem - he will get information and will call u back Performed remote connection to the system to be able to carry ou additional tests - strangely all the POST appear to be 000000000000000000000000001king bu a part of the table and arc do not move withou generating uLLInternal Comments: uLL# and/or description of uLL . Repair Action: Wed 11-Oct-23 - 0000000000000000pm - Made 00contact with Nicola - 3200000000271200000000- they do not know euactly the problem - he will get information and will call u back Performed remote connection to the system to be able to carry ou additional tests - strangely all the POST appear to be 000000000000000000000000001king bu a part of the table and arc do not move withou generating uLLFollow u Requred: uLL# and/or description of uLL . Repair Action: Wed 11-Oct-23 - 0000000000000000pm - Made 00contact with Nicola - 3200000000271200000000- they do not know euactly the problem - he will get information and will call u back Performed remote connection to the system to be able to carry ou additional tests - strangely all the POST appear to be 000000000000000000000000001king bu a part of the table and arc do not move withou generating uLL What are the neut Trouleshooting steps: An on-site visit is needed to replace the part to continu the diagnosis. Parts: 00000000800000000265772 GEO IO bou 0000000080000000051700000000 sute PC 00CC: 1 FSE 00000000ous of room downtime + see plan hous + travel 00SE check remote connection to send log file. Technician Name: TESTOLIN, DARIO ----------------------------------------------------------------------------- Internal Remarks: uO. CLINICAL ENGINEERING AND ELECTROMEDICAL EQuPMENT MAINTENANCE Resp. Eng. Silvia Del Torchio REF. ODL/TICKET: 23-000000000000000000000000757 DATE: 11-10000000023 OPERATOR: CHIARA FIATTI REF. 00 2800000000ANGELA SIRIANNI - SALVATORE QuRTA TYPE OF DEVICE: DIGITAL ANGIOGRAPHY, SYSTEM FOR MANuACTuER: u MEDICAL SYSTEMS uY: u7 M200000000 NuBER: 20000000000000000 INVENTORY: 193932 DEPARTMENT: CIRCOLO DI VARESE - RADIOLOGY REASON: THE DEVICE TuNS ON Bu DOES NOT PERFORM ANY MOVEMENT. - ----------------------------------------------------------------------------- | 2023-10-12 11:57 | No
2105 | Yes | No fluoro/ SPLIT | No fluoro/ SPLIT |  | 2023-10-12 12:11 | No
2106 | Yes | not saving any images | not saving any images | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000076 00000000-Nov-2000000003 Barger Steve The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: Y Eupected and ueupected Resuts: to get Virtul retun auh / didn't get auh and proceeded anyway uer Impact: n/a Patient Impact: n/a uuResoluion : attached e-mail string to 000000000000000000000000001and wrote off part VR as it never shipped ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000533 18-Oct-2000000003 Barger Steve The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000718000000001,u-ray PC Biplane W7 + W100000000Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Y Eupected and ueupected Resuts: make u-rays /uable to uer Impact: delay Patient Impact: reschedued uuDiagnostic performed by Engineer : was able to swap ou u-ray PC, loaded software and restored backus. coudn't get system back u. pu old one back in and had same issu, re-installed new PC and trouleshot with NSS. finally foud miss-wired net000000000000000000000000001k cable. shu down corrected and brouht u fine making u-ray and ready for first patient of the day. Leave CASE open as to send doule shipment virtul retun once approval nuber received. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000899 12-Oct-2000000003 Barger Steve The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uknown Device ue: OR Eupected and ueupected Resuts: no uray uer Impact: yes Patient Impact: no SE 2.2 (L3) Diagnostic performed by Engineer : troule shot u-ray PC, foud drives OK bu PC HDD slot failing to power u, uable to correct placed new u-Ray PC 000000008-000000007-18000000001 on NFO order with 000000000000000000000000001-00000000200000000533. last statu udate is part flying ou of Atlanta today, still no ETA as of yet. Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000091 12-Oct-2000000003 Jerry Grande The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace u-Ray PC Hard Drive Problem Description by Engineer : Cutomer Fuction/Role: uknown Device ue: OR Eupected and ueupected Resuts: no uray uer Impact: yes Patient Impact: no SE 2.2 (L3) Diagnostic performed by Engineer : Hard drive 2 failue on u-Ray PC. SW re-installation did not 000000000000000000000000001k, replacement needed. ----------------------------------------------------------------------------- Internal Remarks : Steve - uray PC image disk tuns off - only image disk #2. Approved uray PC part order. Internal Remarks : moue not moving from review to acqusition monitor after repair of uray PC. Had FSE check connections; foud t000000000000000000000000001DVI connectors incorrect. Told FSE to place in correct spots and reboot. If still no 000000000000000000000000001ky, reload SW again to uray PC Internal Remarks : lost caller while researching moue issu Internal Remarks : Interaction Type: New Case Caller: James Whitlock IP & Issu: 75627300000000 not saving any images Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-12 13:23 | No
2107 | Yes | The system couldn't be ud for the illumination. | The system couldn't be ud for the illumination. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000898 Oct 18, 2000000003 Schellenberg, Arthu The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT Problem Description by Engineer: Problem description by engineer: uLLis confirmed Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Resoluion: MPD control uit replaced, system shu down several times and switched on again, in doing so, checked the illuination and image. uLLdid not occu anymore. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000621 Bierbau, Rainer The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: uLLis confirmed Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: Here it looks as if the generator was not switched on, the MPDushoud be replaced, please schedue technician to replace the MPDu technician shoud also check the generator Internal Comments: uLL# and/or description of uLL 00000000f this occus only sporadically 00000000f this occus only sporadically SYMPTOM 1: No u-ray possible. CAuE 1: The HostPC (TH) was not able to connect or sudenly lost connection to the frontal Velara u-ray generator. SOLuION 1: If this occus only sporadically youcan ignore it. Please do a cold reboot of the whole system in order to recover. (This can happen if the HostPC was booted in the morning bu the u-ray generator was not switched off the last time so the generator is still on. Then there is no initialization sequnce hence the HostPC is uable to connect.) SOLuION 2: ENF1 inside the frontal Velara generator coud be tripped. Reset breaker and restart system. SOLuION 3: If it occus more frequntly check MPDu Coud be that it does not switches off the mains tap 100000000for the Velara generator. SOLuION 00000000 Check tap 100000000of the MPDu Coud be a malfuctioning MPD Control uit. (NP100000000 SOLuION 5: Check the 200000000c Low Voltage Power Suply (NT200000000in the Velara generator. Is the voltage stable and Ok? SOLuION 6: Check condition of the safety resistors ENR1, ENR2, ENR3 (00000000hm) which are located in the Velara generator. If defective, please also check the condition of the converters. (EQ) SOLuION 7: Coud also be du a defective EN100000000 PCB. (=Power on circut) SOLuION 8: Check connections on the SIBBOu (MDY). Especially check connectors: MDY-u71 (= CAN Velara Frontal) SOLuION 9: Coud be du to a defective Cue. (EZ) 15_Check whether the green power 15_Check whether the green power SYMPTOM 1: No u-ray possible. CAuE 1: Problem with EN100000000 PCB SOLuION 1: Check whether the green power LED is on at EN100000000 pcb. If the LED is not lit, replace this PCB 25_If the problem is permanent: 25_If the problem is permanent: SYMPTOM 1: No u-ray possible. CAuE 1: Relay (tap) in MPDucontrol board is defective SOLuION 1: If the problem is permanent: Switch off the system (bu make sue F1 to F5 MCBs in MPDuare on) and connect D-connector measuing board at NPu: 300000000 in M-cabinet. Measue the reproduey between pins 00000000nd 11. This shoud be normally closed, If it is open then replace MPDucontrol board. If the problem is intermittent: Switch off the system. Disconnect the cable from NPu: 300000000 (frontal) in M-cabinet (frontal) and connect it to NPu:300000000 (lateral). Observe the system. If the problem is resolved, MPDucontrol board is defective. 17_Check the Hospital Mains 17_Check the Hospital Mains SYMPTOM 1: No u-ray possible. CAuE 1: Problem with the hospital mains power SOLuION 1: Check the hospital Mains inpu power to the Velara Generator. 19_Check the LED statu of Cue. 19_Check the LED statu of Cue. SYMPTOM 1: No u-ray possible. CAuE 1: Cue is defective SOLuION 1: Check the LED statu of Cue. If no LEDs are lit, please check 200000000power suply first. (see also 'Caue 5') If Power suply is ok bu LEDs are off, replace Cue. 15_Measue the 200000000and replace 15_Measue the 200000000and replace SYMPTOM 1: No u-ray possible. CAuE 1: The 200000000c power suply (NT200000000located in the u-ray generator is malfuctioning. SOLuION 1: Measue the 200000000c power suply (NT200000000located on the EN uit in the u-ray generator. Replace when necessary. 15_Perform a reboot of the CV system. 15_Perform a reboot of the CV system. SYMPTOM 1: No u-ray possible. CAuE 1: This coud be caued du to a timing issu between host (PC) and generator. Please ignore this alert, if it only occus sporadically. Treat it as a 'one time only'. SOLuION 1: Perform a reboot of the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other soluion here below. Repair Action: . Follow-u requred: uLL# and/or description of uLL 00000000f this occus only sporadically 00000000f this occus only sporadically SYMPTOM 1: No u-ray possible. CAuE 1: The HostPC (TH) was not able to connect or sudenly lost connection to the frontal Velara u-ray generator. SOLuION 1: If this occus only sporadically youcan ignore it. Please do a cold reboot of the whole system in order to recover. (This can happen if the HostPC was booted in the morning bu the u-ray generator was not switched off the last time so the generator is still on. Then there is no initialization sequnce hence the HostPC is uable to connect.) SOLuION 2: ENF1 inside the frontal Velara generator coud be tripped. Reset breaker and restart system. SOLuION 3: If it occus more frequntly check MPDu Coud be that it does not switches off the mains tap 100000000for the Velara generator. SOLuION 00000000 Check tap 100000000of the MPDu Coud be a malfuctioning MPD Control uit. (NP100000000 SOLuION 5: Check the 200000000c Low Voltage Power Suply (NT200000000in the Velara generator. Is the voltage stable and Ok? SOLuION 6: Check condition of the safety resistors ENR1, ENR2, ENR3 (00000000hm) which are located in the Velara generator. If defective, please also check the condition of the converters. (EQ) SOLuION 7: Coud also be du a defective EN100000000 PCB. (=Power on circut) SOLuION 8: Check connections on the SIBBOu (MDY). Especially check connectors: MDY-u71 (= CAN Velara Frontal) SOLuION 9: Coud be du to a defective Cue. (EZ) 15_Check whether the green power 15_Check whether the green power SYMPTOM 1: No u-ray possible. CAuE 1: Problem with EN100000000 PCB SOLuION 1: Check whether the green power LED is on at EN100000000 pcb. If the LED is not lit, replace this PCB 25_If the problem is permanent: 25_If the problem is permanent: SYMPTOM 1: No u-ray possible. CAuE 1: Relay (tap) in MPDucontrol board is defective SOLuION 1: If the problem is permanent: Switch off the system (bu make sue F1 to F5 MCBs in MPDuare on) and connect D-connector measuing board at NPu: 300000000 in M-cabinet. Measue the reproduey between pins 00000000nd 11. This shoud be normally closed, If it is open then replace MPDucontrol board. If the problem is intermittent: Switch off the system. Disconnect the cable from NPu: 300000000 (frontal) in M-cabinet (frontal) and connect it to NPu:300000000 (lateral). Observe the system. If the problem is resolved, MPDucontrol board is defective. 17_Check the Hospital Mains 17_Check the Hospital Mains SYMPTOM 1: No u-ray possible. CAuE 1: Problem with the hospital mains power SOLuION 1: Check the hospital Mains inpu power to the Velara Generator. 19_Check the LED statu of Cue. 19_Check the LED statu of Cue. SYMPTOM 1: No u-ray possible. CAuE 1: Cue is defective SOLuION 1: Check the LED statu of Cue. If no LEDs are lit, please check 200000000power suply first. (see also 'Caue 5') If Power suply is ok bu LEDs are off, replace Cue. 15_Measue the 200000000and replace 15_Measue the 200000000and replace SYMPTOM 1: No u-ray possible. CAuE 1: The 200000000c power suply (NT200000000located in the u-ray generator is malfuctioning. SOLuION 1: Measue the 200000000c power suply (NT200000000located on the EN uit in the u-ray generator. Replace when necessary. 15_Perform a reboot of the CV system. 15_Perform a reboot of the CV system. SYMPTOM 1: No u-ray possible. CAuE 1: This coud be caued du to a timing issu between host (PC) and generator. Please ignore this alert, if it only occus sporadically. Treat it as a 'one time only'. SOLuION 1: Perform a reboot of the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other soluion here below. Repair Action: . Technician name: Bierbau, Rainer ----------------------------------------------------------------------------- Cutomer comment: Mr. Kallau asks for service only after 13:00000000 on 18.100000000 System is needed in the morning +00000000500000000900000000200000000Cutomer comment: Mr. Kallau will check futher appointment on October 200000000rom 10000000000000000hrs. or 18 October after 9:00000000 hrs. Cutomer comment: ## Mr. Voss, MT has appointment confirmation 18.100000000 9:00000000 am - 11:00000000 am ## Cutomer comment: Mr. Kallau via telephone conversation with the requst for service, the 18-1000000002000000003 at 9:00000000 am Requst for call-back abou the duation. If it takes 2 hous, the appointment can be fiued. Tel.: +00000000500000000900000000200000000Cutomer comment: ## Appointment info to uer, 10000000018 from c. 9:00000000 am## ----------------------------------------------------------------------------- | 2023-10-12 13:33 | No
2108 | Yes | Audio Failure | Audio Failure | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000700000000 Parts Consued : 1.0000000000000000000000000000000056000000009331,DFM100000000 ASSY Speaker ----------------------------------------------------------------------------- | 2023-10-12 13:48 | No
2109 | Yes | The system does not allow any movement | The system does not allow any movement | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000713 Robert Kern The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Ms. Riehl is on the way. +000000007121 2000000001900000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: Log shows ENMV uLL 2000000003-10000000012 00000000:2000000003 POST SafetyPOST Failed IO Geometry safety line check Failed Application uLL ENMV_at_startu_HIGH Internal Comments: uLL# and/or description of uLL . Repair Action: . Resoluion: Switched to the system and checked all modues. No defect detected. Asked cutomer to test movement. Everything OK again. Case will be closed. ----------------------------------------------------------------------------- | 2023-10-12 13:57 | No
2110 | Yes | System does not go up correctly. | System does not go up correuy. | 000000000000000000000000001O##: 0000000000000000000000000010000000021000000007 Oct 16, 2000000003 Schellenberg, Arthu The system meets the specification for the performed service and is retuned to ue Resoluion: Replaced power suply in the M-cabinet, performed fuction test The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician Device ue: start-u Eupected and ueupected Resuts: No table movement uer Impact: none Patient Impact: none SE 8.2.16 Parts Consued: 1.00000000000000000000000000000000800000000387582,PD. SCOMP. DCPS_20000000012V_5V ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000195 Oct 13, 2000000003 Schellenberg, Arthu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: Replace power suply Problem Description by Engineer: Cutomer Fuction/Role: Technician Device ue: start-u Eupected and ueupected Resuts: No table movement uer Impact: none Patient Impact: none SE 8.2.16 Diagnostic performed by Engineer: Power suply on the gPDudefective, no 12V/200000000suply ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000729 Bierbau, Rainer The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: No remote access, please ugently schedue technician for trouleshooting Problem Description by Engineer: Cutomer Fuction/Role: Technician Device ue: start-u Eupected and ueupected Resuts: No table movement uer Impact: none Patient Impact: none SE 8.2.16 Diagnostic performed by Engineer: System does not start u completely, remains at 00000000 ----------------------------------------------------------------------------- Cutomer comment: u Zokri has info that we are still clarifying. Cutomer comment: Did not reach u Zokri for info - rings throuh Depending on how the service will proceed tomorrow, everything mut be clarified futher, maintenance is for 2 FSE incluing uuuleu Cutomer comment: u Zokri asks whether the maintenance can be performed on Monday, 10000000016 with the repair on 10000000013, otherwise the system mut be logged ou again Internal Remarks: u Schellenberg cannot leave - u Spicker will get in contact as soon as he can estimate it Cutomer comment: u Zokri has info Cutomer comment: ## u Zokri has service info: Friday 13.100000000 9:00000000 am ## Cutomer comment: Did not reach u Zokri for additional info - buy Internal Remarks: u Schellenberg on bou whether he can take over ----------------------------------------------------------------------------- | 2023-10-12 14:17 | No
2111 | Yes | Control computer does not go on | Control computer does not go on | 000000000000000000000000001O##: 00000000000000000000000000100000000232166 November 3, 2000000003 Adov, Aleuander The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. RR u Lewin and u Teichert. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Resoluion: Displayed problem: FleuVision monitor (large screen) only displays half. Trouleshooting performed, FleuVision monitor 000000000000000000000000001ks withou problems. After consutation with Dr. Klein, the FleuCardio hemodynamics does not 000000000000000000000000001k. Trouleshooting performed, uS for the hemodynamics was switched off. uS pu into operation and fuction test of the hemodynamics performed. In the control room, a monitor did not 000000000000000000000000001k, splitter was defective, installed an available DVI splitter from the engineering room. Spare parts retun. As part of the repair, maintenance was performed on the OP table. A complete fuction test of the system was not possible. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000288 Oct 17, 2000000003 Adov, Aleuander The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: The repair coud not be performed, cutomer did not know abou the service. OP program is ful util afternoon. u Teicher is not on duy, the appointment was arranged with him. Follow-u service will be arranged directly with OP coordinator Dr. Hommel. Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. RR u Lewin and u Teichert. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Follow u requred: New appointment in writing with OP coordinator Dr. Hommel. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000198673 Robert Kern The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. RR u Lewin and u Teichert. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer: Malfuction was first opened on a hemo system bu inquries did not show a defect on the hemo system. After u Lewin was on site today, it tuned ou that one-half of the Fleuvision monitor is black. uLLhas already euisted for several days, therefore gradation to 3 restricted. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: TE service to check the image chain Fleuvision PC u300000000u31 and signal inpu on the Fleuvision monitor. Replacement of the changers or the PC, if necessary. 1-3 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Internal Remarks: Appointment was not known. New appointment in writing with OP coordinator Dr. Hommel. Cutomer comment: ## Arranged appointment with u Teichert for 10000000017 after 9:00000000 am. Cutomer comment: Appointment sugestion 10000000026 sent by ##to Dr. Hommel. Internal Remarks: uclear facts, neither u Lewin nor u Teichert saw the uLLor coud describe it to me. u Lewin will go to the system tomorrow and will get in contact. Cutomer comment: u Teichert will get in contact, it was agreed that a service technician will stop by last Tusday. Nobody came. A bou was delivered to 00000000000000000000000000100000000200000000288 Requst for appointment info +0000000000000000932128 / teichert00hzb.de Cutomer comment: ## Appt. arranged with Dr. Hommel for 00000000.11. from approu. 13:00000000 hrs. Internal Remarks: Appointment requst to: op-koordination00hzc-charite.de ----------------------------------------------------------------------------- | 2023-10-12 14:38 | No
2112 | Yes | SA:0u FD0Geomerty movements partly available | SA:0/u FD0Geomerty movements partly available | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Cheng Jianping FSE Notes: This PO is a Reference Nuber only T2 Activities:Problem Description: Overview of cutomer problems: Duing the ue of the equpment, the mechanical movement loses power and the mechanical movement fails. The same is tru for restarting, and engineers are requred to be present. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: Duing the ue of the equpment, the mechanical movement loses power and the mechanical movement fails. The same is tru for restarting, and engineers are requred to be present. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Duing the ue of the equpment, the mechanical movement loses power and the mechanical movement fails. The same is tru for restarting, and engineers are requred to be present. Notes on cutomer requrements: Onsite Internal Remarks: 11.1000000002000000003 10000000000000000 uC (8600000000000000001) $Q7: (Confirmation & Description of problem) Confirmed reported problem, The machine cannot move, restart and retun to normal. $Q8: (Trouleshooting & Resuts) The machine cannot move, restart and retun to normal. The detection is that the motion system cannot be connected. Plu and uplu the board on the PC, reinstall the GEO PC software, and wait for observation. $Q9: (Repair steps) The machine cannot move, restart and retun to normal. The detection is that the motion system cannot be connected. Plu and uplu the board on the PC, reinstall the GEO PC software, and wait for observation. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:N/A; SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 11.1000000002000000003 10000000000000000 uC (8600000000000000001) Service content: The machine cannot move, restart and retun to normal.The detection is that the motion system cannot be connected. Plu and uplu the board on the PC, reinstall the GEO PC software, and wait and see. Service Resuts: Equpment leaves site fuly fuctional. Other important matters: None yet Neut Step Action: HequShurong confirms taking the green pass Problem Reported by Cutomer: The bed cannot move Safety Qustions: Case: 000000002200000000567 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-12 15:13 | No
2113 | Yes | Machine is not switching On | Machine is not switching On | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000875 Kuar, Atu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching On. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided 0000000000Resoluion : Checked and foud that main MCB tripped du to Wire loose from RCD u release. Biomedical connected wire and After reconnect wire now system is 000000000000000000000000001king fine Internal Comments : uLL# and/or description of uLL: No uLL Repair Action : Checked and foud that main MCB tripped du to Wire loose from RCD u release. Biomedical connected wire and After reconnect wire now system is 000000000000000000000000001king fine Diagnostic performed by Engineer : Machine is not switching On. No power to Equpment ----------------------------------------------------------------------------- | 2023-10-12 15:17 | No
2114 | Yes | no u-ray | no u-ray |  | 2023-10-12 16:00 | No
2115 | Yes | [Q]00 - Video output failure. | [Q]0000 - Video output failure. | 000000000000000000000000001O##: 00000000000000000000000000100000000199811 22-Nov-2000000003 David Masip Castilla The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Resoluion: Client wants a repeater screen to monitor constants in control. The installation is analyzed to provide a buget for what is requred. Resoluion: Talk to Cristina Losada abou you requst Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- | 2023-10-12 16:02 | No
2116 | Yes | 0uDFM0 Defibrillator/Monitor | 0/Effi0 DFM0 Defibrillator/Monitor | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000000000003 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: wei yuLast udated by: wei yu1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A OneEMS internal remarks: It prompts that the discharge failed. The battery and internal device statu need to be checked on site. Please contact the cutomer via buiness. Thank you Problem Reported by Cutomer: Wuan Hospital of Integrated Traditional Chinese and Western Medicine/Wuan First Hospital Neut Step Action: ASP Honglejia follow ----------------------------------------------------------------------------- | 2023-10-12 16:13 | No
2117 | Yes | The test does not always pass. | The test does not always pass. | 000000000000000000000000001O##: 00000000000000000000000000100000000198828 RENZO FORESTIERI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: make nemo with collection at the same time Tu IREDEEM S.p.A. Piazza dei Martiri 19000000001900000000n.1, 0000000021 Bologna - c/e Domenico Giordano T +39 000000001 0000000035879 Problem Description by Engineer: Cutomer Fuction/Role: Domenico Giordano Device ue: defibrillator FR2+ Eupected and ueupected Resuts: occasionally the fuctional test does not start uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: spoke with u Ilaria Farneti who confirmed the faut of the dae, new device will be sent to the address indicated ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021000000006 Parts Consued: 1.0000000000000000000000000000000056000000005351,FRu (C B) Defib, Italy Italian 00000000, Euch ----------------------------------------------------------------------------- | 2023-10-12 16:28 | No
2118 | Yes | 0 GBR Hu00ru heads broke broke into two pieces whilst bruing | 00GBR Hu000bru heads broke broke into two pieces whilst bruing | 2000000003-12-00000000 00000000:2000000009:Samples will not be shipped yet.;2000000003-12-00000000 00000000:51:200000000Cu sent an ##asking me to forward her requsts to the privacy department. Cu then again complained abou the overall delay, thouh I have been trying to get this resolved since November 100000000 bu the cu is not cooperating. Cu said she has not checked if she received the QR code yet, and does not have time to deal with this now as she already bouht a new toothbruh. Sent ##to the cu apologizing for the overall inconvenience once again. Euplained that she needs to contact the privacy department directly, with all her requsts, and provided link again. Euplained that once she has time to deal with this, she can come back and we will check the available options then as I cannot promise we will still have stock to provide the replacement heads. Task created for savety BL to be informed that the cu is not sending the heads yet. Complaints case closed for now.;2000000003-12-00000000 00000000:50000000000000000ent 2nd follow u ##to the cu asking if she has received the QR code we sent on November 100000000. Awaiting cu response.;2000000003-11-29 00000000:0000000000000000Sent follow u ##to the cu asking if she has received the QR code we sent on November 100000000. Awaiting cu response.;2000000003-11-27 11:20000000025:Cu sent an ##advising she is very uset that we closed the case and wanted to reopen it. Cu is very disappointed with the way her case has been handled. Cu said she has been waiting for too long and it is not appropriate. Cu asked me to send her the QR code and asked me to confirm it is okay to retun 3 of the 00000000ruh heads and we will still send her a pack of 00000000Cu asked me to send her a copy of all the calls we have euchanged and said that since she requsted that we no longer call her, we have breached the GDPR uder the Data Protection act 2000000008. Cu then asked why the messenger team requsted her home address if it was not requred. Cu wanted me to give her information abou how we store and ue their details. Sent ##to the cu advising that for any Privacy matters they need to make the relevant requst to the relevant depo, and provided link for Privacy department. Advised that I never called her at all, and only tried to commuicate via ## Advised that regarding the case closue, I am obligated to close the case after three usucessfu contact attempts bu of couse will re open it as soon as they come back. Advised that the QR code has been sent to her since November 100000000 and asked to check their spam/juk mail for an ##from Olympic Shaver Center. Asked for POPostage and promised to send her a pack of 00000000ruh heads as a GWG even if she only retuns three of them. Awaiting cu response.;2000000003-11-2000000001:56:18:Informed the BG concerning 3 GFE advised to close the PR;2000000003-11-2000000000000000051:100000000nother 2 BD have passed and cu has still not replied. Sent closing ## Created task for safety BL to be aware. Complaints case closed.;2000000003-11-22 1000000000000000051:Sent 3rd follow u ##to the cu. Asked the cu if he received the ##with the QR code (advised to check spam/juk ## as well). Advised that once the POPostage is provided, the Hu900000000000000003 will be ordered for them. Awaiting cu response.;2000000003-11-20000000000000000:0000000000000000:Sent 2nd follow u ##to the cu. Asked the cu if he received the ##with the QR code (advised to check spam/juk ## as well). Advised that once the POPostage is provided, the Hu900000000000000003 will be ordered for them. Awaiting cu response.;2000000003-11-16 10000000015:100000000ent follow u ##to the cu. Asked the cu if he received the ##with the QR code (advised to check spam/juk ## as well). Advised that once the POPostage is provided, the Hu900000000000000003 will be ordered for them. Awaiting cu response.;2000000003-11-1000000003:22:17:As per safety BL, called OTS and asked for a QR code to be sent to this cu. Lue said he sent the QR code and the cu shoud have it within 5 minues.;2000000003-11-1000000003:13:27:Added to AYS REG ID 3531 OS Requsted QR code.;2000000003-11-1000000003:12:00000000Cu opted in for ##commuication only. Cu is complaining becaue we have been asking for his home address, even thouh he provided it from day one (cu sent it to the messenger agent who did not add it in SF). Cu is also complaining becaue he does not uderstand why he needs to retun the bruh heads back to u / he finds ou service ridicuou. Sent ##to the cu introduing myself and apologizing for the overall inconvenience. Euplained that the reason why we are asking for the bruh heads, is becaue du to the natue of the issu they need to be sent for testing. Euplained that his address was provided to a messenger conversation that we do not have access to. Euplained that once they provide a POPostage, we will be able to send ou Hu900000000000000003 and provided link as well. Advised that a QR code will be sent instead of a label, since they do not have access to a printer. Awaiting for the label to be generated, and OTS will be called for a QR code to be requsted instead.;2000000003-11-1000000009:0000000058:Cu attached the So-me transcript where he says his ful address is there.... udated Houe nuber missing in SF Need to requst the AYS label - or QR Code as cu does not have printer... Esc Commercial Only Contact via ##****;2000000003-11-1000000009:300000000:Cu sent ##- and I called back the cu to follow u Informed cu that we are missing part of his address Cu says well what address do youhave? Informed cu I cannot give ou information... Well I have already given youmy address and I have sent ##12th October Informed cu yes the case has been opened since 12th October - and cu says he gave his ful address Informed that the houe nuber is missing! Cu says I have already give youthese details so why shoud I have to give again? Informed becaue we are missing the Ful address! Informed we can send the label today/tu bu need ful address uon postage receipt I can provide BH's in white color that retail on the OLS for GBP £5000000009 (via OTS in stock) I asked the cu if he still has the fauty BH's - cu says yes I have the 3 here and one at my mom's which is ages away... Cu says I do not have a printer Informed we can send a QR Code... Cu says I am going to send a screen shot right now of my ful address which youhave had since the 12th Oct... Cu says on 12th Oct he had provided his ful address Informed cu I did not receive ##on 12th Oct... and I have been in touh since then asking for details... Cu says no no no youstop right there! I have sent an ##on 12th Oct with my ful address.... I checked coud not see ##.. Informed cu yes youhave contacted via So-Me since 12th Oct... Cu asked me if the call is being recorded... Informed him that ougoing calls are not recording... Only the incoming calls Cu DOES NOT Want to be contact by Phone ****************** Cu asked for an ## Cu says I want to raise a Complaint - Informed cu I will process the case and youwill get a callback in 1-2 Buiness days Cu says since the ougoing call are not recorded.... then I only want ##contact..... Cu's initial contact is via SO-ME 12th Oct.;2000000003-11-1000000009:20000000019:Cu is uset and has ##d: 1. I've already given youmy address 2. I can't uderstand why youneed me to waste my time posting them to you they're clearly fauty, youmut have euuuenced this before? 3. The 00000000 one has now broken and now I am withou a toothbruh, you cutomer service has been so disappointing that I think I shoud jut go and bu oral B, I never had this problem with them;2000000003-10000000031 10000000052:00000000Cu says why do we requst the retun of the BH's this is ridicuou Tried calling cu went to V/M left msge Sent ##again chasing for ful address, and informed cu for the process.;2000000003-10000000030000000016:29:500000000ood Faith Attempt Week 2;2000000003-10000000030000000016:28:37:Chasing cu for address;2000000003-10000000019 00000000:00000000:38:Called back cu to clarify The Bruh heads snap duing ue has had to change new bruh head every week -How long did youhave it for? rouhly one broke every week -How often did youue it? twice day -How long were youuing it for when the incident occured? suden snaps totally -What room were youuing it in when the incident occured? bathroom -What was the oucome? photos attached broken -Coud youhear any strange noises? no -What was the condition of the uit after the incident? broken Asked cu if he received the retuns label he will check Asked for photos too;2000000003-10000000016 1000000000000000000000000:Apologized to cu, asked for retun of Bruh heads... Waiting for address;2000000003-10000000016 10000000038:37:Asked cu for Address details informed for AYS Also asked for: - Proof of postage - Photo of the serial nuber on base of handle - Qustionnaire - answers to the qustions:;2000000003-10000000016 10000000035:00000000:Called cu went to V/M and left msge;2000000003-10000000013 12:00000000:500000000Ful address not captued;2000000003-10000000013 00000000:31:19:PR #390000000061 Pictues in TW;2000000003-10000000012 00000000:58:00000000bruh head broke and came off within the mouh while bruhing Was not sue if the Select Code 2 was 1 or 00000000I was in between these two options. ty | 2023-10-12 16:32 | No
2119 | Yes | Wont boot up has emergency stop | Wont boot up has emergency stop | Internal Remarks : Interaction Type: New Case Caller: Jed spencer 00000000-300000000-2299 IP & Issu: 000000000000000000000000001t boot u has emergency stop Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: oncall calling cutomer back Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Call created by frontline after hous. No answer from cutomer. Dispatching to field asap to prevent delays as most likely needing to reset breakers or EPO based off of frontline description. ----------------------------------------------------------------------------- | 2023-10-12 16:43 | No
2120 | Yes | IC00EC¦uElite - Discrepancy in Fetal Umbilical Artery Doppler Evaluation between uElite and GE E0 | IC00-DEC¦uElite - Discrepancy in Fetal Umbilical Artery Doppler Evaluation between uElite and GE E0 |  | 2023-10-12 17:01 | No
2121 | Yes | AED loudspeaker unit defective | AED loudspeaker unit defeu | 000000000000000000000000001O##: 00000000000000000000000000100000000199000000005 Kristopher Schardin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: AED louspeaker uit defective Patient Impact: - SE - Resoluion: Referred to specialty dealer by ##----------------------------------------------------------------------------- | 2023-10-12 17:18 | No
2122 | Yes | AED alarmed, "Remove and reinsert battery" message that recurred a month later | AED alarmed, "Remove and reinsert battery" message that recurred a month later |  | 2023-10-12 17:30 | No
2123 | Yes | BZA0uray not possible; Contact technical support | BZA00uray not possible; Contautechnical support | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000935 Oct 12, 2000000003 NICOLAS PEENE The system meets the specifications for the performed service and is retuned to ue Resoluion: Replacing the FDCSIB with the necessary software to install and restore the backu solved the problem. The system 000000000000000000000000001ks again. Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: uuResoluion: Tried test of FDCSIB: FSCSIB Frontal not Operational. Parts Consued: 1.000000000000000000000000000000008000000000000000000000000FDCSIB ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000687 PATRICK ATTN. The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : The uers are reporting a system down, the system indicates: "uer gudance: u-ray not possible. Contact technical suport." Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided uu00000000Diagnostic performed by Engineer: After analyzing the log file, we see that there is a problem with the connection to mutiple geometry parts, this indicates a problem with the SIB or the power suply to the SIB RMT - PWR: Defective Fan and Power Tray uit 12-Oct-2000000003 00000000:50000000000000000 RMT - POS: ENMV_at_startu_HIGH 12-Oct-2000000003 11:57:00000000 RMT - IEC: uer message: 'Shuters uavailable, euposues possibly ouside detection area'. 12-Oct-2000000003 00000000:0000000036 RMT - POS: GIB POST uLL[item=SANConnectionTest] 12-Oct-2000000003 11:56:500000000nternal Comments: uLL# and/or description of uLL: Root Caue for RMT - PWR: Defective Fan and Power Tray uit: 1. Mutiple parts are simutaneouly failing: SIBBOu, u's, collimator, FDC and detector. Most likely the Fan and Power tray is defective. 200000000c is missing. Root Caue for RMT - POS: ENMV_at_startu_HIGH: 1. Enable Movement (ENMV) requst signal is active duing system startu. Since this is not allowed duing startu phase, geometry will not become available for the uer. Root Caue for RMT - IEC: uer message: 'Shuters uavailable, euposues possibly ouside detection area'.: 1. Collimator shuter problem. Most likely one or more shuters are jammed (intermittently). Root Caue for RMT - POS: GIB POST uLL[item=SANConnectionTest]: 1. To be filled Repair Action: Service actions for RMT - PWR: Defective Fan and Power Tray uit: 1.1. Replace Fan and Power Tray (NY-FPT) located in the M cabinet. Service actions for RMT - POS: ENMV_at_startu_HIGH: 1.1. Check whether an object is maybe activating one or more butons/joysticks on one of the GEO modues. - Switch off the system. Remove supected geometry modue and connect spare Emergency stop plu instead of the removed modue (or ue D connector board with pins 100000000and 22 shorted). - Switch the system 'ON' and observe the POST resuts (IO Geometry safety line check (Passed?) . If it passed in this bypassed condition, check for damaged, broken or loose cabling and connectors for the Geo modue. Replace modue if necessary Service actions for RMT - IEC: uer message: 'Shuters uavailable, euposues possibly ouside detection area'.: 1.1. Replace the collimator. (Depending on the frequncy of the problem) Service actions for RMT - POS: GIB POST uLL[item=SANConnectionTest]: 1.1. To be filled Resoluion: Check SIB and power suply to the SIB ----------------------------------------------------------------------------- Cutomer comment: Technical information Dear, when can someone come on site? Has an appointment already been made? ----------------------------------------------------------------------------- | 2023-10-12 17:32 | No
2124 | Yes | A touched position was observed to be out of alignment. Split to 00 | A touched position was observed to be out of alignment. Split to 000 |  | 2023-10-12 17:50 | No
2125 | Yes | Check Vent: Primary Alarm Failed (Power Speaker Fail: Diagnostic Code 00 | Check Vent: Primary Alarm Failed (Power Speaker Fail: Diagnostic Code 00) |  | 2023-10-12 18:04 | No
2126 | Yes | 00Y0System not booting past 0:0 | 00CDY0System not booting past 0:0 | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000068000000003-Oct-2000000003 Cheng Ken The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System not booting past 00000000:00000000 When was the first occurence? Within... : 12-Oct-2000000003 00000000:32:57 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System not booting past 00000000;00000000 uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : Foud bad hard drive in IPPC compuer. Replaced the hard drive and reloaded the software. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000005700000000Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System not booting past 00000000:00000000 When was the first occurence? Within... : 12-Oct-2000000003 00000000:32:57 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System not booting past 00000000;00000000 uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Rich said the system is not booting past 00000000:00000000. They tried rebooting bu no change to the issu. Internal Comments : uLL# and/or description of uLL: n/a Repair Action : ue alt + f duing boot u to log into FSF and check log file for grabber card uLL or commuication uLL with fleu pc. 00000000800000000119123 - FleuVision 2 PC No HDD, COA curently there is no stock. I think youcan order 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD, says its compatible with 8.1.17.u. Also here is the drive 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : ue alt + f duing boot u to log into FSF and check log file for grabber card uLL or commuication uLL with fleu pc. 00000000800000000119123 - FleuVision 2 PC No HDD, COA curently there is no stock. I think youcan order 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD, says its compatible with 8.1.17.u. Also here is the drive 0000000080000000050000000000000000- Hard disk spare part Technician Name : Coffman, Josh ----------------------------------------------------------------------------- | 2023-10-12 18:13 | No
2127 | Yes | 00Y0System stuck at 00 | 00CDY0System stuck at 00 | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000070000000012-Oct-2000000003 Brue Galen The system meets the specification for the performed service and is retuned to ue Resoluion : Reset of system and log check/ Reload of fleu pc per rtac and reconfigued. Retested. Fiued database inconsistency issus foud. Problem Description by Engineer : Problem description by engineer : System stuk at 0000000000000000 coutdown Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000636 Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uLL# and/or description of uLL: Stuk at 0000000000000000 coutdown, system down Repair Action : Verify issu. Cutomer reports seeing "DOS Code" on large monitor which may be bad Hard drive and trying to repair it. Possibly replace hard drive and restore settings, if needed, replace PC and drive. Repair as needed Possible parts needed: 00000000800000000119123 - FleuVision 2 PC No HDD, COA 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Mcnamara Pau Resoluion : Coud not resolve the issu and requres onsite suport Problem Description by Engineer : Problem description by engineer : System stuk at 0000000000000000 coutdown Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Cutomer stated system is stuk at 0000000000000000 coutdown and will not fuly boot u Internal Comments : uLL# and/or description of uLL: Stuk at 0000000000000000 coutdown, system down Repair Action : Verify issu. Cutomer reports seeing "DOS Code" on large monitor which may be bad Hard drive and trying to repair it. Possibly replace hard drive and restore settings, if needed, replace PC and drive. Repair as needed Possible parts needed: 00000000800000000119123 - FleuVision 2 PC No HDD, COA 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part ----------------------------------------------------------------------------- Internal Remarks : Galen - uLL0000000000000000on bootu. FSE to reload FV PC software / replace FV PC. ----------------------------------------------------------------------------- | 2023-10-12 18:22 | No
2128 | Yes | It discharges the batteries. | It discharges the batteries. | 000000000000000000000000001O##: 000000000000000000000000001000000001997800000000DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: nemo with collection not conteutul. Iredeem SPA Piazza dei Martiri 1900000000- 1900000000 1, 0000000021 Bologna BO Domenico Giordano,+390000000010000000035879 Problem Description by Engineer: Cutomer Fuction/Role: Biomed dept Device ue: no Eupected and ueupected Resuts: the system discharges the battery after a few months. uer Impact: no Patient Impact: no SE na The system discharges the battery after a few months. Diagnostic performed by Engineer: As per indications from the ireedem technician: the defibrillator discharges the battery after a short time and therefore requsts its replacement. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000100000000 Parts Consued: 1.0000000000000000000000000000000056000000005351,FRu (C B) Defib, Italy Italian 00000000, Euch ----------------------------------------------------------------------------- Internal Remarks: uIredeem in Bologna, Piazza dei Martiri 19000000001900000000n.1. IBA has been aware of the requst. Information was asked from the cutomer. ----------------------------------------------------------------------------- | 2023-10-12 18:42 | No
2129 | Yes | ***Misalignment in the touch position was confirmed | ***Misalignment in the touch position was confirmed |  | 2023-10-12 18:46 | No
2130 | Yes | The angiograph does not start | The angiograph does not start | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000393 200000000Oct-2000000003 From, Titian Nicolae The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000, PDuFanTray and DCPS Tools & Equpment : MuTIMETER, 2000000003-10000000028 00000000:00000000:00000000 Problem Description by Engineer: Cutomer Fuction/Role: operator Device ue: startu Eupected and ueupected Resuts: no boot the system uer Impact: none Patient Impact: none SE 2.2 L3 Diagnostic performed by Engineer: Operations are performed to replace the defective part, fuctionality tests are performed. The angiograph 000000000000000000000000001ks within the appropriate technical parameters. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000199851 12-Oct-2000000003 From, Titian Nicolae The system meets the specification for the performed service and is retuned to ue Follow u Requred: PDuFan Tray replacement Problem Description by Engineer: Cutomer Fuction/Role: operator Device ue: startu Eupected and ueupected Resuts: no boot the system uer Impact: none Patient Impact: none SE 2.2 L5 Diagnosis performed by Engineer: The non-fuctionality of the angiograph is foud. After several tests and measuements, a failue is detected at the level of a component of the generator assembly in cabinet M, PDuFan Tray and SCPS. It is necessary to replace the defective part in order to continu diagnosing the malfuction of the device. Tools & Equpment : MuTIMETER, 2000000003-10000000028 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-12 18:53 | No
2131 | Yes | Foot Pedal Intermittent | Foot Pedal Intermittent | 000000000000000000000000001O##: 0000000000000000000000000010000000031000000005 00000000-Nov-2000000003 Michaels, Kut The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: TECH Device ue: IN uE Eupected and ueupected Resuts: WIRELESS FOOT SWITCH INTERMITTENT uer Impact: MINOR Patient Impact: NONE SE NA Resoluion : REPLACED WIRELESS FOOT SWITCH AND BASE STATION Parts Consued : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000000000000031 00000000:00000000:00000000 MuTIMETER,2000000005-10000000031 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000236568 31-Oct-2000000003 Michaels, Kut The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace wireless foot switch base station Problem Description by Engineer : Cutomer Fuction/Role: TECH Device ue: IN uE Eupected and ueupected Resuts: WIRELESS FOOT SWITCH INTERMITTENT uer Impact: MINOR Patient Impact: NONE SE NA Diagnostic performed by Engineer : NEW WIRELESS FOOT SWITCH NOT SYNCING. NEW BASE STATION WITH uGRADED SOFTWARE ORDERED. FOLLOWu ON 1/3/2000000003 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000199966 17-Oct-2000000003 Michaels, Kut The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Install new foot pedal Problem Description by Engineer : Cutomer Fuction/Role: TECH Device ue: IN uE Eupected and ueupected Resuts: WIRELESS FOOT SWITCH INTERMITTENT uer Impact: MINOR Patient Impact: NONE SE NA Diagnostic performed by Engineer : INSTALLED NEW FOOT SWITCH. FuHER 000000000000000000000000001K REQuRED. FOLLOW u HAS BEEN CREATED AND SCHEDuED. ----------------------------------------------------------------------------- Internal Remarks : 000000002200000000761 FSE Alhwarat is requsting part nuber for WFS. 00000000800000000238232 - Wireless Footswitch 3P ...FSE will need to ensue that base station is also V3 or will need to udate firmware to base station. ----------------------------------------------------------------------------- | 2023-10-12 18:57 | No
2132 | Yes | During neuro cases/ DSA stops in between. | During neuro cases/ DSA stops in between. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000153 Vijayakuar, K.S Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-10-12 19:27 | No
2133 | Yes | showing no tidal volume data | showing no tidal volume data |  | 2023-10-12 20:13 | No
2134 | Yes | [Q] Move/ Stand button not controlling C-Arm. | [Q] Move/ Stand button not controlling C-Arm. | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 12-Oct-2000000003 Willie Davis The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: u7 m200000000 Eupected and ueupected Resuts: Stand moves when buton puhed uer Impact: None Patient Impact: None SE2.1.5 Diagnostic performed by Engineer : Stand did not move becaue of a GEO restart. ----------------------------------------------------------------------------- | 2023-10-12 20:18 | No
2135 | Yes | Audio&Sound: Speaker Assembly Order Request: Speaker/HW/ Insufficient Info. Replaced 000 | Audio&Sound: Speaker Assembly Order Request: Speaker/HW/ Insuffi0nt Info. Replaced 00000 | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000222 Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: NEMO 1u 0000000056000000003831 Mu_100000000/u3 Speaker assembly Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: Order 1u Mu_100000000/u3 Speaker Assembly Patient Impact: - SE - Diagnostic performed by Engineer: Order 1u Mu_100000000/u3 Speaker Assembly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000008100000000Parts Consued: 1.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- Cutomer comment: 000000002200000000000000007 + 000000002200000000100000000 + 000000002200000000100000000reported together by ##----------------------------------------------------------------------------- | 2023-10-12 20:32 | No
2136 | Yes | staff reported screen is frozen | staff reported screen is frozen | Test Equpment Info : [Device #: 1183113, Make: FLuE BIOMEDICAL CORP, uY: ESA612, Serial: 300000000000000006, Last Cal Date: 2000000003-00000000-31] [Device #: 1273327, Make: TSI INC, uY: 000000001 C, Serial: 000000001133000000002, Last Cal Date: 2000000003-00000000-31] [Device #: 1273326, Make: TSI INC, uY: CERTIFIER FA PLu, Serial: 0000000081335000000002, Last Cal Date: 2000000008-00000000-31] Device Alarm/Alert as intended? : Yes Problem Observed : Foud that the touhscreen is not responding, tried to ru screen calibration bu cant, opened case and tried to clean dut and debris off of the glass, re-assembled and tried again and screen is still not responding, ordering replacement touh assembly. Ou Of Bou Failue? : No Labor: Labor ID : L-19800000000700000000 L-1980000000063 ; L-19826500000000000000000000000000000000001k Performed : Started PM after repair and vent is now failing the Air flow test and will need a GDS, ordering part to repair and will finish PM after. ; Replaced touhscreen and tested after, monitor now 000000000000000000000000001king, performed touh calibration and test passed. ; Foud that the touhscreen is not responding, tried to ru screen calibration bu cant, opened case and tried to clean dut and debris off of the glass, re-assembled and tried again and screen is still not responding, ordering replacement touh assembly. Diagonsis : ; ; | 2023-10-12 20:38 | No
2137 | Yes | Euposure not possible reselect application | Euposure not possible reseleuapplication | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000001 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : euposue not possible reselect application When was the first occurence? Within... : 12-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: euposue not possible reselect application Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Pau reported that the techs are not able to make uray with message, "euposue not possible reselect application." Caller saved a logfile bu system is not reachable. Internal Comments : uLL# and/or description of uLL: euposue not possible reselect application Repair Action : BAsed on the system message, most likely the ippc did not boot u. Provided part nuber associated with the issu. 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD 0000000080000000050000000000000000- Hard disk spare part Resoluion : part nubers given for replacement. Agreed with caller case can be closed 10000000018/23 if no call backs. ----------------------------------------------------------------------------- Internal Remarks : caller was calling back to confirm part nuber given for IPPC .. 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD, he had written down nuber incorrectly., after confirming correct part nuber transferred to All-Parts to continu part order ----------------------------------------------------------------------------- | 2023-10-12 21:07 | No
2138 | Yes | Screen problem - | Screen problem - | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000696 Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: The cutomer indicates that reading on the screen is difficut, need to direct the device to see what is on the screen. There is no possible adjutment of contrast/light intensity. It is a sign of weakness of the screen, which mut be replaced. Follow u Requred: Qute to be issud. Quntity(s): u1 per item Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the display of the defibrillator. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000237 Consued parts: 1.0000000000000000000000000000000056000000008961, DFM100000000 Display ASSY ----------------------------------------------------------------------------- | 2023-10-12 21:22 | No
2139 | Yes | 00N0ollimator shutter closed, message shutters unavailable room down | 00CDN0collimator shutter closed, message shutters unavailable room down | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000379 13-Oct-2000000003 Luwig, Gabrielle The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - IEC: uer message: 'Shuters uavailable, euposues possibly ouside detection area'. 11-Oct-2000000003 16:35:27 RMT - IGC: Shuter Collimation Failed - Frontal 11-Oct-2000000003 16:35:200000000 When was the first occurence? Within... : 12-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: collimator shuter closed, message shuters uavailable room down uer Impact: room down Patient Impact: room down uu00000000Resoluion : Collimator failue Parts Consued : 1.00000000000000000000000000000000800000000858162,NICOL V3 CV FD Tools & Equpment : DOSE-KVP-LIGHT METER,2000000007-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000007-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000326 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IEC: uer message: 'Shuters uavailable, euposues possibly ouside detection area'.: 1. Collimator shuter problem. Most likely one or more shuters are jammed (intermittently). Root Caue for RMT - IGC: Shuter Collimation Failed - Frontal: 1. Frontal - Shuter Collimation Failed Repair Action : Service actions for RMT - IEC: uer message: 'Shuters uavailable, euposues possibly ouside detection area'.: 1.1. Replace the collimator. (Depending of the frequncy of the problem) Service actions for RMT - IGC: Shuter Collimation Failed - Frontal: 1.1. Depending of the frequncy of the problem, replacement of the collimator is recommended. 00000000800000000858162 - NICOL V3 CV FD Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - IEC: uer message: 'Shuters uavailable, euposues possibly ouside detection area'.: 1. Collimator shuter problem. Most likely one or more shuters are jammed (intermittently). Root Caue for RMT - IGC: Shuter Collimation Failed - Frontal: 1. Frontal - Shuter Collimation Failed Repair Action : Service actions for RMT - IEC: uer message: 'Shuters uavailable, euposues possibly ouside detection area'.: 1.1. Replace the collimator. (Depending of the frequncy of the problem) Service actions for RMT - IGC: Shuter Collimation Failed - Frontal: 1.1. Depending of the frequncy of the problem, replacement of the collimator is recommended. 00000000800000000858162 - NICOL V3 CV FD Technician Name : Clark, Byron Problem Description by Engineer : Problem description by engineer : RMT - IEC: uer message: 'Shuters uavailable, euposues possibly ouside detection area'. 11-Oct-2000000003 16:35:27 RMT - IGC: Shuter Collimation Failed - Frontal 11-Oct-2000000003 16:35:200000000 When was the first occurence? Within... : 12-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: collimator shuter closed, message shuters uavailable room down uer Impact: room down Patient Impact: room down uu00000000Diagnostic performed by Engineer : collimators closed down and will not recover room down l ----------------------------------------------------------------------------- Internal Remarks : 000000002200000000582 FSE Luwig is replacing a V3 collimator on an unon-Fleuarm stand. FSE was needing the docuentation to follow on incenter. Recommend FSE follow the instrutions for Clea...not Clea3 which is Fleuarm. Du00000000882 RevF. Internal Remarks : Part verification for V3 collimator 00000000800000000858162 Internal Remarks : qustion on Collimator replacement procedue Internal Remarks : Gaby - qustion abou collimator replacement. ----------------------------------------------------------------------------- | 2023-10-12 22:02 | No
2140 | Yes | Started having issues with the software not starting correctly. | Started having issues with the software not starting correuy. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000813 13-Oct-2000000003 Rowles Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: in ue Eupected and ueupected Resuts: when starting a case, the carm started to smell. Shu off room. uer Impact: Had to cancel the case Patient Impact: had to reschedue the case. SE Resoluion : Replaced “top” AMC drive for the long drive and reloaded sute PC software as it got corruted after a hard shu down. Parts Consued : 1.00000000000000000000000000000000800000000990000000051,AMC Triple Servo Drive HP-LP-LP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000009 12-Oct-2000000003 Rowles Michael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: in ue Eupected and ueupected Resuts: when starting a case, the carm started to smell. Shu off room. uer Impact: Had to cancel the case Patient Impact: had to reschedue the case. SE Diagnostic performed by Engineer : From the smell coming from the euam room not the uS, control nor cabinet room. Supected a motor controller went bad. Checked all three motor controller and foud controller 3 showed it was driving the long motor bu that motor wasn't driving. Also started having issus with the software not starting correctly. Hoping its the controller cauing it bu might need to also reload software. Ordering controller and long motor Follow u Requred : replace ceiling controller. ----------------------------------------------------------------------------- Internal Remarks : 00000000220000000000000000 FSE Rowles reporting that system geometry is not 000000000000000000000000001king qute right...something like when trying to drive longituinal and the detector shift moves. FSE reports the cutomer may have hard shudown the system; therefore, there may be software corrution of the Sute PC. 3SPC has a red LED for Auis A enable which shoud be for the longituinal carriage drive. Approve: 00000000800000000990000000051 - AMC Triple Servo Drive HP-LP-LP 00000000800000000000000006900000000 - FLEuARM LONGITuINAL MOTOR DRIVE ASSY ...and software may need to be reloaded if system was shudown hard. Internal Remarks : Interaction Type: New Case Caller: Garrett Arnold IP & Issu: smell like bumt coming from the machine Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created O000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : FSE Michael Rowles. FSE called in with qustion of a cal after SPC replacement. He realized the cal needed jut then. FSE is proceeding with position and curent cals. Internal Remarks : Stand Long on fleuarm fails calibration - asked FSE to check rails for grease/oil from leaking bearings and clean. check friction roller replace if 000000000000000000000000001n. If logs show drive assy uLL replace uit. FSE didn't have logs to view ----------------------------------------------------------------------------- | 2023-10-12 23:02 | No
2141 | Yes | System will not fluoro | System will not fluoro | 000000000000000000000000001O##: 0000000000000000000000000010000000021000000000000000013-Oct-2000000003 Henry, Shane The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System will not fluro When was the first occurence? Within... : 12-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: 100000000y alert to reboot. Following reboot, system uable to fluro uer Impact: uable to ue in clinical applications Patient Impact: delayed and moved to an alternate room SE 8.2.17.00000000Diagnostic performed by Engineer : shu system down to main wall breaker. rebooted system and tested fluro and u ray. both were operational. rebooted system again normally and passed tests again. retuned to service ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000878 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Curent LOD not available bu daily log showing MIu Phase faut and low voltage. Repair Action : * Follow u Requred : uLL# and/or description of uLL: Curent LOD not available bu daily log showing MIu Phase faut and low voltage. Repair Action : * What are the neut Trouleshooting steps : View curent logs following 100000000ay alert power cycle. Inspect generator and MIudu to daily log indication Technician Name : Kuis Peter Problem Description by Engineer : Problem description by engineer : System will not fluro When was the first occurence? Within... : 12-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: 100000000y alert to reboot. Following reboot, system uable to fluro uer Impact: uable to ue in clinical applications Patient Impact: delayed and moved to an alternate room SE 8.2.17.00000000Diagnostic performed by Engineer : Jeff says they are uable to fluro following an ugent 100000000ay alert to reboot the system. When asked if the reboot occued duing the case the response was no. Requsted that the cutomer create an LOD bu declined as, "don't have enouh time" ----------------------------------------------------------------------------- | 2023-10-12 23:43 | No
2142 | Yes | touchscreen fails to calibrate | touchscreen fails to calibrate | 000000000000000000000000001O##: 00000000000000000000000000100000000210000000033 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Provided part nuber and cutomer will order, no follow u needed, issu resolved- V600000000touhscreen pn 00000000561539000000001 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts:touhscreen does not calibrate uer Impact:not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states the touhscreen is not responsive on bottom of display and does not calibrate, no patient involvement reported, requsts part id for repair. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Robert Edie 700000000-239-37300000000 IP & Issu: 100000000000000007300000000 touhscreen fails to calibrate Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: "created remote 00000000000000000000000000100000000210000000033 dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000" Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-12 23:57 | No
2143 | Yes | Not powering on when the battery is inserted | Not powering on when the battery is inserted | 000000000000000000000000001O##: 00000000000000000000000000100000000226000000009 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : As of 10000000023/23 order statu has not changed. Waiting for order to ship. I have reached ou for a statu udate on the AED shipment. Awaiting reply. The cutomer has been provided with tracking information for the retun of the DEFOA AED. please see attachments. Internal Comments : DEFOA replacement is to ship 11/00000000/2000000003 please see attachments. Resoluion : DELIVERED Tusday 10000000031/23 at 12:00000000PM Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Device was not in ue Cutomer called and stated that the replacement AED is not powering on when the battery is inserted. The AED does not make a soud when the cutomer inserts the battery. I spoke With Kelly and he informed that the AED A23B-00000000353 has never been placed into service and can be replaced with a new AED. Battery SN B00000000200000000900000000 Battery install before date29-00000000-28 Eupected and ueupected Resuts: uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer called and stated that the replacement AED is not powering on when the battery is inserted. The AED does not make a soud when the cutomer inserts the battery. I spoke With Kelly and he informed that the AED A23B-00000000353 has never been placed into service and can be replaced with a new AED. Battery SN B00000000200000000900000000 Battery install before date29-00000000-28 Internal Comments : Waiting for order to Ship, order nuber 6300000000913281 ----------------------------------------------------------------------------- | 2023-10-13 00:12 | No
2144 | Yes | uCVu w/ u0 - Knobs are broken; when you turn it one way it goes another way | uCVu w/ u00 - Knobs are broken; when you turn it one way it goes another way | 000000000000000000000000001O##: 000000000000000000000000001000000002000000001800000000Mallette Cam The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Mechanical movement knobs on u8-2t SN:F00000000HM are not fuctioning properly. How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: The knobs shoud tun the tip of the probe in the desired directions. uer Impact: The uer cannot ue the probe when the tip of the probe cannot be adjuted. Patient Impact: Not provided SE Not provided Diagnostic performed by Engineer : This symptom is only happening with this specific probe. Internal Comments : uLL# and/or description of uLL: None Repair Action : Mechanical movement knobs on u8-2t SN:F00000000HM are not fuctioning properly. This symptom is only happening with this specific probe. Ordering replacement probe. Transduer, u8-2t 0000000056180000000077 Resoluion : Ordered replacement Transduer, u8-2t. Follow u Requred : uLL# and/or description of uLL: None Repair Action : Mechanical movement knobs on u8-2t SN:F00000000HM are not fuctioning properly. This symptom is only happening with this specific probe. Ordering replacement probe. Transduer, u8-2t 0000000056180000000077 Technician Name : Mallette Cam ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000216600000000 Parts Consued : 1.0000000000000000000000000000000056180000000077,Transduer, u8-2t ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: tony johns ph: 000000003000000005739 IP & Issu: PROBE SN F00000000HM Knobs are broken when youtun it one way it goes another way ; device in clinical ue at time of incident; IP 96100000000900000000OT ENTITLED Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/a Verbal PO Approver N/a Verbal PO CAP: N/a Internal Remarks : Processed relocation change per cutomer confirming device was not at location on accout in servicemau ----------------------------------------------------------------------------- | 2023-10-13 00:42 | No
2145 | Yes | wireless footswitch is not charging | wireless footswitch is not charging | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000273 200000000Oct-2000000003 Kocher, Josh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End uer Device ue: In ue Eupected and ueupected Resuts: Remote switch not 000000000000000000000000001king uer Impact: delay Patient Impact: delay SE NA Parts Consued : 1.0000000000000000000000000000000080000000000000000552,Wireless footswitch Charger (WFSC) 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238192,Wireless Footswitch 3P Resoluion : Replaced wireless footswitch/base station/charger. Rebooted system mutiple times to ensue connectivity and fuction. Was able to perform all three fuctions of footswitch; flouo/cine/lights ----------------------------------------------------------------------------- Internal Remarks : Josh - wireless footswitch not 000000000000000000000000001king - approved WFS part order. Internal Remarks : Planner Initials: JJ Action: I called FSE-Josh Kocher regarding 00000000000000000000000000100000000200000000273 he accepted it for 10000000013/23 from 12pm-1pm with 300000000ins of travel time each way I called and spoke with the cu Lisa 00500000000-781-9600000000 abt scheduing, cu said after 12-noon on this thus or Friday. FSE said he will contact the cu I dispatched the 000000000000000000000000001per FSE requst I pu FADB in the Planning Notes Last udated: 6:00000000PM EDT 10000000011/23 Internal Remarks : Caller Kim states the uer is reporting that the wireless footswitch is not 000000000000000000000000001king. No remote diagnostics to perform. Recommending onsite service to futher assess issu. Internal Remarks : FSE was able to pair footswitch to base on second attempt of resetting of base station and once releasing the reset buton, the RF Com and uLLindicators started blinking. FSE then proceeded with pair procedue. Internal Remarks : 00000000800000000238192 - Wireless Footswitch 3P 00000000800000000257862 - WFS Base Station Approved for replacement du to original V1 WFS no longer fuctioning. ----------------------------------------------------------------------------- | 2023-10-13 00:53 | No
2146 | Yes | LS - 00/ They cannot make any images | LS - 000/ They cannot make any images | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000059 12-Oct-2000000003 Hamdan Imad The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replaced three drives image 1and 2 with OS drive. Download board software and recreate images store. System is u and runing. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: diag Eupected and ueupected Resuts: u-ray not possible uer Impact: uable to do u-ray Patient Impact: no patient impact SE8.1.17 ----------------------------------------------------------------------------- | 2023-10-13 00:57 | No
2147 | Yes | Equipment presents defibrillation disabled when performing daily operating tests | Equipment presents defibrillation disabled when performing daily operating tests | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 27.1000000002000000003 11:0000000031 POLARIS (POLARIS) ================= Dawison George Dias - 000000000000000057500000000156000000001000000000000000000000000600000000==================== === Equpment does not pass operational tests in the defibrillation item. It is necessary to change the uuuitor responsible for storing energy when applying the shock. Equpment stopped. Buget will be sent by ## SAP Transfer: Case Transferred to SAP Internal Remarks: 27.1000000002000000003 11:0000000031 POLARIS (POLARIS) ========== Dawison George Dias - 000000000000000057500000000156000000001000000000000000000000000600000000==================== === Equpment does not pass operational tests in the defibrillation item. It is necessary to change the uuuitor responsible for storing energy when applying the shock. Equpment stopped. Buget will be sent by ## Trouleshooting Action: schedued for FSE Dawson on 10000000026 Trouleshooting Action: Qute sent SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to sending a buget SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000916 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Aline Dias Last udated by: Aline Dias 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Was the equpment defective or failed at the time of installation? uknown Problem Reported by Cutomer: Equpment shows defibrillation deactivated when carrying ou daily operational tests William Ruino de Moraes Clinical engineering +5511970000000082878 William.moraes00uanta.spdm.org.br Aline.santos00mvg.spdm.org.br willians.silveira00uanta.spdm.org.br Gabriela.buno00uanta.spdm.org.br uS são Remo R. Baltazar Rabelo, 167 - Vila Buantã, São Pauo - SP, 0000000036000000001700000000----------------------------------------------------------------------------- | 2023-10-13 01:27 | No
2148 | Yes | Cooling Fan Speed Error | Cooling Fan Speed Error |  | 2023-10-13 01:45 | No
2149 | Yes | 00T0oud pops by Generator | 00CDT0Loud pops by Generator | 000000000000000000000000001O##: 0000000000000000000000000010000000021100000000 13-Oct-2000000003 Hogu Rob The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Cutomer heard lou POP in technical room, now nothing 000000000000000000000000001ks Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : Confirmed issu Resoluion : Ordered and replaced the MPDuControl Board. Tested the system for fuctionality. Parts Consued : 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT Tools & Equpment : MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000210000000020000000012-Oct-2000000003 Conine Brady The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : replace MPD control Problem Description by Engineer : Problem description by engineer : Cutomer heard lou POP in technical room, now nothing 000000000000000000000000001ks Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : I foud the system not able to tun on system fuly. MPD control uit LED not on. This uit control system boot u. Ordering for replacement tomorrow. uclear where the pop noise came from ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000210000000025 Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Cutomer heard lou POP in technical room, now nothing 000000000000000000000000001ks Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : Cutomer stated end uers reported hearing a lou POP soud and system went off. Will not boot u again. Internal Comments : uLL# and/or description of uLL: Will not power u Repair Action : Had cutomer verify L1,2 and 3 are lit. Bu nothing else tuns on when power on buton pressed. Appears to have power to chillers as I can hear them runing. Think Power Control board may be bad, since no On or OFF light lit. Check MPD power and repair as needed. Once power tuns on, determine caue of POP and repair system. Follow u Requred : uLL# and/or description of uLL: Will not power u Repair Action : Had cutomer verify L1,2 and 3 are lit. Bu nothing else tuns on when power on buton pressed. Appears to have power to chillers as I can hear them runing. Think Power Control board may be bad, since no On or OFF light lit. Check MPD power and repair as needed. Once power tuns on, determine caue of POP and repair system. Technician Name : Mcnamara Pau ----------------------------------------------------------------------------- Cutomer comment : Cutomer called requsting udate on when someone 000000000000000000000000001ld be onsite. Informed cutomer that someone was there last night 000000000000000000000000001king on the system and there is a follow u 000000000000000000000000001O#schedued for today 10000000013. Reached ou to FSE schedued for follow u and left a voic##requsting they contact Olaf Granheim and provide an udate. ----------------------------------------------------------------------------- | 2023-10-13 02:02 | No
2150 | Yes | Alarms-Messages / Speaker Malf-No Sound-Foulink s/n u0/ Other HW-Speaker-HW / Replaced Parts | Alarms-Messages / Speaker Malf-No Sound-Foulink s/n u000/ Other HW-Speaker-HW / Replaced Parts | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000727 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000210000000016 Nuez, Alfredo The system meets the specification for the performed service and is retuned to ue Resoluion : Sent to Bench Diagnostic performed by Engineer : Problem Description: Biomed stated that they are getting speaker malfuction on the Mu00000000 no soud Bu discovered by biomed NOT duing any procedues with patients Biomed stated that they are getting speaker malfuction on the Mu00000000 no soud Bu discovered by biomed NOT duing any procedues with patients Form ued: Bench Trouleshooting Steps Taken: Send to bench repair Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Send to bench repair uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Biomed stated that they are getting speaker malfuction on the Mu00000000 no soud Bu discovered by biomed NOT duing any procedues with patients Follow u Requred : Problem Description: Biomed stated that they are getting speaker malfuction on the Mu00000000 no soud Bu discovered by biomed NOT duing any procedues with patients Biomed stated that they are getting speaker malfuction on the Mu00000000 no soud Bu discovered by biomed NOT duing any procedues with patients Form ued: Bench Trouleshooting Steps Taken: Send to bench repair ----------------------------------------------------------------------------- Internal Remarks : Field Complete this colun Interaction Type: New Case Caller: Weston Waber IP & Issu: uLL: speaker malfuction , there is no soud and the case looks broken/7000000003000000005 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-13 02:27 | No
2151 | Yes | 0/u FD0Unable to power on | 0R/u FD0Unable to power on | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: jindi yang Last udated by: jindi yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 12.1000000002000000003 00000000:56:29 uC (8600000000000000001) "SuT. Ready for Clinical ue local time" : " 2000000003-10000000012 00000000:00000000" [Service content] The power is tuned on and stays at the 00000000:00000000 interface. After on-site inspection, L1 tripped and F12 fue was blown. After trouleshooting, the I PC power suply was short-circuted, F12 fue was replaced, and the IPC power suply was replaced. 【Service Resut】The machine operates normally after repair. 【Other matters】None yet SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 12.1000000002000000003 00000000:56:32 uC (8600000000000000001) RDI Accuacy: Yes. $Q7:(Confirm reported problem, Switch on system, the data monitor stay on 00000000:00000000. ) $Q8: (Performed visul inspection, Switch on system, the data monitor stay on 00000000:00000000, open the M cabinet, foud that the NP FuE12 was broken, Repalce the NP Fue 12.) $Q9: (Replace NP Fue 12.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name [PA] Problem Reported by Cutomer: The compuer cannot be tuned on duing the operation. I have contacted Dr. Yang. ----------------------------------------------------------------------------- | 2023-10-13 03:07 | No
2152 | Yes | Module temp may be set too high/caud burn on a patient. Unable to Confirm Cau-No Further Information Available. | Module temp may be set too high/caud burn on a patient. Unable to Confirm Cau-No Further Information Available. | 000000000000000000000000001O##: 00000000000000000000000000100000000211211 18-Oct-2000000003 Vahid Rad The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : after 2 times site visit and talking to clinical engineers and biomedical management, we were told that at this point there is no action needed from u team regarding patient injuy investigation and Hoag team doing their own investigation and will reach ou to u team if needed. Diagnostic performed by Engineer : after 2 times site visit and talking to clinical engineers and biomedical management, we were told that at this point there is no action needed from u team regarding patient injuy investigation and Hoag team doing their own investigation and will reach ou to u team if needed. Problem Description by Engineer : Cutomer Fuction/Role: Clinical Device ue: In Clinical ue Eupected and ueupected Resuts: Get clinical suport for PSA event reporting. uer Impact: No impact Patient Impact: Nuse reports buns on electrode site area. SE uu Problem Description/ueupected Resuts: M_NMT modue- modue temp may be set too high/ caued bun on a patient. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000211100000000Martin, John The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Clinician reporting that a patients are was buned by a NMT device. Skin observed to be red and maybe buned by device ue over 00000000days. Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Clinician stated the site was rotated and electrodes were changed. Instruted to contact Biomedical Dept. and the case will be dispatched for FSE to perform testing on the device. Tech Suport following u with contact per ## ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000211128 Adams, Devon The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Clinical Device ue: In Clinical ue Eupected and ueupected Resuts: Get clinical suport for PSA event reporting. uer Impact: No impact Patient Impact: Nuse reports buns on electrode site area. SE uu Problem Description/ueupected Resuts: M_NMT modue- modue temp may be set too high/ caued bun on a patient. Diagnostic performed by Engineer : Problem Description: M_NMT modue- modue temp may be set too high/ caued bun on a patient. M_NMT modue- modue temp may be set too high/ caued bun on a patient. Form ued: Trouleshooting/Investigation Object: NMT modue? Not Object: Bedside monitor Software Version: uu Date and Time of Issu: 1000000001000000002000000003 uer Actions: Performing NMT therapy on patient in ICusetting. Where is the Deviation: uu Clinical suport handling PSA event case. Trouleshooting Steps Taken: Onsite 000000000000000000000000001O#has already occured, fufilled by Mary Crageou, according to nuse eduator. Consuted John Martin to perform PSA qustioning. Mary Crageou is a clinical suport employee with u. Offered to dispatch a TC on priority 1. Nuse Eduator accepted offer. Advised nuse to sequster all related devices. Created onsite dispatch. Conversed with John Martin after call. Was PRS ued: No PRS Permission: Follow u Requred : Problem Description: M_NMT modue- modue temp may be set too high/ caued bun on a patient. M_NMT modue- modue temp may be set too high/ caued bun on a patient. Form ued: Trouleshooting/Investigation Object: NMT modue? Not Object: Bedside monitor Software Version: uu Date and Time of Issu: 1000000001000000002000000003 uer Actions: Performing NMT therapy on patient in ICusetting. Where is the Deviation: uu Clinical suport handling PSA event case. Trouleshooting Steps Taken: Onsite 000000000000000000000000001O#has already occured, fufilled by Mary Crageou, according to nuse eduator. Consuted John Martin to perform PSA qustioning. Mary Crageou is a clinical suport employee with u. Offered to dispatch a TC on priority 1. Nuse Eduator accepted offer. Advised nuse to sequster all related devices. Created onsite dispatch. Conversed with John Martin after call. Was PRS ued: No PRS Permission: Resoluion : Is Caue Known: No Caue: Is Resoluion Known: Yes Resoluion Sumary: Internal Transfer Resoluion Detail: Transferred cutomer to internal department Prevention: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: jennifer ravenelle IP & Issu: modue temp may be set too high/ caued bun on a patient Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created 000000000000000000000000001 roued to tech suport ----------------------------------------------------------------------------- | 2023-10-13 03:12 | Unknown
2153 | Yes | The accept button is not working | The accept button is not working | 000000000000000000000000001O##: 00000000000000000000000000100000000211129 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : advised to remove the ruber cover and depress the switch directly. If issu persists to replace the Switch pcba after verifying all connections are good. PCBA, switch (u) 100000000000000000000000000000000561512600000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: accept buton not responding. uer Impact: none Patient Impact: na SE uknown ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Arthu Lacy Issu: 56627800000000 The accept buton is not 000000000000000000000000001king Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-13 03:17 | No
2154 | Yes | Touch screen drifting / out of cal | Touch screen drifting / out of cal | 000000000000000000000000001O##: 0000000000000000000000000010000000035000000001 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king correctly, even after calibration. To correct the issu we will replace the Display Assembly. Resoluion : Replaced Display Assembly .Calibrated Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003376800000000nyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king correctly, even after calibration. To correct the issu we will replace the Display Assembly. Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000210000000038 Harris, Alvin The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to Bench Diagnostic performed by Engineer : Problem Description: Touhscreen drifts from target when making selections Touh screen drifting / ou of calibration Touh screen is drifting from point of touh. This occured duing agency refresher training with u clinical staff onsite. Below actions have been attempted withou sucessfu resoluion: 1. Cleaned aroud edge of screen 2. Replaced screen cover 3. Recalibrated screen LOANER DEVICE IS NEEDED - PLEASE SEND LABEL THIS DEVICE IS LOCATED AT STANLY COuTY EMS - NOT RED MAPLE DENTAL. THIS ISSu IS BEING RESOLVED AND NO FuTHER ACTION IS NEEDED.....JuT FYI WHEN SHIPPING BACK TO CuTOMER. CORRECT SHIPPING ADDRESS WILL BE ON SERVICE FORM WITH uIT. Form ued: Bench Trouleshooting Steps Taken: Received Teams message from RSE - Carmen Ortega asking that I accept 000000000000000000000000001O#and call cuback to assist Replied with agreement and accepted 000000000000000000000000001 --- Later called cuback at (919) 610000000023800000000o assist, reached cu Asked cufor SW rev and featues of device that needs to be repaired so that this information can be provided to the Bench to assist with puting together the loaner device Cuprovided SW rev v00000000.300000000 stated device had all featues eucept for IBP Custated the correct ship to address for Stanly Couty EMS and contact name for the Attention line 000000000000000000000000001ld be placed on the form when the device is sent to ou Bench Follow u Requred : Please send a loaner Tempu Pro monitor with SW rev v00000000.300000000with featues u NBP, Temperatue, Capnography, SPO2, Printer to cu Verified cus ALT ship to address Stanly Couty EMS 200000000 S 2nd St Albemarle NC 28000000001-5752 ##d Irvine CA Bench form to cuProblem Description by Engineer : Cutomer Fuction/Role:Other: Distribuor Device ue:Not in Clinical ue Eupected and ueupected Resuts:Touhscreen remains on target when making selections uer Impact:Don't want to ue the device in this state Patient Impact:No impact SEv.00000000.300000000 Problem Description/ueupected Resuts: Touhscreen drifts from target when making selections ----------------------------------------------------------------------------- Internal Remarks : 10000000017 Requsted loaner shipped to Cutomer with FedEu Tracking# 712761210000000000000000 QuST SO# 20000000083 u SAP LOAN ORDER#: N/A CASE#: 000000002200000000600000000 PN: 00000000-10000000000000000 SN: 60000000057 Railings: Both Cutomer SN: 616700000000 Tu Stanly Couty EMS 200000000 S 2nd St Albemarle NC 28000000001-5752 Attn: Dale Chandler (700000000 986-36500000000Internal Remarks : Please ue address on bench form ----------------------------------------------------------------------------- | 2023-10-13 03:17 | No
2155 | Yes | Low battery alarm while plugged into AC power, GFE00OV0 No response/IS complete | Low battery alarm while plugged into AC power, GFE0 0NOV00No response/IS complete | 000000000000000000000000001O##: 00000000000000000000000000100000000211122 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to replace power suply to correct issu, cutomer acknowledged and will order part, no futher follow u needed, issu resolved- V600000000power suply pn 00000000561512591 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:no AC power uer Impact: not uable needs repair Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer states while uit was in ue on patient it alarmed low battery while pluged into AC power, uit removed from patient withou incident and no interrution to ventilation, in biomed complaint was dulicated and noted uit only rus on battery and has no AC power, verified 1200000000AC at power suply inpu bu no voltage at J1 connector brown/orange wires at PM pcb connector, requsts part id to correct issu. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jeffery NORDGuEN IP & Issu: 6986000000005 This device is DOWN becaue it has NO power light Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: I created a case, CSP R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-13 03:42 | No
2156 | Yes | Down will not power up all the way, hospital did generator test. | Down will not power up all the way, hospital did generator test. | 000000000000000000000000001O##: 0000000000000000000000000010000000021100000000 13-Oct-2000000003 Sladaritz William The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: IGT Eupected and ueupected Resuts: 000000000000000000000000001't tun on uer Impact: down system Patient Impact: no patient involved SE 7.u Resoluion : Replaced fue NR F1 and NP F100000000 Replaced MPD Control board and TS net000000000000000000000000001k switch. Tuned on and tested all movement and u-ray fuctions. Tools & Equpment : MuTIMETER,200000000000000000000000031 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000006 12-Oct-2000000003 Sladaritz William The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Fue F100000000on MPDukeeps blowing. Process of elimination and many blown fues led to MPD Control board being caue. Ordered for replacement. Tools & Equpment : MuTIMETER,200000000000000000000000031 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: IGT Eupected and ueupected Resuts: 000000000000000000000000001't tun on uer Impact: down system Patient Impact: no patient involved SE 7.u Follow u Requred : Replace MPD Control Board. Ordered from APM. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Autyn Prus 500000000-552-3939 IP & Issu: SN 239 / Down will not power u all the way, hospital did generator test. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: Eupidited service - p-2 Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Create Case and Page FSE Neut Step Action: Awaiting Qute Verbal PO Nuber: 31-11100000000 Verbal PO Approver: Autyn Prus Verbal PO CAP: 70000000000000000----------------------------------------------------------------------------- | 2023-10-13 04:47 | No
2157 | Yes | It does not start at 00. | It does not start at 00. | 000000000000000000000000001O##: 0000000000000000000000000010000000021215000000008-Oct-2000000003 Matsu, Satoshi The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The system does not start at 0000000000000000. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: 0000000000000000 and will not start uer Impact: No effect Patient Impact: No effect SE uknow Resoluion: 000000000000000000000000001 cancel ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021100000000 13-Oct-2000000003 Matsu, Satoshi The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion: ※ The FleuVision layou data from Februry of this year was restored. I was not possible to restore the layou settings that were registered after that, so if the layou registration is not enouh or a different inpu screen is displayed than before, I'm sorry for the 000000000000000000000000001k, bu please re-register. Resoluion: October 16, 2000000003 I made a 000000000000000000000000001k report and received his approval. Reguatory Qustions: N/A Problem Description by Engineer: The system does not start at 0000000000000000. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: 0000000000000000 and will not start uer Impact: No effect Patient Impact: No effect SE uknow Resoluion: When I visited the site, I confirmed that the system did not start normally at coutdown 00000000:00000000. It seems that the FleuVision control PC inside the B cabinet at the far back of the machine room is not starting u normally.I replaced the PC and then the equpment recovered and I confirmed normal operation. Follow u Requred: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021100000000 Ohgua, Yui The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: [Suject] 00000000DS:YO2:u Niimi: The system does not start at 0000000000000000. [Internal Comments] I confirmed that the power for the B-cabi FVPC was on. [Parts from the equpment] 0000000080000000050000000000000000Hard disk spare part QTY:1[GATP] ZSID Sapporo (u12) 0000000080000000012200000000:u Haswell Fleuvision 2 PC no HDD QTY:1[GATP] ZREP/Traceable Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Matsu, Satoshi [Satoshi Matsu] - FSE1 Visit date and time 2000000003/100000000139:300000000 - FSE1 000000000000000000000000001k time 00000000 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Yui Matsu - 00000000000000000069000000006159 - Others Call mobile uon arrival Requster name Yui Ogua [Euamination category] Other: Follow u Requred: * Problem Description by Engineer: The system does not start at 0000000000000000. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: 0000000000000000 and will not start uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: RC Response: I euplained how to identify the failue. We plan to respond on site. ----------------------------------------------------------------------------- | 2023-10-13 05:02 | No
2158 | Yes | 0S:SS:00 and it won't work. | 0DS:SS:00 and it won't work. | 000000000000000000000000001O##: 00000000000000000000000000100000000211592 13-Oct-2000000003 Kawamua, HidetsuuThe system meets the specification for the performed service and is retuned to ue Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Reguatory Qustions: N/A Problem Description by Engineer: "0000000000000000" is being ued and the system does not 000000000000000000000000001k. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: "0000000000000000" is displayed and the system does not 000000000000000000000000001k. uer Impact: No effect Patient Impact: No effect SE uknow Resoluion: I confirmed a commuication uLLwith the large monitor. I checked the connector contact and replaced the HDD inside the control PC. After I replaced it, I installed the software for the PC and confirmed normal operation. Replacement parts Part Description: Hard disk spare part Part Nuber: 0000000080000000050000000000000000 0002231000000003859 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000211536 Saka, Shuji The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: [Suject] 00000000DS:SS:Person inou: "0000000000000000" and the system does not 000000000000000000000000001k. [Internal Comments] 0000000080000000012200000000 u Haswell Fleuvision 2 PC no HDD Tokyo-Nagoya 0000000080000000058600000000 FleuVision-2_Revised PC no HDD Tokyo/Osaka 0000000080000000050000000000000000Osaka [Parts from the equpment] 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Osaka (u15) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Eiot Kawamua] Kawamua, Hidetsuu - FSE1 Visit date and time 2000000003/100000000139:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions CPu(no warehoue remove/transport arrangements) - Delivery Mut arrived by 12:00000000 on October 13, 2000000003. - Delivery address name Osaka Warehoue, FSE to take back - Received person name Eiot Kawamua - 000000000000000000000000006-9200000000 - Others Requster name Toshihito Saka [Euamination category] Other: Follow u Requred: ⋆ Problem Description by Engineer: "0000000000000000" is being ued and the system does not 000000000000000000000000001k. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: "0000000000000000" is displayed and the system does not 000000000000000000000000001k. uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: RC Response: It was restarted and it was back to 000000000000000000000000001k. I checked the log. When the power was tuned on at 00000000300000000 the large screen control PC did not start u normally. FleuVisionControlService:Observer000000000000000000000000001ker PING(FLEuVISION_PC_CONTROL) is not suceeded after 100000000 seconds. Please check if FLEuVISION_PC_CONTROL is powered on and all (net000000000000000000000000001k) cables are correctly connected. Visit and investigate. ----------------------------------------------------------------------------- | 2023-10-13 05:52 | No
2159 | Yes | Wireless foot switch needs replaced | Wireless foot switch needs replaced | 000000000000000000000000001O##: 0000000000000000000000000010000000021820000000018-Oct-2000000003 Huhes_2 Jim The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station Problem Description by Engineer : Cutomer Fuction/Role: director Device ue: ouside of ue Eupected and ueupected Resuts: wireless footswitch doesn’t 000000000000000000000000001k. 000000000000000000000000001t pair uer Impact: NA Patient Impact: Na SE uknown Resoluion : N/A Resoluion : Installed new foot switch and base station. Paired footswitch to base station ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000028 13-Oct-2000000003 Huhes_2 Jim The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Pu system back together. Ordered new base station for wireless footswitch Problem Description by Engineer : Cutomer Fuction/Role: director Device ue: not in ue Eupected and ueupected Resuts: wireless footswitch 000000000000000000000000001’t pair uer Impact: na Patient Impact: na SEuknown Diagnostic performed by Engineer : Tried nuerou times to pair footswitch and was usucessfu ----------------------------------------------------------------------------- Internal Remarks : 00000000220000000067 FSE Huhes is requsting a part nuber for an uWFS. Recommend: 00000000800000000238232 - Wireless Footswitch 3P ...FSE will need to udate the base station to V3 to be compatible if needed. ----------------------------------------------------------------------------- | 2023-10-13 06:08 | No
2160 | Yes | M0A / u00/ HeartStart uL Defibrillator/Monitor / Battery cannot be charged. | M00A / u000 / HeartStart uL Defibrillator/Monitor / Battery cannot be charged. | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: The First Affiliated Hospital of Nanchang uiversity Neut Step Action: Contacted FSE Hung uinzhang for buiness follow-u Problem Reported by Cutomer: 00000000/HeartStart uL Defibrillator/Monitor Neut Step Action: Contacted FSE Hung uinzhang again for buiness follow-u Safety Qustions: Case: 00000000220000000016 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-13 06:23 | No
2161 | Yes | system boot fail | system boot fail | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000052 00000000-Nov-2000000003 K.J. Sug The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts: part transfer uer Impact: no Patient Impact: no SE8 Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 Resoluion : For part transfer ----------------------------------------------------------------------------- | 2023-10-13 06:57 | No
2162 | Yes | QAA0/u FD0Startup fail | QAA0R/u FD0Startup fail | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 12.1000000002000000003 1000000008:23 uC (86000000008900000000 "RDI Accuacy: YES" $Q7: (Confirmed reported problem, Startu fail) $Q8: (Check the hostpc and os disk, foud the interface board of the hard disk defect.) $Q9: (Bypass the interface board of the hard disk.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools (NA ) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: NA T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine coud not enter the application interface normally and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine cannot enter the application interface normally Remote diagnosis and test resuts: power ouage and restart have no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the Host PC starts normally, try to clean and plu in the internal board, 2. Check whether the BIOS battery is normal 3. ue PC Diagnostic Tool for diagnostic testing 00000000Check whether the software is normal, try to ue Ghost to restore the system 5. Check M Check whether the DCPS of the cabinet is normal. 6. Check whether the internal commuication of the system is normal. 7. Check the machine uLLand repair the faut. If it cannot be repaired, replace the corresponding spare parts. Recommended spare part uuand name: 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD & 0000000080000000050000000000000000 Hard disk spare part Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 00000000220000000011 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 12.1000000002000000003 1000000008:200000000uC (86000000008900000000 Service Content: The system cannot start. Check the host pc and bypass hard disk bou, and the device will retun to normal. Service resuts: After repair, the equpment is normal. Others: none SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: QAA00000000L/u uuuFD200000000 ----------------------------------------------------------------------------- | 2023-10-13 07:28 | No
2163 | Yes | 0 - u0 - Quote for metallic paddles | 00- u000- Quote for metallic paddles | BuQ&R Feedback: No Reason: Corrective Maintenance. Safety Qustions: Case: 00000000220000000058 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: uyni Edwirges. Last udated by: uyni Edwirges. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? N/A, Problem Reported by Cutomer: Qute for metallic paddles Internal Remarks: Technical validation ----------------------------------------------------------------------------- | 2023-10-13 07:43 | No
2164 | Yes | The AC light of the device almost turns on, but there is no response. The screen is black and there is no startup sound | The AC light of the device almost turns on, but there is no response. The screen is black and there is no startup sound | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/HeartStart uL Defibrillator/Monitor/ Neut Step Action: Teacher Tu 13890000000099785 called to uge me, bu no one has contacted me so far. I have ##d ASP Neut Step Action: Yibin Mine Emergency Hospital Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: ##d to ASP Siboru for buiness follow-u Safety Qustions: Case: 000000002200000000000000002 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A OneEMS internal remarks: The AC light of the device almost tuns on, bu there is no response when it is tuned on. The screen is black and there is no startu soud. Buiness follow-u is requred. ----------------------------------------------------------------------------- | 2023-10-13 07:57 | No
2165 | Yes | 0u FD0Space low | 0/u FD0Space low | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/u uuuFD200000000Safety Qustions: Case: 0000000022000000000000000000000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-13 08:02 | No
2166 | Yes | Live uuunot available on fleu vision | Live uuunot available on fleu vision | 000000000000000000000000001O##: 00000000000000000000000000100000000218100000000 17-Oct-2000000003 Baaso, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : installed new NY3 PDuin B cabinet. verified operation and retuned uit to service. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: missing fleuvision inpus uer Impact: med Patient Impact: none SE 1.2 Parts Consued : 1.00000000000000000000000000000000800000000111891,PDuSingle Phase Distribuion uit SP 2.0000000000000000000000000000000080000000000000000900000000,PDuFues Oupus ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021100000000 13-Oct-2000000003 Ramirez, Carlos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech: Device ue: in ue uon arrival Eupected and ueupected Resuts: replacement of parts uer Impact: no patient involved Patient Impact: none SEuknown Follow u Requred : Perform parts replacement. Continu with diagnosis. Parts shipped not foud. Resoluion : Actions taken: no 000000000000000000000000001k performed. System not available to me. System was in ue on arrival ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021100000000000000003-Oct-2000000003 Baaso, Scott The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: Cath lab Eupected and ueupected Resuts: duing power u live and ref feed did not show u on Fleuvision main monitor uer Impact: high Patient Impact: high SE1.2 Diagnostic performed by Engineer : after replacement of footswitch on prior 000000000000000000000000001O# system was powered u, bu the live, ref 1 and coronary tools feed for the main fleuvision monitor did not show u. trouleshot issu with assistance from NSS suport. identified that B cabinet PDuwas not produing 5v and 200000000oupus. this was preventing the transmitters from sending the 3 video feeds to the fleuvision pc. opened follow u 000000000000000000000000001k requst, researched and ordered PDufor overnight delivery. arranged for FSE to be on site to install 10000000013. udated staff and leadership. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Ming Wang IP & Issu: 75920000000093 & wanting fse to come ou today or tomorrrow Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: verified contact info , Created AA and set eupectations Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Planner Notes Planner Initials: TB Action Taken: The planner collaborated with FSE Scott Baaso to follow u on a cutomer call back for a rescheduing requst, FSE Scott Baaso rejected the rescheduing stating that he has already spoken to the listed cutomer: Ming Wang to let him know that 00000000000000000000000000100000000218100000000 has already been schedued with the room and scheduing director "Emily" for the assigned time/date in PST. The planner contacted the cutomer: Ming Wang to relay the FSE's remarks. Oucome: Rescheduing rejected Neut Steps: N/A Last udated: 10000000016/23 5PM EST Internal Remarks : Planner Initials: GM Action Taken: Follow u activity for Scott Baaso for 1000000002100PM (SAT) Oucome: dispatched Internal Remarks : Planner Initials: GM Action Taken: FSE Scott Basso requsted call open and dispatched for 100000000130000000000M Oucome: dispatched Internal Remarks : Planner Initials: GM Action Taken: FSE Scott Baaso requsted call open and dispatched for 1000000001200:300000000M Oucome: dispatched Internal Remarks : FSE Scott Baaso. FSE onsite and not getting live/ref on Fleuvision. It is showing on rabbit ears. He tracked cabling from splitter and then no power 3NY has no 200000000or 5V. Internal Remarks : Planner Initials: GM Action Taken: FSE Scott Baaso requsted call transferred to Carlos Ramirez for 100000000130000000000M Oucome: dispatched ----------------------------------------------------------------------------- | 2023-10-13 08:27 | No
2167 | Yes | uL+ system language has changed to English and cannot be ud normally | uL+ system language has changed to English and cannot be ud normally | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000182 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: ECR:00000000/HeartStart uL+ Defibrillator/Monitor OneEMS internal remarks: The system languge has changed to English and cannot be ued normally. Buiness follow-u is requred. ----------------------------------------------------------------------------- | 2023-10-13 08:57 | No
2168 | Yes | Euhalation flow sensor defective. | Euhalation flow sensor defeu. | 000000000000000000000000001O##: 0000000000000000000000000010000000021200000000 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : Euhalation flow sensor defective. Requsted cutomer go for equpment level ugrade as this uY is EOL and no part suport available. ----------------------------------------------------------------------------- | 2023-10-13 10:02 | No
2169 | Yes | 0 JPN Hu00The nozzle was broken while cleaning between his teeth | 00JPN Hu0/0 The nozzle was broken while cleaning between his teeth | 2000000003-10000000012 00000000:13:28:Produt safety accident PSI5500000000; 2000000003-10000000012 00000000:53:21:1：N/A 2: I received a power line to ask abou a produt safety incident. 3：N/A 00000000Safety was juged based on the hearing of the gudelines.I've finished hearing abou a produt safety incident. I confirmed that the AED was fauty du to the problem, so I let him know that the AED will be replaced.He uderstood, so I will close this.; 2000000003-10000000012 00000000:0000000032:A_ occurence date 2000000003/10000000000000000 B_ produt puchase date 2000000002/00000000/00000000 C_ manuactuer's code 2100000000300000000 D_ I actully heard abou the part nuber Hu8516/00000000 E_ code change date Withou 1. Symptoms overview Is broken 2. Please tell me the details of the sitution when this happened.(Please also tell me abou the sitution if there is any injuy or injuy to the wheren) While cleaning the back teeth, it got broken on the teeth. 3. When did the produt start to be ued? Since aroud Juy 2000000002 00000000What is the condition of the produt after the problem occured?Can the appearance of the produt in particuar be confirmed to be fauty or damaged?Please attach the images.(If the problem occus with an accessory, please also tell me the statu of the accessory.) Withou 5. Please specifically tell me the frequncy of ue of the produt in qustion (if the problem occus with the accessory, please also tell me the frequncy of ue of that accessory). Once a day 6. When the problem occured, where (in what room) was the produt ued?(If the problem occus with an accessory, please also tell me abou that accessory.) Wash wash 7. Where was the produt stored? Wash wash 8_ (if applicable) Where did the produt charge? Wash wash Did youcontact 38_euternal (other than the puchased store)?Please provide the details. Withou; 2000000003-10000000012 00000000:00000000500000000_ハンドルにノズル did younotice before cleaning the teeth even when the battery was fitted? No 2. Did younotice it in the oral cavity (did youfeel it)? Yes 3. Injuy? No 00000000Concern? No 5_ Comment (Safety or Qulity jugement) Safety jugement; 2000000003-10000000000000000 00000000:52:00000000Masako Kojima (u Masako Kojma) | 2023-10-13 10:32 | No
2170 | Yes | System failure/Starts in service mode | System failure/Starts in service mode | 000000000000000000000000001O##: 00000000000000000000000000100000000357000000009 Katori, Masato The system does not meet the specification for the performed service and is not retuned to ue Resoluion: [000000000000000000000000001k details] December 000000002000000003 u Hanano of Fuua Electronics Technical Service contacted u to cancel the repair and retun the produt. [Simutaneou FCO 000000000000000000000000001k] There are no FCOs that requre 000000000000000000000000001k on this equpment. [Special Note] December 5, 2000000003 We sent the equpment and 000000000000000000000000001k report to u Hanano of Fuua Electronics Technical Service. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit the 000000000000000000000000001k report on paper? <YES> Resoluion: [Special Note] December 6, 2000000003 We received notification from Fuua Electronics Technical Service regarding the receipt of the device. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Fuua Denshi Technical Service Device ue: Not for clinical ue Eupected and ueupected Resuts: Cannot be started and ued in service mode uer Impact: None Patient Impact: None SE2.00000000.21 Diagnostic performed by Engineer: [Confirmation details] Althouh we were not able to confirm the phenomenon of power failue, we confirmed that when the power is tuned on, it starts in service mode. When we performed an operation check, we foud that the General System Test item displayed In Progress and did not proceed futher.(Confirmed that the problem was improved by replacing the Processor PCA) ----------------------------------------------------------------------------- | 2023-10-13 10:42 | No
2171 | Yes | 0/FD0Startup failed | 0R/FD0Startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 15.1000000002000000003 23:29:00000000 uC (8600000000900000000) "SuT. Ready for Clinical ue local time" : " 2000000003-1000000001000000008:00000000" Service content: The device cannot enter the system after being tuned on. It is diagnosed as Host PC. OS hard disk failue, replace the hard disk and restore system backu. Repair resuts: After repair, the equpment retuned to normal operation SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 15.1000000002000000003 23:29:16 uC (8600000000900000000) "RDI Accuacy:Yes" $Q7:(Confirmed reported problem. The failue occured in system control, The system cannot be accessed after the device is powered on. $Q8:(Checked logs and foud NA Checked cables connections on IPPC and HOST PC was ok, bypass the HOST PC OS disk was not OK, and the system cannot enter the FSF mode, restore the Ghost Image was failed. the HOST PC OS disk was defect) $Q9:(Replaced The Hard disk spare part (uu0000000080000000050000000000000000 Performed Restore a Ghost Image T2 sugestion effective.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. $Q11: (Calibrated Tools Name $Calibrated Tools name (NA); S/N (NA); Neut Calibration date (NA)) SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine coud not enter the application interface normally and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine cannot enter the application interface normally Remote diagnosis and test resuts: power ouage and restart have no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the Host PC starts normally, try to clean and plu in the internal board, 2. Check whether the BIOS battery is normal 3. ue PC Diagnostic Tool to perform diagnostic tests 00000000Check whether the software is normal, try to ue Ghost to restore the system 5. Check M Check whether the DCPS of the cabinet is normal. 6. Check whether the internal commuication of the system is normal. 7. Check the machine uLLand repair the faut. If it cannot be repaired, replace the corresponding spare parts. Recommended spare part uuand name: 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD & 0000000080000000050000000000000000 Hard disk spare part Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000512 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD 0000000080000000050000000000000000*1 200000000Wang Shengkai-86-15000000008986515 Normal shipment Neut Step Action: PPD 00000000800000000500000000000000001 200000000Wang Shengkai-86-15000000008986515 Normal shipment SAP Transfer: Case Transferred to SAP T2 Activities:Repair Action: 0000000080000000050000000000000000Problem Reported by Cutomer: 00000000R/FD200000000 ----------------------------------------------------------------------------- | 2023-10-13 11:22 | No
2172 | Yes | SA:0/FD0uray intermittent problem | SA:0K/FD0uray intermittent problem | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 13.1000000002000000003 12:5000000003 uC (86000000009567) Only on job training Internal Remarks: 15.1000000002000000003 00000000:00000000:29 uC (86000000009567) only on job training Internal Remarks: 16.1000000002000000003 11:00000000:33 uC (8600000000800000000)$Q1: (Y) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem.uray intermittent problem.) $Q8: (Performed system test, the Q is defective Need to replace the Q) $Q9: (Replace the Q) $Q100000000 (Test and inspection data refer to: other fse PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: (N/A) ^Curent Tue installed date: (N/A) ^Curent Tue uage: (N/A) Neut Step Action: PPD 00000000217800000000100000000 1. Ou of stock. ugent order has been arranged. uu Yu -86-1330000000010000000011 Euternal Remarks: 2000000001000000002000000003 00000000:2000000001 uC (86000000009567) only on the job Neut Step Action: uu Yu Internal Remarks: 17.1000000002000000003 13:10000000000000000TC (860000000095300000000$Q1: (Y) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem.uray intermittent problem.) $Q8: (Performed system test, the Q is defective Need to replace the Q) $Q9: (Replace the Q) $Q100000000 (Test and inspection data refer to:other fse. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: (N/A) ^Curent Tue installed date: (N/A) ^Curent Tue uage: (N/A) Euternal Remarks: 17.1000000002000000003 13:1000000001 uC (860000000095300000000 Service content: Check the reason for failue to eupose, replace Q and euchange test tank. Service resut: The device is in normal ue. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Repair Action: 00000000217800000000100000000 uit dig. kV mA control SAP Transfer: Case Transferred to SAP Internal Remarks: 13.1000000002000000003 12:5000000006 uC (86000000009567) Only on job training Internal Remarks: 00000000.11.2000000003 10000000000000000:200000000uC (86000000008277)"RDI Accuacy:YES" $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem, no u-ray intermittent) $Q8: (check logfile and euchange Q1and Q2 test TANK uLLtransferred from Q1to Q2 replace Q1 bu problem still occued then do ghost system ） $Q9: (replace Q1 and do ghost system restore backu epu system was ok) $Q100000000 (Test and inspection data refert to PA Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: N/A. ^Curent Tue installed date: NA Internal Remarks: 2000000001000000002000000003 00000000:20000000000000000TC (86000000009567) only on the job Internal Remarks: 17.1000000002000000003 0000000000000000:00000000 uC (8600000000800000000)$Q1: (Y) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem.uray intermittent problem.) $Q8: (Performed system test, the Q is defective Need to replace the Q) $Q9: (Replace the Q) $Q100000000 (Test and inspection data refer to: other fse. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: (N/A) ^Curent Tue installed date: (N/A) ^Curent Tue uage: (N/A) Problem Reported by Cutomer: SA:00000000K/FD200000000 Internal Remarks: 17.1000000002000000003 0000000010000000012 uC (8600000000800000000)$Q1: (Y) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem.uray intermittent problem.) $Q8: (Performed system test, the Q is defective Need to replace the Q) $Q9: (Replace the Q) $Q100000000 (Test and inspection data refer to: other fse. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: (N/A) ^Curent Tue installed date: (N/A) ^Curent Tue uage: (N/A) Euternal Remarks: 16.1000000002000000003 0000000000000000:00000000uC (86000000009567) only on job training Internal Remarks: 16.1000000002000000003 0000000000000000:51 uC (86000000009567) only on job training Euternal Remarks: 16.1000000002000000003 11:00000000:300000000uC (8600000000800000000) Service content: Check the reason why it cannot be euposed, replace Q and euchange test tank. Service resut: The device is in normal ue. SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000552 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: FSE uu Yu 1330000000010000000011 called. On the way, he went to answer the cutomer's call and asked RSE to call him back later. Neut Step Action: Cutomers who have retuned to visit cannot see the resuts. Internal Remarks: 00000000.11.2000000003 12:57:00000000uC (8600000000900000000)$Q7: (Confirmed reported problem, no u-ray intermittent) $Q8: (check logfile and euchange Q1and Q2 test TANK uLLtransferred from Q1to Q2 replace Q1 bu problem still occued then do ghost system ） $Q9: (replace Q1 and do ghost system restore backu epu system was ok) $Q100000000 (Test and inspection data refert to PA Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: N/A. ^Curent Tue installed date: NA Euternal Remarks: 00000000.11.2000000003 12:57:00000000uC (8600000000900000000) Service content: Check the reason why it cannot be euposed, replace Q and euchange test tank. Service resut: The device is in normal ue. Euternal Remarks: 17.1000000002000000003 0000000000000000:00000000uC (8600000000800000000) Service content: Check the reason why it cannot be euposed, replace Q and euchange test tank. Service resut: The device is in normal ue. Euternal Remarks: 17.1000000002000000003 0000000010000000000000000 uC (8600000000800000000) Service content: Check the reason why it cannot be euposed, replace Q and euchange test tank. Service resut: The device is in normal ue. SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine was intermittently uable to emit rays and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine is intermittently uable to emit rays Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the high-voltage commuication is normal. 2. Check whether the MPD control board is normal. 3. Check whether the IPPC is normal. 00000000Check the machine uLLand perform faut repairs. Recommended spare parts uuand designation: NA Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 15.1000000002000000003 00000000:00000000:26 uC (86000000009567)only on job training SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 10000000000000000:17 uC (86000000008277) Service content: Check the reason why it cannot be euposed, replace Q and euchange test tank.The test wiring faut still occured intermittently. After the system was restored, the equpment retuned to normal. Service resut: The equpment retuned to normal and all fuctions are working normally. ----------------------------------------------------------------------------- | 2023-10-13 11:37 | No
2173 | Yes | Individual images with foot switch not possible. Filming is not possible like this | Individual images with foot switch not possible. Filming is not possible like this | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000001 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Sw Conny +000000001812960000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: No release via foot switch possible anymore. Asked cutomer to check connections and cables. They are all OK. Log shows no defect. Asked whether youcan replace the foot switch youself, does not 000000000000000000000000001k becaue youdo not see the connection at all? Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: Schedue a service for the replacement of the foot switch. 1-2 hous of 000000000000000000000000001k There is a service in the neighboring lab on Monday anyway. Cutomer is asking for this case to be planned in addition and to be addressed first. Technician name: Robert Kern ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000212635 Krub, Horst Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Sw Conny +000000001812960000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Cutomer comment: ### Date with Sw. Conny Monday, 10000000016, 9:00000000 am ----------------------------------------------------------------------------- | 2023-10-13 12:02 | No
2174 | Yes | Problem with wi-fi pedal. | Problem with wi-fi pedal. | 000000000000000000000000001O##: 00000000000000000000000000100000000212700000000 October 13th, 2000000003 Chevalley Christophe The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Sebastien Kase Device ue: Yes Eupected and ueupected Resuts: No connection with Wi-Fi foot switch. uer Impact: No. Patient Impact: No. SE 8.2.17 Resoluion: Disassembly of the pedal and reset of the pedal. Disassembled the table and paired the pedal with the WiFi base. Pedal test. Reset the system database to zero. Final test. In order. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000Chevalley Christophe The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: Diagnosis of Wi-Fi pedal. Problem Description by Engineer: Cutomer Fuction/Role: Sebastien Kase Device ue: Yes Eupected and ueupected Resuts: No connection with Wi-Fi foot switch. uer Impact: No. Patient Impact: No. SE 8.2.17 Diagnostic performed by Engineer: Contacted the cutomer. On wi-fi pedal, red Wi-Fi connection LED. Go on site for a dignostic ----------------------------------------------------------------------------- | 2023-10-13 12:17 | No
2175 | Yes | The AED beeped once and continued to require self-test | The AED beeped once and continued to require self-test | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000588 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: I confirmed that the "Cannot ue" message and the three-beep alarm souded when I performed a battery test on the AED.I confirmed that the same alarm occured when the test battery was ued. I verified the operation record of the AED and foud that a self test detected an uLLin the high voltage element inside the AED, which failed the test.There was a record of an AED failue, so I replaced the AED. "I apologize for the inconvenience this has caued. " Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "October 23, 2000000003 Fuua Denshi Hiroshima Sales Report to the on-site Person in Charge and I also sent the equpment." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The AED beeps once when it is normally stored. The self test does not end. He re-removed the battery and re-tested the AED, and then a self test was requred. The actul consuable part is not a sustitue uit. LOT Adut pad Y000000001822-00000000 Battery 91875P ----------------------------------------------------------------------------- | 2023-10-13 12:22 | No
2176 | Yes | Monitor needs to be adjued for alarming correctly. No Product Malfunction-Configuration. Config changes. | Monitor needs to be adjued for alarming correuy. No ProduuMalfunuon-Configuration. Config changes. | Test Equpment Info : [Device #: 12277700000000Make: PRONK TECHNOLOGIES INC, uY: SC-5 SIMCuE, Serial: 5281, Last Cal Date: 2000000003-00000000-15] [Device #: 1722266, Make: PRONK TECHNOLOGIES INC, uY: Ou-2 OuSim Fleu, Serial: 000000009, Last Cal Date: 2000000003-00000000-00000000] [Device #: 7000000007000000007, Make: PRONK TECHNOLOGIES INC, uY: SL-8 SIMSLIM, Serial: 95300000000Last Cal Date: 2000000003-00000000-00000000] Device Alarm/Alert as intended? : No Problem Observed : Monitor needs to be adjuted for alarming correctly Ou Of Bou Failue? : No Labor: Labor ID : L-19830000000017 000000000000000000000000001k Performed : Change profile settings to how PACuprefers it. Diagonsis : Change profile settings | 2023-10-13 12:31 | No
2177 | Yes | Therapy Receptacle damaged. - | Therapy Receptacle damaged. - | 000000000000000000000000001O##: 0000000000000000000000000010000000056600000000 Buak Tezel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Diagnostic. Eupected and ueupected Resuts: Therapy Receptacle is broken. uer Impact: No. Patient Impact: No. SE 2.00000000.21 Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000000000000000000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000003-16 00000000:00000000:00000000 MuTIMETER, 200000000000000002-16 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle Resoluion: The broken part of the device was replaced with a new one. All tests and controls were carried ou according to the procedue specified in the service docuent. No problems were observed. The device was delivered to the uer in 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000212900000000 Buak Tezel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Diagnostic. Eupected and ueupected Resuts: Therapy Receptacle is broken. uer Impact: No. Patient Impact: No. SE 2.00000000.21 Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000000000000000000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000003-16 00000000:00000000:00000000 MuTIMETER, 200000000000000002-16 00000000:00000000:00000000 Follow u Requred: N/A Resoluion: Since the part has been moved to the 000000000000000000000000001O#nubered 0000000000000000000000000010000000056600000000, all operations of the device have been closed in the relevant 000000000000000000000000001O# All tests and controls of the device have been carried ou and it has been delivered to the uer in 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000212855 Buak Tezel The system does not meet the specification for the performed service and is not retuned to ue Resoluion: Duing the checks, it was determined that the Spoon Inpu Socket of the device was broken du to uer-related reasons. The fauty Therapy Receptacle part mut be replaced with a new one. According to the faut detection, 0000000056000000008971 MuT BE OFFERED TO CERRAHPAŞA TRAINING RESEARCH HOSPITAL. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Diagnostic. Eupected and ueupected Resuts: Therapy Receptacle is broken. uer Impact: No. Patient Impact: No. SE 2.00000000.21 Follow u Requred: Duing the checks, it was determined that the Spoon Inpu Socket of the device was broken du to uer-related reasons. The fauty Therapy Receptacle part mut be replaced with a new one. According to the faut detection, it shoud be OFFERED TO 0000000056000000008971 - CERRAHPAŞA TRAINING RESEARCH HOSPITAL. ----------------------------------------------------------------------------- | 2023-10-13 12:52 | No
2178 | Yes | QA 000000- No audio alarm - G0Goldway G0Patient Monitor - Changed speaker | QA 00000000 - No audio alarm - G0Goldway G0Patient Monitor - Changed speaker |  | 2023-10-13 12:56 | No
2179 | Yes | QA 000000 "temperature is too high" alarm shows | QA 00000000 - "temperature is too high" alarm shows |  | 2023-10-13 12:57 | No
2180 | Yes | QA 000000 Device disabled during u | QA 00000000- Device disabled during u |  | 2023-10-13 13:14 | No
2181 | Yes | The AED beeped three times | The AED beeped three times | Internal Remarks: Failue content 2 3-beep soud Adut pad: C00000000723-00000000 Adut pad: C00000000723-00000000 Battery: 96887P ----------------------------------------------------------------------------- | 2023-10-13 13:42 | No
2182 | Yes | Plates are damaged | Plates are damaged | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000273 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: no Eupected and ueupected Resuts: Damaged plates uer Impact: no Patient Impact: no SE na The plates are damaged. Diagnostic performed by Engineer: Heard from the cutomer: plates are damaged. Qute will follow. Follow u Requred: Emac, Via del Commercio, 5F 16167, Genoa. C\A Silvia Scala. 000000000000000029000000001. ----------------------------------------------------------------------------- Internal Remarks: Good morning, This is to requst estimate and delivery times for: - 0000000056300000000991 water-resistant paddle assembly for defibrillator u uDFM100000000 s/n CN3260000000072. uEmergency Department P-1 Hospital Savona Waiting for you kind feedback, Kind regards Chiara D'Antuno Cutomer Care - Technical Office ________________________________ ----------------------------------------------------------------------------- | 2023-10-13 14:28 | No
2183 | Yes | 0DM0M0）/Unable to discharge | 0/DM0M00/Unable to discharge | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: uable to discharge Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: liuuiaohuLast udated by: liuuiaohu1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: SN#72797000000003726 provided by the cutomer, uY nuber DM100000000(M200000000 was not foud in the system Neut Step Action: ##d ASP Honglejia for follow u ----------------------------------------------------------------------------- | 2023-10-13 14:48 | No
2184 | Yes | display monitor on fleuvision intermittently shuts off during case (found during pm) | display monitor on fleusion intermittently shuts off during case (found during pm) | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 16-Nov-2000000003 Ortiz, Romy The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.000000000000000000000000300000000000000006935721,19" Color LCD Monitor ER (DC19-LER) 1.000000000000000000000000300000000000000006935721,19" Color LCD Monitor ER (DC19-LER) Problem Description by Engineer : Cutomer Fuction/Role: radtech Device ue: ouside of ue Eupected and ueupected Resuts: intermittent issu on display monitor of fleuvision uer Impact: no impact Patient Impact: no impact SE 8.2 Resoluion : Replace 3DRA and hemo monitor on fleuvision Re configue display Check connections Endorse to uer for check System monitoring ----------------------------------------------------------------------------- | 2023-10-13 14:48 | No
2185 | Yes | Requests a new plug test becau it is worn and not working | Requests a new plug test becau it is worn and not working | 000000000000000000000000001O##: 0000000000000000000000000010000000021000000003 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000213218 RENZO FORESTIERI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Francesco Germinal Device ue: defibrillator udfm100000000 Eupected and ueupected Resuts: suply plu test uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Spoke with Francesco Germinal, he requsts a new plu test becaue it is 000000000000000000000000001n and not 000000000000000000000000001king, the part will be sent to the address indicated Follow u Requred: make nemo withou collection at the same time send to: Vito Fazi Hosp. Muatore Squre Dea pavilion Heart Center 3rd floor c/o Francesco Germinal tel 0000000032335300000000/3280000000086283 ----------------------------------------------------------------------------- | 2023-10-13 15:47 | No
2186 | Yes | Machine not getting on | Machine not getting on | 000000000000000000000000001O##: 0000000000000000000000000010000000028122000000009-Oct-2000000003 Ramesh, S The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine not getting on When was the first occurence? Within... : 13-Oct-2000000003 15:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine not getting on uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.3000000005 Resoluion : Cutomer complaint regarding uable to switch on the system, reached site and observed and diagnosed foud front panel assembly was fauty so the same ordered and foud Singapore part so waited, mean while observed earth to neural voltages are abnormal and seen some rodent activities in all cabinets & condole room also so the same advised to cutomer to correct it abd fiued the received part and loaded software foud system botted normal Parts Consued : 1.0000000000000000000000000000000080000000000000000321,PD.ASSY.REARPANEL BOuED 1.0000000000000000000000000000000080000000000000000300000000,PD.ASSY.FRONTPANEL BOuED ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000213900000000 15-Oct-2000000003 Ramesh, S The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Machine not getting on When was the first occurence? Within... : 13-Oct-2000000003 15:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine not getting on uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.3000000005 Diagnostic performed by Engineer : Cutomer complaint regarding system not switching on, reached site & observed , diagnosed and foud PDufront panel assembly is fauty so the same ordered and was in Singapore nee to replace the when arrived Follow u Requred : Need to order parts Tools & Equpment : MuTIMETER,200000000000000009-21 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000038700000000,PD.SCOMP.uD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000212726 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Machine not getting on When was the first occurence? Within... : 13-Oct-2000000003 15:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine not getting on uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.3000000005 Diagnostic performed by Engineer : FSE visit requred for futher checks on GPDuin mcabinet Follow u Requred : uLL# and/or description of uLL: Machine not getting on Repair Action : FSE visit requred for futher checks on GPDuin mcabinet Technician Name : Vijayakuar, K.S Resoluion : FSE visit requred for futher checks on GPDuin mcabinet Internal Comments : uLL# and/or description of uLL: Machine not getting on Repair Action : FSE visit requred for futher checks on GPDuin mcabinet ----------------------------------------------------------------------------- | 2023-10-13 15:48 | No
2187 | Yes | Therapy delivery test fail | Therapy delivery test fail | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000335 Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Resoluion : Problem solved by cleaning the Paddle plates. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Therapy delivery test fail ----------------------------------------------------------------------------- | 2023-10-13 16:02 | No
2188 | Yes | SA:0/u FD0cannot startup | SA:0R/u FD0cannot startup | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 15.1000000002000000003 00000000:13:300000000uC (8600000000600000000"SuT. Ready for Clinical ue local time" : " 2000000003-1000000001000000000000000000000000" Service content (requred): Replacement of M cabinet back panel real panel Service resut (requred): The device is runing normally Other important matters: None SAP Transfer: Case Transferred to SAP Internal Remarks: 15.1000000002000000003 00000000:13:37 uC (8600000000600000000 "RDI Accuacy: Yes" $Q7: (Confirmed reported problem. Cannot boot.) $Q8: (Replace the real panel, the device retuned to normal.+The real panel problem.) $Q9: (Replace the real panel. uuof the part: 0000000080000000000000000321.) $Q100000000 (Test and inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (FLuE 287C) S/N (23350000000017) Neut Calibration date (200000000000000007/100000000 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: SA:00000000R/u uuuFD200000000Neut Step Action: PPD0000000080000000000000000321 1 Arranged to go to the hospital on 18H FSE: Zheng Kai-86-18100000000000000000000000033 T2 Activities:Problem Description: Overview of cutomer problems: Cannot boot Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: IPPC and host PC do not have 2200000000, and the switch indicator light does not light u Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: 1. Check the power suply startu circut Recommended spare parts uuand designation: PPD uu0000000080000000000000000321 PD.ASSY.rearpanel boued 200000000to Hosp Safety Qustions: Case: 000000002200000000953 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-13 16:07 | No
2189 | Yes | uLLsays avtive button has been detected, release button for startup. | uLLsays avtive button has been deteud, release button for startup. | Internal Remarks : Interaction Type: New Case Caller: Kathy Boddie 7573880000000022 IP & Issu: 82000000008929/uLLsays avtive buton has been detected, release buton for startu. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Qute Reason for Entitlement Claim: n/a Action Taken: Requsted qute Neut Step Action: Awaiting Qute Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-13 17:17 | No
2190 | Yes | CATH LAB WONT BOOT UP | CATH LAB WONT BOOT UP | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000500000000 October 13th, 2000000003 Ameu Abderrahim The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: u FD100000000 Eupected and ueupected Resuts: Fleuvision PC (Grabber board) issu PC was replaced and reloaded software. uer Impact: No. Patient Impact: No. SE 8.1.3000000005 Resoluion: Fleuvision PC (Grabber board) issu PC was replaced and reloaded software. Parts Consued: 1.0000000000000000000000000000000080000000012200000000, u Haswell Fleuvision 2 PC no HDD ----------------------------------------------------------------------------- Internal Remarks: RSE: ReseAn kiernan asked to cancel remote 00000000000000000000000000100000000200000000538. He said that he coud not open the case (he didn't have the priviledges). He spoke to the Cu. Cu's their room is down, they wanted to see if anything coud be done to resolve it. Onsite 000000000000000000000000001was already created. Internal Remarks: Abderrahim - system boots to 0000000000000000 fleuvision frame grabber card uLL. Approved fleuvision PC part order. Internal Remarks: Cu Dany called again wanting remote suport. Said He called earlier bu coudn't get remote suport as it was before 8 AM. Adv'd an onsite was already created. Cu insisted on remote suport. Created another Remote 000000000000000000000000001 transferred to RSE. Internal Remarks: Interaction Type: New Case Caller: Dany Maheu IP & Issu: 6100000000500000000CATH LAB 000000000000000000000000001'T BOOT u Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: ENTITLED Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 CALLED ON CALL Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-13 17:17 | No
2191 | Yes | M0A / u00/ The defibrillator does not turn on anymore, they tested the battery and it works. | M00A / u000 / The defibrillator does not turn on anymore, they tested the battery and it works. |  | 2023-10-13 17:32 | No
2192 | Yes | [Q] Wireless footswitch not operational | [Q] Wireless footswitch not operational | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 23-Nov-2000000003 Glyn Allen The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: replace footswitch du to lost connection uer Impact: Patient Impact: SE Resoluion : Replaced footswitch and base station du to lost connection issus. Tested and all 000000000000000000000000001king as eupected. System retuned to clinical ue. Parts Consued : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station ----------------------------------------------------------------------------- Cutomer comment : From: Lay, Abigail Sent: Thusday, October 19, 2000000003 9:100000000M To: Muray, Brian <brian.murayucom>; Planning uI <planning.uiucom>; Watson, David <david.watsonucom> Suject: RE: Absence Hi Brian, No problem I have altered planning and added you sickness. Kind Regards Abigail Lay Long Term Planner www.u.co.u u ELECTRONICS u LIMITED Registered Office: Ascent 1, Aerospace Bouevard, Farnborouh, Gu00000000uW. Registered in England No 00000000897 www.u.co.u ________________________________________ The information contained in this message may be confidential and legally protected uder applicable law. The message is intended solely for the addressee(s). If youare not the intended recipient, youare hereby notified that any ue, forwarding, dissemination, or reprodution of this message is strictly prohibited and may be ulawfu. If youare not the intended recipient, please contact the sender by retun e-mail and destroy all copies of the original message. From: Muray, Brian <brian.murayucom> Sent: Thusday, October 19, 2000000003 8:59 AM To: Planning uI <planning.uiucom>; Lay, Abigail <abigail.layucom>; Watson, David <david.watsonucom> Suject: Absence Hi Abigail, I am uavailable to 000000000000000000000000001k today and tomorrow du to a chest infection. I've teut Fergal abou the service this morning, please move 00000000000000000000000000100000000899000000007 to Tusday 21st of November , If we have no FSE availability to cover the corrective 0000000000000000000000000010000000020000000000000000 on Friday, then move this to Thusday 23rd of November. Kind regards, Brian. Internal Remarks : Brian Muray spoke with Fergal Doherty and confirmed Friday 200000000h October for corrective 000000000000000000000000001k to be completed (AM). ----------------------------------------------------------------------------- | 2023-10-13 17:42 | No
2193 | Yes | Alarm problem | Alarm problem | 000000000000000000000000001O##: 00000000000000000000000000100000000213382 CHEN, Cheng The system meets the specifications for the performed service and is retuned to ue. Resoluion: Reconfiguation of the alarm at the windows configuation panel in Soud. First place the defaut louspeakers then disconnect the screen and pu back the screen. The louspeakers do annouce the alarms. System operational. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: no more auio alarm on the central station. ----------------------------------------------------------------------------- | 2023-10-13 18:08 | No
2194 | Yes | [Q] THE TABLE MOVES WITHOUT CONTROL - Split | [Q] THE TABLE MOVES WITHOUT CONTROL - Split |  | 2023-10-13 18:10 | No
2195 | Yes | Low tidal volume, GFE00CT0 Response received/IS complete | Low tidal volume, GFE0 0Ou0Response received/IS complete |  | 2023-10-13 18:10 | No
2196 | Yes | [Q] CV u uLLFLUOSTR-IMAGEDS | [Q] CV u uLLFLUOSTR-IMAGEDS | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000973 Sulivan Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : remote alert of CV u: uLLFLuSTR-IMAGEDS When was the first occurence? Within... : 16-Oct-2000000003 23:58:00000000 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: M2M/Radar Device ue: No patient involved Eupected and ueupected Resuts: uer msg; WARNING; Flu storage not poss. Image disk problem. uer Impact: possible loss of storage space Patient Impact: none SE 8.2.3000000005 Resoluion : system does not meet the criteria for dispath rue of 00000000PPENDING ALERTS IN 2 DAYS closing case Internal Comments : uLL# and/or description of uLL: remote alert Repair Action : The following alarm has been generated by asset CV_722000000006-297: Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog: Flu storage not poss. Image disk problem Description=alert rue version RMA uuu200000000-----WAIT FOR 00000000PPENDING ALERTS IN 2 DAYS!! KEEP THIS M2M CASE OPEN SO M2M CAN ADD THE APPENDING ALERTS! IF NO APPENDING ALERT IS RECEIVED THEN CLOSE THIS CASE. DO NOT DISPATCH FSE IF FAIL RATE IS LESS THAN 00000000-----------------------------------------------------------Duing POST of the IP-PC Image Disk has failed. Order 2 Image Disks. Save all local euams to PACS or other archive. Replace Image Disks.------------------------------------------------------------(IF FSE is onsite THEN ignore this alert. Alerts are filtered when u is in FSF-mode and Service-mode. To ensue alerts are filtered ou when FSE is doing PM or CM let FSE activate Service-mode via the DataMonitor as long as he is onsite) Diagnostic performed by Engineer : Rse to monitor for additional alert fillowing alert rue gudance ----------------------------------------------------------------------------- | 2023-10-13 18:37 | No
2197 | Yes | II not responding | II not responding | Internal Remarks : called cintact twice .left vmail ,created on site attached lod to case level ----------------------------------------------------------------------------- | 2023-10-13 18:37 | No
2198 | Yes | No energy is coming to the device. | No energy is coming to the device. | 000000000000000000000000001O##: 00000000000000000000000000100000000215318 23-Oct-2000000003 Eu Gok The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: When the device is opened, F100000000fue in MPDuis ou of ue. How often is the issu occuring? : First occurence Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved. Eupected and ueupected Resuts: System is not switched on. uer Impact: The system is down. Patient Impact: No impact. SE 000Resoluion: The MPDupreviouly foud to be defective was replaced with a new one. A test was made and the system was delivered in active. Parts: *1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000215000000001 Buak Yediyildiz The system does not meet the specification for the performed service and is not retuned to ue. Internal Comments: uLL# and/or description of the uLL When the device is opened, F100000000fue in MPDuis ou of ue. Repair Action: u Erhan was met. He told me that MPD F100000000fue was thrown when the device was opened. MPD control uit will be changed. Follow u Requred: uLL# and/or description of the uLL When the device is opened, F100000000fue in MPDuis ou of ue. Repair Action: u Erhan was met. He told me that MPD F100000000fue was thrown when the device was opened. MPD control uit will be changed. What are the neut trouleshooting steps: Disconnect NP-u100000000 and NP-u100000000 and remove fue F11. When the fue F100000000still blows, then replace the MPDucontrollor Technician Name: Buak Yediyildiz Problem Description by Engineer: Problem description by engineer: When the device is opened, F100000000fue in MPDuis ou of ue. How often is the issu occuring? : First occurence Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved. Eupected and ueupected Resuts: System is not switched on. uer Impact: The system is down. Patient Impact: No impact. SE 000Diagnostic performed by Engineer: I talked to u Erhan. He told me that MPD F100000000fue was thrown when the device was opened. MPD control uit will be changed. ----------------------------------------------------------------------------- | 2023-10-13 18:52 | No
2199 | Yes | Speaker malfunction - no sound. Cuomer Ordered Part: Assembly/PCB/Card-Speaker-HW. | Speaker malfunuon - no sound. Cuomer Ordered Part: Assembly/PCB/Card-Speaker-HW. | 000000000000000000000000001O##: 00000000000000000000000000100000000215200000000Mottley, Sheldon The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Is Caue Known: Yes Caue: PartsID for speaker Is Resoluion Known: No Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Parts/Suplies Information Provided Part ID and Price: 00000000261000000009000000001 Speaker Assembly uD 85.00000000 Prevention: Follow u Requred : Beth Israel Deaconess Med Ctr Inc 00000000 Brookline Ave Boston, MA 00000000215-500000000 ATTN: RICHARD PELLETIER Clinical Engineering Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: PartsID for speaker uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: PartsID for speaker Diagnostic performed by Engineer : Problem Description: PartsID for speaker PartsID for speaker Speaker does not 000000000000000000000000001k device is uable to produe soud Form ued: Parts Trouleshooting Steps Taken: Beth Israel Deaconess Med Ctr Inc00000000 Brookline Ave Boston, MA 00000000215-500000000ATTN: RICHARD PELLETIER Clinical Engineering Was there a “speaker malfuction” INOP message? (No)Was the speaker able to produe soud, despite the uLLmessage. (No)Coud youplease confirm, there was no health impact to a patient at any time du to the described issu? (No Health Impact) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021500000000 Parts Consued : 1.00000000000000000000000000000000261000000009000000001,IV-MP5 CBL Speaker Assembly ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Richard Pelletier +16176328325 IP & Issu: 6100000000900000000Parts ID Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: R000000000000000000000000001created and roued call to RSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-13 18:58 | No
2200 | Yes | Main monitor froze and started working after reboot | Main monitor froze and started working after reboot | 000000000000000000000000001O##: 000000000000000000000000001000000002000000000000000023-Oct-2000000003 Jonathan, Lopez The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Problem Description by Engineer : Cutomer Fuction/Role: rad tech Device ue: none Eupected and ueupected Resuts: not booting uer Impact: none Patient Impact: none SE8.1.300000000Resoluion : replaced fleuvision PC and performed software load fuction test passed ----------------------------------------------------------------------------- Internal Remarks : 000000002200000000992 Jonathan Lopez 000000000000000000000000001king with FSE on u FD2000000008.1.3000000005 where Fleu PC had failed and FSE ordered new one. FSE went from a Haswell to a Nehalem. Now FSE cant see the Fleu PC in DBS. Asked FSE for old Haswell nuber and was able to find good part nuber. FSE is going to reorder the correct fleu pc. Advised FSE to try the hard drive to see if it 000000000000000000000000001k, else he will have to reload software. Internal Remarks : 000000002200000000992 Jonathan Lopez FSE calling back after puting back old FV PC and old hard drive. FSE says that the system gets to 00000000:00000000. FSE says FV PC does not show in the DBS. Asking FSE to take a pictue of DBS. Had FSE go into open profile to make sue he had backu of FV on the E drive. He did. Had FSE go back to closed profile and now everything is good. Having FSE reboot and see if it all comes backu. Internal Remarks : Johnathan Lopez - uuhelp Internal Remarks : Johnathan Lopez - replaced fleu pc, bu still stuk at 0000000000000000. New logs show that after new fleu PC still has grabber uLL. 850000000000000000000000001 Failed to initialize all grabber cards: DgInit failed: 3. [FREE_TEuT] [SEVERITY] Fatal uLL [SEVERITY] uLL Sugest to have FSE ue fresh new disk, re download board SW. Internal Remarks : Interaction Type: Cutomer Call Back Caller: JAMIE B IP & Issu: ccb to check statu of dispatch Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: advised the FSE will be onsite 10000000016 001pm and added note to FSE 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-13 19:38 | No
2201 | Yes | System constantly rebooting | System constantly rebooting | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Problem Description by Engineer : u Intrepid Defibrillator for repair Resoluion : uit failed duing electrical safety test. System constantly rebooting Traced failue to the SOM PCB Replaced the SOM PCB reloaded software. Device passed operational check. Parts Consued : 1.0000000000000000000000000000000056000000001600000000,Intrepid PCCRD System on Modue (SOM) ----------------------------------------------------------------------------- | 2023-10-13 19:47 | No
2202 | Yes | Does not allow fluoroscopy | Does not allow fluoroscopy | 000000000000000000000000001O##: 00000000000000000000000000100000000216100000000October 16, 2000000003 Meira, Fernando The system meets the specification for the performed service and is retuned to ue. Tools & Equpment: DOSE-KVP-LIGHT METER, 2000000000000000000000000000000000000000000000000:00000000 Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal 13-Oct-2000000003 15:30000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No futher information is available as per the cutomer. Eupected and ueupected Resuts: u-rays cannot be made. uer Impact: The equpment cannot be ued. Patient Impact: No futher information is available as per the cutomer. SE 8.2.17.00000000Resoluion: Generator uSC controller replacement. Software load, calibration and check for correct operation. See 000000000000000000000000001000000002261500000000or part consuption and Prontoforms. Follow u Requred: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002261500000000ctober 18, 2000000003 Meira, Fernando The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.00000000000000000000000000000000210000000010000000000000000SC Certeray Tools & Equpment: DOSE-KVP-LIGHT METER, 2000000000000000000000000000000000000000000000000:00000000 Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal 13-Oct-2000000003 15:30000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No futher information is available as per the cutomer. Eupected and ueupected Resuts: u-rays cannot be made. uer Impact: The equpment cannot be ued. Patient Impact: No futher information is available as per the cutomer. SE 8.2.17.00000000Resoluion: Generator uSC controller replacement. Software load, calibration and check for correct operation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000215635 Adrián Sancho López The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal 13-Oct-2000000003 15:30000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No futher information is available as per the cutomer. Eupected and ueupected Resuts: u-rays cannot be made. uer Impact: The equpment cannot be ued. Patient Impact: No futher information is available as per the cutomer. SE 8.2.17.00000000Diagnostic performed by Engineer: The event log of the system is checked and we see uer messages of failed fluroscopy. The equpment is detecting the u-ray tue as not adapted, the uSC uit of the generator coud be failing, it is recommended to replace it. Internal Comments: uLL# and/or description of uLL Root Caue for RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal: 1. Front - u-ray Tue 000is missing in Adaptation data Root Caue for RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal: 1. Front - uSC coud be defective Repair Action: Service actions for RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal: 1.1. Check whether correct 000of Frontal u-ray tue is configued in FSF/PSC. If the 000is not present or incorrect, please configue the correct Frontal 000and perform Frontal Tue adaptation. Reboot the system and look for the udated tue 000in the additional info of logfile with Description = Adaptation Data. Service actions for RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal: 1.1. If configuation of Frontal tue 000is not solving the problem, replace Frontal uSC and perform necessary Generator calibration for the uSC replacement. Follow u Requred: uLL# and/or description of uLL Root Caue for RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal: 1. Front - u-ray Tue 000is missing in Adaptation data Root Caue for RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal: 1. Front - uSC coud be defective Repair Action: Service actions for RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal: 1.1. Check whether correct 000of Frontal u-ray tue is configued in FSF/PSC. If the 000is not present or incorrect, please configue the correct Frontal 000and perform Frontal Tue adaptation. Reboot the system and look for the udated tue 000in the additional info of logfile with Description = Adaptation Data. Service actions for RMT - uGEN: u-ray Tue 000is missing in Adaptation data of Certeray - Frontal: 1.1. If configuation of Frontal tue 000is not solving the problem, replace Frontal uSC and perform necessary Generator calibration for the uSC replacement. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-10-13 19:48 | No
2203 | Yes | The pressure values ​​remain fiued | The pressure values ​​remain fiued | 000000000000000000000000001O##: 0000000000000000000000000010000000021630000000016-Oct-2000000003 DAMIANI, MAuIZIO The system meets the specification for the performed service and is retuned to ue. Internal Comments: The checks of the performance of the machine were carried ou completing the maintenance. The data are reported in the Connected form (0000000000000000000000000010000000083000000008) Problem Description by Engineer: Cutomer Fuction/Role: FSE u Laua Bastoni Device ue: ventilation Eupected and ueupected Resuts: The pressue valus remain fiued uer Impact: no Patient Impact: no SE na Resoluion: Checked log withou detecting uLL. Carried ou flow and pressue checks withou detecting any anomalies Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000003-15 00000000:00000000:00000000 FLOW METER, 200000000000000001-23 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000213373 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: The pressue valus remain fiued. Schedue Mauizio Damiani on 16.10000000023. Problem Description by Engineer: Cutomer Fuction/Role: FSE u Laua Bastoni Device ue: ventilation Eupected and ueupected Resuts: The pressue valus remain fiued uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Colleagu Laua Bastoni reports that pressue valus remain fiued ----------------------------------------------------------------------------- | 2023-10-13 20:02 | No
2204 | Yes | The Stand could not move_SPLIT | The Stand could not move_SPLIT |  | 2023-10-13 20:34 | No
2205 | Yes | Device won't power on; bad power cord | Device won't power on; bad power cord | Test Equpment Info : [Device #: 1273326, Make: TSI INC, uY: CERTIFIER FA PLu, Serial: 0000000081335000000002, Last Cal Date: 2000000008-00000000-31] [Device #: 1273327, Make: TSI INC, uY: 000000001 C, Serial: 000000001133000000002, Last Cal Date: 2000000003-00000000-31] [Device #: 1183113, Make: FLuE BIOMEDICAL CORP, uY: ESA612, Serial: 300000000000000006, Last Cal Date: 2000000003-00000000-31] Device Alarm/Alert as intended? : N/A Problem Observed : Started trouleshooting device and foud that it will not power on, pluged in and foud that the AC light and battery charging light are not on, removed cord and ohmed ou, power cord is bad and needs replaced, tried a known good cord and uit is now 000000000000000000000000001king and comes on, ac and Battery lights now 000000000000000000000000001king. opening case and ordering replacement Ou Of Bou Failue? : No Labor: Labor ID : L-1982900000000 000000000000000000000000001k Performed : Started trouleshooting device and foud that it will not power on, pluged in and foud that the AC light and battery charging light are not on, removed cord and ohmed ou, power cord is bad and needs replaced, tried a known good cord and uit is now 000000000000000000000000001king and comes on, ac and Battery lights now 000000000000000000000000001king. opening case and ordering replacement Diagonsis : Not Answered | 2023-10-13 20:34 | No
2206 | Yes | Request a new plug test becau it is worn and not working | Request a new plug test becau it is worn and not working | 000000000000000000000000001O##: 000000000000000000000000001000000002236800000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000213100000000RENZO FORESTIERI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Francesco Germinal Device ue: defibrillator udfm100000000 Eupected and ueupected Resuts: suply plu test uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Spoke with Francesco Germinal, he requsts a new plu test becaue it is 000000000000000000000000001n and not 000000000000000000000000001king, the part will be sent to the address indicated Follow u Requred: make nemo withou collection at the same time send to: Vito Fazi Hosp. Muatore Squre Dea pavilion Heart Center 3rd floor c/e Francesco Germinal tel 0000000032335300000000 ----------------------------------------------------------------------------- | 2023-10-13 20:48 | No
2207 | Yes | Requests a new plug test becau it is worn and not working | Requests a new plug test becau it is worn and not working | 000000000000000000000000001O##: 00000000000000000000000000100000000220000000000000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000213300000000RENZO FORESTIERI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Francesco Germinal Device ue: defibrillator udfm100000000 Eupected and ueupected Resuts: suply plu test uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Spoke with Francesco Germinal, he requsts a new plu test becaue it is 000000000000000000000000001n and not 000000000000000000000000001king, the part will be sent to the address indicated Follow u Requred: make nemo withou collection at the same time send to: Vito Fazi Hosp. Muatore Squre Dea pavilion Heart Center 3rd floor c/e Francesco Germinal tel 0000000032335300000000 ----------------------------------------------------------------------------- Internal Remarks: Good morning Attached is requst for technical intervention for repair of the following electromedical app. u DEFIBRILLATOR uY uDFM 100000000 SN CN 326100000000500000000 REPORTED FAuT: suply defibrillator test plu 989800000000171271 For futher information please contact coordinator Francesco Germinal 0000000032335300000000 Waiting for feedback. Sincerely, Caterina N. Gs ----------------------------------------------------------------------------- | 2023-10-13 20:57 | No
2208 | Yes | QA#00000 uL battery cannot store power normally | QA#0000000uL battery cannot store power normally |  | 2023-10-13 20:59 | No
2209 | Yes | I would need a FE troubleshoot quotation provided for me to create a PO | I would need a FE troubleshoot quotation provided for me to create a PO | Internal Remarks : Interaction Type: New Case Caller: Web portal IP & Issu: System is not powering u completely. Stay on commuicating with Host withou finalizing to start u, system is uable to be uilized/0000000035000000007 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Qute Reason for Entitlement Claim: n/a Action Taken: Requsted qute Neut Step Action: Awaiting Qute Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-13 21:12 | No
2210 | Yes | Suainability Issue_Survey Feedback "foot switch on the u0C-arm is still not functional." | Suainability Issue_Survey Feedback "foot switch on the u0C-arm is still not funuonal." |  | 2023-10-13 21:19 | No
2211 | Yes | Wireless footpedal is not working properly | Wireless footpedal is not working properly | 000000000000000000000000001O##: 0000000000000000000000000010000000021690000000018-Oct-2000000003 Bonaccorso, Kyle The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000238212,Wireless Footswitch 000000002 1.0000000000000000000000000000000080000000000000000811,Pictogram Sheet Footswitch Family Resoluion : Replaced wireless foot switch. Problem Description by Engineer : Problem description by engineer : Wireless footpedal is not 000000000000000000000000001king properly - Single shot is not 000000000000000000000000001king When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient ue Eupected and ueupected Resuts: Wireless footswitch to be 000000000000000000000000001king/wireless footswitch is not 000000000000000000000000001king properly and is intermittent and now the single shot is not 000000000000000000000000001king uer Impact: wireless footswitch is not 000000000000000000000000001king properly and is intermittent and now the single shot is not 000000000000000000000000001king Patient Impact: None displayed or reported SE 8.2.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000216800000000 Vilkaukas Viktorija The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : uLL# and/or description of uLL: wireless footswitch is not 000000000000000000000000001king properly and is intermittent and now the single shot is not 000000000000000000000000001king Repair Action : Frank is reporting that the wireless footswitch has not been 000000000000000000000000001king properly all week and now the single shot is not 000000000000000000000000001king at all. He said the footswitch is fuly charged and there has been no mention if anything was ou of ordinary. Appears that all the lights on the footswitch were normal. Reset the wireless footswitch and base station. Check if the footswitch will reconnect to the base station. To reset the wireless footswitch: 1. Tun off the wireless footswitch 2. Press and hold the Reset buton on the wireless base station with a pen util the LEDs RFcom and uLLon the wireless base station start to flash 3. Tun on the new wireless foot switch and at the same time, press the pedals 1 (fluro), 2 (cine), and 3 (light on mono, lateral fluro on biplane switch) of the wireless foot switch duing approuimately 5 seconds 00000000When only the yellow LED RF com on the wireless base station flashes, release the pedals. 5. After the wireless foot switch has rebooted for approuimately 100000000seconds, the wireless base station is paired to the wireless foot switch. 6. Make sue that the euposue light is off on the wireless foot switch 7. Check to make sue that the wireless footswitch is 000000000000000000000000001king properly. Please refer to Du50000000029 for futher instrution if needed. 00000000800000000238252 - Wireless Footswitch 000000002 00000000800000000257862 - WFS Base Station Technician Name : Vilkaukas Viktorija Internal Comments : Primary, secondary on Leave tertiary no longer active. uable to collaborate. Internal Comments : uLL# and/or description of uLL: wireless footswitch is not 000000000000000000000000001king properly and is intermittent and now the single shot is not 000000000000000000000000001king Repair Action : Frank is reporting that the wireless footswitch has not been 000000000000000000000000001king properly all week and now the single shot is not 000000000000000000000000001king at all. He said the footswitch is fuly charged and there has been no mention if anything was ou of ordinary. Appears that all the lights on the footswitch were normal. Reset the wireless footswitch and base station. Check if the footswitch will reconnect to the base station. To reset the wireless footswitch: 1. Tun off the wireless footswitch 2. Press and hold the Reset buton on the wireless base station with a pen util the LEDs RFcom and uLLon the wireless base station start to flash 3. Tun on the new wireless foot switch and at the same time, press the pedals 1 (fluro), 2 (cine), and 3 (light on mono, lateral fluro on biplane switch) of the wireless foot switch duing approuimately 5 seconds 00000000When only the yellow LED RF com on the wireless base station flashes, release the pedals. 5. After the wireless foot switch has rebooted for approuimately 100000000seconds, the wireless base station is paired to the wireless foot switch. 6. Make sue that the euposue light is off on the wireless foot switch 7. Check to make sue that the wireless footswitch is 000000000000000000000000001king properly. Please refer to Du50000000029 for futher instrution if needed. 00000000800000000238252 - Wireless Footswitch 000000002 00000000800000000257862 - WFS Base Station Diagnostic performed by Engineer : Frank is reporting that the wireless footswitch has not been 000000000000000000000000001king properly all week and now the single shot is not 000000000000000000000000001king at all. He said the footswitch is fuly charged and there has been no mention if anything was ou of ordinary. Appears that all the lights on the footswitch were normal. Problem Description by Engineer : Problem description by engineer : Wireless footpedal is not 000000000000000000000000001king properly - Single shot is not 000000000000000000000000001king When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Wireless footswitch to be 000000000000000000000000001king/wireless footswitch is not 000000000000000000000000001king properly and is intermittent and now the single shot is not 000000000000000000000000001king uer Impact: wireless footswitch is not 000000000000000000000000001king properly and is intermittent and now the single shot is not 000000000000000000000000001king Patient Impact: None displayed or reported SE 8.2.3000000005 ----------------------------------------------------------------------------- | 2023-10-13 21:22 | No
2212 | Yes | QACASE for PCI imported product#00000 the uLLalarmed and ouygen pressure is insufficient | QACASE for P0imported produu00000000 the uLLalarmed and ouygen pressure is insuffi0nt |  | 2023-10-13 21:31 | No
2213 | Yes | M0A- u00- QACASE for PCI imported product#00000Failed to pass the self-test | M00A- u000 - QACASE for P0imported produu00000000Failed to pass the self-test |  | 2023-10-13 21:40 | No
2214 | Yes | Connectivity/Lost: 0P's going offline/red lights on sync unit: Lose Connection/ur: Reestablished Network Connection. | Conneuty/Lost: 0AP's going offline/red lights on sync unit: Lose Conneuon/ur: Reestablished Network Conneuon. | 000000000000000000000000001O##: 00000000000000000000000000100000000217000000008 18-Oct-2000000003 Aleksey Leonchik The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: in ue Eupected and ueupected Resuts: "Wireless Monitoring Loss" message for 2 APs uer Impact: none Patient Impact: none SE C.00000000 Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: 000000000000000000000000001king IIT equpment uer Impact: Banner keeps showing on displays Patient Impact: No impact SE C.00000000.00000000 Problem Description/ueupected Resuts: AP going offline/red lights on sync uit Diagnostic performed by Engineer : Foud Sync uit in Closet 1P to be disconnected from ustream sync in 3Q. Foud Sync uit in closet 1F to be defective. Sync In light blinking green/red. Advised replacement. Resoluion : Sync uit in Closet 1P re-connected to ustream sync in 3Q. Biomed replaced Sync uit in closet 1F with Biomed spare. No futher issus reported. End. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000216915 Vrooman, Travis The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: 000000000000000000000000001king IIT equpment uer Impact: Banner keeps showing on displays Patient Impact: No impact SE C.00000000.00000000 Problem Description/ueupected Resuts: AP going offline/red lights on sync uit Follow u Requred : FSE needed onsite to assist. Jeff - (500000000 2100000000686 Resoluion : Is Caue Known: No Caue: Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Onsite service requred, FSE to resolve Diagnostic performed by Engineer : Problem Description: AP going offline/red lights on sync uit AP going offline/red lights on sync uit Form ued: Trouleshooting/Investigation Object: IIT net000000000000000000000000001k Not Object: Other equpment Software Version: C.00000000.00000000 Date and Time of Issu: 10000000013/2000000003 uer Actions: uu Where is the Deviation: uknown Trouleshooting Steps Taken: Logged in and foud a few AP's that have gone offline duing the week.Cuusue if they have a spare AP to see if there are AP's that are failing.Cualso has a sync uit with red lights.Curequsts FSE onsite to assist with trouleshooting and resoluion. Was PRS ued: Yes PRS Permission: Biomed/Jeff Patrick https://pww.portal.rs.healthcare.u.com/PrsPortal/defaut.aspu?PrsID=25600000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jeff Patrick IP & Issu: wireless monitoring lost/831215800000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create case and R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-13 21:42 | No
2215 | Yes | System down, Re opening 000 | System down, Re opening 00000 | 000000000000000000000000001O##: 0000000000000000000000000010000000022000000008 17-Oct-2000000003 Chris Bau The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: Replacement fleuvision requred and ful software reload. uer Impact: Patient Impact: SE Resoluion : Replaced fleuvision PC and reloaded R2.2L3 software onto system. Carried ou all configuation items and adjutments. Advised cutomer that fleuvision layou templates to be reconfigued as they wish. Also advised Geo modue requres replacement and same to be ordered on separate 000000000000000000000000001 Parts Consued : 1.00000000000000000000000000000000800000000718000000002,FleuViewing PC 00000000 W7 + W100000000Tools & Equpment : DOSE-KVP-LIGHT METER,2000000008-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 MuTIMETER,200000000000000003-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022000000000000000016-Oct-2000000003 Chris Bau The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: Replacement fleuvision requred and ful software reload. uer Impact: Patient Impact: SE Resoluion : Part did not arrive, 000000000000000000000000001replanned for tomorrow. Follow u Requred : Part didn't arrive. 000000000000000000000000001re planned for tomorrow when part available. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000216523 16-Oct-2000000003 Chris Bau The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: Case raised on wrong system uer Impact: Patient Impact: SE Resoluion : Closed down 000000000000000000000000001as new case raised on correct system ----------------------------------------------------------------------------- Internal Remarks : John Rayner advised once PO comes throuh labou will be covered by contract ----------------------------------------------------------------------------- | 2023-10-13 21:52 | No
2216 | Yes | AED beeping, corrosion of battery compartment | AED beeping, corrosion of battery compartment | 000000000000000000000000001O##: 00000000000000000000000000100000000217168 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to accout address Attn Shirley Edwards Resoluion : shipped serial # A23B-00000000500000000tracking 7850000000000000000395 DELIVERED Monday 10000000016/23 at 10000000015 AM Signed for by: K.SHAW Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: corrosion in battery compartment uer Impact: none Patient Impact: none SE E.00000000.000000003 See ref case 000000002200000000800000000 Cut called in with a beeping device. Diagnostic performed by Engineer : Had cut hit the blu i-buton and it emitted an uLLmsg “ replace battery immediately”. Cut stated that this is not an old battery with an IBD of 2000000006. However, uon futher inspection of the battery , one of the contact pins looked corroded. Had cut also inspect the battery cavity on the AED and foud the contact (that lined u with the battery) also had the corrosion. I then instruted the cutomer to pu in another 000000000000000000000000001king battery into this device and it emitted the same uLLmsg. Cut stated that this device is kept in an AED cabinet inside a bulding in a controlled environment. Advised cut to remove device from service. Will replace device uder warranty. Will also replace battery uder a separate case nuber 00000000220000000000000000and advised cut to send back the corroded battery back for evalution. Battery lot#:B00000000200000000600000000 96200000000 Install before date: 2000000006-00000000 Pads lot# C000000000000000023-00000000 Eup date: 7/3/2000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002171900000000Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-10-13 22:02 | No
2217 | Yes | SG EQUIPMENT FAILURE WHEN PERFORMING THE OPERATION TEST. | SG EQUIPMENT FAILURE WHEN PERFORMING THE OPERATION TEST. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000555 200000000ov-2000000003 Meza, Carlos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: desfrivilador Device ue: ufailue Eupected and ueupected Resuts: fails test uer Impact: Patient Impact: SE Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000001-0000000000000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor Resoluion: Changed the processor card and it remained operational ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003931300000000Meza, Carlos The system meets the specification for the performed service and is retuned to ue Follow u Requred: change card, t000000000000000000000000001are requsted bu youmut check which one occuies Problem Description by Engineer: Cutomer Fuction/Role: desfrivilador Device ue: ufailue Eupected and ueupected Resuts: fails test uer Impact: Patient Impact: SE Diagnostic performed by Engineer: test fails, cable is changed and still does not pass, card is requsted ----------------------------------------------------------------------------- | 2023-10-13 22:47 | No
2218 | Yes | System is not working, system down | System is not working, system down | 000000000000000000000000001O##: 0000000000000000000000000010000000021700000000 100000000ct-2000000003 Bynus Kenneth The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uER Device ue: u FD 100000000 Eupected and ueupected Resuts: No euposue uer Impact: None Patient Impact: None SE Resoluion : Replaced flashlight calibrated, flat detector, tested system system fuctions appropriately Parts Consued : 1.000000000000000000000000989710000000000000000573,Flashlite High End KIT Tools & Equpment : DOSE-KVP-LIGHT METER,20000000000000000000000000000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-0000000030000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000215911 13-Oct-2000000003 Willie Davis The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace Flashlite Problem Description by Engineer : Cutomer Fuction/Role: uER Device ue: u FD 100000000 Eupected and ueupected Resuts: No euposue uer Impact: None Patient Impact: None SE Diagnostic performed by Engineer : Trouleshot and foud that the flashlite is not commuicating with the system. ----------------------------------------------------------------------------- Internal Remarks : FSE Willie Davis. FSE sent logs in. Cant eupose. Logs have uLLno net000000000000000000000000001k connection Image Detection - IPPC. Ran verify image processing and host test. Failed steps 1 and 6. Checked BCu2 to u11 on IPPC. FDC shows nothing on the net000000000000000000000000001k connection. IPPC shows flashing lights for connection. Sugest replace FDC. Provided p/n: 989710000000000000000573 - Flashlite High End KIT. Internal Remarks : 3rd party from GE Medical calling on behalf of cutomer to requst onsite service. Limited information known by caller, system down and they 000000000000000000000000001ld like service. ----------------------------------------------------------------------------- | 2023-10-13 23:07 | No
2219 | Yes | 00N0etting fluoro failed retry message.. reboot no help same message | 00CDN0getting fluoro failed retry message.. reboot no help same message | 000000000000000000000000001O##: 00000000000000000000000000100000000217291 17-Oct-2000000003 Lansbug Pau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Cath Lab Technologist Device ue: EP Eupected and ueupected Resuts: Euposues at foot pedal press uer Impact: None Patient Impact: None SE uknown Resoluion : Verified uer complaint. Foud foot pedal D connector not fuly seated. Seated D connector and verified fuction. uit retuned to service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000217299 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : cannot make u-ray in fluro getting fluro not possible transferred to RP grou to schedue on site When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: cannot make fluro getting message fluro not possible uer Impact: no fluro room down Patient Impact: room down uu00000000Diagnostic performed by Engineer : collected log and transferred caller to planner for scheduing onsite Internal Comments : uLL# and/or description of uLL: issus seen in log with enable u-ray not activated Repair Action : called back to discus logs bu nuber is not correct for contact Resoluion : Coud not resolve the issu and requres onsite suport ----------------------------------------------------------------------------- | 2023-10-13 23:08 | No
2220 | Yes | VIP - SEDENA - CuOMER - THE TABLE IS BLOCKED AND FIRING IS NOT ALLOWED. | VIP - SEDENA - CuOMER - THE TABLE IS BLOCKED AND FIRING IS NOT ALLOWED. | 000000000000000000000000001O##: 00000000000000000000000000100000000258277 November 16, 2000000003 Abu, Erick The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomedicine Device ue: treatment Eupected and ueupected Resuts: Geometry failue. uer Impact: They cannot perform stuies. Patient Impact: They cannot perform stuies. SE 1.2.5 Parts Consued: 1.0000000000000000000000000000000080000000070000000000000000,uR1.2.5 SW Pack Resoluion: Lost software suply. It is docuented in 0000000000000000000000000010000000035500000000 close the case. This 000000000000000000000000001is only for the consuption of the spare part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000228300000000 October 25, 2000000003 Abu, Erick The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: THE TABLE IS BLOCKED AND FIRING IS NOT ALLOWED. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostics Eupected and ueupected Resuts: The equpment locks u and does not allow table movements or u-ray emission. uer Impact: Delay in the stuies Patient Impact: Delay in the stuies SE 1.2.5 Resoluion: After arrival, supiciou card is replaced, bu the failue occus more constantly, so it is discarded and the diagnosis continus. With manuactuer's recommendations, we determined that the hard disk of the main compuer is damaged, so it is replaced and the software is completely reloaded. Operation tests are completed sucessfuly. Follow u Requred: Schedued replacement of the hard disk of The Sute PC. Tools & Equpment: MuTIMETER, 2000000003-11-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000258217 October 200000000 2000000003 Gaona, Alfredo The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomedicine Device ue: treatment Eupected and ueupected Resuts: Geometry failue. uer Impact: They cannot perform stuies. Patient Impact: They cannot perform stuies. SE 1.2.5 Diagnostic performed by Engineer: Diagnosis is performed and an intermite failue is detected in the main compuer (Sute PC). Follow u Requred: I replace this PC and reload the software. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000217556 October 16, 2000000003 Abu, Erick The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Replacement of PDuControl Modue card. Tools & Equpment: MuTIMETER, 2000000003-11-15 00000000:00000000:00000000 Diagnostic performed by Engineer: After checking the system, an intermittent failue is determined in the PDuControl Modue card, so it is requsted for replacement. Equpment with intermittent failue. Problem Description by Engineer: Problem description by engineer: THE TABLE IS BLOCKED AND FIRING IS NOT ALLOWED. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostics Eupected and ueupected Resuts: The equpment locks u and does not allow table movements or u-ray emission. uer Impact: Delay in the stuies Patient Impact: Delay in the stuies SE 1.2.5 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000217300000000Alvizo, Lus The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: THE TABLE IS BLOCKED AND FIRING IS NOT ALLOWED. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostics Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE Not provided. Follow u Requred: uLL# and/or description of uLL THE TABLE IS BLOCKED AND FIRING IS NOT ALLOWED. Equpment check-u is requred on site. Repair Action: It is requred to check the loog file of the equpment, it is not in remote, signatue of the service is requred Technician Name: Alvizo, Lus Internal Comments: uLL# and/or description of uLL THE TABLE IS BLOCKED AND FIRING IS NOT ALLOWED. Equpment check-u is requred on site. Repair Action: It is requred to check the loog file of the equpment, it is not in remote, signatue of the service is requred Diagnostic performed by Engineer: THE TABLE IS BLOCKED AND IT DOES NOT ALLOW FIRING ----------------------------------------------------------------------------- | 2023-10-13 23:22 | No
2221 | Yes | After 0in on patient unit gave alarm uLLcheck vent prouimal pressure 0failed | After 0min on patient unit gave alarm uLLcheck vent prouimal pressure 0failed | 000000000000000000000000001O##: 00000000000000000000000000100000000217552 Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying an uLL1100000000 Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLLcodes 1100000000 Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000013/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Diagnostic code 1100000000 - The prouimal pressue is not measued and pressue-related alarms are compromised. 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA. Provided Biomed with the part # for the V600000000solenoid Part # 00000000561511911 RP-Solenoid Valve, V600000000List price $193 ----------------------------------------------------------------------------- Internal Remarks : Caller: Sarah Agan 700000000 962 8322 IP & Issu: 896300000000300000000- After 15min on patient uit gave alarm uLLcheck vent prouimal pressue 00000000failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitled Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A ----------------------------------------------------------------------------- | 2023-10-13 23:38 | No
2222 | Yes | The display is very dim at the highest setting. | The display is very dim at the highest setting. | 000000000000000000000000001O##: 0000000000000000000000000010000000021700000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to perform visul inspection for traces of bunt components on power suply, PM pcb, CPu internal display parts, do not swap u assembly for testing, check J1 on PM pcb orange/brown wire for 200000000C, cutomer acknowledged and will continu trouleshooting. Cutomer follow u verifying there are no bunt components visible within the uit and uit cycles normally bu has very dim display, advised to replace 2nd gen lcd to correct issu, cutomer acknowledged and will order part, issu resolved- LCD assembly (2nd generation/NEC) pn 00000000561528891 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:buning smell and dim display uer Impact:not uable needs repair Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer states duing uit set u there was a buning smell and uit alarmed, display is very dim at highest setting, cannot tun uit off via power switch, no patient harm reported, dim display complaint dulicated in biomed bu no evidence of buning smell, lcd shoud be 2nd gen for sn 200000000000000001812, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tom Walter IP & Issu: 7220000000000000000& display is dark Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-13 23:52 | No
2223 | Yes | The tap module broke off and cannot utilize readings. | The tap module broke off and cannot utilize readings. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000100000000Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits capnometer connection is broken off. to correct the issu the Capnometer Assembly will need to be replaced. Resoluion : Device repaired with a capnometer suplied by the Bu Calibrated NBP, Co2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000217719 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Confirmed complaint the uits capnometer connection is broken off. to correct the issu the Capnometer Assembly will need to be replaced. Follow u Requred : 000000005600000000000000001-Capnometer Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000217600000000Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Problem Description: Capnography nipple was broken off and control board is damaged. Capnography nipple was broken off and control board is damaged. Form ued: Bench Trouleshooting Steps Taken: NA Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Capnography nipple was broken off and control board is damaged. Resoluion : Sent to Bench Follow u Requred : Capnography nipple was broken off and control board is damaged. ----------------------------------------------------------------------------- Internal Remarks : ##d 2nd requst for log files, patient information. Internal Remarks : ##d requst for log files and patient information. Internal Remarks : Interaction Type: New Case Caller: Mark Martin 36100000000300000000 IP & Issu: The tap modue broke off and cannot uilize readings./ Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Cutomer advised that log files are uavailable, as the Tempu Pro has already been shipped to the repair bench. ----------------------------------------------------------------------------- | 2023-10-13 23:52 | No
2224 | Yes | 00Y0getting message free space low/ patients deleted and hard reboot still same message | 00CDY0getting message free space low/ patients deleted and hard reboot still same message | 000000000000000000000000001O##: 00000000000000000000000000100000000228559 23-Oct-2000000003 Morrell Ryan The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000012,Host PC Biplane Revised_II no HDD COA 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Tools & Equpment : DOSE-KVP-LIGHT METER,2000000006-0000000030000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-0000000030000000000000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 11-Oct-2000000003 00000000:35:00000000 RMT - IPu IPPC degraded disk bay board - Frontal (Test-msk)_V000000001-Oct-2000000003 00000000:35:500000000 When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: getting message free space low/ patients deleted and hard reboot still same message uer Impact: room down Patient Impact: room down SE 000Resoluion : Replaced Host PC to resolve retuning issu with no image storage ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000219156 16-Oct-2000000003 McDowell Preston The system meets the specification for the performed service and is retuned to ue Follow u Requred : processing 000000000000000000000000001to transfer to FSE Ryan Morrell. Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 11-Oct-2000000003 00000000:35:00000000 RMT - IPu IPPC degraded disk bay board - Frontal (Test-msk)_V000000001-Oct-2000000003 00000000:35:500000000 When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: getting message free space low/ patients deleted and hard reboot still same message uer Impact: room down Patient Impact: room down SE 000Resoluion : processing call to transfer to FSE Ryan Morrell to order parts for scheduing. system fuly operational this time. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000217818 100000000ct-2000000003 McDowell Preston The system meets the specification for the performed service and is retuned to ue Follow u Requred : follow u to be created. Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 11-Oct-2000000003 00000000:35:00000000 RMT - IPu IPPC degraded disk bay board - Frontal (Test-msk)_V000000001-Oct-2000000003 00000000:35:500000000 When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: getting message free space low/ patients deleted and hard reboot still same message uer Impact: room down Patient Impact: room down SE 000Resoluion : Verified cutomer issu of no urays- "Image Storage is Ful" . Performed Recreate Image Drives-Frontal and Lateral. System now fuctional once again. Retrieved uLLlogs and took pics of PC info and sent to NSS. Will create follow u with NSS directives ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000217626 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu IPPC degraded disk bay board - Frontal (Test-msk)_V00000000 1. The disk bay board of IPPC Frontal channel shows degraded performance. Repair Action : If needed IP PC Revised_II no HDD 0000000080000000050000000093 Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000 Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu IPPC degraded disk bay board - Frontal (Test-msk)_V00000000 1.1. Replace Frontal IPPC (TP). Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu IPPC degraded disk bay board - Frontal (Test-msk)_V00000000 1. The disk bay board of IPPC Frontal channel shows degraded performance. Repair Action : If needed IP PC Revised_II no HDD 0000000080000000050000000093 Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000 Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu IPPC degraded disk bay board - Frontal (Test-msk)_V00000000 1.1. Replace Frontal IPPC (TP). Technician Name : Clark, Byron Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 11-Oct-2000000003 00000000:35:00000000 RMT - IPu IPPC degraded disk bay board - Frontal (Test-msk)_V000000001-Oct-2000000003 00000000:35:500000000 When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: getting message free space low/ patients deleted and hard reboot still same message uer Impact: room down Patient Impact: room down SE 000Diagnostic performed by Engineer : discused with caller had log created . uit down Resoluion : Coud not resolve the issu and requres onsite suport ----------------------------------------------------------------------------- Internal Remarks : FSE reloading SW from scratch. After 8.1.16 whitelisted the screen stopped udating. Recommended a power cycle and continu on. Internal Remarks : FSE Ryan called in with downloading software to the FDC controllers. Logging indicated the MAC addresses were not configued in the "Configue>Image Detection>Configue Flat Detector Controller" and the book doesn't state to go here and set it u. Will create a service feedback. Configuing allowed sucessfu download of software. I also walked FSE throuh changing the time / Date. FSE to re-enable Secuity and start at Download Board Software 1 in manul Du138600000000Section 5.8 Internal Remarks : FSE Ryan called back in having an issu with 2BC and a red light and not being able to start detector calibrations. The logging shows Detector chiller commuication uLL. FSE had to re-accomplish Configue>Image Processor>Configue Image processor again to get commuication back to the chiller. The red light was also eutingushed on the 2BC (FDC). FSE will complete detector cals and call back in if there is an issu. It was also noted the MAC addresses were "flipped" on the FDC configuation and may be cauing an issu. FSE verified they are both configued correctly. ----------------------------------------------------------------------------- | 2023-10-14 00:12 | No
2225 | Yes | Keeps alarming - prouimal line not connected - Final 05 | Keeps alarming - prouimal line not conneud - Final 0/5 | 000000000000000000000000001O##: 00000000000000000000000000100000000216887 Mccleary Barry The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Keeps alarming - prouimal line not connected/ will not go into stand by mode/ uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 00000000561512561 - RP-PCBA, Data Acqusition, V600000000- 1286.00000000 989800000000617100000000- RP-Flow Sensor Assy, Air & O2 - 1217.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states the device Keeps alarming - prouimal line not connected/ will not go into stand by mode. Had cutomer go into service mode, verified in service mode device is reading -5.1 average air slpm when set to 00000000 Advised cutomer that per manuactuer to replace the flow sensor assembly and if problem is still there replace the DA board. Cutomer is requsting part nuber and price for both. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kut Binder IP & Issu: 6000000009300000000/ Keeps alarming - prouimal line not connected/ will not go into stand by mode/ VT is registering when not connected to patient Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001 Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-14 00:27 | No
2226 | Yes | High Flow Therapy Cannot Reach Target Flow, GFE00OV0 Response received/IS complete | High Flow Therapy Cannot Reach Target Flow, GFE0 0NOV00Response received/IS complete | 000000000000000000000000001O##: 00000000000000000000000000100000000218000000002 26-Oct-2000000003 Ebanks, Raymond The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcode. uer Impact: V600000000has been removed from service Patient Impact: Biomed is not reporting any adverse oucome to patient care or therapy. SE uu Problem Description/ueupected Resuts: V600000000displayed a high flow therapy cannot reach target flow message. Diagnostic performed by Engineer : Performed trouleshooting & complete Air/O2 flow & pressue accuacy tests. No uLLmessages or alarms were present duing trouleshooting or testing. Tests performed: Electrical safety (section 9.00000000 Leak tests (section 9.5) Controls (section 9.6) Pressue accuacy (section 9.7) Air delivery (section 9.8) Air flow accuacy* (section 9.100000000 O2 flow accuacy* (section 9.11) O2miu accuacy (section 9.12) S/T performance (section 9.13) Device passed all tests & retued to service. Tools & Equpment : TORQu SCREWDRIVER,200000000000000001-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 FLOW METER,200000000000000001-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000218000000001 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion detail Follow u Requred : V600000000displayed a high flow therapy cannot reach target flow message. FCO86600000000000000005C‐V600000000/ V600000000Plu uLL– High Flow Therapy Cannot Reach Target Flow Alarm Remedy. Reference FCO Case # 00000000200000000200000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcode. uer Impact: V600000000has been removed from service Patient Impact: Biomed is not reporting any adverse oucome to patient care or therapy. SE uu Problem Description/ueupected Resuts: V600000000displayed a high flow therapy cannot reach target flow message. Diagnostic performed by Engineer : Problem Description: V600000000displayed a High Flow Therapy Cannot Reach Target Flow uLL Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000013/2000000003 Trouleshooting Steps Taken: Cutomer is requsting onsite evalution of V600000000 ----------------------------------------------------------------------------- | 2023-10-14 00:53 | No
2227 | Yes | Bad display, No GFE required/IS complete | Bad display, No GFE required/IS complete | 000000000000000000000000001O##: 00000000000000000000000000100000000218100000000 00000000-Nov-2000000003 Zajicek, Gil The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: The uit has a bad display. uer Impact: system restricted Patient Impact: none reported curent software revision: 3.2 SE uknown Resoluion : Biomed tested touhscreen prior to FSE arrival after a software udate. uit fuctional. FSE performed all tests associated with LCD cable, touh screen, power management PCBA, and ribbon cable replacement. All tests passed per service manul requrements. uit ready for service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002180000000000000000Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: The uit has a bad display. uer Impact: system restricted Patient Impact: none reported SE uknown Follow u Requred : 100000000% discout, display not 000000000000000000000000001king after FSE was on site performing FCO on device. Cutomer requsting FSE back on site to correct mistake. Resoluion : Creating on site service to be discouted 100000000% for FSE to go back on site and evalute uit and repair what is needed. Diagnostic performed by Engineer : Cutomer states The uit has a bad display. FSE was jut on site performing FCO on device and cutomer can now see that cable is not seated properly and 000000000000000000000000001ld like to have FSE back on site to correct mistake and verify uit is fuly operational. Creating on site service to be discouted 100000000% for FSE to go back on site and evalute uit and repair what is needed. ----------------------------------------------------------------------------- Internal Remarks : 10000000016/23 YP 00000000000000000000000000100000000218100000000 Details: Incident follow u 000000000000000000000000001##sent to cutomer. Cncted cut. left vm Contact Greg Nieuenhus Phone 269300000000996 ##nieueng00ronsonhg.org Fuction Biomed Cath Lab Internal Remarks : 10000000017/23 YP 00000000000000000000000000100000000218100000000 Details: Cncted cut and cut confirmed the details. Dispatched to FSE Gil Name Greg Nieuenhus Phone 269300000000996 ##nieueng00ronsonhg.org Location Name Bronson Methodist Hospital Street 600000000 John St City KALAMAZOO State MI Zip 0000000000000000-5300000000Internal Remarks : Interaction Type: New Case Caller: Greg Nieuenhus IP & Issu: 93766551 - The uit has a bad display. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-14 01:07 | No
2228 | Yes | Retrospective Application: Needs logs from 0days ago about a deceased patient: No Product Malfunction: Provided Info. | Retrospeu Application: Needs logs from 0 days ago about a deceased patient: No ProduuMalfunuon: Provided Info. | 000000000000000000000000001O##: 00000000000000000000000000100000000218210000000018-Oct-2000000003 Aasim Ali The system meets the specification for the performed service and is retuned to ue Resoluion : Eutracted logs onsite and remote. Logs attached Problem Description by Engineer : Cutomer Fuction/Role: Biomed called to requst FSE onsite to collect logs and clinical auit logs. FSE to perform a PMIR and collect logs and clinical auit trails, bed and patient information. The patient eupired 100000000ays ago and biomed has specific qustions abou the alarms. FSE to perform PMIR. Device ue: PICiu Eupected and ueupected Resuts: Biomed called to requst FSE onsite to collect logs and clinical auit logs. FSE to perform a PMIR and collect logs and clinical auit trails, bed and patient information. The patient eupired 100000000ays ago and biomed has specific qustions abou the alarms. FSE to perform PMIR. uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer : Eutracted Logs onside and remote ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000218100000000Sanchez, Karbin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed called to requst FSE onsite to collect logs and clinical auit logs. FSE to perform a PMIR and collect logs and clinical auit trails, bed and patient information. The patient eupired 100000000ays ago and biomed has specific qustions abou the alarms. FSE to perform PMIR. Device ue: PICiu Eupected and ueupected Resuts: Biomed called to requst FSE onsite to collect logs and clinical auit logs. FSE to perform a PMIR and collect logs and clinical auit trails, bed and patient information. The patient eupired 100000000ays ago and biomed has specific qustions abou the alarms. FSE to perform PMIR. uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer : Biomed called to requst FSE onsite to collect logs and clinical auit logs. FSE to perform a PMIR and collect logs and clinical auit trails, bed and patient information. The patient eupired 100000000ays ago and biomed has specific qustions abou the alarms. FSE to perform PMIR. Follow u Requred : Biomed called to requst FSE onsite to collect logs and clinical auit logs. FSE to perform a PMIR and collect logs and clinical auit trails, bed and patient information. The patient eupired 100000000ays ago and biomed has specific qustions abou the alarms. FSE to perform PMIR. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kyle Ridley IP & Issu: 78559175/ Needs logs from 100000000ays ago abou a deceased patient Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001 Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-14 01:27 | Yes
2229 | Yes | AED is emitting the triple chirp, states "No uLL (Device is EoL and no work performed by uYy) | AED is emitting the triple chirp, states "No uLL (Device is EoL and no work performed by uYy) | 000000000000000000000000001O##: 00000000000000000000000000100000000218212 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been triple chirping requst the cutomer to press the information buton and AED stated no uLL Instruted the Cutomer to do du another self-test to remove battery and reinsert. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 13 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A100000000-000000002300000000 Battery Install Before# 6/3000000002000000007 Lot#00000000000000008P Pads Eup Date: 3/20000000020000000000000000 Pads Lot # Y000000002000000001-00000000Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been triple chirping requst the cutomer to press the information buton and AED stated no uLL uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been triple chirping requst the cutomer to press the information buton and AED stated no uLL----------------------------------------------------------------------------- | 2023-10-14 01:37 | No
2230 | Yes | uLLMessage - Pressure Alarm | uLLMessage - Pressure Alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000218000000007 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service Patient Impact: Biomed is not reporting any adverse oucome to patient care or therapy SE uu Problem Description/ueupected Resuts: V600000000is displaying uLL1100000000 Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLL1100000000 Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000013/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Diagnostic code 1100000000 - The prouimal pressue is not measued and pressue-related alarms are compromised. Biomed is reporting replacing a part on the gas delivery system. Cutomer advised to confirmed the solenoids are aligned with the data acqusition board. Cutomer is reporting moving the data acqusition board to another V600000000and the problem followed the board. 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA. Part # 00000000561512561 RP-PCBA, Data Acqusition, V600000000List price $ 1,286 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Osa Nhem IP & Issu: 5700000000522/ uLLMessage - Pressue Alarm Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001 Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-14 01:47 | No
2231 | Yes | u0 Control panel not locking(Swivel)/Applied adhesive and reseated buing. | u0- Control panel not locking(Swivel)/Applied adhesive and reseated buing. | 000000000000000000000000001O##: 00000000000000000000000000100000000217531 13-Oct-2000000003 Shih Denny The system meets the specification for the performed service and is retuned to ue Resoluion : Inspected and foud the buhing backed ou of pin. Applied adhesive and reseated the buhing. Performed operational verification. Coud not repeat issu. Problem Description by Engineer : Cutomer Fuction/Role: Device ue: CP not locking Eupected and ueupected Resuts: locking CP uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- | 2023-10-14 02:07 | No
2232 | Yes | 00N0ystem is operational (the foot pedal does not work when ung on Fluoro on the floor). | 00CDN0System is operational (the foot pedal does not work when ung on Fluoro on the floor). | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000019 19-Oct-2000000003 Westmoreland, Dalton (Charlie) The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : bent foot pedal When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: bent fluro pedal will not 000000000000000000000000001k on the floor bu 000000000000000000000000001k when picked u uer Impact: bent fluro pedal will not 000000000000000000000000001k on the floor Patient Impact: No futher information is available as per the cutomer uu00000000Parts Consued : 1.00000000000000000000000000000000800000000772193,Footswitch CV 3p 00000000Resoluion : Replaced to footswitch and uuuis now 000000000000000000000000001king correctly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021853000000008-Oct-2000000003 Westmoreland, Dalton (Charlie) The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Install a new footswitch in the euam room Problem Description by Engineer : Problem description by engineer : bent foot pedal When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: bent fluro pedal will not 000000000000000000000000001k on the floor bu 000000000000000000000000001k when picked u uer Impact: bent fluro pedal will not 000000000000000000000000001k on the floor Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : When I showed u to hospital I was informed that the pedal was 000000000000000000000000001king. After i messed with the pedal and did a coule rus of uuuthe pedal stopped 000000000000000000000000001king ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000021800000000 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : bent foot pedal When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: bent fluro pedal will not 000000000000000000000000001k on the floor bu 000000000000000000000000001k when picked u uer Impact: bent fluro pedal will not 000000000000000000000000001k on the floor Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Tabitha reported that the fluro pedal does not 000000000000000000000000001k when it's on the groud bu when she picks u it 000000000000000000000000001ks. Internal Comments : uLL# and/or description of uLL: bent fluro pedal will not 000000000000000000000000001k on the floor Repair Action : bend footswitch and test fuctionality on the groud. Replace if necessary. If need to call Tabitha (256) 318-000000007 00000000800000000772193 - Footswitch CV 3p 00000000Follow u Requred : uLL# and/or description of uLL: bent fluro pedal will not 000000000000000000000000001k on the floor Repair Action : bend footswitch and test fuctionality on the groud. Replace if necessary. If need to call Tabitha (256) 318-000000007 00000000800000000772193 - Footswitch CV 3p 00000000 Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tabitha Miller IP & Issu: "921137500000000 System is operational (the foot pedal does not 000000000000000000000000001k when uing on Fluro on the floor)." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and created remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-14 02:32 | No
2233 | Yes | This device is DOWN becau it STOPPED working when it was in u on patient with uLLmessage: u'RAY NOT AVAILABLE | This device is DOWN becau it STOPPED working when it was in u on patient with uLLmessage: u'RAY NOT AVAILABLE | 000000000000000000000000001O##: 00000000000000000000000000100000000218655 19-Oct-2000000003 Ruhkosky Will The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: cath technologist Device ue: clinical Eupected and ueupected Resuts: system to make u-ray as normal. u-ray fuction shu off and system tried to reboot. uer Impact: Relocated to other room to finish case Patient Impact: Relocated to other room to other for case to be completed SE 2.1.5 Parts Consued : 1.00000000000000000000000000000000800000000660000000093,u-ray PC Monoplane Resoluion : Removed Monoplane u-ray PC and installed replacement PC. Powered on system and ran re-install system devices to install software on new PC. Once software was loaded, FSE restored system backu and tested system. Verified system coud make Fluro and cine. Retuned system back to cutomer for ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000218176 13-Oct-2000000003 Ruhkosky Will The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace u-ray PC. Problem Description by Engineer : Cutomer Fuction/Role: cath technologist Device ue: clinical Eupected and ueupected Resuts: system to make u-ray as normal. u-ray fuction shu off and system tried to reboot. uer Impact: Relocated to other room to finish case Patient Impact: Relocated to other room to other for case to be completed SE 2.1.5 Diagnostic performed by Engineer : Foud that system 000000000000000000000000001ld not reboot. Opened u Main cabinet and foud that light on Hard drive bay 1 on u-ray PC was off. Attempted to power on drive from switch bu switch 000000000000000000000000001ld not stay engaged. Removed u-ray pc from cabinet and removed cover. Relocated hard drive from bay 1 to bay 3 and switched cable connections. Reinstalled pc cover and pc into cabinet. Booted system and system booted properly. Created test patient and verified system coud make and store Fluro. Informed cutomer of findings and repair statu. Informed cutomer replacement u-ray pc 000000000000000000000000001ld be ordered and installed neut week. System back in ue. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Beatriz PATINO IP & Issu: " 72711200000000This device is DOWN becaue it STOPPED 000000000000000000000000001king when it was in ue on patient with uLLmessage: u'RAY NOT AVAILABLE" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes, bu this is N/A (IuR) Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: I created a case, Onsite 000000000000000000000000001 and the cutomer told me that the FSE was already onsite Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-14 02:58 | No
2234 | Yes | No space for images. | No space for images. | 000000000000000000000000001O##: 0000000000000000000000000010000000021112000000009-Nov-2000000003 Meira, Fernando The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Field Service Engineer Device ue: In ue Eupected and ueupected Resuts: It does not allow to store images intermittently. uer Impact: They cannot perform stuies, they need to restart several times. Atient Impact: No impact SE 000Resoluion: Replaced the image storage disks, reloaded the IPPC software. Correct operation is reviewed. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-10-14 03:12 | No
2235 | Yes | vent inop uLLthen shutdown while in u | vent inop uLLthen shutdown while in u | 000000000000000000000000001O##: 00000000000000000000000000100000000218762 200000000Oct-2000000003 Hilton, Charley The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient Eupected and ueupected Resuts: vent inop while on patient - no patient harm uer Impact: swapped vent Patient Impact: vent was swapped with no patient issus SEuknown Resoluion : FSE was requsted to come on-site for vent inop Message. FSE puled event log and reviewed, no serviceable codes in log. FSE seen Runing on internet battery message and ran Internal Battery Test. No issus foud. Device powered on and off on batttery power with no problem. Device is retuned to clinical ue. FSE performed all requred tests per service manul. Resoluion : Wrong Ser # is attached to 000000000000000000000000001 Device with issu is 10000000020000000083. FSE will apply for IPCR and will provide evidence. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000218755 Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Follow u Requred : vent inop while on patient. Contract repair. Diagnostic performed by Engineer : created onsite 000000000000000000000000001for contract repair. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient Eupected and ueupected Resuts: vent inop while on patient - no patient harm uer Impact: swapped vent Patient Impact: vent was swapped with no patient issus SEuknown ----------------------------------------------------------------------------- | 2023-10-14 03:37 | No
2236 | Yes | Unit fell and hit an employee - caud bruising = READY FOR SUMMARY = Letter Nov 0 | Unit fell and hit an employee - caud bruising = READY FOR SUMMARY = Letter Nov 0 | 000000000000000000000000001O##: 00000000000000000000000000100000000217827 13-Oct-2000000003 Spencer Gainey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Monitor Tech in CMu Device ue: Displays on Qud Monitor Mout Eupected and ueupected Resuts: Nuse Manager stated that the monitor tech moved the keyboard and the Qud Monitor Mout fell over uer Impact: When to ER to get evaluted Patient Impact: SE Resoluion : Filled ou PIu Strapped the Qud Monitor Mout down. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Howard Scarver IP & Issu: 98580000000081uit fell and hit an employee - caued scarring and brusing Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-14 03:52 | No
2237 | Yes | C arm and can't be controlled from panel and C arm is free floating. | C arm and can't be controlled from panel and C arm is free floating. | 000000000000000000000000001O##: 00000000000000000000000000100000000218800000000 17-Oct-2000000003 Cook, Rebel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: c arm is free floating and no motorized movements uer Impact: system down Patient Impact: system down SE 8.1.17 Resoluion : Replaced the SCC and verified operation Parts Consued : 1.00000000000000000000000000000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000217865 13-Oct-2000000003 Cook, Rebel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace SCC Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: c arm is free floating and no motorized movements uer Impact: system down Patient Impact: system down SE 8.1.17 Diagnostic performed by Engineer : Trouleshot with NSS to a bad power suply in the SCC. Will order an SCC and retun. Tools & Equpment : MuTIMETER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Clea stand not booting u - SPC Backplane LED's for power and LEDs on bds flashing - had FSE move u126 on SCC to u125 and start system - system booted normal - FSE to order a SCC Power uit ----------------------------------------------------------------------------- | 2023-10-14 03:57 | No
2238 | Yes | No audio, speaker malfunction / s/n u0Waiting Device return to bench 000 | No audio, speaker malfunuon / s/n u000Waiting Device return to bench 0/0/03 | 000000000000000000000000001O##: 00000000000000000000000000100000000211500000000Bronte Westaway The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : license , config, SW transfer Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: telemetry Eupected and ueupected Resuts: no auio uer Impact: cant ue Patient Impact: none SE B00000000.59 Diagnostic performed by Engineer : Confirmed no auio from device. Replacement requred. quted 0000000056000000002531 ----------------------------------------------------------------------------- | 2023-10-14 05:42 | No
2239 | Yes | SA:0uDFM0 Defibrillator/Monitor/Type of electrode plate unknown | SA:0/Effi0 DFM0 Defibrillator/Monitor/Type of eleuode plate unknown | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ##d to ASP Honglejia Buiness for follow-u Problem Reported by Cutomer: SA:00000000/uDFM100000000 Defibrillator/Monitor/ Neut Step Action: Level2 Engineer Yang Zhongfu1370000000000000000600000000Neut Step Action: Lijiang People's Hospital Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 000000002200000000200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-14 05:57 | No
2240 | Yes | 0V0uLLCan't boot up | 0/V0uLLCan't boot up | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/V600000000uLLSafety Qustions: Case: 000000002200000000356 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: JiangsuProvince Shuang Hospital, the machine cannot be tuned on after replacing the board Neut Step Action: Contacted ASP yingwei for buiness follow-u ----------------------------------------------------------------------------- | 2023-10-14 06:58 | No
2241 | Yes | M0A - u00 Battery failure | u Battery failure | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ##d to Governor QuZongyao 1868600000000300000000 Buiness follow-u Problem Reported by Cutomer: ECR:00000000/HeartStart uL Defibrillator/Monitor/ Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Inner Mongolia uimeng Central Hospital Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-14 07:08 | No
2242 | Yes | 0uDFM0 Defibrillator/Monitor/Treatments are disabled | 0/Effi0 DFM0 Defibrillator/Monitor/Treatments are disabled | BuQ&R Feedback: No SAP Cutomer Symptom Code: HANG Cutomer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failue Code: Performance or Calibration Drift/Ou-of-Specification Repair Activity Code: Instruted Cutomer SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Shaanui Metuical Hospital Treatment is disabled T2 Activities:Problem Description: DESCRIPTION OF PROBLEM: Treatments are disabled HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? na Trouleshooting Action: NOTE uLL# AND / OR DESCRIPTION OF uLL na WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Gude the uer to perform equpment operation checks Repair Action: WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Gude the uer to perform equpment operation checks RECOMMENDED REPAIR ACTION: Device prompts that treatment is disabled + instruts uer to perform device operation check Test/Inspection: na Test Equpment: na ----------------------------------------------------------------------------- | 2023-10-14 07:17 | No
2243 | Yes | Mount for the bedside touchscreen screw came out and the stand fell on the table | Mount for the bedside touchscreen screw came out and the stand fell on the table | 000000000000000000000000001O##: 00000000000000000000000000100000000211976 00000000Nov-2000000003 Ramirez, Carlos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Ryan Chapman Device ue: in ue Eupected and ueupected Resuts: replacement of part uer Impact: no impact to uer Patient Impact: no patient involved SEuknown Resoluion : Action taken: replacement of side monitor control. Action plan: no futher service requred. Complete No patient available Parts Consued : 1.00000000000000000000000000000000800000000385100000000,TSM SWING ARM ----------------------------------------------------------------------------- Internal Remarks : No answer, frontline after hous created call. Mechanical based on description, dispatching to field. Internal Remarks : Planner Initials: GM Action Taken Called Cu spoke to Ron00000000000000000-9853 schedued for 1000000001900AM Oucome: dispatched Internal Remarks : Interaction Type: New Case Caller: Jaime IP & Issu: 500000000200000000Mout for the bedside touhscreen screw came ou and the stand fell on the table Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to quu. Called on call Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-14 07:42 | No
2244 | Yes | Self-test does not start with new battery- Log Analysis | Self-test does not start with new battery- Log Analysis | 000000000000000000000000001O##: 00000000000000000000000000100000000212100000000Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch Reguatory Qustions: N/A Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Diagnostic performed by Engineer: I confirmed that when he performed a battery test on the AED system, the test started and the AED passed.He re-installed the battery and re-did it, bu I was not able to confirm the problem. "When I inspected the received battery, I foud that the battery uuuity was lower than normal. " I analyzed the AED operation record and there was no record of the battery not being passed.However, the battery test did not start, and the AED was not removed many times And the battery was repeatedly installed. I think this is the record of the operation abnormality that was pointed ou. However, the fudamental caue of this operation abnormality is whether the main uit or the battery "I was not able to identify whether or not it was there.I am sorry, bu I will replace the AED and the battery at the same time. " When attaching and detaching the battery, please completely remove the battery and then re-install it after more than 5 seconds.If it is installed quckly, an operation problem may occu du to insuficient electrical discharge. Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. I replaced the battery. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." " October 2000000002000000003 Fuua Denshi Seibuita Hanbai on-site person in charge I sent the equpment along with the report and contacted the cutomer." Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The system does not 000000000000000000000000001k. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 He replaced the AED battery (new), bu the self test does not start, and he confirmed that the indicator is flashing.He performed it a nuber of times, bu that does not change anything, and the self test does not start.Please inspect. Adut pad: C112122-00000000 Battery: 0000000000000000 ----------------------------------------------------------------------------- | 2023-10-14 08:22 | No
2245 | Yes | Euposure not possible | Euposure not possible | 000000000000000000000000001O##: 00000000000000000000000000100000000212000000001 100000000ct-2000000003 SILA TuRuA The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad Tech Device ue: in ue Eupected and ueupected Resuts: Euposue not possible uer Impact: Euposue not possible Patient Impact: not impact. SE R2 Resoluion : Euposue not possible cleaning and verify collimator frontal check and test system ok. ----------------------------------------------------------------------------- | 2023-10-14 08:32 | No
2246 | Yes | 0/IntelliVue Patient Monitor Mu0/Speaker failure (GFE0ENT Dec0 F/U= Dec0 Case in progress Dec0 | 0N/IntelliVue Patient Monitor Mu00Speaker failure (GFE0SENT Dec0) F/U= Dec0) Case in progress Dec0 | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000783 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Neut Step Action: PH Li Yi 18515200000000200000000called and the equpment is at GungzhouZheuian Medical Equpment Co., Ltd. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: Retun visit sucessfu SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000G/IntelliVu Patient Monitor Mu00000000/ T2 Activities:Problem Description: The cutomer's description of the problem: The uLLmessage of speaker failue. The cutomer observes the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observes the uLLmessage displayed on the monitor: na T2 Remote Diagnosis and Testing: na Repair Action: T2 Recommended on-site repair steps: Check the motherboard and speakers T2 Recommended spare parts uuand description: 0000000056000000006331 Cutomer demand notes: na ----------------------------------------------------------------------------- | 2023-10-14 08:52 | No
2247 | Yes | Poor battery contact (Provided information to cuomer) | Poor battery contau(Provided information to cuomer) | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: Instruted Cutomer Neut Step Action: A23B-00000000900000000 The cutomer informs that the end uer is China Water Conservancy and Hydropower Research Institue in Beijing Problem Reported by Cutomer: 00000000G/AED/ Neut Step Action: This newly installed equpment has not yet been insued. Please do not add Onsite labor. Please inform C1 if youneed to dispatch labor. Safety Qustions: Case: 000000002200000000823 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- DESCRIPTION OF PROBLEM: Poor battery contact WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Boot u POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? Chiller Service/3rd Party Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Poor battery contact WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? N\A WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Gude cutomers to replace batteries Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: Gude cutomers to replace batteries Service Content: If the battery contact is poor, the device will restart self-test repeatedly and gude the cutomer to replace the battery. Service resuts: After remote service, the equpment resues normal operation Test/Inspection: NA Test Equpment: NA Internal Remarks: Waiting for the test resuts of the repair applicant SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-14 09:07 | No
2248 | Yes | 00efibrillator is not functioning | 000Defibrillator is not funuoning | 000000000000000000000000001O##: 000000000000000000000000001000000002199100000000nanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : 0000000056000000007191 DFM100000000 Display ASSY 2.1 Ordered. Diagnostic performed by Engineer : Defibrillator is not fuctioning Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Follow u Requred : 0000000056000000007191 DFM100000000 Display ASSY 2.1 Ordered. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000219918 Parts Consued : 1.0000000000000000000000000000000056000000007191,DFM100000000 Display ASSY 2.1 ----------------------------------------------------------------------------- | 2023-10-14 10:22 | No
2249 | Yes | System not coming up. No activity on main monitors. | System not coming up. No auty on main monitors. | 000000000000000000000000001O##: 00000000000000000000000000100000000218125 100000000ct-2000000003 Botts, Andrew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Ouside of ue Eupected and ueupected Resuts: C-arm 000000000000000000000000001ld not move and system 000000000000000000000000001ldn’t boot back u uer Impact: Down Patient Impact: Delays SENA Diagnostic performed by Engineer : System was off uon arrival. Checked and reseated cabling to PCs and Ethernet switches. Coud not dulicate issu after mutiple reboots. Discused uLL in log files with NSS and it was determined that this was most likely an isolated issu bu will take futher action if the issu reoccus. ----------------------------------------------------------------------------- Internal Remarks : Case: 000000002200000000622 FSE: Botts, Andrew Intermittent boot u issu. FSE states Site has been having Air Conditioning issus. Reviewed logs; instruted FSE to perform the following soluions: Caue 1 The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. Soluion Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start u of the system. If so, replace GEO IPC. Caue 2 Bad cabling and/or connections. Soluion 1 Check the cable RC:u25 to SCC1:u22. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. Soluion 2 Tighten all the connectors of GeoIPC and SCC1 uit. Additionally, FSE to speak with cutomer to not ue system util Environmental Conditions on-site are u to u specification. ----------------------------------------------------------------------------- | 2023-10-14 10:32 | No
2250 | Yes | Geometry movement fault | Geometry movement fault | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Overview of cutomer problems: Mechanical movement uLL Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Mechanical movement uLL Remote diagnostics and test resuts: Consider IPC startu or commuication abnormalities PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Check whether there are abnormal uLL in the post statu, whether ipc starts normally, whether there are any abnormalities in synqnet commuication and PCu and futher trouleshoot based on on-site uLLreports. Recommended spare part uuand name: N/A Notes on cutomer requrements: onsite T2 Activities:Repair Action: PPD,00000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS, 200000000to hosp, the repair person will collect it SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 100000000000000002000000003 12:5000000003 uC (8600000000900000000) Service Content: 1. Check the equpment failue, emergency stop buton activity, and emergency stop circut. 2. Replace the PCu If the problem still cannot be solved, recheck the circut. 3. If there is a virtul connection in the emergency stop circut, reconnect the line and the equpment will retun to normal. Service resuts: The device is powered on normally and the perspective euposue is normal. SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022000000002300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang Zhu Last udated by: Weiqiang Zhu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD 000000002300000000325000000001. Shipped 200000000 Cao Peng-86-1810000000062800000000 Internal Remarks: 19.1000000002000000003 0000000039:00000000uC (8600000000900000000) $Q7: (Confirmed the report problem that the emergency stop buton activity. ) $Q8: (1. Checked the logfile, 1.check the emergency Buton, 2.check the emergency circut.3.reconncet The cable, replace the PCu not OK, 00000000check the cable again, foud the cable is lossen) $Q9: (Reconnect the emergency buton, the resuts passed. Device retuns to ful service. ) $Q100000000 Test and inspection data refer to PA. Device retuns to ful service. $Q11: Calibrated Tools:N/A. Problem Reported by Cutomer: The bed and frame cannot move duing sugery ----------------------------------------------------------------------------- | 2023-10-14 11:43 | No
2251 | Yes | Failure to start | Failure to start | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 100000000000000002000000003 10000000015:15 uC (860000000062500000000 RDI accuacy? Yes. $Q7: (Confirmed reported problem, the system can't startu) $Q8: (check the HostPC not startu, the power suply is defective. ) $Q9: (replace the power suply) $Q100000000 Test and inspection data refer to: PA. Device retuned to ful fuctionality. $Q11: N/A Problem Reported by Cutomer: QAA00000000R/u uuuFD200000000 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 100000000000000002000000003 10000000000000000:51 uC (860000000062500000000 "SuT. Ready for Clinical ue local time" : " 2000000003-1000000001000000000000000000000000 Service content: The system cannot be started and the host power suply has no oupu. Replace the power suply. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. SAP Transfer: Dispatched for Onsite Suport Neut Step Action: PPD Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000 1 Arranged to go to the hospital 200000000ous a day, FSE Zhang Yue-86-18656130000000022 T2 Activities:Problem Description: Overview of cutomer problems: Start failue Trouleshooting Action: Key faut codes or phenomena: Host pc cannot start PRS TR statu: Doesn't make sense Repair Action: Recommended on-site repair steps: " 1. Go to the site and check whether the host pc power suply is normal. 2. Replace host pc" Recommended spare parts uuand designation: "ppd/0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD/200000000to hospital/site1" Safety Qustions: Case: 000000002200000000755 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-14 13:28 | No
2252 | Yes | fluoroscopy not allowed | fluoroscopy not allowed | 000000000000000000000000001O##: 0000000000000000000000000010000000021338000000006-Oct-2000000003 Jamil Hossain The system meets the specification for the performed service and is retuned to ue Resoluion : No issu observe in system. Hospital electrical panel and changeover switch not 000000000000000000000000001king properly. Hospital maintenance team rectified the issu Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: NA uer Impact: NA Patient Impact: NA SE ----------------------------------------------------------------------------- | 2023-10-14 13:42 | No
2253 | Yes | geometry movement issues | geometry movement issues | 000000000000000000000000001O##: 0000000000000000000000000010000000032700000000 Anand Pishe The system meets the specification for the performed service and is retuned to ue Resoluion : Informed magnu table engineer abou issu. Problem of magnu table resolved by Magnu table engineer Problem Description by Engineer : Cutomer Fuction/Role: Biomedical engineer Device ue: No Eupected and ueupected Resuts: Magnu table movement problem uer Impact: No Patient Impact: No SE 1.2.5 Diagnostic performed by Engineer : magnu table movement issu ----------------------------------------------------------------------------- | 2023-10-14 13:47 | No
2254 | Yes | [Q] Geo module mechanically defective | [Q] Geo module mechanically defeu | 000000000000000000000000001O##: 0000000000000000000000000010000000022000000003 Oct 27, 2000000003 Krub, Horst The system meets the specification for the performed service and is retuned to ue Resoluion: On cutomer requst: Connection of a geometry control modue from cutomer stock. Fuction test OK. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Problem Description by Engineer: Cutomer Fuction/Role: Medical Technology Device ue: no Eupected and ueupected Resuts: Mech damaged uer Impact: no Patient Impact: no SE 8.1.17 ----------------------------------------------------------------------------- Cutomer comment: Confirmed by ## Cutomer comment: ## Arranged appointment with Sabine for October 27, 8:00000000 hrs. ## Cutomer comment: Call to the cutomer Please arrange appointment with u Jug tomorrow. Cutomer comment: Rs. with u Jug, appointment 000000000000000000000000001ld be possible on October 27, 8:00000000-10000000000000000 hrs., we will get in touh when we have held a meeting with u Krub. Internal Remarks: Via u Baugarten / PH KVA for 1u 00000000800000000151715 GEO MODuE WP Kit - in euchange 2 hous of 000000000000000000000000001k Cutomer comment: Scheduing possible on Monday with u Jug. ----------------------------------------------------------------------------- | 2023-10-14 14:28 | No
2255 | Yes | ulead failure | ulead failure | 000000000000000000000000001O##: 00000000000000000000000000100000000219755 Gnanasekar, D The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : 0000000056000000009000000001 DFM100000000 ASSY Therapy Diagnostic performed by Engineer : ulead failue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Follow u Requred : 0000000056000000009000000001 DFM100000000 ASSY Therapy ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000219823 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy ----------------------------------------------------------------------------- | 2023-10-14 16:12 | No
2256 | Yes | 00he paddle/electrode connection is broken. | 000The paddle/eleuode conneuon is broken. | 000000000000000000000000001O##: 0000000000000000000000000010000000027700000000 November 1000000002000000003 Alejandro Rodrigo Mauier Córdoba The system meets the specification for the performed service and is retuned to ue. Tools & Equpment: DEFIBRILLATOR ANALYZER, 200000000000000000000000020000000000000000:00000000:00000000 Reguatory Qustions: N/A Parts Consued: 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle 1.000000000000000000000000000000005600000000871,DFM100000000 Therapy Collar ASSY Resoluion: Therapy connector and protection collar are replaced. Fuctional parameters are checked. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Clinical ue Eupected and ueupected Resuts: Therapy connector is damaged. uer Impact: NO Patient Impact: NO SE 2.00000000.21 ----------------------------------------------------------------------------- | 2023-10-14 16:27 | No
2257 | Yes | ETCO2 | ETCO2 | 000000000000000000000000001O##: 0000000000000000000000000010000000021500000000 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to Bench Follow u Requred : EMS site Manager requst Tempu Pro: Etco2 checked for accuacy and fuction. Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Site Manager Device ue:Not in Clinical ue Eupected and ueupected Resuts: Cutomer requst to have this uits Etco2 checked for accuacy and fuction uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Cutomer requst to have this uits Etco2 checked for accuacy and fuction Diagnostic performed by Engineer : Problem Description: Cutomer requst to have this uits Etco2 checked for accuacy and fuction Cutomer callback Tempu Pro: Cutomer requst to have this uits Etco2 checked for accuacy and fuction Form ued: Bench Trouleshooting Steps Taken: EMS site Manager requst Tempu Pro: Cutomer requst to have this uits Etco2 checked for accuacy and fuction. It was believed that ETCO2 did not coincide with patient condition. No failue or uLLreported. Cutomer is qustioning accuacy. Cutomer requst to send uit to bench to have the Etco2 evaluted/repaired. Created Bench 000000000000000000000000001 ##d cutomer bench repair form, and requsted shipping label. ----------------------------------------------------------------------------- | 2023-10-14 18:32 | No
2258 | Yes | [Q] C-arm not moving with side panel | [Q] C-arm not moving with side panel | 000000000000000000000000001O##: 00000000000000000000000000100000000325688 00000000-Nov-2000000003 Head Kelly The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000230000000022113,potm.uit longituinal Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: no Eupected and ueupected Resuts: intermittent slippage of c-arm; uer Impact: no Patient Impact: no SE n/a Resoluion : tested alignment with laser alignment tool and level; replaced c-arm longituinal pot; calibrated long pot; completed curent calibrations; tested movement of c-arm in longituinal position; system fuctioning properly & retuned to cutomer for ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000251310000000000000000-Nov-2000000003 Riley Earl The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Lead Tech Device ue: In ue Eupected and ueupected Resuts: When driving c-arm longituinally, it 000000000000000000000000001ld not move intermittently. uer Impact: Inconvenienced Patient Impact: None SE R1.2.5 Diagnostic performed by Engineer : Diagnostics: Checked logfile and foud uLLin longituinal movement relating to Pot. Resoluion : Resoluion: We changed ou the c-arm longituinal Pot, calibrated it, and checked operation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000215000000006 00000000-Nov-2000000003 Head Kelly The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : follow u Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: no Eupected and ueupected Resuts: intermittent slippage of c-arm; uer Impact: no Patient Impact: no SE n/a Diagnostic performed by Engineer : uable to access system du to emergent patient ue; reschedued for tomorrow; no 000000000000000000000000001k performed today ----------------------------------------------------------------------------- | 2023-10-14 18:33 | No
2259 | Yes | Live monitor has no signal | Live monitor has no signal | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 27.1000000002000000003 00000000:17:12 uC (8600000000000000003) RDI Accuacy: Yes $Q7: (Confirm reported problem, device has no problem, live monitor can't light u) $Q8: (performed visul inspection and check cable foud live monitor can't light u) $Q9: (reseat cable of signal inpu. ) $Q100000000 (Test and inspection refer to PA, Device in ful fuctionality) $Q11: Calibrated Tools name (NA) Euternal Remarks: 26.1000000002000000003 00000000:00000000:57 uC (8600000000000000003) Service content: The live screen sometimes goes black, bu it becomes normal after reprocessing the split screener. Service Resut: Equpment is fuly fuctional. Other important matters: None yet. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 1 The machine needs maintenance 2 There is no signal on the main screen. It still doesn’t 000000000000000000000000001k after restarting. 3 The muhroom head on the console is broken SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Overview of cutomer problems: No signal on main screen Detailed description of cutomer problem or faut phenomenon: No signal on real-time screen Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Check video wiring and monitor DVI 5V Notes on cutomer requrements: Please contact the cutomer for on-site repair Safety Qustions: Case: 00000000220000000023 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-14 19:43 | No
2260 | Yes | [Q] VIP-IMSS-CLIENT-POLYGRAPH FAILURE, THE APPLICATION DOES NOT START. | [Q] VIP-IMSS-CLIENT-POLYGRAPH FAILURE, THE APPLICATION DOES NOT START. |  | 2023-10-14 20:32 | No
2261 | Yes | Intrepid sn CN00 D-SV00 | Intrepid sn CN0000D-SV000 |  | 2023-10-14 20:48 | No
2262 | Yes | SKH ¦ DigitalDiagnost 0hest / Emergency ¦ A AND E resupanel 0disconnected from the ceiling fortitude | SKH ¦ DigitalDiagnost 0Chest / Emergency ¦ A AND E resupanel 0 disconneud from the ceiling fortitude | 000000000000000000000000001O##: 00000000000000000000000000100000000219930000000000000000ct-2000000003 Sebastian, Nithin The system meets the specification for the performed service and is retuned to ue Resoluion : Confirm faut was reported Visited site and verified holder damaged ( loosen and came ou) Managed to get new spare. Replaced new carriage holder and tested ok Problem Description by Engineer : Cutomer Fuction/Role: na Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SEna Cable carriage holder came ou from the ceiling rail ----------------------------------------------------------------------------- | 2023-10-14 20:48 | No
2263 | Yes | Angiography continues showing a memory full message even when deleting images from previoueuams. it is impacting on... | Angiography continues showing a memory full message even when deleting images from previoueuams. it is impaung on... | BuQ&R Feedback: No SAP Cutomer Symptom Code: ACQP Cutomer Symptom Code: Imaging/Patient Recording/Acqusition Problem Reason: Corrective Maintenance Engineer Failue Code: Electrical/Electronic Failue Repair Activity Code: Instruted Cutomer Problem Reported by Cutomer: Angiograph continus to display memory ful message even deleting images from previou euams. It is impacting the performance of procedues and runing the risk of locking the device duing a high-risk procedue. 000000000000000000000000001king No one was hut Lais Sector: Life Image 71 991225582 maria.heloisa00apvida.com.br ID: CV00000000300000000 Is the u uuuFD100000000uY the equpment? It belongs to HOSPITAL ANTONIO PRuENTE DA BAHIA, and is located at Avenida Antonio Carlos Magalhaes 200000000 Salvador-BA 0000000000000000-700000000 T2 Activities:Problem Description: Problem description by engineer : Disk Ful Trouleshooting Action: What steps are taken duing the system diagnosis : Video Connection & Remote Connection uLL# and/or description of uLL: Image Disk Ful Repair Action: ► Checked the IPPC post and no uLLwas foud; ► Database consistency test pointed to 87 failues. They were repaired by the system, bu as there were many of them, we decided to rebuld the database; ► After rebulding the database, tuning the equpment off and on and testing, the equpment showed the correct nuber of available images and did not give an uLLin the database consistency tests; ► Left 1 week in testing with no new failues. The Lais client is aware of the closue of the call as resolved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK Test Equpment: N/A OneEMS internal remarks: Problem description by engineer : Disk Ful What steps are taken duing the system diagnosis : Video Connection & Remote Connection uLL# and/or description of uLL: Image Disk Ful Repair Action "Euternal" : ► Checked the IPPC post and no uLLwas foud; ► Database consistency test pointed to 87 failues. They were repaired by the system, bu as there were many of them, we decided to rebuld the database; ► After rebulding the database, tuning the equpment off and on and testing, the equpment showed the correct nuber of available images and did not give an uLLin the database consistency tests; ► Left 1 week in testing with no new failues. The Lais client is aware of the closue of the call as resolved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK Test Equpment: N/A Eutra Internal comment field : CSR Marcel Ogassawara SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Renata Lima Last udated by: Renata Lima 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A ----------------------------------------------------------------------------- | 2023-10-14 22:08 | No
2264 | Yes | Q/O Wireless Footswitch failed during FCO | Q/O Wireless Footswitch failed during FCO | 000000000000000000000000001O##: 00000000000000000000000000100000000218100000000 300000000Oct-2000000003 Vanmeter Timothy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: flouo Eupected and ueupected Resuts: footswitch FCO failed uer Impact: had to ue wired footswitch Patient Impact: none SE 2.1 Parts Consued : 1.00000000000000000000000000000000800000000257862,WFS Base Station Diagnostic performed by Engineer : WFS FCO failed, had to replace WFS base ----------------------------------------------------------------------------- | 2023-10-15 01:02 | No
2265 | Yes | FCO0PM board not compatible with after-market battery OEM letter | FCO0 PM board not compatible with after-market battery OEM letter | Internal Remarks : 10000000013/23 YP 000000000000000000000000001000000002181900000000 Details: Cncted cut for qutes. cut not reachable. left vm Contact Alonso Lares Phone 00000000-531-9731 ##alonso.lares00e.com ----------------------------------------------------------------------------- | 2023-10-15 01:22 | No
2266 | Yes | Device failed the PM due to capnometer port being physically broken | Device failed the PM due to capnometer port being physically broken | 000000000000000000000000001O##: 0000000000000000000000000010000000035000000009 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Device was repaired by parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits CO2 connector is physically damaged. Additionally the uits Charging port is physically damaged. To correct both issus the CO2 assembly will need to be replaced and the Front Case Assembly needs to bee ordered. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000218538 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits CO2 connector is physically damaged. Additionally the uits Charging port is physically damaged. To correct both issus the CO2 assembly will need to be replaced and the Front Case Assembly needs to bee ordered. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 000000005600000000000000001-Capnometer 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000323695 Harris, Alvin The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to Bench --- ##d Irvine CA Bench form to cuProblem Description by Engineer : Cutomer Fuction/Role:Other: Chief/Officer/Equpment Manager Device ue:Not in Clinical ue Eupected and ueupected Resuts:NA uer Impact:Can't perform capnometer measuements Patient Impact:No impact SEuknown, cuis not with the device Problem Description/ueupected Resuts: Capnometer port broken/Needs repair Diagnostic performed by Engineer : Problem Description: Capnometer port broken/Needs repair Device failed the PM du to capnometer port being physically broken Annul PM 2000000003 Form ued: Bench Trouleshooting Steps Taken: Received Teams message from Kim Jordan asking that I accept 000000000000000000000000001O#and call cuback to assist Replied with agreement and accepted 000000000000000000000000001 --- Foud case, onsite service 000000000000000000000000001O# and bench repair 000000000000000000000000001O#were all created on 10000000013/23, yet the remote service 000000000000000000000000001O#was opened today, 11/1/23 --- Later called cuback at (817) 00000000-000000003 and reached cu Qustioned cuon the fact that his original case was created on 10000000013/23 Cueuplained that he already had his device boued u with a shipping label and a cutomer created Bench repair form at his local Fed Eu location waiting to be shipped to u Custated he had been waiting for a case nuber, called in, and foud a case had been created Custated he did not want to also wait for a Bench repair form to be ##d to him Apologized to cufor the strange sequnce of events and delay in his being able to send his device to ou Bench Verified cus ship to address Azle Fire Department 900000000 Lakeview Dr AZLE Tu 76000000000000000037200000000 ATT: Lee Godbold/Case 000000002200000000926 Informed cuthat I 000000000000000000000000001ld still ##him the official Bench repair form in the event that it was requsted and inform the Bench that he is sending the device in Confirmed nothing else is needed from technical suport ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Lee Godbold 817300000000900000000 IP & Issu: 61000000005/000000002200000000926/ CCB for repair Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-15 02:27 | No
2267 | Yes | Image disk full | Image disk full | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 19.1000000002000000003 00000000:19:100000000uC (8600000000600000000) Faut description: Cannot see throuh and eupose Service content: The log indicates that IPPC cannot operate normally. uplu and insert the IPPC board and memory modue. Service resuts: After repair, the equpment retuned to normal operation. Internal Remarks: 19.1000000002000000003 00000000:19:13 uC (8600000000600000000) RDI Accuacy:YES $Q7: (Confirm reported problem & No perspective or euposue) $Q8: (If the IPPC fails to ru properly, remove and reinsert the IPPC board and memory modue) $Q9: (If the IPPC fails to ru properly, remove and reinsert the IPPC board and memory modue after maintaining the system normal ) $Q100000000 (Test and inspection data refer to PA, Device retun to ful fuctionality) $Q11: Calibrated Tools name:N/A For Mainland China CV / uR, CT engineer SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000366 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: QAA00000000R/u uuuFD200000000T2 Activities:Problem Description: Overview of cutomer problems: Image disk ful Detailed description of cutomer problem or faut phenomenon: No euposue, only perspective Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: The cutomer requres the engineer to be present Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended to check the logs and check the IPPC hard disk smart data If IPPC does not start, it is recommended to replace the bios battery and plu in and ou the board If it has been started, it is recommended to bypass the hard disk bou, download IPPCsoftware again, and recreate imagedisk Notes on cutomer requrements: Please contact the cutomer for on-site repair ----------------------------------------------------------------------------- | 2023-10-15 06:52 | No
2268 | Yes | Startup failed | Startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: NTD: The cutomer said that someone in the department contacted the engineer and refued to transfer. The hard drive is broken NNY Neut Step Action: Notify FSE Ni Zhiqiang for the time being 18667936622 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 16.1000000002000000003 00000000:12:31 uC (86000000006300000000 Service content: The system cannot start, the hard disk modue is damaged, replace the hard disk modue. Service resuts: After repair, the equpment retuned to normal operation. "SuT. Ready for Clinical ue local time" : " 2000000003-10000000015 18:300000000 Internal Remarks: 17.1000000002000000003 00000000:0000000035 uC (86000000006300000000 "RDI Accuacy: YES " $Q7: ( Confirmed reported problem & The system can't boot u ) $Q8: ( Check the boot device, no hard disk & The OS disk is defective ) $Q9: ( Replace the OS disk uu000000080000000050000000000000000and reinstall the OS ans APP ) $Q100000000 ( Test and inspection data refer to PA & Device retuned to ful fuctionality ) $Q11: ( Calibrated Tools: N/A ) SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine coud not enter the application interface normally and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine cannot enter the application interface normally Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the Host PC starts normally, try to clean and plu in the internal board, 2. Check whether the BIOS battery is normal 3. ue PC Diagnostic Tool for diagnostic testing 00000000Check whether the software is normal, try to ue Ghost to restore the system 5. Check M Check whether the DCPS of the cabinet is normal. 6. Check whether the internal commuication of the system is normal. 7. Check the machine uLLand repair the faut. If it cannot be repaired, replace the corresponding spare parts. Recommended spare part uuand name: 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD & 0000000080000000050000000000000000 Hard disk spare part Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Case: 000000002200000000668 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD0000000080000000050000000000000000*1, 0000000080000000050000000000000000*1, has been shipped, please send 000000000000000000000000001k FSE: LiuChao -86-13979783177 Neut Step Action: Negotiating for labor dispatch Problem Reported by Cutomer: 00000000R/FD200000000 Neut Step Action: NTD: The cutomer said that someone in the department had contacted the engineer and refued to transfer. The hard drive is broken NNN SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-15 11:42 | No
2269 | Yes | System runs slowly | System runs slowly | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 2000000001000000002000000003 12:00000000200000000TC (86000000008200000000 "RDI Accuacy:Yes" $Q7: (Confirmed reported problem, the system did not bootu, and The startu program stops at 0000000000000000) $Q8: (maybe the software corruted or operation system disk issu make the startu program failed, and reinstall the software failed) $Q9: (replaced operation system disk and restore the system) $Q100000000 (Test and inspection data refer to PA, performing the cardiac scanning, perfectedimages display on monitors, Device retuned to ful fuctionality) $Q11:NA For Mainland China CV / uR, CT engineers only.Please fill in: ^Curent Tue installed date: ^Curent Tue uage: scan second_ Euternal Remarks: 15.1000000002000000003 00000000:00000000:23 uC (86000000008200000000 Service content: The device cannot be tuned on intermittently. Replace the system hard disk and install the software. Service resuts: After maintenance, it was confirmed that all fuctions were operating normally and the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD spare parts 0000000080000000050000000012 and 0000000080000000050000000000000000arrange 12H to the hospital, Yu Jiangmin -86-18630000000000000000181 Safety Qustions: Case: 000000002200000000700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: FSE Yu Jiangmin CV/u-ray(18630000000000000000181) T2 Activities:Problem Description: Overview of cutomer problems: System crash Detailed description of cutomer problem or faut phenomenon: Slow boot, system lag Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: n/a Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Try replacing the hard drive or host Recommended spare parts uuand designation: PPD uu0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II and 0000000080000000050000000012 12H to Hosp Notes on cutomer requrements: Please contact the cutomer for on-site repair SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-15 12:27 | No
2270 | Yes | Intermittent startup failed | Intermittent startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: FSE Zhao Hongqun has been dispatched 1537697000000005 Euternal Remarks: 19.1000000002000000003 13:50000000021 uC (8600000000800000000) Service content: The host tuns on a black screen, replace the host hard disk slot Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000795 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine coud not enter the application interface normally and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: Read uLLRate: 200000000 (3), Reallocated Sectors Cout: 200000000 (00000000, Reallocation Event Cout: 200000000 (00000000, Curent Pending Sector Cout: 200000000 (6) Remote diagnosis and test resuts: Restart can restore PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check whether the Host PC hard disk is normal, try to bypass or format and then restore the system. 2. Check whether the IPPC is normal. 3. Check the machine for uLL and repair the faut. If it cannot be repaired, replace the corresponding spare parts. Recommended spare part uuand name: 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD & 0000000080000000050000000000000000 Hard disk spare part Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Problem Reported by Cutomer: 00000000L/FD200000000 Internal Remarks: 19.1000000002000000003 13:500000000200000000TC (8600000000800000000) "RDI Accuacy:yes" $Q7: (Confirmed reported problem &Intermittent startu failed) $Q8: (duing startu, have blu screen, 000000000000000000000000000000007A, os disk have problem) $Q9: (change os strip on host pc) $Q100000000 (Test and inspection data refer to: PA & Device retuned to ful fuctionality) $Q11: Calibrated Tools name (N/A) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: N/A SAP Transfer: Case Transferred to SAP Neut Step Action: NTD: The cutomer contacted site1 and refued to transfer. Intermittent failue to power on NNN Limited labor times, there are still labor times available, and 000000000000000000000000001kers can be dispatched normally ----------------------------------------------------------------------------- | 2023-10-15 14:02 | No
2271 | Yes | Unit did not deliver shock | Unit did not deliver shock | 000000000000000000000000001O##: 00000000000000000000000000100000000220000000013 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Placed Material 000000000000000000000000001for euchange Tempu LS Follow u Requred : Please Tu COLuBIA RIVER FIRE AND RESCu Attention: Jeff Lockhart 100000000 S 12th St SAINT HELENS OR 97000000001-1825 Problem Description by Engineer : Cutomer Fuction/Role:Sales Device ue:Not in Clinical ue Eupected and ueupected Resuts: "Duing a cardiac arrest, ou team pressed the shock buton and althouh it recorded a shock, it did not deliver any energy." uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: "Duing a cardiac arrest, ou team pressed the shock buton and althouh it recorded a shock, it did not deliver any energy." Diagnostic performed by Engineer : Problem Description: "Duing a cardiac arrest, ou team pressed the shock buton and althouh it recorded a shock, it did not deliver any energy." Cutomer callback Tempu LS: "Duing a cardiac arrest, ou team pressed the shock buton and althouh it recorded a shock, it did not deliver any energy." Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 100000000100000000000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Accout Manager called to report Tempu LS: "Duing a cardiac arrest, ou team pressed the shock buton and althouh it recorded a shock, it did not deliver any energy. This was verified/witnessed by mutiple people. It 000000000000000000000000001ked normally all of the other times. Device has been taken ou of service." The following information was requsted from cutomer: *Log and Rescu Files: Attached to case. *Date of event? 100000000100000000000000003 *Brief description of the event? Tempu LS: "Duing a cardiac arrest, ou team pressed the shock buton and althouh it recorded a shock, it did not deliver any energy. This was verified/witnessed by mutiple people. It 000000000000000000000000001ked normally all of the other times. Device has been taken ou of service." *Operator of the produt? Paramedic *Did the uit display any uLLmessages or warnings before the incident? No *Was the Patient or uer injued or caued any harm? uu *Patient Oucome? uu *Age of Patient? uu *Patient Gender? uu Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000225625 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-10-15 22:47 | No
2272 | Yes | O0ressure sensor range uLL(EC 0) | O0pressure sensor range uLL(EC 00) | Internal Remarks : 10000000016/23 YP 000000000000000000000000001000000002200000000700000000 Attempt1 Details: spoke to cutomer Helen for qutes. Transferred to manager Karen who has to review the qutes. left vm Contact Helen Bobo Phone 615-396-5651 ##helen.bobo00scension.org Internal Remarks : 10000000018/23 YP 000000000000000000000000001000000002200000000700000000 Attempt2 Details: Cncted cut in the below nos. Transferred to Karen. Karen stated that Christopher is the correct P.O.C. christopher.asberry00scension.org Ph# 615396000000001 Director of Respiratory Contact Helen Bobo Phone 615-396-5651 ##helen.bobo00scension.org Internal Remarks : 10000000018/23 YP 000000000000000000000000001000000002200000000700000000 Attempt3 Details: Cncted the below no.s. Cutomer not reachable. left vm. sent qutes ##to the cutomer Name: Christopher Asberry Phone: 615-396-0000000000000000Mainline ## christopher.asberry00scension.org Internal Remarks : 10000000019/23 YP 000000000000000000000000001000000002200000000700000000 Attempt00000000 Details: Contacted cut for qutes. cut not reachable. Sent ##to cut if they want u to proceed with the repair Contact Christopher Asberry Phone (615) 396-0000000000000000 ##christopher.asberry00scension.org ----------------------------------------------------------------------------- | 2023-10-16 02:57 | No
2273 | Yes | uLL0 related to the OVP circuit failure. | uLL00 related to the OVP 0cuit failure. | Internal Remarks : 10000000018/23 YP 00000000000000000000000000100000000220000000071 Attempt2 Details: Cncted the below no.s. Cutomer not reachable. left vm. sent qutes ##to the cutomer Name: Christopher Asberry Phone: 615-396-0000000000000000Mainline ## christopher.asberry00scension.org Internal Remarks : 10000000017/23 YP 00000000000000000000000000100000000220000000071 Attempt1 Details: spoke to cutomer Helen for qutes. Transferred to manager Karen who has to review the qutes. left vm Contact Helen Bobo Phone 615-396-5651 ##helen.bobo00scension.org Internal Remarks : 10000000018/23 YP 00000000000000000000000000100000000220000000071 Cut sent ##stating that biomed team will look into servicing these uits. If they find that the repairs are beyond their scope, they will coordinate any service calls requred Contact Christopher Asberry Phone (615) 396-0000000000000000 ##christopher.asberry00scension.org Internal Remarks : 10000000018/23 YP 00000000000000000000000000100000000220000000071 Attempt00000000 Details: Contacted cut for qutes. cut not reachable. Sent ##to cut if they want u to proceed with the repair Contact Christopher Asberry Phone (615) 396-0000000000000000 ##christopher.asberry00scension.org ----------------------------------------------------------------------------- | 2023-10-16 02:57 | No
2274 | Yes | [Q] 0S:SS: The equipment does not work properly. | [Q] 0DS:SS: The equipment does not work properly. | 000000000000000000000000001O##: 00000000000000000000000000100000000386257 100000000ov-2000000003 Kodera, Norihiro The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of radiology Device ue: In ue Eupected and ueupected Resuts: Device does not start uer Impact: None Patient Impact: None SER2.2.3 Resoluion: We determined that the PC uit in the machine room was broken and replaced it (Fleuviewing PC for Fleuspot). After replacement, normal operation was confirmed. Part nuber〈00000000800000000718000000002〉/Part name〈Fleuviewing PC for Fleuspot〉/Serial〈531611000000000000000037〉/Quntity〈1〉 Parts Consued: 1.00000000000000000000000000000000800000000718000000002,FleuViewing PC 00000000 W7 + W100000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000291983 00000000-Nov-2000000003 Kuota, Yuchiro The system meets the specification for the performed service and is retuned to ue Resoluion: I was able to recover after receiving emergency treatment. Reguatory Qustions: N/A Resoluion: YES Same-day reguar report Problem Description by Engineer: Cutomer Fuction/Role: Radiation u Nakajima Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE2.2 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000238800000000 27-Oct-2000000003 Kuota, Yuchiro The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The device does not start u properly. Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: The device does not start u properly. uer Impact: Test canceled Patient Impact: Test canceled SE uknow Resoluion: YES Same-day reguar report Resoluion: No 000000000000000000000000001k done du to initial defective parts. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000221565 17-Oct-2000000003 Kuota, Yuchiro The system meets the specification for the performed service and is retuned to ue Follow u Requred: N/A Diagnostic performed by Engineer: YES Normal report on the same day Reguatory Qustions: N/A Problem Description by Engineer: The device does not start u properly. Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: The device does not start u properly. uer Impact: Test canceled Patient Impact: Test canceled SE uknow Resoluion: The Fleu spot PC was temporarily repaired and restored. We plan to replace the parts as soon as they arrive. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000221552 Saka, Shuji The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: The device does not start u properly. Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: The device does not start u properly. uer Impact: Test canceled Patient Impact: Test canceled SE uknow Diagnostic performed by Engineer: RC Contents: The monitor in the operation room displays the u logo. The screen in the euamination room is displaying normally. Even after rebooting, the sitution remains the same. Visit suvey. Internal Comments: [Suject] 00000000DS:SS:u Ishimoto: The device does not start u properly. [Internal Comments] It seems that there is no video oupu from FleuViewing PC for FleuSpot. [Ou of stock parts] 00000000800000000577362:FLEuVIEWING PC 00000000 mDP QTY:1 ZREP/Traceable Tokyo (u11) [Dispatch] 1 RC -----------------FSE registration information---- ------------ -Visit FSE1 Name [Yuchiro Kuota] Kuota, Yuchiro -FSE1 Visit date and time 2000000003/1000000001611:300000000-FSE1 000000000000000000000000001king hous 2 [Shipping instrutions] ------- ----------Shipping instrution #1----------------- -Delivery conditions ASAP (transportation arrangement) -Delivery date ASAP -Destination name Local delivery -Receipt Person's name: Yuchiro Kuota -0000000000000000005100000000-15300000000-Other Mobile phone call uon arrival -Requster's name: Toshihito Saka [Test category] Other Follow u Requred: * ----------------------------------------------------------------------------- | 2023-10-16 06:47 | No
2275 | Yes | The power on button does not work. | The power on button does not work. | 000000000000000000000000001O##: 0000000000000000000000000010000000022100000000 Watanabe, Atsuhi The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: RC Response: I euplained how to do it: OK Euplanation regarding failue locations: OK Recurence issu.FSE Kajiwara is schedued to respond on site today at 15:00000000 uder the euisting case 000000002200000000500000000 00000000000000000000000000100000000211899. Problem Description by Engineer: The power on buton does not 000000000000000000000000001k. He said that he had previouly consuted with an FSE abou a response, and the problem had 000000000000000000000000001sened, and this morning he had to operate it abou 100000000times before the power coud be tuned on. It was a consutation abou an early response. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: Cannot operate the power on buton. He said that he had previouly consuted with an FSE abou a response, and the problem had 000000000000000000000000001sened, and this morning he had to operate it abou 100000000times before the power coud be tuned on. It was a consutation abou an early response. uer Impact: No effect Patient Impact: No effect SE 8.2.3000000005 ----------------------------------------------------------------------------- | 2023-10-16 07:52 | No
2276 | Yes | M0A - M0A - The battery does not charge | M00A - M00A - The battery does not charge | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Qingyun District People's Hospital, Baoding City, Hebei Province ----------------------------------------------------------------------------- | 2023-10-16 08:47 | No
2277 | Yes | System is not switching on | System is not switching on | 000000000000000000000000001O##: 00000000000000000000000000100000000221887 16-Oct-2000000003 Ramesh, S The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,200000000000000009-21 00000000:00000000:00000000 Follow u Requred : Need to order part Problem Description by Engineer : Problem description by engineer : uray generator is not switching ON. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided Resoluion : Cutomer complaint regarding system not switching on , reached site observed , diagnosed and foud mains interface uit was fauty so the same ordered and fiued and system uder observation ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000229731 21-Oct-2000000003 Ramesh, S The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uray generator is not switching ON. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided Resoluion : Cutomer complaint regarding duing switching on foud buning smell from technical room, reached site and foud MIuis fauty so the same ordered and replaced foud same issu so futher diagnosis foud uSC is fauty so the same reordered & replaced and did callibration fiud system was normal Parts Consued : 1.00000000000000000000000000000000210000000020000000091,uSC Certeray 1.00000000000000000000000000000000210000000019971,Mains Interface uit (MIu Certeray Tools & Equpment : MuTIMETER,200000000000000009-21 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000221251 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : uray generator is not switching ON. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided Internal Comments : uLL# and/or description of uLL: uray generator is not switching ON. Repair Action : MIuneed to replace Diagnostic performed by Engineer : uray generator is not switching ON. MIuinpu is ok bu no oupu and inside MIuShort circut and buning foud Follow u Requred : uLL# and/or description of uLL: uray generator is not switching ON. Repair Action : MIuneed to replace What are the neut Trouleshooting steps : MIuneed to replace 00000000210000000019971 Mains Interface uit Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-10-16 09:13 | No
2278 | Yes | Installation screen is displayed. | Installation screen is displayed. | 000000000000000000000000001O##: 00000000000000000000000000100000000231260000000000000000ov-2000000003 Martinic, Chris The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: In ue Eupected and ueupected Resuts: System was rebooted and it booted u to a net000000000000000000000000001k uLLand an installation screen with cogs. At the moment all the control room monitors are showing loading screens. The FleuVision monitor is ok. uer Impact: System is not uable, has to reschedue patients. Patient Impact: No impact SE 1.2.5.00000000Resoluion : Replaced fleuviewing PC, configued, licensed, tested operation pass. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000222000000000000000016-Oct-2000000003 Martinic, Chris The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: chief radiographer Device ue: coronary care Eupected and ueupected Resuts: no uer Impact: nil Patient Impact: nil SE Resoluion : Fleuview pc replaced, reinstalled software 1.2.5 , licensing server not available, placed new installed hard disk in original compuer hardware, performance tests passed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000221950000000016-Oct-2000000003 Frigon, Christopher The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: In ue Eupected and ueupected Resuts: System was rebooted and it booted u to a net000000000000000000000000001k uLLand an installation screen with cogs. At the moment all the control room monitors are showing loading screens. The FleuVision monitor is ok. uer Impact: System is not uable, has to reschedue patients. Patient Impact: No impact SE 1.2.5.00000000Diagnostic performed by Engineer : ued PC diagnostics tool to confirm the hard drive is not booting on the fleuviewing PC. Du to not being able to generate a new license file for the system the hard drive from the new PC was installed into one of the 000000000000000000000000001king slots of the euisting fleuviewing PC, and software loaded to the new hard disk. A follow u has been created to install the new PC when a new license can be generated and retun the defective PC. Resoluion : Performed a fuctional test of the system and retuned the system to clinical ue with the euplanation that at a later date an engineer will need to retun to swap ou the fleuviewing PCs. Follow u Requred : Replace the PC when wisekey is back u and runing with PC that was ordered and get me license file for the system. Parts Consued : 1.00000000000000000000000000000000800000000660000000013,FleuViewing PC 00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002218600000000Christian De Candia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Technician requred to: - Check FleuSpot PC in B cabinet, see if powering on etc. - Reload S/W if possible - Replace PC is needed Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: In ue Eupected and ueupected Resuts: System was rebooted and it booted u to a net000000000000000000000000001k uLLand an installation screen with cogs. At the moment all the control room monitors are showing loading screens. The FleuVision monitor is ok. uer Impact: System is not uable, has to reschedue patients. Patient Impact: No impact SE 1.2.5.00000000Diagnostic performed by Engineer : Saved logs files and determined some uLL with the FleuViewing PC for the FleuSpot. Supecting that there might be an issu with the boot u of the PC, or some software problems. This will requre a technician on site. ----------------------------------------------------------------------------- | 2023-10-16 09:33 | No
2279 | Yes | 0/UNIQ FD0Cannot fluoro | 0R/UNIQ FD0Cannot fluoro | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Neut Step Action: 000000000000000000000000001king with FSE Zhang Jingyu13967122176 commuicating Internal Remarks: 31.1000000002000000003 15:31:00000000 uC (860000000063600000000 RDI Accuacy: YES " $Q7: (Confirmed reported problem, Cannot fluro- free image space low, system can't euposue&fluro) $Q8: (Check system log, foud uLL"Image Disk uLL(Write uLL"; bu check the disk smart data, no uLL so need to reseat image disks and recreate image store. ) $Q9: (Reseat image disks and recreate image store.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11:N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage: N/A Internal Remarks: 18.1000000002000000003 1000000008:00000000 uC (86000000006573) Only on job training Problem Reported by Cutomer: 00000000R/uIQ FD200000000 SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 16.1000000002000000003 11:55:00000000uC (860000000063600000000SuT. Ready for Clinical ue local time" : " 2000000003-10000000016 00000000:300000000 T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine was uable to perform perspective euposue and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine cannot see throuh the euposue Remote diagnostics and test resuts: NA PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check whether the foot brake is normal. 2. Check whether there is an uLLin the high voltage. 3. Check whether the image chain is normal. 00000000Check the machine uLLand perform faut repairs. Recommended spare part uuwith name: 00000000800000000219713 FOOTSWITCH PACKED Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000691 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: NTD: The cutomer refued to transfer and said that he coud not see throuh YNY Neut Step Action: Temporarily dispatch 000000000000000000000000001ker Zhang Euternal Remarks: 31.1000000002000000003 15:31:00000000 uC (860000000063600000000Service content (requred): uable to perform perspective euposue; check image disk connection and rebuld image storage. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: none. Neut Step Action: FSE Golden Bulding 1575782100000000 Call and add 300000000OJT ----------------------------------------------------------------------------- | 2023-10-16 10:28 | No
2280 | Yes | [Q]Images do not display on the general purpose monitor in the euamination room. | [Q]Images do not display on the general purpose monitor in the euamination room. | 000000000000000000000000001O##: 00000000000000000000000000100000000222336 16-Oct-2000000003 Konishi, Hiroto The system meets the specification for the performed service and is retuned to ue Resoluion: I made a 000000000000000000000000001k report on October 16, 2000000003 and received his approval. Reguatory Qustions: N/A Problem Description by Engineer: Images from the general-pupose monitor in the euamination room do not display. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: Images are not displayed on the general-pupose monitor in the euamination room. uer Impact: No effect Patient Impact: No effect SE 000Resoluion: The power adapter for the DVI converter was fauty. After I replaced the power adapter, I checked the images and there was no problem. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000222175 Noda, Kotaro The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Images from the general-pupose monitor in the euamination room do not display. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: Images are not displayed on the general-pupose monitor in the euamination room. uer Impact: No effect Patient Impact: No effect SE 000Follow u Requred: * Internal Comments: Simpl physical [Suject] 00000000DS:KN:u Ogawa: Images do not display on the general-pupose monitor in the euamination room. [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Hirokachi Konishi] Konishi, Hiroto L1 Engineer - FSE1 Visit date and time 2000000003/1000000001617:00000000 - FSE1 000000000000000000000000001k time 2 - Name of dispatched suervisor [Manawa T know] Tachinami, Makoto ¦ L2 { 2900000000} (Level 2) Dispatched suervisor [Euamination category] Normal euamination [Dispatched suervision] Dispatched suervisor Case_ L1 FSE_ [Hirokachi Konishi] Konishi, Hiroto ¦ L00000000{ 393 } L2 FSE_ [Manawa T know] Tachinami, Makoto ¦ L2 { 2900000000} L1 FSE_ This system was done uder the suervision of the L2FSE, and it is necessary for the L2FSE (engaged in another operation) to approve the parts order, manage technical escalation, and receive final approval before starting ue on the patient. L2 FSE_ The role of dispatch suervisors is to approve all parts orders, manage all technical escalations, provide gudance and instrutions for service 000000000000000000000000001k to the L1FSE who has been assign, and review and approve the 000000000000000000000000001k review and approval of the FSE. The final approval of the content of the case by the L2 FSE and the close (CNF) means that the repair of the S000000000000000000000000001and the recorded content (incluing all recorded comments) were officially approved. Diagnostic performed by Engineer: Images are not displayed on the general pupose monitor in the euamination room.It appears that "Ni signal" is displayed. It looked like there was an Imagenics switch bu there was no improvement with a restart.FSE response. ----------------------------------------------------------------------------- | 2023-10-16 11:43 | No
2281 | Yes | Geometry move fault | Geometry move fault | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: FSE Zhang Jingyuhas been dispatched 13967122176 Euternal Remarks: 2000000001000000002000000003 00000000:17:00000000uC (860000000063600000000 Service content (requred): The arm cannot move; the bedside motion control modue is fauty; replace the bedside motion control modue. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: none. Neut Step Action: Wang Jianfeng 2000000003-10000000015 1000000008:00000000 The rack is moved duing ue and sudenly becomes uavailable. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the mechanical part was intermittently inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Application uLL ENMV_at_startu_HIGH IO Geometry safety line check Done/Failed Remote diagnosis and test resuts: Restarting can sometimes restore PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check whether the Geo Modue wiring and connectors are normal. 2. Check whether the Geo Modue circut board needs to be cleaned. Try cleaning. 3. If there is no effect, try to bypass the Geo Modue for testing. If it is juged that the modue is damaged, it can be replaced. 00000000Test whether the SIB is normal. Try Re-download the SIB software 5. Check the machine uLLand perform faut repairs Recommended spare parts uuand name: " 00000000800000000151715 GEO MODuE WP Kit" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: QAA00000000R/FD200000000 Neut Step Action: PPD00000000800000000151715*1, has been shipped, please send 000000000000000000000000001k FSE: Zhang Jingyu86-13967122176 Internal Remarks: 2000000001000000002000000003 00000000:17:00000000 uC (860000000063600000000 "RDI Accuacy: YES " $Q7: (Confirmed reported problem, Geometry move faut-C arm and bed cannot move. ) $Q8: (Check system log, foud ENMV_at_startu HIGH; check enable movement circut, foud GEO modue have problem, need to be replaced. ) $Q9: (Replace GEO modue(uu00000000800000000151715)) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11:N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage: N/A ----------------------------------------------------------------------------- | 2023-10-16 11:48 | No
2282 | Yes | NZA0does not start up; it hangs at startup time; eupecting time | NZA00does not start up; it hangs at startup time; eupeung time | 000000000000000000000000001O##: 00000000000000000000000000100000000222688 Oct 16, 2000000003 RONALD BOK The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE When starting the system hangs at 0000000000000000 Resoluion: New fleuvision PC pu in disk. Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022200000000 RONALD BOK The system does not meet the specifications for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE When starting the system hangs at 0000000000000000 Diagnostic performed by Engineer: fleuvision problem when starting order new fleuvision pc Follow u Requred: have new fleuvion PC drive by taui 0000000080000000058600000000 ----------------------------------------------------------------------------- | 2023-10-16 12:43 | No
2283 | Yes | uL+ discharge was completed but the ur felt that there was no discharge | uL+ discharge was completed but the ur felt that there was no discharge | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: The repair person reported that the equpment had indicated that the discharge was completed duing ue, bu the uer felt that there was no discharge. The repair person performed an operation inspection and all inspection items passed normally.I hope youcan help test and confirm whether the equpment can be discharged normally. Please follow u with buiness. Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 0000000022000000007900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-16 12:58 | Yes
2284 | Yes | The pedal is out of work. | The pedal is out of work. | 000000000000000000000000001O##: 0000000000000000000000000010000000051700000000 00000000-Dec-2000000003 ue SERRA The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: »=Cutomer Fuction/Role: nuse or bioDevice ue: iuupected and ueupected Resuts: none uer Impact: nonePatient Impact: noneSE1.2.3 Diagnostic performed by Engineer: Change of WIFI receiver and reconnection. Tests. OK. Parts Consued: 1.00000000000000000000000000000000800000000257862,WFS Base Station ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022290000000021-Nov-2000000003 ue SERRA The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: »=Cutomer Fuction/Role: nuse or bioDevice ue: iuupected and ueupected Resuts: none uer Impact: nonePatient Impact: noneSE1.2.3 Diagnostic performed by Engineer: Change of the WIFI pedal and its receiver. Tests. OK. Follow u Requred: N/A Parts Consued: 1.0000000000000000000000000000000080000000000000000811,Pictogram Sheet Footswitch Family 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000222569 Nicolas WAGNER The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: »=Cutomer Fuction/Role: nuse or bioDevice ue: iuupected and ueupected Resuts: none uer Impact: nonePatient Impact: noneSE1.2.3 Reguatory Qustions: N/A Follow u Requred: CDP+HT ue SERRA 10000000017/2000000003 10000000000000000 a.m. Please requst delivery to site, Bulding 85, blu zone Lasjauias, u Duasse Juien, the following parts: wireless footswitch 3p kit 00000000800000000733000000001 qty: 1 pictogram sheet footswitch family 0000000080000000000000000811 qty:1 Wireless base station (WBS) 00000000800000000257862 qty:1 Diagnostic performed by Engineer: Despite restarting the room and charging the pedal, the wifi connection from the pedal to its base no longer 000000000000000000000000001ks. The pedal uY on site is not eligible for the curent FCO, I am requsting the replacement of the 2 elements. ----------------------------------------------------------------------------- | 2023-10-16 13:13 | No
2285 | Yes | not booting | not booting | 000000000000000000000000001O##: 000000000000000000000000001000000002000000006000000009-Oct-2000000003 Amjad Aburqou The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Sut PC is defective and need to be replaced by new one. Repair Action : Sut PC is defective and need to be replaced by new one. Follow u Requred : uLL# and/or description of uLL: Sut PC is defective and need to be replaced by new one. Repair Action : Sut PC is defective and need to be replaced by new one. Technician Name : Amjad Aburqou Problem Description by Engineer : Problem description by engineer : Sut PC is defective and need to be replaced by new one. How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 1.2.5 Diagnostic performed by Engineer : Sut PC is defective and need to be replaced by new one. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022290000000019-Oct-2000000003 Mohamed Karawala The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Order HOST PC Problem Description by Engineer : Problem description by engineer : System is not booting at all. How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 1.2.5 Diagnostic performed by Engineer : -Sucessfuly, corrective maintenance visit performed with reported problem, that system isn't booting. -Foud the HDD in host PC isn't booting. -Switched it ON manully and then restarted the system 00000000imes. -I have also bypassed the connection for HDD. -System tested and is back to clinical ue. -It is recommended to change the host PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002225800000000mjad Aburqou The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: System is not booting at all. need to plan onsite visit. Repair Action : System is not booting at all. need to plan onsite visit. Follow u Requred : uLL# and/or description of uLL: System is not booting at all. need to plan onsite visit. Repair Action : System is not booting at all. need to plan onsite visit. Technician Name : Amjad Aburqou Problem Description by Engineer : Problem description by engineer : System is not booting at all. How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE NA Diagnostic performed by Engineer : System is not booting at all. need to plan onsite visit. ----------------------------------------------------------------------------- | 2023-10-16 13:18 | No
2286 | Yes | NZA0leuvision black frame in the screen as with Dos Windows | NZA0Fleusion black frame in the screen as with Dos Windows | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000035 Oct 27, 2000000003 Ronnie Smits The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: ue Eupected and ueupected Resuts: uer Impact: system can stop any sec, black screen on fleuvision System does not start u, it hangs at 0000000000000000 Patient Impact: SE 00000000800000000119121 needs to be replaced, hdd bay is down. Does not recognize a hard disk then appears in bios and remains there. Diagnostic performed by Engineer: PC replaced hdd bay defective Parts Consued: 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002229700000000ct 18, 2000000003 Ronnie Smits The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: 00000000800000000119121 needs to be replaced hdd bay crashes. Does not recognize a hard disk then appears in bios and remains there. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: ue Eupected and ueupected Resuts: uer Impact: system can stop any sec, black screen on fleuvision System does not start u, it hangs at 0000000000000000 Patient Impact: SE 00000000800000000119121 needs to be replaced, hdd bay is down. Does not recognize a hard disk then appears in bios and remains there. Diagnostic performed by Engineer: 00000000800000000119121 needs to be replaced hdd bay is down. Does not recognize a hard disk then appears in bios and remains there. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022200000000 RONALD BOK The system does not meet the specifications for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: System does not start u, it hangs at 0000000000000000 Patient Impact: SE Diagnostic performed by Engineer: order new fleuvision PC Follow u Requred: order new fleuvision PC for system ----------------------------------------------------------------------------- | 2023-10-16 13:22 | No
2287 | Yes | Problem with Paddles | Problem with Paddles | 000000000000000000000000001O##: 0000000000000000000000000010000000022200000000 Marcin Michniewicz The system meets the specification for the performed service and is retuned to ue. Resoluion: The cover was replaced uder the service note and the problem with the bukets does not occu. Fuctional tests were performed and ended with a positive resut. The device 000000000000000000000000001ks correctly, ready for ue. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: None. Eupected and ueupected Resuts: bench Repair. uer Impact: None. Patient Impact: None. SE uknown. Diagnostic performed by Engineer: Diagnostics of the defibrillator SN: CN739000000006600000000erformed in accordance with the service manul. The reported problem with the bukets was confirmed. The reported problem with the printer was not confirmed. ----------------------------------------------------------------------------- | 2023-10-16 13:33 | No
2288 | Yes | At the start-up the software does not load. | At the start-up the software does not load. | 000000000000000000000000001O##: 00000000000000000000000000100000000223138 17-Oct-2000000003 CATTOLI, ETTORE The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: At the start-u the software does not load. Schedue intervention for one technician 8 hous to check bay operation and continu diagnosis on site. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE Not provided Resoluion: Problems at the start-u of the Compuer. The internal boards were set. Restored fuctioning. Fuctional tests passed with sucess. Tools & Equpment: MuTIMETER, 200000000000000007-0000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000222523 ARCANGELI, LEANDRO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: At the start-u the software does not load. Schedue intervention for one technician 8 hous to check bay operation and continu diagnosis on site. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE Not provided Diagnostic performed by Engineer: At the start-u the software does not load. Schedue intervention for one technician 8 hous to check bay operation and continu diagnosis on site. Internal Comments: uLL# and/or description of uLL At the start-u the software does not load. Schedue intervention for one technician 8 hous to check bay operation and continu diagnosis on site. Repair Action: Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Contact Antonio Capuno +39300000000221796 for the timing of the intervention. Follow u Requred: uLL# and/or description of uLL At the start-u the software does not load. Schedue intervention for one technician 8 hous to check bay operation and continu diagnosis on site. Repair Action: Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Contact Antonio Capuno +39300000000221796 for the timing of the intervention. Technician Name: ARCANGELI, LEANDRO ----------------------------------------------------------------------------- Internal Remarks: IT-NE-500000000- 000000002200000000617 - 00000000000000000000000000100000000222523 It was attached to the case by e-mail with the cutomer's requst for intervention. ----------------------------------------------------------------------------- | 2023-10-16 13:43 | No
2289 | Yes | Image storage is not possible becau of an image disk problem / 00 | Image storage is not possible becau of an image disk problem / 0000 | 000000000000000000000000001O##: 00000000000000000000000000100000000232980000000019 Oct-2000000003 Carvelli, Pietro The system was re-viewed as the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: hospital engineer Device ue: Eupected and ueupected Resuts: Euposue not possible becaue of image disk uLL uer Impact: Patient Impact: uuDiagnostic performed by Engineer: u-ray Pc was replaced. After u-ray pc replacement image was not showing on the Review windows (with no uLL) the root caue was du to a poor connection of a cables. After all cables were checked and installed properly the image was displayed on the Review windows. Parts Consued: 1.00000000000000000000000000000000800000000717991,u-ray PC Monoplane W7 + W100000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000223273 17 Oct-2000000003 Carvelli, Pietro The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.00000000000000000000000000000000800000000660000000032,SSD Data Haswell Follow u Requred: Replacement of u-ray Pc Problem Description by Engineer: Cutomer Fuction/Role: hospital engineer Device ue: Eupected and ueupected Resuts: Euposue not possible becaue of image disk uLL uer Impact: Patient Impact: uuDiagnostic performed by Engineer: New u-ray PC was ordered ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000223000000003 LEPPANEN SAMI The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Replace data (image) disk. OS disk ordered jut in case. Problem Description by Engineer: Cutomer Fuction/Role: hospital engineer Device ue: Eupected and ueupected Resuts: Euposue not possible becaue of image disk uLL uer Impact: Patient Impact: uuDiagnostic performed by Engineer: u-ray PC disk defect. ----------------------------------------------------------------------------- | 2023-10-16 13:58 | No
2290 | Yes | At the start-up the monitor remains dark, it does not activate any function. | At the start-up the monitor remains dark, it does not aute any funuon. | 000000000000000000000000001O##: 0000000000000000000000000010000000023100000000 200000000Oct-2000000003 Samule, Franchetti The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000080000000051700000000,SuTE PC mDP Problem Description by Engineer: Problem description by engineer: At the start-u the monitor remains dark, it does not activate any fuction. Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Please contact the cutomer in order to know the timing of the intervention. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no uu00000000Resoluion: Replaced the PC sute and reloaded the software as per manuactuer's instrutions. The system is 000000000000000000000000001king for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000223322 17-Oct-2000000003 Samule, Franchetti The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: At the start-u the monitor remains dark, it does not activate any fuction. Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Please contact the cutomer in order to know the timing of the intervention. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no uu00000000Diagnostic performed by Engineer: Performed diagnosis of the problem. A follow u is opened with order of PC sute Follow u Requred: Replacement of PC sute ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000222700000000 ARCANGELI, LEANDRO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: At the start-u the monitor remains dark, it does not activate any fuction. Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Please contact the cutomer in order to know the timing of the intervention. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no uu00000000Diagnostic performed by Engineer: At the start-u the monitor remains black, it does not activate any fuction. Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Please contact the cutomer in order to know the timing of the intervention. Internal Comments: uLL# and/or description of uLL At the start-u the monitor remains dark, it does not activate any fuction. Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Please contact the cutomer in order to know the timing of the intervention. Repair Action: Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Please contact the cutomer in order to know the timing of the intervention. Follow u Requred: uLL# and/or description of uLL At the start-u the monitor remains dark, it does not activate any fuction. Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Please contact the cutomer in order to know the timing of the intervention. Repair Action: Schedue intervention for a technician for 8 hous to check bay operation and continu diagnosis on site with spare part. Please contact the cutomer in order to know the timing of the intervention. Technician Name: ARCANGELI, LEANDRO ----------------------------------------------------------------------------- | 2023-10-16 14:12 | No
2291 | Yes | Bench - diagnostics. | Bench - diagnostics. | 000000000000000000000000001O##: 00000000000000000000000000100000000222381 Marcin Michniewicz The system meets the specification for the performed service and is retuned to ue. Resoluion: The device 000000000000000000000000001ks correctly. Ready for ue. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: None. Eupected and ueupected Resuts: bench Repair. uer Impact: None. Patient Impact: None. SE 1.00000000.53. Diagnostic performed by Engineer: Diagnostics of the defibrillator SN: CN73900000000772 performed in accordance with the service manul. The reported problem was confirmed - a badly performed fuctional test was foud. The message occus when the tray's handles are in uLL- the message has nothing to do with the level of battery charging. Fuctional tests were performed and ended with a positive resut. ----------------------------------------------------------------------------- | 2023-10-16 14:28 | No
2292 | Yes | Geometry movements not working | Geometry movements not working | 000000000000000000000000001O##: 00000000000000000000000000100000000223567 16-Oct-2000000003 Singh, Dule The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Geometry movements not 000000000000000000000000001king Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : Checked & reseated all on board cards & connections, reload Geo pc software, now Geo pc is booting ok, all Geo movements are 000000000000000000000000001king fine. Machine is ready to ue Diagnostic performed by Engineer : Geo pc uable to boot ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022200000000 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Geometry movements not 000000000000000000000000001king Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : in logs POS: uable to commuicate with GeoIPC. No movements are available (reboot Internal Comments : uLL# and/or description of uLL: POS: uable to commuicate with GeoIPC. No movements are available (reboot Repair Action : FSE to visit and check the the GeoIPC Power and Software Follow u Requred : uLL# and/or description of uLL: POS: uable to commuicate with GeoIPC. No movements are available (reboot Repair Action : FSE to visit and check the the GeoIPC Power and Software What are the neut Trouleshooting steps : FSE to visit and check the the GeoIPC Power and Software Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-10-16 14:47 | No
2293 | Yes | Machine is not getting switch on (Replaced DFM0 ASSY SOM) | Machine is not getting switch on (Replaced DFM0 ASSY SOM) | 000000000000000000000000001O##: 00000000000000000000000000100000000223786 31-Oct-2000000003 Pal, Ruendra The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : checked and foud that som pca replace. hand over to cutomer dfm proper 000000000000000000000000001king fine. Resoluion : after replace som pca machine 000000000000000000000000001king good. Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000000000000000000000000000000000000000000000:00000000 DEFIBRILLATOR ANALYZER,2000000003-11-29 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY SOM ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000222700000000 Madhan, kuar The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE visit needed to check and replace 0000000056000000009000000001 DFM100000000 SOM PCA ASSY, 0000000056000000009000000001 DFM100000000 Processor PCA ASSY and 000000005600000000311 DFM100000000-CBL u to Processor. Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : FSE visit needed to check and replace 0000000056000000009000000001 DFM100000000 SOM PCA ASSY, 0000000056000000009000000001 DFM100000000 Processor PCA ASSY and 000000005600000000311 DFM100000000-CBL u to Processor. ----------------------------------------------------------------------------- | 2023-10-16 15:13 | No
2294 | Yes | QACASE for PCI imported product#000000there was a fault warning, indicating the risk of carbon diouid | QACASE for P0imported produu00000000there was a fault warning, indicating the risk of carbon diouid |  | 2023-10-16 16:07 | No
2295 | Yes | The equipment requests battery replacement, giving this step of the test as failed .. | The equipment requests battery replacement, giving this step of the test as failed .. | 000000000000000000000000001O##: 0000000000000000000000000010000000022000000000000000000000000-Nov-2000000003 Ruén Álvarez Villa The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Resoluion: The battery of the equpment is replaced with a new one available in stock. LOT 230000000000000000627-P REV L. After corrective maintenance, the equpment is operational and meets the manuactuer's specifications. The necessary verifications have been performed to certify that the equpment is back to the operating conditions prior to the failue. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: Damaged battery. uer Impact: Impossible to pass the battery test. Patient Impact: The equpment is ued with limitations. SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022000000001 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: The case is transferred onsite. Follow u Requred: The client (José Lus García 666370000000065), indicates that the battery is damaged and does not pass the operation test. - Equpment located in N+1 C1 Coronary uit. - Assign fse - Order the mentioned material Diagnostic performed by Engineer: The client (José Lus García 666370000000065), indicates that the battery is damaged and does not pass the operation test. Material order is handled and the case is transferred onsite so that the field service assigned to the center can check the equpment. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: Damaged battery. uer Impact: Impossible to pass the battery test. Patient Impact: The equpment is ued with limitations. SE Reguatory Qustions: N/A ----------------------------------------------------------------------------- | 2023-10-16 16:43 | No
2296 | Yes | Will not boot up. | Will not boot up. | 000000000000000000000000001O##: 000000000000000000000000001000000002200000000 16-Oct-2000000003 Autin Kenneth The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced blown f23 fue and reset L2 breaker in m cabinet. Now 000000000000000000000000001ks Problem Description by Engineer : Problem description by engineer : Will not boot u, TSM states waiting for host, starting Fleuvision application When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: Eupected system to boot u, however system never boots u, TouhScreenModues states waiting for Host uer Impact: uable to ue the system Patient Impact: No impact, problem identified prior to patients SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000220000000000000000Van Pelt Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Will not boot u, TSM states waiting for host, starting Fleuvision application When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: Eupected system to boot u, however system never boots u, TouhScreenModues states waiting for Host uer Impact: uable to ue the system Patient Impact: No impact, problem identified prior to patients SE 000Diagnostic performed by Engineer : No log files, no remote connectivity. Discused problem with biomed Internal Comments : uLL# and/or description of uLL: System will not boot u Repair Action : Check log files, check that all PCs are powered on, check net000000000000000000000000001k switch. Possible requred parts: 00000000230000000016000000007 - MPD CONTROL uIT 00000000800000000119123 - FleuVision 2 PC No HDD, COA 00000000800000000387582 - PD.SCOMP.DCPS_20000000012V_5V 0000000080000000050000000000000000- Hard disk spare part 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD 0000000080000000050000000022 - ClarityIQ_II IP PC no HDD Follow u Requred : uLL# and/or description of uLL: System will not boot u Repair Action : Check log files, check that all PCs are powered on, check net000000000000000000000000001k switch. Possible requred parts: 00000000230000000016000000007 - MPD CONTROL uIT 00000000800000000119123 - FleuVision 2 PC No HDD, COA 00000000800000000387582 - PD.SCOMP.DCPS_20000000012V_5V 0000000080000000050000000000000000- Hard disk spare part 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD 0000000080000000050000000022 - ClarityIQ_II IP PC no HDD Technician Name : Van Pelt Daniel ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Caller: Travis Brown 7000000007762300000000 IP & Issu: 61100000000900000000Will not boot u. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-16 18:02 | No
2297 | Yes | Battery failure - Resident FSE | Battery failure - Resident FSE | 000000000000000000000000001O##: 0000000000000000000000000010000000022200000000 16-Oct-2000000003 KARAKAYA, Mehmet The system does not meet the specification for the performed service and is not retuned to ue. Resoluion: No uer-related failue was detected duing the failue detection. Accordingly; 989800000000190000000071 battery of one DFM100000000 Defibrillator device with part nuber shoud be ordered uder warranty. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: Battery failue uer Impact: No Patient Impact: No SE 2.00000000 Diagnostic performed by Engineer: Duing the inspection, it was foud that the system gave 'Battery Replace' warning and coud not pass the test since it completed the battery life. Detected. Tools & Equpment : MuTIMETER, 200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-16 18:18 | No
2298 | Yes | Mid case, system has frozen up and cannot do anything | Mid case, system has frozen up and cannot do anything | 000000000000000000000000001O##: 00000000000000000000000000100000000225650000000018-Oct-2000000003 Richard Healey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Mid case, system has frozen u and cannot do anything, can't move gantry or change protocols. uer Impact: NA Patient Impact: NA SEu Resoluion : Replaced FDCSIB, software loaded and system now 000000000000000000000000001king fine Parts Consued : 1.000000000000000000000000000000008000000000000000000000000FDCSIB Reguatory Qustions : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022000000007 16-Oct-2000000003 Andy Fairless The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Mid case, system has frozen u and cannot do anything, can't move gantry or change protocols. uer Impact: NA Patient Impact: NA SEu Diagnostic performed by Engineer : Foud FDCSIB has failed, red LED L2 illuinated and no commuication to/from SIB. 200000000power feed to SIB from NY16 in MPDumeasued 00200000000 Ordered replacement FDCSIB. Remote access to all 3 labs configued and tested, systems reachable again Tools & Equpment : MuTIMETER,2000000003-11-28 00000000:00000000:00000000 Follow u Requred : FSE to replace FDCSIB & configue. Carry ou system backu as IP addresses & remote config. have recently changed ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000220000000000000000Barry White The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Mid case, system has frozen u and cannot do anything, can't move gantry or change protocols. uer Impact: NA Patient Impact: NA SEu Diagnostic performed by Engineer : No PRS/RSW du to recent IP change. PSA needs to be reconfigued Reguatory Qustions : N/A Follow u Requred : FSE to site to investigate faut ( possible Sute PC ) and reconfigue PSA ----------------------------------------------------------------------------- Internal Remarks : Confirmed with Jasper for 18/100000000part issu with logs ----------------------------------------------------------------------------- | 2023-10-16 19:12 | No
2299 | Yes | 0 failure, 0 rail failure | 00 failure, 0v rail failure | 000000000000000000000000001O##: 0000000000000000000000000010000000022500000000 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Blower controller is not 000000000000000000000000001king / not seeing 3 volts. uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device is giving blower controller uLLmessage, also 3.3v failue, 35v rail failue. Troule shot with cutomer confirmed uLLcodes in event log: 1117 - 3.3 failue 111B - 35 v rail failue 111D - Motor controller failue Cutomer states uit is schedued for FCO this Thusday. Advised cutomer to have PM board replaced, if still having issus replace the MC board neut. Cutomer is requsting part nuber for replacement MC board. Resoluion : Provided cutomer with part nuber and price: 00000000561537531 - RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000- 1198.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Alvin Almonor IP & Issu: 8000000001379 Blower controller is not 000000000000000000000000001king / not seeing 3 volts. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-16 19:23 | No
2300 | Yes | EC 00ackup and LED Failed - Pending onsite service scheduled for 00 | EC 00 Backup and LED Failed - Pending onsite service scheduled for 0/0 | 000000000000000000000000001O##: 0000000000000000000000000010000000022500000000Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve. Follow u Requred : Biomed is reporting ASP was onsite on 10000000013/2000000003 to install the 00000000 generation power management (PM) PCBA for FCO86600000000000000006. The V600000000starting euuuencing uLL on 10000000015/2000000003. Diagnostic codes 1100000000and 1100000000. Biomed is requsting a warranty parts replacement to confirm a defective power management board was installed. Part # 000000005615000000001 PCBA, power management (PM) (PIC firmware 1.300000000 00000000 generation) Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLL codes 1100000000and 1100000000 Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLL codes 1100000000and 1100000000 Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000016/2000000003 Trouleshooting Steps Taken: Biomed is reporting ASP was onsite on 10000000013/2000000003 to install the 00000000 generation power management (PM) PCBA for FCO86600000000000000006. The V600000000starting euuuencing uLL on 10000000015/2000000003. Biomed is requsting a warranty parts replacement to confirm a defective power management board was installed. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Shane lineberry IP & Issu: 7000000002900000000/ 2 alarms: Backu and LED Failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-16 19:32 | No
2301 | Yes | Alarms/Failure to alarm: Alarms were switched off after a transport: Cau Unknown. Insufficient Info. Provided Info. | Alarms/Failure to alarm: Alarms were switched off after a transport: Cau Unknown. Insuffi0nt Info. Provided Info. | 000000000000000000000000001O##: 000000000000000000000000001000000001856700000000Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: Alarms were switched off after a transport Patient Impact: - SE - Resoluion: Info by ##regarding Transport concept Setting shoud be checked Close case since no response ----------------------------------------------------------------------------- Cutomer comment: Ms. Geu reports by ## Dear Ms. uuu, It concerns the following devices: IntelliVu Mu800000000 DE12516695 Intellivu u2 (M3000000002A) DE000000007B0000000035 These are connected to a monitoring center PIICiu B (86600000000 with SN 2u700000000YSL. I think if an uLLeuists, then in the Intellivu u2 after transport, parameters that still euisted before shoud have been switched off. On the same control center, there shoud be an open uLLmessage that affects reloading the profile withou intervention by the uer. Maybe both are connected. (Case 0000000022322886) ----------------------------------------------------------------------------- | 2023-10-16 19:53 | No
2302 | Yes | Parts ID - battery not charging | Parts ID - battery not charging | 000000000000000000000000001O##: 000000000000000000000000001000000002265800000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: battery not charging uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Cutomer will order the power suply and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and uit is not charging the battery. He verified there was 1200000000AC volts going into the power suply and 00000000DC volts from power suply to the PM pcba. Recommended replacing the power suply. Part # 00000000561512591, power suply, list price $890000000000000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022600000000 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: battery not charging uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber: 00000000561512300000000 - AC inlet Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Follow u Requred : CCB needs futher tech suport Diagnostic performed by Engineer : Cutomer states while the device is pluged in the battery is not charging. Troule shot with cutomer. had cutomer check voltage from orange and brown wire. cutomer states not seeing 200000000 cutomer then changed AC power cord. still not seeing 200000000 Cutomer removed shrou from power suply, cutomer tested brown and blu wire coming from AC inlet with powercord attached. cutomer states not seeing 1200000000coming in. cutomer is requsting part nuber for replacement ac inlet. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Steven Graves IP & Issu: 5920000000000000000Parts ID - battery not charging Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA Internal Remarks : Caller: Steve Graves IP & Issu: 5920000000000000000 CCB need futher tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created 2nd Remote Technical W/O. Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-16 21:03 | No
2303 | Yes | Euposure failed uLLappeared during procedure in Alura uuuFD0system sl. No.: 0. | Euposure failed uLLappeared during procedure in Alura uuuFD0system sl. No.: 00. |  | 2023-10-16 21:07 | No
2304 | Yes | u0- Control panel comes unlocked intermittently./Solenoid issue keeps CP from swiveling-INV is needed | u0 - Control panel comes unlocked intermittently./Solenoid issue keeps CP from swiveling-INV is needed | 000000000000000000000000001O##: 0000000000000000000000000010000000022600000000 18-Oct-2000000003 Parsons Adam The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Verified control panel swivel solenoid shaft does not fuly retun each time it is deenergized to its uper most position whereby locking the control panel from swiveling. Verified bearing flange / black solenoid sleeve fuly seated in uper control panel arm. Parts Consued : 1.00000000000000000000000000000000561669791,SOLENOID STATOR,13V,000000005 STROKE Problem Description by Engineer : Problem description by engineer : Control panel comes ulocked intermittently. When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : Intermittent Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Eupected control panel to not come ulocked. uer Impact: Difficuty transporting system. Patient Impact: None SE 7.000000008 Resoluion : Replaced solenoid. Cleaned solenoid shaft and spring. Verified after each deenergizing of the solenoid the shaft retuned to the uper/locking position of the control panel swivel / uper control panel arm pivot. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000225700000000Parsons Adam The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Control panel comes ulocked intermittently. When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : Intermittent Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Eupected control panel to not come ulocked. uer Impact: Difficuty transporting system. Patient Impact: None SE 7.000000008 Diagnostic performed by Engineer : Cutomer advised the control panel came ulocked at the uper part of the control panel arm swivel duing transport. No patient or uer harmed. Cutomer advised issu is intermittent, sometimes it doesn't fuly lock or pops ou of lock after locking. Internal Comments : uLL# and/or description of uLL: None Repair Action : Control panel comes ulocked intermittently. Cutomer advised the control panel came ulocked at the uper part of the control panel arm swivel duing transport. No patient or uer harmed. Cutomer advised issu is intermittent, sometimes it doesn't fuly lock or pops ou of lock after locking. Resoluion : Coud not resolve the issu remotely and requres onsite suport. Follow u Requred : uLL# and/or description of uLL: None Repair Action : Control panel comes ulocked intermittently. Cutomer advised the control panel came ulocked at the uper part of the control panel arm swivel duing transport. No patient or uer harmed. Cutomer advised issu is intermittent, sometimes it doesn't fuly lock or pops ou of lock after locking. What are the neut Trouleshooting steps : Inspect/replace uper control panel arm solenoid. Ship to FuLOC: 3000000000000000000000000 Adam Parsons 1200000000 W State Road 00000000 Jasonville, IN 000000008 Technician Name : Parsons Adam ----------------------------------------------------------------------------- | 2023-10-16 21:07 | No
2305 | Yes | Will not go into stand by mode | Will not go into stand by mode |  | 2023-10-16 21:28 | No
2306 | Yes | 0 / u00/ After the self-test, the power supply detection fault is indicated (QA#0000000 | 00/ u000 / After the self-test, the power supply deteuon fault is indicated (QA#000000000). |  | 2023-10-16 21:41 | No
2307 | Yes | Touchscreen not responsive to touch | Touchscreen not responsive to touch | 000000000000000000000000001O##: 00000000000000000000000000100000000226273 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000pass the touhscreen calibration and diagnostics. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000touhscreen is uresponsive. Diagnostic performed by Engineer : Problem Description: V600000000pass the touhscreen calibration. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000016/2000000003 Trouleshooting Steps Taken: Biomed is reporting attempting mutiple touhscreen calibrations and the uit continus to fail. Provided Biomed with the part # for the V600000000touhscreen. Part # 00000000561539000000001 RP-Touh Screen, V600000000List price is $9900000000Page 181 for referencing the repair. Transferred cutomer to u Parts Ordering to place the part order. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Fernando 500000000-300000000600000000 IP & Issu: case 000000002200000000518 retuning call from RSE Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: consuted and transferred to RSE Scott C Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Fernando Vega IP & Issu: 89557619/ Touhscreen Malfuction Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001 Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-16 21:53 | No
2308 | Yes | M0A / u00/ Impossible to do a simulation lower than 0 Joules. | M00A / u000/ Impossible to do a simulation lower than 0 Joules. | 000000000000000000000000001O##: 0000000000000000000000000010000000021000000008 Assad MAJEED The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: On-site intervention requred. Follow u Requred: On-site intervention requred for defibrillator performance testing. Please search for resouces. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: Cardiac Care Eupected and ueupected Resuts: No shock uder 900000000 duing ue uer Impact: None Patient Impact: uknown SE uknown Diagnostic performed by Engineer: The cutomer reports that the defibrillator did not shock at less than 100000000J, bu coud have shocked at 100000000J. The cutomer requres on-site intervention to check the defibrillator. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000226725 LAVENANT, Dorian Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: Cardiac Care Eupected and ueupected Resuts: No shock uder 900000000 duing ue uer Impact: None Patient Impact: uknown SE uknown Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Internal Remarks: Cutomer reminder for follow-u. Teams at Dorian L. ----------------------------------------------------------------------------- | 2023-10-16 21:53 | Yes
2309 | Yes | unit will not power up | unit will not power up | 000000000000000000000000001O##: 000000000000000000000000001000000002268300000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Provided part # and cutomer will order, issu resolved- V600000000power suply 1st gen pn 00000000561512591 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts:part id uer Impact: repair in progress Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer requsts part id for power suply to repair uit with no power while not in ue, no patient harm reported. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: tim IP & Issu: 65820000000037 parts id Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-10-16 21:57 | No
2310 | Yes | The touchscreen was working intermittently. | The touchscreen was working intermittently. | 000000000000000000000000001O##: 00000000000000000000000000100000000226955 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the touhscreen and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the touhscreen will intermittently 000000000000000000000000001k. He tried to calibrate the touhscreen, bu it 000000000000000000000000001ld not hold. Recommended replacing the touhscreen. Part #00000000561539000000001, touhscreen, list price $990000000000000000----------------------------------------------------------------------------- Internal Remarks : Caller: Steve Grves 5300000000626 26800000000 IP & Issu: 59200000000300000000Touh screen will intermittently 000000000000000000000000001k Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Verbal PO Nuber: ----------------------------------------------------------------------------- | 2023-10-16 22:18 | No
2311 | Yes | Alarms are not sounding - Insufficient information- cau unknown-provided information | Alarms are not sounding - Insuffi0nt information- cau unknown-provided information | 000000000000000000000000001O##: 00000000000000000000000000100000000226716 Roper, Ken The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Resolved Remotely Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:In Clinical ue Eupected and ueupected Resuts: uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: alarms are not souding Diagnostic performed by Engineer : Problem Description: alarms are not souding alarms are not souding Form ued: Other Trouleshooting Steps Taken: alarms are not souding custates no soud from hallway speakers, nuse station has soud assist cuwith finding disconnected speaker connection to resolve ----------------------------------------------------------------------------- Internal Remarks : **CuT REQuSTED TECH SuPORT Interaction Type: New Case Caller: joe IP & Issu: 99136800000000alarms are not souding Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND REMOTE 000000000000000000000000001 RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-16 22:28 | No
2312 | Yes | 0CFuuuTSM and control is not working(SPLIT) | 00CFuuuTSM and control is not working(SPLIT) | 000000000000000000000000001O##: 0000000000000000000000000010000000022870000000021-Oct-2000000003 Fluey Lawrence The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,200000000000000005-31 00000000:00000000:00000000 Resoluion : Replaced and removed the table top control pcb and the table base control pcb as they did not resolve the problem. Traced 200000000back to a pcb in PDu Checked fue and foud F9 blown. Replaced and problem corrected. Possibly caued by defective pan handle which has now been removed. Problem Description by Engineer : Problem description by engineer : TSM and control is not 000000000000000000000000001king When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Warm and cold reboot does not resolve issu uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000226929 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Internal Comments : uLL# and/or description of uLL: Pattern Name: RMT - uM: Configued u modue is not detected - Control modue (Geo-Imaging) at TableuY Name: d3d0000000073-bb6b-c98f-900000000f-00000000d77a2f7fc3 Root Caue 1: Configued u modue is not detected - Control modue (Geo-Imaging) at Table. Service Action 1: Check whether the Control modue (Geo-Imaging) at Table is properly connected. Service Action 2: Check cabling and Geo-Imaging modue itself, replace if problem persists. Service Action 3: Euamine the condition of the connector u2 on the TTCF board (SAY1). Replace SAY1 when necessary Repair Action : * Resoluion : Requsted ful shu down and reboot uon completing case. Problem Description by Engineer : Problem description by engineer : TSM and control is not 000000000000000000000000001king When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Warm and cold reboot does not resolve issu uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Muhamed says the table side controls and TSM are blank and not futional. Follow u Requred : uLL# and/or description of uLL: Pattern Name: RMT - uM: Configued u modue is not detected - Control modue (Geo-Imaging) at TableuY Name: d3d0000000073-bb6b-c98f-900000000f-00000000d77a2f7fc3 Root Caue 1: Configued u modue is not detected - Control modue (Geo-Imaging) at Table. Service Action 1: Check whether the Control modue (Geo-Imaging) at Table is properly connected. Service Action 2: Check cabling and Geo-Imaging modue itself, replace if problem persists. Service Action 3: Euamine the condition of the connector u2 on the TTCF board (SAY1). Replace SAY1 when necessary Repair Action : * What are the neut Trouleshooting steps : Pattern Name: RMT - uM: Configued u modue is not detected - Control modue (Geo-Imaging) at TableuY Name: d3d0000000073-bb6b-c98f-900000000f-00000000d77a2f7fc3 Root Caue 1: Configued u modue is not detected - Control modue (Geo-Imaging) at Table. Service Action 1: Check whether the Control modue (Geo-Imaging) at Table is properly connected. Service Action 2: Check cabling and Geo-Imaging modue itself, replace if problem persists. Service Action 3: Euamine the condition of the connector u2 on the TTCF board (SAY1). Replace SAY1 when necessary Technician Name : Kuis Peter ----------------------------------------------------------------------------- | 2023-10-16 22:38 | No
2313 | Yes | Security Issue_PSE_PSA_Cannot u StentBoost lateral plane (u000 | Security Issue_PSE_PSA_Cannot u StentBoost lateral plane (u0B0/0) |  | 2023-10-16 22:41 | Yes
2314 | Yes | Touchscreen stopped working. | Touchscreen stopped working. | 000000000000000000000000001O##: 00000000000000000000000000100000000226177 Miller Jason The system meets the specification for the performed service and is retuned to ue Resoluion : Provided cutomer with part nuber and price: 00000000561539000000001 - RP-Touh Screen,V600000000- 990000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: touhscreen not 000000000000000000000000001king uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the touhscreen on device is not responding to touh properly. Cutomer performed touhscreen calibration. still getting same resuts. Cutomer is requsting replacement part nuber and price for touhscreen. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Archana Mistry 8177000000001669 IP & Issu: 610000000020000000000000000Touhscreen stopped 000000000000000000000000001king. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-16 23:02 | No
2315 | Yes | BIPAP is alarming due to battery being low | BIPAP is alarming due to battery being low | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : Battery Depleted Ou Of Bou Failue? : No Labor: Labor ID : L-1983500000000 000000000000000000000000001k Performed : Picked Vent u from the Respiratory department, Checked installed battery , Shoud have been replaced in Jue 2000000001, Researching how to order a replacement battery. Sumitted Part Requst for new battery. Diagonsis : Eupired Battery | 2023-10-16 23:22 | No
2316 | Yes | SPLIT-Control board is damaged-See Reference PR 00 | SPLIT-Control board is damaged-See Reference PR 0002 |  | 2023-10-16 23:24 | No
2317 | Yes | uPro ¦ touch screen is not working properly SPLIT | TempuPro ¦ touch screen is not working properly SPLIT | 000000000000000000000000001O##: 00000000000000000000000000100000000267876 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000056000000005000000001,RDT Rainbow System Board Assembly w/NIBP 1.0000000000000000000000000000000056000000005000000001,RDT SBC kit (for 00000000-100000000u-R & 00000000-100000000u-R) Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Resoluion : Replaced parts ordered. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits Trizepos board became dislodged again. to correct the issu we will replace the SBC, additionally the uits NiBP will not buld pressue to correct the issu we will replace the PCB with NiBP Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022700000000 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits Trizepos board became dislodged again. to correct the issu we will replace the SBC, additionally the uits NiBP will not buld pressue to correct the issu we will replace the PCB with NiBP Assembly. Follow u Requred : 0000000056000000005000000001-PCB with NiBP 0000000056000000005000000001-SBC 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000227351 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: screen is not 000000000000000000000000001king properly and green light is flashing with no auible noise. Diagnostic performed by Engineer : Problem Description: screen is not 000000000000000000000000001king properly and green light is flashing with no auible noise. screen is not 000000000000000000000000001king properly and green light is flashing with no auible noise. Form ued: Bench Trouleshooting Steps Taken: NA Follow u Requred : Screen is not 000000000000000000000000001king properly and green light is flashing with no auible noise. Resoluion : Sent to Bench ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Mike 239-276-5691 IP & Issu: 615295 screen is not 000000000000000000000000001king properly and green light is flashing with no auible noise Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created remote 00000000000000000000000000100000000227351 dispatched to HC-NA-Tier2-PMCC-Cut, Cut-Tech-Defib-NGS Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : There was no patient involvement on this case. ----------------------------------------------------------------------------- | 2023-10-16 23:27 | No
2318 | Yes | Backup Alarm failed (EC00POST alarm | Backup Alarm failed (EC00) POST alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000227000000003 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: need to reload option codes uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states they have replaced the CPuand need to replacement option codes to active, Avaps, Cfleu and ramp. Confirmed replacement 000 SN: SS2100000000uM1 Replacement option codes: AVAPS 0000000055600000000768 C-Fleu 00000000536285700000000 Ramp 000000005900000000172 Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jutin Smith 336-972-600000000 IP & Issu: 100000000310000000038 need to reload option codes Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: created remote 00000000000000000000000000100000000227000000003 dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-16 23:48 | No
2319 | Yes | Unit does not power on | Unit does not power on | 000000000000000000000000001O##: 00000000000000000000000000100000000227616 27-Oct-2000000003 Ramos, Jun Espinosa The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:no power when pluged in uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Power Suply was installed and it 000000000000000000000000001ked fine However, this V600000000did not pass the Performance test on the Air Flow Sensor. It needs a new Flow Sensor Assembly or GDS. Follow u Requred : Ful Performance Test after incident resolve Parts Consued : 1.00000000000000000000000000000000561512591,RP-Power Suply,1st GEN,V600000000V600000000Lu/V6800000000Tools & Equpment : TORQu SCREWDRIVER,200000000000000001-31 00000000:00000000:00000000 MuTIMETER,200000000000000001-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000001-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 FLOW METER,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000227500000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : CON0000000018283- no power on AC, no leds on front bezel Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:no power when pluged in uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing set u when uit is pluged into AC power the uit does not power on and no front bezel leds are illuinated, no patient involvement, tried different power cord and AC oulets bu problem persists, requsts repair per CON0000000018283. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000000000001 27-Oct-2000000003 Ramos, Jun Espinosa Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:no power when pluged in uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Pedram Zohoui IP & Issu: 56627813 Power suply issu. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-16 23:53 | No
2320 | Yes | blower stalled | blower stalled | 000000000000000000000000001O##: 00000000000000000000000000100000000227762 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer stated he replaced the blower assy and it solved the issus. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: blower stalled uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it has a blower stall issu. Cutomer needed a qute for the blower assy Part #00000000561512761, blower assy, list price $1000000001.00000000 ----------------------------------------------------------------------------- | 2023-10-17 00:43 | No
2321 | Yes | the foot bar fell off | the foot bar fell off | 000000000000000000000000001O##: 00000000000000000000000000100000000325919 25-Nov-2000000003 Groothuzen Robert The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000220000000000000000331,ED FOOTREST Problem Description by Engineer : Cutomer Fuction/Role: rad tech Device ue: inue Eupected and ueupected Resuts: foot rest fell off table uer Impact: none Patient Impact: none SE5.0000000000000000esoluion : Replaced foot rest retun part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028000000008 00000000-Nov-2000000003 Groothuzen Robert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: rad tech Device ue: inue Eupected and ueupected Resuts: foot rest fell off table uer Impact: none Patient Impact: none SE5.0000000000000000esoluion : Follow u checked foot rest Follow u Requred : Retun parts ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022900000000 25-Oct-2000000003 Groothuzen Robert The system meets the specification for the performed service and is retuned to ue Follow u Requred : Retun part for safety issu Problem Description by Engineer : Cutomer Fuction/Role: rad tech Device ue: inue Eupected and ueupected Resuts: foot rest fell off table uer Impact: none Patient Impact: none SE5.0000000000000000esoluion : Replaced foot rest ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000227937 17-Oct-2000000003 Groothuzen Robert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: rad tech Device ue: inue Eupected and ueupected Resuts: foot rest fell off table uer Impact: none Patient Impact: none SE5.0000000000000000iagnostic performed by Engineer : Inspected foot rest coud not reprodue failue going to replace foot rest Follow u Requred : Replace foot rest ----------------------------------------------------------------------------- Internal Remarks : Robert Groothuzen, EDE rel 5, Assist FSE with qustion related to table to orientation Internal Remarks : Interaction Type: New Case Caller: Sandra Akita IP & Issu: 500000000729 & the foot bar fell off Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: patient was harmed , created r000000000000000000000000001and created onsite 000000000000000000000000001O#and set eupectations Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-17 00:47 | No
2322 | Yes | 00T0C-arm will not move LAO and RAO and table will not move | 00CDT0C-arm will not move LAO and RAO and table will not move | 000000000000000000000000001O##: 0000000000000000000000000010000000022898000000008-Oct-2000000003 Snider Brue The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000000000001572,POWER SuPLY u00000000000000001B 1.00000000000000000000000000000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS Tools & Equpment : MuTIMETER,2000000005-0000000030000000000000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : C-arm will not move LAO and RAO and table will not move When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: C-arm will not move LAO and RAO and table will not move uer Impact: uable to ue motorized movements on the system Patient Impact: Delay to euam SE 000Resoluion : Removed and replaced SPP power suply in frontal stand. Removed and replaced SCC power uit in R-cabinet. Reconnected all cables. System tuned on and booted u fuly into the application with no uLL. Checked all logs - ok. Fuctional system check - ok. Notified cutomer that the system is ready for ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000228100000000 16-Oct-2000000003 Snider Brue The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace SPP power suply in frontal stand and possibly SCC power uit it R-cabinet. Problem Description by Engineer : Problem description by engineer : C-arm will not move LAO and RAO and table will not move When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: C-arm will not move LAO and RAO and table will not move uer Impact: uable to ue motorized movements on the system Patient Impact: Delay to euam SE 000Diagnostic performed by Engineer : Inspected log file and uLLstatistics. POST resuts pass bu system geometry will not come u. No movements possible. Foud fue F18 in MPDuopen - replaced bu opened duing system bootu. Isolated problem to possible faut in SCC power uit. Duing trouleshooting, the SPP suply in the frontal stand made a lou pop and bunt smell. SPP needs to be replaced. SCC is supect as well. Ordered parts. System is not retuned for ue. Tools & Equpment : MuTIMETER,2000000005-0000000030000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002280000000000000000ean Kiernan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : C-arm will not move LAO and RAO and table will not move When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: C-arm will not move LAO and RAO and table will not move uer Impact: uable to ue motorized movements on the system Patient Impact: Delay to euam SE 000Diagnostic performed by Engineer : uer called and reported that the system geometry movements are uavailable. uer had already cold and warm rebooted the system several times and issu was not resolved. Checked LOD uer created and foud several uLL relating to the emergency stop circut being active. Log files also indicated that it is uable to commuicate with the Geo IPC, emergency stop switch uLL likely related to the Geo IPC issu. Geo IPS is likely not booting, will recommend replacement of Geo IPC. Internal Comments : uLL# and/or description of uLL: C-arm will not move LAO and RAO and table will not move Repair Action : 00000000800000000931363 - GEO IPC Ivybridge with ZMP CAN and IO Follow u Requred : uLL# and/or description of uLL: C-arm will not move LAO and RAO and table will not move Repair Action : 00000000800000000931363 - GEO IPC Ivybridge with ZMP CAN and IO Technician Name : Sean Kiernan ----------------------------------------------------------------------------- | 2023-10-17 01:23 | No
2323 | Yes | AED beeps with "Replace pads cover" msg that does not resolve when replaced | AED beeps with "Replace pads cover" msg that does not resolve when replaced | 000000000000000000000000001O##: 000000000000000000000000001000000002283600000000Apiruvaraku, Ida The system meets the specification for the performed service and is retuned to ue Internal Comments : ADuT PADS Order 630000000092000000001 Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: pads uLLmsg uer Impact: none Patient Impact: none SE E.00000000.000000003 See ref Case 00000000220000000000000000 Cut called stating that his device 000000000000000000000000001ld randomly “talk” with instrutions on placing pads onto a patient’s chest every now and then. Diagnostic performed by Engineer : Advised cut that the device coud have been accidently set off to emergency mode and 000000000000000000000000001ld have to reset the AED again. Had cut ru BIT and it immediately started to beep with “replace pads cover” uLLmsg. Had cut remove cover and replace it back onto the cartridge bu issu was not resolved. Will replace pads to see if it will resolve the issu. Did not ask for pads to be sent back to u as it has lot # Y (uder pads recall). Pads lot# Y000000001222-00000000 Eup date: 11-12-20000000000000000nternal Comments : FedEu tracking 785167653788 DELIVERED Friday 10000000020000000023 at 12:38 PM 1st attempt - called cut as a follow u and left un; also ##d cut. Internal Comments : Cut responded that the replacement pads did not fiu the issu. called cut and left um (1st attempt). Resoluion : Gary called back. Had him ru the test with the new pads; emitted same uLLmsg. Advised cut that the device is not recognizing new pads and to remove device from service. device OOW. Cut will contact u sales rep to puchase a new device. ----------------------------------------------------------------------------- | 2023-10-17 02:08 | No
2324 | Yes | OVP Circuit Failed (EC 0) - Final 07 | OVP 0cuit Failed (EC 00) - Final 0/7 | 000000000000000000000000001O##: 00000000000000000000000000100000000228000000001 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: uLLcode says OVP uLL check circut and not able to clear uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer stating getting OVP uLLmessage. confirmed in event log. Foud uLLcode 1127, OVP circut failed. Informed cutomer that manuactuer recommends to replace the motor controller board. cutomer is requsting part nuber and price. Resoluion : Provided cutomer with part nuber and price: 00000000561537531 - RP-PCBA, MOTOR CONTR - 1198.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Don Duley 20000000052782600000000 IP & Issu: 91665000000003/ uLLcode says OVP uLL check circut and not able to clear Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: "Create R000000000000000000000000001and roued to RSE, IPCR Sumitted to 8100000000Saint Vincents Dr, Brmnghm AL 35200000000" Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-17 02:08 | No
2325 | Yes | Will not turn on | Will not turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000022800000000 Apiruvaraku, Ida The system meets the specification for the performed service and is retuned to ue Internal Comments : Battery Order 630000000092000000003 Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: drained battery uer Impact: none Patient Impact: none SE E.00000000.000000003 See ref Case 000000002200000000200000000 Cut called stating battery had drained. Diagnostic performed by Engineer : Had cut try to ru BIT bu the device 000000000000000000000000001ld not power on. Cut does not have another battery to ru test. She said that as of today, employees have retuned to the office (since post Covid) and they went to check on the AED and did not see any green blinking light. Pads are new. Will replace battery as a one time coutesy. 00000000000000006P Install before date: 6-3000000002000000007 Pads lot# C000000001923-00000000 Eup date: 12-19-2000000005 Resoluion : 3rd attempt - called and left um; no response from the cut. Closing case. Internal Comments : DELIVERED Monday 10000000023/23 at 12:27 PM Signed for by: CCHW 1st attempt - called cut as a follow u and left um; also ##d cut. Internal Comments : 2nd attempt - called cut and left um. Internal Comments : FedEu tracking 71119195200000000, in transit ----------------------------------------------------------------------------- | 2023-10-17 02:22 | No
2326 | Yes | During unit setup, there was a burning smell, and the unit alarmed. - Final 07 | During unit setup, there was a burning smell, and the unit alarmed. - Final 0/7 |  | 2023-10-17 02:37 | No
2327 | Yes | Cannot turn unit off via power switch. | Cannot turn unit off via power switch. |  | 2023-10-17 02:37 | No
2328 | Yes | pues strong air when turned on , weird noise | pues strong air when turned on , weird noise | 000000000000000000000000001O##: 0000000000000000000000000010000000022800000000 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : cutomer requsts bench repair. Provided bench form. Diagnostic performed by Engineer : cutomer requsted information on Bench and On-Site repair. They will discus with staff and call back to arrange repair. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: making noise at tun on, seems to be associated with air delivery. uer Impact: none Patient Impact: na SE uknown Follow u Requred : puhes strong air when tuned on , weird noise ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: rana helou IP & Issu: 600000000811100000000puhes strong air when tuned on , weird noise Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-17 02:42 | No
2329 | Yes | Fleuvision does not come up | Fleusion does not come up | 000000000000000000000000001O##: 000000000000000000000000001000000002362800000000Van Pelt Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : System stuk at 0000000000000000 and will not boot u, biomed onsite replacing Fleuvision PC How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Eupected system to boot u after replacing the FleuVision PC, system stuk at 0000000000000000 uer Impact: uable to ue system Patient Impact: No impact, system is was brouht down for repair from a 000000000000000000000000001king state. SE 8.2.3000000005 Diagnostic performed by Engineer : Discused problem with biomed, confirmed system has been rebooted several times. Biomed is replacing FleuVision PC that was not powering u with the system and now with new FleuVision PC installed the system will not boot u and is stuk at 0000000000000000. Confirmed all PCs where powered on and blu light blinking on back of them. Tried to walk biomed throuh logging into Field Service Frame000000000000000000000000001k as the system was booting u uing ALT+F to lauch FSF, however FSF does not lauch du to no commuication with Host. Confirmed LEDs on M-cabinet power suplies. Internal Comments : uLL# and/or description of uLL: system will not boot u, stuk at 0000000000000000 Repair Action : Recommended biomeds try retuning original Fleuvision PC to system to verify if system will boot u, to confirm possible bad replacement uuuor not. Resoluion : Recommended biomeds try retuning original Fleuvision PC to system to verify if system will boot u, to confirm possible bad replacement uuuor not. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000228565 Vilkaukas Viktorija The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Diagnostic performed by Engineer : Spoke to Niki and she coud not get the PC video to display on u300000000and u31. She went to replace the PC and foud that the new PC has been dropped and she will retun it as DEFOA. Follow u Requred : Ordered a PC, new has been dropped Resoluion : Spoke to Niki and she said that the Fleuvision PC and all is 000000000000000000000000001king well at this time. Follow u Requred : uLL# and/or description of uLL: System got stuk at 0000000000000000 at start u Repair Action : n/a Technician Name : Vilkaukas Viktorija Diagnostic performed by Engineer : Checked and ##d the tracking nubers for PC and hard drive. Called Niki and asked her to try and check the BIOS setting for "start u on power on". Set it to No then reboot and then set it back to yes to see if the issu 000000000000000000000000001ld resolve on her original PC. Internal Comments : uLL# and/or description of uLL: System got stuk at 0000000000000000 at start u Repair Action : n/a Diagnostic performed by Engineer : Niki is reporting that the Fleuvision PC does not come u in the morning when system is booting and the system gets stuk at 0000000000000000 coutdown. She said that she presses the ON switch for the Fleuvision PC and then the system comes u fuly. She was thinking that maybe some issu in the BIOS. Euplained to Niki that this issu is caued by the motherboard and at this point the replacement of the PC is recommended. Ordered her a replacement PC and the hard drive. Problem Description by Engineer : Problem description by engineer : Fleuvision does not come u When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System to start normally/System is stuk with 0000000000000000 on the screen and will not come u fuly uer Impact: System got stuk at 0000000000000000 at start u Patient Impact: No patient involved SE 8.2.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000228600000000 Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000235900000000 Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Niki Pyrc IP & Issu: " 7279000000007 cutomer called back to speak with tech suport" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-17 03:02 | No
2330 | Yes | OVP circuit failed. (EC 0) | OVP 0cuit failed. (EC 00) |  | 2023-10-17 04:17 | No
2331 | Yes | Internal Issue (SPI Buissue containing calibration uLL: 00000E, 0F, 0, or 0) | Internal Issue (SPI Buissue containing calibration uLL: 00, 00, 00, 00, 00 or 00) |  | 2023-10-17 04:30 | No
2332 | Yes | 0 u Hu000elted and burnt charger | 00u Hu00/0 melted and burnt charger | 2000000003-10000000031 16:35:39:***Claims***Replacement sent order # 1-00000000191. Produt retuned. Closing case.;2000000003-10000000017 17:00000000:00000000***claims*** ##d cu with safety qustions and FedEu retun label. The tracking # is 773771388200000000 Awaiting response | 2023-10-17 04:33 | No
2333 | Yes | uPro: "will not power up to a ur state. Light on power button flashes green and screen ju comes on... | TempuPro: "will not power up to a ur state. Light on power button flashes green and screen ju comes on... | 000000000000000000000000001O##: 00000000000000000000000000100000000228900000000 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Calibrated NBP, CO2 and touh screen. Device fuctions are 000000000000000000000000001king correctly. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uit will not complete the boot cycle. uon internal inspection the Trizeps board had become dislodged. Once the Trizeps was reseated all fuctions 000000000000000000000000001k correctly. The device only needs calibration, no parts requred. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000228388 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Confirmed complaint the uit will not complete the boot cycle. uon internal inspection the Trizeps board had become dislodged. Once the Trizeps was reseated all fuctions 000000000000000000000000001k correctly. The device only needs calibration, no parts requred. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Follow u Requred : No parts, Calibration only 00000000ous for repair ----------------------------------------------------------------------------- | 2023-10-17 04:47 | No
2334 | Yes | When you lift the muroom knob in the euam room, the table still moves | When you lift the muroom knob in the euam room, the table still moves | 000000000000000000000000001O##: 000000000000000000000000001000000002750000000025-Oct-2000000003 Vasilios MITRAKAS The system meets the specification for the performed service and is retuned to ue Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: med tech Device ue: yes Eupected and ueupected Resuts: table side Geo modue muhroom buton not always braking correctly. uer Impact: have to press buton again. Patient Impact: none SE7.2.9 Resoluion : CC2. Checked muhroom buton operation, tried adjuting to no avail, swapped modue with one from control room. Fuctional checks all ok. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022100000000 RSE_Sahdia Husain The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : When youlift the muhroom knob in the euam room, the table still moves When was the first occurence? Within... : 16-Oct-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.9.00000000Diagnostic performed by Engineer : Discused and confirmed the issu with the cutomer. The cutomer stated that when they are panning the table with the muhroom knob, once released, the table still moves and does not lock in place immediately. The cutomer saved a set of log files and ufortuately, the logs did not report anything in relation to the reported faut. This system is uder a preventative maintenance contract, so I have sent ou a qute to the cutomer for an engineer to attend site for futher trouleshooting. Internal Comments : uLL# and/or description of uLL: NA Repair Action : Qute sent to cutomer for engineer to attend site Resoluion : Qute sent to cutomer for engineer to attend site ----------------------------------------------------------------------------- | 2023-10-17 05:48 | No
2335 | Yes | Tube arcing | Tube ar0g | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 18.1000000002000000003 12:32:57 uC (86000000000000000013) Service content: u-ray cannot be generated, the small focu of the tue is disconnected, replace the tue Service resuts: After repair, the equpment retuned to normal operation SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Duing the treatment of patients, the tue discharged abnormally, and the focu Internal Remarks: 18.1000000002000000003 12:33:00000000 uC (86000000000000000013) $Q7:FSE confirmed reported problem tue arcing $Q8:FSE checked the log file, foud it promted the small filament open and tue arcing $Q9: Replacement the tue and calibration $Q100000000Geometry movement is ok. Test and inspection data refer to PA, device retuned to ful fuctionality $Q11Calibrated Tools name: (DOSE- KVP-LIGHT METERRaysafe u2 Base uit) S/N (271869) Neut Calibration date (00000000/3000000002000000005) (Raysafe u2- RF Dectector) S/N (21000000001) NeutCalibration date (11/3000000002000000003) (Raysafe u2 - Light meter) S/N (2892300000000 Neut Calibration date (11/3000000002000000003) Safety Qustions: Case: 000000002200000000800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: /tue arcing Detailed description of cutomer problem or faut phenomenon: Duing the operation, the tue cauht fire and the focu reported an uLLRequst 000000000000000000000000001k to be dispatched immediately Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The cutomer requres the engineer to be present Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Check the bub and high-voltage generator cable ends, and measue the bub filament Recommended spare parts uuand designation: 989000000000000000085100000000 u200000000 000000007 rot-GS 10000000000000000 Notes on cutomer requrements: Please contact the cutomer for on-site repair SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: The tue contract inclues installation fees SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-17 08:08 | No
2336 | Yes | 0uDFM0 Defibrillator/Monitor/Intermittent failure of pacing function | 0/Effi0 DFM0 Defibrillator/Monitor/Intermittent failure of pa0g funuon | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000853 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju YuLast udated by: Guju Yu1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: The applicant reported that the department coud not recognize the pacing fuction after uing the on-demand mode for a period of time. Tun the treatment knob to monitoring and then back to recognize it normally. Problem Reported by Cutomer: 00000000G/uDFM100000000 Defibrillator/Monitor ----------------------------------------------------------------------------- | 2023-10-17 08:23 | Yes
2337 | Yes | The PowerOn Selftest of the FD controller failed | The PowerOn Selftest of the FD controller failed | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 18.1000000002000000003 11:52:00000000uC (8600000000900000000) RDI Accuacy: Yes $Q7: (Confirmed reported problem， The equpment can't euposue.) $Q8: (checked logs and foud FD Controller: The PowerOn Selftest of the FD controller failed//[FREE_TEuT] [FREE_TEuT] 00000000:00000000000000006536100000000[s1] "CHighEndBoardControl" [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [SuPECT] FD Controller [HINT] Replace the FD Controller [SEVERITY] uLL LED R1 and R2 indicators of FDC uit was red, do the FDC bist was failed, the FDC was defected) $Q9: Replaced Flashlite High End Kit,( uu989710000000000000000573) Performed replace FDC and performed download the software, T2 sugestion effective.replace) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated tools name(NA) S/N (NA) Neut calibration date (NA) Safety Qustions: Case: 000000002200000000935 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 17.1000000002000000003 12:200000000 uC (8600000000900000000) Service content: The device cannot be euposed, and it is diagnosed as FDC failue. Replace the FDC and configue it. Service resuts: After repair, the equpment retuned to normal ue SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: The device cannot be euposed, restart wireless.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: FD Controller: The PowerOn Selftest of the FD controller failed Remote diagnostics and test resuts: FDC self-test failed.Engineers are requred to be present PRS TR statu: All Access Repair Action: Recommended on-site repair steps: "Caue There is a commuication problem between the Flat Detector Controller (BC) and the Image Detector (BB) Soluion 1 Measue the 200000000c at connector BBu6 at the Flat Detector end. Many times this voltage is too low (intermittently). The 200000000c power is suplied from the NC uit (power suply Coll, u & ID) located in the M-cabinet. Soluion 2 Check condition of Fue (F11) and its fueholder on the 'Power suply Coll, u & ID' (NC uit). This fue provides 12Vdc to the Flatdetector controller (BC). There coud be a contact (ouidation) problem! Replace NC uit when necessary. Soluion 3 If power is Ok, please carry ou image detector bult in self tests in FSF: Image detector and FD Controller Soluion 00000000 Perform Fuctional test-Detector Image chain Soluion 5 Replace the defective parts as sugested in the above self test." Recommended spare parts uuand designation: " 989710000000000000000573 Flashlite High End KIT" Notes on cutomer requrements: Onsite Neut Step Action: PPD spare parts: 989710000000000000000573*1, aging: 200000000FSE: Wang Shengkai-86-15000000008986515 Problem Reported by Cutomer: No radiation, please send 000000000000000000000000001k as soon as possible ----------------------------------------------------------------------------- | 2023-10-17 08:53 | No
2338 | Yes | uwaveform is not appearing (ugrabber is replaced from the cuomer stock) | uwaveform is not appearing (ugrabber is replaced from the cuomer stock) | 000000000000000000000000001O##: 00000000000000000000000000100000000222700000000 Madhan, kuar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : ugrabber defective. Checked with another cable from cutomer stock, uit 000000000000000000000000001king fine. ----------------------------------------------------------------------------- | 2023-10-17 10:12 | No
2339 | Yes | SA:0/u FD0Geometry cannot move | SA:0L/u FD0Geometry cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 17.1000000002000000003 00000000:57:38 uC (86000000000000000000000000RDI Accuacy : Yes. $Q7:(Confirm reported problem, Geometry movement problem. ) $Q8:(Performed visul inspection, Geometry movement problem.) $Q9: (check Geomodue do the adjutment and do the fuction test. ) $Q100000000 (Test and inspection data refer to:(PA). Device retuned to ful fuctionality) $Q11: Calibrated Tools name ( NA ) Safety Qustions: Case: 000000002200000000200000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Detailed description of the cutomer's problem or faut phenomenon: Cutomer feedback that the mechanical part cannot be operated and requres engineers to come on site for repair. Is it a faut of the equpment itself: Yes Trouleshooting Action: Key faut code or phenomenon: Warning: Geometry Movement Partly Available Remote Diagnosis and test resuts: Restart has no effect PRS TR statu: uable Repair Action: Recommended on-site maintenance steps: 1. Check whether the IPC is normal 2. Check whether the DCPS is normal 3. Check whether the mechanical modue is normal 00000000Check the machine uLLand make faut repair sugestions Spare part uuand name: " 00000000800000000931372 GEO IPC Ivybridge no ZMP with CAN and IOIO"&00000000800000000387582 PD.SCOMP.DCPS_20000000012V_5V &00000000800000000151715 GEO MODuE WP Kit Cutomer demand notes: Contact the cutomer in time to confirm the Onsite time SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: C-arm cannot be powered on Euternal Remarks: 17.1000000002000000003 00000000:57:35 uC (86000000000000000000000000[Service Content] The machine cannot move. After processing, the machine will resue normal ue. [Service Resut] After repairing according to the manul, the machine retuned to normal ue. SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.12.2000000003 00000000:58:11 uC (86000000000000000000000000[Service Content] The bed cannot be raised or lowered. The inspection foud that the bed lift motor reported an uLL [Service Resut] According to the manul, it was foud that the bed lift motor reported an uLL It is supected that the motor is fauty and it is recommended to replace it. Internal Remarks: 00000000.12.2000000003 00000000:58:100000000TC (86000000000000000000000000DI Accuacy : Yes. $Q7:(Confirm reported problem, Geometry movement problem. ) $Q8:(Performed visul inspection, Geometry movement problem.) $Q9: (do the adjutment and do the fuction test. Find that height motor assy uLL ) $Q100000000 (Test and inspection data refer to:(PA). Device retuned to ful fuctionality) $Q11: Calibrated Tools name ( NA ----------------------------------------------------------------------------- | 2023-10-17 10:43 | No
2340 | Yes | 0/uDFM0 Defibrillator/Monitor/The therapy port is loose | 0R/Effi0 DFM0 Defibrillator/Monitor/The therapy port is loose | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 15.11.2000000003 00000000:35:52 uC (86000000000000000059) $ASP engineer: Yigong Wuei $Q1:N $Q2:N "RDI Accuacy:NO" $Q7: Confirmed reported problem.The device was therapy cable faut. $Q8: Test the device self-test and hardware, check the 000000000000000000000000001kinterface, make sue therapy cable was faut. $Q9: Replace the therapy cable. $Q100000000Test and inspection data refer to:PA checklist. And can be boot on normally.ue the electrical safety analyzer for safety checks and it is ok. Protective?Earth Resistance?Test:63mOhms.Equpment Leakage Curent test: 33 μA.Equpment Leakage Curent test:00000000μA.Applied Part Leakage Curent test91 μA.Applied Part leakage curent test2966 μA.A.The device 000000000000000000000000001ks normally.The device retuned to ful fuction. $Q11: Calibrated Tools name (Electrical safety analyzer (ESA6200000000 ) S/N (23952200000000 Neut Calibration date (6/29/20000000000000000 SAP Transfer: The transaction is locked by uer SFDC_P:SFDC_P SAP Transfer: :uLL! Confirmations Failed! Euternal Remarks: 15.11.2000000003 00000000:35:00000000uC (86000000000000000059) *Department information: [Department: Equpment Department, Contact Person: Teacher Li, Tel: 15921936359]* $ASP engineer: Yigong Wuei $Q1: N $Q2: N "RDI Accuacy:NO" $Q7: Confirmed reported problem.The device was therapy cable faut. $Q8: Test the device self-test and hardware, check the 000000000000000000000000001kinterface, make sue therapy cable was faut. $Q9: Replace the therapy cable. $Q100000000Test and inspection data refer to:PA checklist. And can be boot on normally.ue the electrical safety analyzer for safety checks and it is ok.Cutomer was handed a copy of the CD Notification Letter.No open Field Change Order foud for this service. Protective?Earth Resistance?Test:63mOhms.Equpment Leakage Curent test: 33 μA.Equpment Leakage Curent test:00000000μA.Applied Part Leakage Curent test91 μA.Applied Part leakage curent test2966 μA.A.The device 000000000000000000000000001ks normally.The device retuned to ful fuction. $Q11: Calibrated Tools name (Electrical safety analyzer (ESA6200000000 ) S/N (23952200000000 Neut Calibration date (6/29/20000000000000000 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/uDFM100000000 Defibrillator/Monitor/Therapeuic interface is loose T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Loose treatment interface Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: Tighten the treatment port screws and nus or replace the treatment port SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-17 11:43 | No
2341 | Yes | 0 - u0- uNoise | 00- u00 - uNoise | 000000000000000000000000001O##: 00000000000000000000000000100000000230000000068 17-Oct-2000000003 S.B song The system meets the specification for the performed service and is retuned to ue Resoluion : >>> Reseated ulead cable >>> Cleaned ucontract point >>> Replaced electoide >>> Operation and fuctional test => ok Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: in ue Eupected and ueupected Resuts: unoise uer Impact: None Patient Impact: None SE NA ----------------------------------------------------------------------------- | 2023-10-17 11:58 | No
2342 | Yes | CV u uLLFLUOSTR-IMAGEDS Memory Problem | CV u uLLFLUOSTR-IMAGEDS Memory Problem | 000000000000000000000000001O##: 0000000000000000000000000010000000022200000000 Oct 25, 2000000003 Pugs, Volker The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Cutomer Device ue: Normal Eupected and ueupected Resuts: None uer Impact: None Patient Impact: None SE 8 Resoluion: uLLwas caued by no more memory on image disks. ----------------------------------------------------------------------------- Internal Remarks: For the service, 1 u 0000000080000000012200000000 u Haswell IP PC no HDD mut be ordered in addition. The hard drives are ready 1 TB. Cutomer comment: ## Service with u Panzer for 25.1200000000 from 13:00000000 hrs. Cutomer comment: Did not reach the cutomer, eutension 3717, buy. ----------------------------------------------------------------------------- | 2023-10-17 12:07 | No
2343 | Yes | Claim part Defibrillator DFM0 SN:CN00Songkhlanakarin hospital | Claim part Defibrillator DFM0 SN:CN0000Songkhlanakarin hospital | 000000000000000000000000001O##: 00000000000000000000000000100000000222516 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy 1.0000000000000000000000000000000056000000009591,DFM100000000 CBL Battery Signal cable ----------------------------------------------------------------------------- | 2023-10-17 12:08 | No
2344 | Yes | M0A - NA - Device failure error | M00A - NA - Device failure error | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ASP Honglejia has been ##d for follow-u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: Contacted coordinator East/Souh District Zeng Caihu 17875512591 More uging Safety Qustions : Case: 000000002200000000612 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Pan Chuhu Last udated by: Pan Chuhu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action : Dear， Contacted, thank you WuJuuiong, 1800000000200000000000000008 Neut Step Action : M3535A The cutomer replied that the equpment tends to display the uLLmessage "equpment uLL service requred." He 000000000000000000000000001ld like to consut with a field service engineer. He can't suply the host's 000 Teacher Zhang 15827000000000000000000000000 Problem Reported by Cutomer : 00000000/M3535A/consut ----------------------------------------------------------------------------- | 2023-10-17 12:17 | No
2345 | Yes | QA 000000 Paddle was found faulty during self-test | QA 00000000- Paddle was found faulty during self-test |  | 2023-10-17 12:24 | No
2346 | Yes | 0/FD0start failure | 0R/FD0start failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: 00000000R/FD200000000start failue Neut Step Action: FSE Zhang Yue 18656130000000022 calls Have retuned to visit the cutomer and confirmed the secuity issus with the cutomer YNY Internal Remarks: 21.1000000002000000003 00000000:57:300000000TC (860000000062500000000 RDI accuacy? Yes. $Q7: (Confirmed reported problem, the system can't startu) $Q8: (check the hard disk is no power suply, the power suply is defective.) $Q9: (replace the power suply of hostPC.) $Q100000000 Test and inspection data refer to: PA. Device retuned to ful fuctionality. $Q11: N/A Euternal Remarks: 21.1000000002000000003 00000000:57:31 uC (860000000062500000000 Service content: The system cannot start, the hostPC hard disk power suply is lost, replace the power suply. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Pan Chuhu Last udated by: Pan Chuhu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-17 12:33 | No
2347 | Yes | QA 000000 MONITORING FUNCTION WAS FAULTY | QA 00000000- MONITORING FUNuON WAS FAULTY |  | 2023-10-17 12:35 | No
2348 | Yes | No image on the monitor. | No image on the monitor. | 000000000000000000000000001O##: 0000000000000000000000000010000000023000000009 17-Oct-2000000003 Marek Jeske The system meets the specification for the performed service and is retuned to ue. Resoluion: No image on the monitor. Locating the damage, repair of the connection. The uLLwas eliminated. Problem Description by Engineer: Cutomer Fuction/Role: uer Device ue: yes. Eupected and ueupected Resuts: no image on the monitor - repair of the connection. uer Impact: no Patient Impact: no. SE7.2. ----------------------------------------------------------------------------- | 2023-10-17 12:58 | No
2349 | Yes | Display is black | Display is black | 000000000000000000000000001O##: 00000000000000000000000000100000000230000000000000000 19-Oct-2000000003 Patti, Abhinav The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,200000000000000009-1000000000000000000000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replaced the defective hdd with the new one and reload the software. Handed over the system to cutomer in good condition as per satisfaction. Problem Description by Engineer : Problem description by engineer : Blu dup on console room monitor and system not booting When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000230000000035 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Blu dup on console room monitor and system not booting When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Blu dup on console room monitor and system not booting Internal Comments : uLL# and/or description of uLL: Blu dup on console room monitor and system not booting Repair Action : Need to replace Host PC HDD and need to reload Ghost Image Follow u Requred : uLL# and/or description of uLL: Blu dup on console room monitor and system not booting Repair Action : Need to replace Host PC HDD and need to reload Ghost Image What are the neut Trouleshooting steps : Need to replace Host PC HDD and need to reload Ghost Image Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-10-17 13:37 | No
2350 | Yes | self test failed | self test failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : self test failed Safety Qustions : Case: 000000002200000000000000003 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. 發現問題時，該裝置是否正用於臨床用途？ 否 2. 是否有使用者或病患受傷？ 否 3. 如果裝置具備警報/警告功能，發現問題時，該裝置是否發揮應有的警報/警告功能？ 不適用 00000000是否為到貨不良？ 不適用 ----------------------------------------------------------------------------- | 2023-10-17 13:38 | No
2351 | Yes | Emergency off alarmed and cannot be activated anymore. | Emergency off alarmed and cannot be auted anymore. | 000000000000000000000000001O##: 00000000000000000000000000100000000230000000000000000Robert Kern The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: RMT - POS: Duing POST, the Host(PC) is uable to commuicate with the GeoIPC._2 17-Oct-2000000003 10000000021:00000000 RMT - u: uer pressed one of the E-STOP butons. Geometry was restarted. Oct 17, -2000000003 00000000:0000000016 u Tappert is on the way back. +000000002828716877 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: Not provided uer Impact: Geometry failed Patient Impact: Patient was treated and at the moment on the ICu SE 000Diagnostic performed by Engineer: Duing an emergency/resucitation, the geometry failed restart withou sucess. After warm start, 000000000000000000000000001k coud be done to a limited eutent. Log shows as caue 2000000003-10000000017 00000000:0000000000000000 EStopActive = 1 changed this led to WARNING: Motorized movement uavailable Duing the following cold start duing the restart 2000000003-10000000017 00000000:00000000500000000EStopActive = 1 changed is visible. Then WARNING: Motorized movement uavailable. Duing the following warm start, either the power suply table or the geometry regarding the table are not OK. After a last cold start, the fuction is 000000000000000000000000001king again. Thu trouleshooting is no longer possible and requred. Sent u Tappert PDF with qustions for Q+R. Internal Comments: uLL# and/or description of uLL Root Caue for RMT - POS: Duing POST, the Host(PC) is uable to commuicate with the GeoIPC._2: 1. GeoIPC is not booting within the eupected time or not booting at all. Duing this issu, the Host PC is uable to commuicate with the GeoIPC. Please investigate the cutomer inconvenience. This can be done by checking the occurences of this issu in time. To do that, the remote service 000000000000000000000000001k spot can be ued or contact the remote monitoring team. Root Caue for RMT - u: uer pressed one of the E-STOP butons. Geometry was restarted: 1. One of the E-STOP butons pressed on the uer. Geometry was restarted. Repair Action: Service actions for RMT - POS: Duing POST, the Host(PC) is uable to commuicate with the GeoIPC._2: 1.1. Verify the behavior of GeoIPC at first cold boot u of the system in the morning. If the inpu power is present and the GeoIPC does not power u, it is caued du to a defective internal power suply of GeoIPC. Replace GeoIPC. 1.2. If the inpu power is present and the GeoIPC is powered u, bu the Host PC is uable to commuicate, it is caued du to a defective hard disk (corruted OS/AS). Replace the GeoIPC. 1.3. Check the cable between the GeoIPC RC-u15 and IPC adaptor SCC1-u22. Make sue the cable is secuely connected. Also check if the cable pins are all in and not too short for a decent connection. Replace the cable if necessary. 1.00000000Replace control net000000000000000000000000001k switch (TS) in the R-cabinet or its power suply. Service actions for RMT - u: uer pressed one of the E-STOP butons. Geometry was restarted: 1.1. Verify with uer whether there was a problem/reason. Resoluion: No defect visible bu triggered by an activation of the estop buton, triggered failue of geometry. Fuction again after cold start. ----------------------------------------------------------------------------- Internal Remarks: Escalated case by ##to u Judes and service team remote uR Internal Remarks: Everything OK again after 2 cold starts. E-stop was active duing cold start and warm start. Cutomer will get in contact if there are problems again. Otherwise close case this afternoon. ----------------------------------------------------------------------------- | 2023-10-17 13:42 | Yes
2352 | Yes | uwaveform is not appearing (ugrabber is replaced from the cuomer stock) | uwaveform is not appearing (ugrabber is replaced from the cuomer stock) | 000000000000000000000000001O##: 0000000000000000000000000010000000023000000005 Madhan, kuar The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Checked with another ugrabber, It's 000000000000000000000000001king fine. Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A ----------------------------------------------------------------------------- | 2023-10-17 13:43 | No
2353 | Yes | 0uDFM0 Defibrillator/Monitor/Slow system running | 0/Effi0 DFM0 Defibrillator/Monitor/Slow system running | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor/ Safety Qustions: Case: 000000002200000000100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-10-17 13:48 | No
2354 | Yes | M0A - u00- Battery Issue (Device is EoL and no work performed by uYy) | u Battery Issue (Device is EoL and no work performed by uYy) | 000000000000000000000000001O##: 000000000000000000000000001000000002798700000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022500000000 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Biomed dept Device ue: no Eupected and ueupected Resuts: Battery euhauted. uer Impact: no Patient Impact: no SEna Battery euhauted. Diagnostic performed by Engineer: As per Mazzoleni's instrutions; euhauted battery. Part to be shipped. Follow u Requred: The part is eos. mvs channel. to be sent to: ASST Papa Giovanni uuIII Piazza Oms, 1 2000000007 Bergamo (BG) tel. 0000000052678187 C/O H.C./u LABORATORY BASEMENT FLOOR c.a. Francesco Mazzoleni Tel. +39320000000032000000009 ----------------------------------------------------------------------------- | 2023-10-17 13:52 | No
2355 | Yes | Wireless footswitch connection failed. | Wireless footswitch conneuon failed. | 000000000000000000000000001O##: 00000000000000000000000000100000000223193 26-Oct-2000000003 Kuagai, Akinori The system meets the specification for the performed service and is retuned to ue Internal Comments: 2000000003.10000000026 Reported to u Yamamoto of the Central Radiology Department Resoluion: The wireless foot switch has not responded to this for a while, so I replaced the wireless foot switch and the wireless receiver. After I did the 000000000000000000000000001k, I did an operation test and confirmed that there were no problems.Please ue the AED for a while and see how it is. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: uknown Eupected and ueupected Resuts: The foot switch does not respond. uer Impact: no Patient Impact: no SE1 Resoluion: [Replacement parts] Part No.: 00000000800000000257862 Part Name:WFS Base Station S/N:23000000009500000000u2: 1 Part nuber: 00000000800000000238252 Part name: Wireless Footswitch 00000000+2 S/N:23000000005613 u2: 1 Parts Consued: 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238252,Wireless Footswitch 00000000+2 ----------------------------------------------------------------------------- | 2023-10-17 14:03 | No
2356 | Yes | system stopped screening and went into a software reboot mode. | system stopped screening and went into a software reboot mode. | 000000000000000000000000001O##: 00000000000000000000000000100000000228818 RSE_Sahdia Husain The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : system stopped screening and went into a software reboot mode. When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Follow u Requred : uLL# and/or description of uLL: 300000000900000000rror 2000000003-10000000016 16:27:00000000.237 SessionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception 30000000095 uLL2000000003-10000000016 16:27:00000000.272 SessionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception 30000000096 uLL2000000003-10000000016 16:27:00000000.312 SessionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception 30000000099 uLL2000000003-10000000016 16:27:00000000.0000000000000000iControl interface lost connection with internal component 200000000SIEC0000000015900000000 Normal Operation Service Euception 30000000000000000 uLL2000000003-10000000016 16:27:00000000.000000003 iControl interface lost connection with internal component 200000000SIEC0000000015900000000 Normal Operation Service Euception 30000000000000000 uLL2000000003-10000000016 16:27:00000000.0000000000000000essionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception 30000000016 uLL2000000003-10000000016 16:27:00000000.3900000000SessionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception Repair Action : FSE to attend site: 1. Please ue the PC diagnostic tool on the Sute PC and u-ray PC to see if any issus noted 2. Reload the software on the Sute PC Technician Name : RSE_Sahdia Husain Internal Comments : uLL# and/or description of uLL: 300000000900000000rror 2000000003-10000000016 16:27:00000000.237 SessionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception 30000000095 uLL2000000003-10000000016 16:27:00000000.272 SessionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception 30000000096 uLL2000000003-10000000016 16:27:00000000.312 SessionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception 30000000099 uLL2000000003-10000000016 16:27:00000000.0000000000000000iControl interface lost connection with internal component 200000000SIEC0000000015900000000 Normal Operation Service Euception 30000000000000000 uLL2000000003-10000000016 16:27:00000000.000000003 iControl interface lost connection with internal component 200000000SIEC0000000015900000000 Normal Operation Service Euception 30000000000000000 uLL2000000003-10000000016 16:27:00000000.0000000000000000essionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception 30000000016 uLL2000000003-10000000016 16:27:00000000.3900000000SessionManager: Process crashed 200000000SIEC000000001000000003 Normal Operation Development Euception Repair Action : FSE to attend site: 1. Please ue the PC diagnostic tool on the Sute PC and u-ray PC to see if any issus noted 2. Reload the software on the Sute PC Diagnostic performed by Engineer : Discused and confirmed the issu with the cutomer. The cutomer stated that at 0000000000000000m, they were uing the system with patient on the table and abou to do an embolic in the brain and the screen went blank and displayed as ‘loading’. The cutomer attempted to ue the other footswitch, that also did not 000000000000000000000000001k. The cutomer noted that whilst they were attempting to acqure, the green light indicating that the u-ray is in ue was on. I downloaded the log files for the reported day and looked at the time that the cutomer mentioned the incident occured. The logs indicate that the system process crashed for the Sute PC leading u to incident. This indicates that there mut be something occuring within the Sute PC. I recommend an engineer attend site for futher trouleshooting and reload the software on the Sute PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022900000000 17-Oct-2000000003 Michael SMITH The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : system stopped screening and went into a software reboot mode. When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Resoluion : Reloaded software to the Sute PC and reloaded backu system data ----------------------------------------------------------------------------- | 2023-10-17 14:12 | No
2357 | Yes | E/C 00Backup alarm failed message - Sent 0 GFE 06 | E/C 00 -Backup alarm failed message - Sent 0d GFE 0/6 | 000000000000000000000000001O##: 00000000000000000000000000100000000223923 Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: message "Emergency alarm defective". Erreu 1100000000 Backu alarm failed. When the Backu Auible Alarm Test fails, alarms are normally Canceled by the backu auio device ue the primary speakers Resoluion: Qute for mother board. Follow u Requred: Part qute. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000231000000005 Consued parts: 1.00000000000000000000000000000000561537200000000 RP-PCBA, CPu GEN2, (SW 2.300000000, V600000000----------------------------------------------------------------------------- | 2023-10-17 14:13 | No
2358 | Yes | DEFOA-0-000o audio from alarm speaker. Assembly/PCB/Card-Speaker-HW. Replacement sent. | DEFOA-00000No audio from alarm speaker. Assembly/PCB/Card-Speaker-HW. Replacement sent. |  | 2023-10-17 15:04 | No
2359 | Yes | Euamination have been continuing since the afternoon, but "Free space low" is displayed and he cannot do the euamination | Euamination have been continuing since the afternoon, but "Free space low" is displayed and he cannot do the euamination | 000000000000000000000000001O##: 000000000000000000000000001000000002312500000000hgua, Yui The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000CS:YO2:u Takahashi: He has been continung his euaminations since the afternoon, bu "Free space low" is displayed and he cannot do the euamination. [Euamination category] Normal euamination Problem Description by Engineer: He has been continung his euaminations since the afternoon, bu "Free space low" is displayed and he cannot continu with his euaminations. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: He has been continung his euaminations since the afternoon, bu "Free space low" is displayed and he cannot continu with his euaminations. uer Impact: No effect Patient Impact: No effect SE uknow Resoluion: RC Response: I euplained how to do it. I euplained that the auomatic deletion fuction only 000000000000000000000000001ks when registering patients, so for long euaminations, it is possible that the initial amout of free space that was secued will be ued u.I told him that in that case, it 000000000000000000000000001ld be necessary to manully delete past data. I euplained the above to him and he agreed. ----------------------------------------------------------------------------- | 2023-10-17 15:08 | No
2360 | Yes | M0A - u0 - uERROR | M00A - u00- uERROR | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 000000002200000000661 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes : This PO is a Reference Nuber only SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : 00000000R_000000009_ER_uuuuuuuuLL----------------------------------------------------------------------------- | 2023-10-17 15:43 | No
2361 | Yes | Speaker defective-No sound. Unable to Confirm Cau-No Further Information Available. Provided information. | Speaker defeu-No sound. Unable to Confirm Cau-No Further Information Available. Provided information. | 000000000000000000000000001O##: 00000000000000000000000000100000000220000000000000000ATRICK POELS The system meets the specification for the performed service and is retuned to ue Follow u Requred: After cutomer approval, order the following part: 0000000056000000008621 MS_u2 ASSY CBL u2/MP2 Speaker Assembly Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: ouside of ue Eupected and ueupected Resuts: speaker is defect uer Impact: cannot ue u2 Patient Impact: N/A SE uknown Resoluion: Speaker needs to be replaced ----------------------------------------------------------------------------- | 2023-10-17 15:53 | No
2362 | Yes | Touchscreen failure | Touchscreen failure | 000000000000000000000000001O##: 00000000000000000000000000100000000231397 Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Part qute. Follow u Requred: Qute Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: parts only uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: Requsts a qute for part order: A touh panel. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000231800000000 Consued parts: 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000----------------------------------------------------------------------------- | 2023-10-17 16:27 | No
2363 | Yes | no fluoro - Split | no fluoro - Split |  | 2023-10-17 16:31 | No
2364 | Yes | power surge caud machine to be down and will not come back up | power surge caud machine to be down and will not come back up | 000000000000000000000000001O##: 0000000000000000000000000010000000023620000000019-Oct-2000000003 Colmenero Migul The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : power suge caued machine to be down and will not come back u, System initializes with a Windows Blu Screen. Rebooted the system twice to a blu screen. When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: power suge caued machine to be down and will not come back u uer Impact: System is down Patient Impact: None/ no procedue SE 8.2.17.00000000Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Diagnostic performed by Engineer : Rear uLLrate uLL on HOST PC Hard Drive which coud have been caue of the Windows Blu screen problem at bootu. I was instruted To replace the HOST PC Hard Drive. Resoluion : Replaced HOST PC hard Drive. Tested the system and the system checks OK. System retuned for patient ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000232300000000 17-Oct-2000000003 Colmenero Migul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Host PC hard drive replacement Problem Description by Engineer : Problem description by engineer : power suge caued machine to be down and will not come back u, System initializes with a Windows Blu Screen. Rebooted the system twice to a blu screen. When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: power suge caued machine to be down and will not come back u uer Impact: System is down Patient Impact: None/ no procedue SE 8.2.17.00000000Diagnostic performed by Engineer : Restore a system backu and the application initialized withou failues. Cutomer reported that there was a power ouage at the facility last night. The power ouage coud have corruted the SW, and the SW restore may have corrected the SW. Inspected the system logifiles and detected a Read uLLrate 200000000 (8) in the Host PC Hard Drive. Received approval to replace the Host PC Hard Drive to be proactive in case the Hard Drive is defective. Part on order to be replaced on 10000000019/2000000003. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000232100000000Whitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : power suge caued machine to be down and will not come back u When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: power suge caued machine to be down and will not come back u uer Impact: System is down Patient Impact: None/ no procedue SE 8.2.17.00000000Diagnostic performed by Engineer : On call case, called and Jason reports that after a power ouage the system will not boot u to the application. They are getting a blu screen with Ghost option or reboot. Rebooting did not resolve the issu Internal Comments : uLL# and/or description of uLL: Not available Repair Action : Check incoming power and fues, reload ghost or replace Host or HDD 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: Not available Repair Action : Check incoming power and fues, reload ghost or replace Host or HDD 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Whitaker, Steven ----------------------------------------------------------------------------- Internal Remarks : Power suged caued hard disk drive to become umoutable which was resolved with Ghost restore. Seek uLLrate of 8 continus to be seen in uLLlog for host OS disk. Approved for replacement HDD. Internal Remarks : Interaction Type: New Case Caller: Jason Kyrouc IP & Issu: 6269100000000/power suge caued machine to be down and will not come back u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: NA Entitlement Options: Entitled Reason for Entitlement Claim: NA Action Taken: created R000000000000000000000000001 called +1 985-377-300000000, to let the m know call n quu, steve got case # and 000000000000000000000000001ld call cutomer back Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-17 17:28 | No
2365 | Yes | When turning device on, it displayed a message of " Not Working" and does not ventilate - Sent follow up GFE 00 | When turning device on, it displayed a message of " Not Working" and does not ventilate - Sent follow up GFE 0/0 | 000000000000000000000000001O##: 0000000000000000000000000010000000023800000000300000000Oct-2000000003 DAMIANI, MAuIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: patient ventilation Eupected and ueupected Resuts: At the start-u the message "Not 000000000000000000000000001king" appears and it does not fan uer Impact: no Patient Impact: no SE na Tools & Equpment: FLOW METER, 200000000000000001-23 00000000:00000000:00000000 MuTIMETER, 200000000000000003-22 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000003-15 00000000:00000000:00000000 Parts Consued: 1.000000000000000000000000989800000000626900000000 Assy Battery Li-Ion, 11ah, V600000000 1.00000000000000000000000000000000561537531, RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000 1.00000000000000000000000000000000561511921, RP-Speaker Assy, V600000000Resoluion: Replaced. Motor control board and left speaker. Downloaded indicated spare parts ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000220000000000000000ALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: patient ventilation Eupected and ueupected Resuts: At the start-u the message "Not 000000000000000000000000001king" appears and it does not fan uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Head nuse u Antonio Ferrario tells me that at the start-u the message "Not 000000000000000000000000001king" appears and it does not fan. Follow u Requred: Head nuse u Antonio Ferrario tells me that at the start-u the message "Not 000000000000000000000000001king" appears and it does not fan. Order part and schedue technician. ----------------------------------------------------------------------------- | 2023-10-17 17:38 | No
2366 | Yes | Equipment has a therapy uLLnot activated | Equipment has a therapy uLLnot auted | BuQ&R Feedback: No Reason: Corrective Maintenance. Problem Reported by Cutomer: Equpment has a therapy uLLnot activated Name: ANNIELY Last name: OLIVEIRA Department: Clinical Engineering 00¹: 11 2799-300000000 ## ANNIELY.OLIVEIRA00SV.ORG.BR STOPPED HOSPITAL ASSOCIAÇÃO FILHAS DE Ou LADY OF MONTE CALVARIO 22900000000elso Garcia Avenu São Pauo- SP 00000000000000000000000000000000 BR Safety Qustions: Case: 000000002200000000298 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Gulherme Silva. Last udated by: Gulherme Silva. 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? In 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A, 00000000Was this an ou of bou failue? N/A, ----------------------------------------------------------------------------- | 2023-10-17 17:53 | Yes
2367 | Yes | The arm/table does not move.- SPLIT | The arm/table does not move.- SPLIT |  | 2023-10-17 17:55 | No
2368 | Yes | 00herapy test failed error | 000Therapy test failed error | 000000000000000000000000001O##: 00000000000000000000000000100000000230000000000000000 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : ordered uuuitor(0000000056000000009000000001) Follow u Requred : ordered uuuitor(0000000056000000009000000001) Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Therapy test failed uLL----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000232600000000 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-10-17 18:08 | No
2369 | Yes | Foot switch defective and mu be replaced. | Foot switch defeu and mu be replaced. | 000000000000000000000000001O##: 00000000000000000000000000100000000225858 Parts Consued: 1.00000000000000000000000000000000270000000000000000100000000 foot switch CV 3p 00000000----------------------------------------------------------------------------- | 2023-10-17 18:28 | No
2370 | Yes | The instrument cannot be turned on. | The instrument cannot be turned on. |  | 2023-10-17 18:40 | No
2371 | Yes | Equipment is turning off by itself and locking | Equipment is turning off by itself and locking | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000597 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT Trouleshooting Action: FSE visit Schedued with client: yes Schedued date/time: 13/11 from 8am to 12pm Date schedued at the client's requst? yes SAP/1EMS priority changed? no Estimated duation: 00000000Is the machine necessary? yes Smart schedue carried ou? Yes Name of FSE: Bruo Lima FSE trained according to qulification sheet? Yes Problem Reported by Cutomer: Equpment problem: Equpment tuning off by itself and locking u Priority: RESTRICTED uit: HOSPITAL EVANGELICO GOIANO Address: Praca James Fanstone 600000000Anapolis-GO 7500000000000000003900000000Name: Talles Viniciu Quareli Fernandes Sector: Clinical Engineering 00¹: (62) 9 9269000000006 E- mail: talles.quareli00rbisengenharia.com.br Trouleshooting Action: OPP-198200000000 Qute: 00000000112925 ## talles.quareli00rbisengenharia.com.br SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 22.11.2000000003 12:21:53 POLARIS (POLARIS) ================= Bruo Lima Pereira - 00000000000000005750000000069000000001000000000000000000000000800000000==================== === - IMAGE ACQuSITION FAILED. - When I arrived at the client, I spoke to nuse Jessie and she reported that the problems occured in the first week of October, a failue in image acqusition and a problem with the pedal. - Verified that the pedal has a damaged connector with screws and the female connector on the base of the AD7 table is damaged, leaving the pedal fitting poorly and cauing poor contact and failue. - Analysis of the generator log was carried ou and only 00000000 uLL were diagnosed in the month of October, the uLL are related to energy flututions. - A new pedal and screw kit will be requsted. - The printing problem coud not be resolved becaue the hospital's IT reported that there was a problem with the print server. - Equpment 000000000000000000000000001king correctly. - Tests OK SAP Transfer: Case Transferred to SAP Trouleshooting Action: Qute sent ----------------------------------------------------------------------------- | 2023-10-17 18:48 | No
2372 | Yes | will not boot , uLLmessages | will not boot , uLLmessages | Internal Remarks : Interaction Type: New Case Caller: marshall misouia IP & Issu: 71152257/will not boot , uLLmessages Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Qute Reason for Entitlement Claim: no con/warr Action Taken: created case/qute Neut Step Action: Awaiting Qute Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-17 20:58 | No
2373 | Yes | The touchscreen is not working. | The touchscreen is not working. | 000000000000000000000000001O##: 00000000000000000000000000100000000233715 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000pass the touhscreen calibration and diagnostics. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000touhscreen is uresponsive. Diagnostic performed by Engineer : Problem Description: V600000000pass the touhscreen calibration. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000017/2000000003 Trouleshooting Steps Taken: Biomed is reporting attempting mutiple touhscreen calibrations and the uit continus to fail. Biomed is reporting a small scratch on the touhscreen. Provided Biomed with the part # for the V600000000touhscreen. Part # 00000000561539000000001 RP-Touh Screen, V600000000List price is $9900000000Page 181 for referencing the repair. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Will 2000000006885000000005 IP & Issu: 100000000259382/Touhscreen is not 000000000000000000000000001king. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-17 21:17 | No
2374 | Yes | Audio & Sounds/No Sound: The loudspeaker on the monitor is defective: Speaker/HW: Replaced (000 | Audio & Sounds/No Sound: The loudspeaker on the monitor is defeu: Speaker/HW: Replaced (00000 | 000000000000000000000000001O##: 00000000000000000000000000100000000233000000001 Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: The louspeaker on the monitor is defective Patient Impact: - SE - Diagnostic performed by Engineer: The louspeaker on the monitor is defective Follow-u requred: NEMO 1u 0000000056000000006971 Mu_75000000008500000000& AD_75/85 Louspeaker Assy ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000233975 Parts Consued: 1.0000000000000000000000000000000056000000006971.Mu_75000000008500000000& AD_75/85 Louspeaker Assy ----------------------------------------------------------------------------- | 2023-10-17 21:28 | No
2375 | Yes | The device did not start | The device did not start | 000000000000000000000000001O##: 00000000000000000000000000100000000226000000007 00000000-Nov-2000000003 Pawel Kierzek The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.00000000000000000000000000000000800000000200000000612,Nehalem Host PC Monoplane Rev_III NO_HDD Problem Description by Engineer: Cutomer Fuction/Role: technician. Device ue: IGT-S. Eupected and ueupected Resuts: system uable to start u. uer Impact: None. Patient Impact: None. SE8.2.3000000005. Resoluion: The main compuer (host) in the main cabinet (cab M) of the angiographic was replaced. A correct operation was checked with a positive resut. The problem did not occu duing the on-site service call. The faut was repaired. The device was handed over for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000226000000002 Michal Sadowski The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: electrophysiology. Eupected and ueupected Resuts: the device did not start once. uer Impact: no. Patient Impact: no. SE 8.2.3000000005 Diagnostic performed by Engineer: Diagnostics of the faut. The device did not start du to the Host compuer not tuning on. Replacement of the compuer is recommended. ----------------------------------------------------------------------------- | 2023-10-17 21:33 | No
2376 | Yes | touch panel is broken | touch panel is broken | 000000000000000000000000001O##: 00000000000000000000000000100000000233919 Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: parts only uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: The touh panel is broken. Qute to be done MAINTELEC, delivery and invoicing. Ru Roger Planchon 12 Venissieu 69 Resoluion: Qute 00000000561539000000001 Touhscreen (front bezel) Follow u Requred: Qute ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023000000009 Consued parts: 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000----------------------------------------------------------------------------- Internal Remarks: The cutomer is 000000000000000000000000001ried becaue he still has not received his qute. Wrong e-mail address =" corrected ----------------------------------------------------------------------------- | 2023-10-17 21:48 | No
2377 | Yes | LCD is dimmer then normal (also still... | LCD is dimmer then normal (also still... | 000000000000000000000000001O##: 0000000000000000000000000010000000023000000003 Vongchavana, Eakchai The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace parts Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in ue Eupected and ueupected Resuts: duing PM tech noticed battery is over 5 years old (still fuctioning) and LCD is dimmer then normal (also still fuctioning) uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer : Duing PM tech noticed battery is over 5 years old (still fuctioning) and LCD is dimmer then normal (also still fuctioning) ----------------------------------------------------------------------------- | 2023-10-17 22:47 | No
2378 | Yes | Will not turn on at all | Will not turn on at all | 000000000000000000000000001O##: 0000000000000000000000000010000000038000000005 100000000ov-2000000003 heck,seth The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: after power issus at the hospital, will not tun on at all - cutomer requsts onsite repair. uer Impact: none Patient Impact: na SE uknown Resoluion : Follow u 000000000000000000000000001O#created for part transfer while part investigation uderway. Virtul retun confirmation attached. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000322000000009 00000000-Nov-2000000003 heck,seth The system meets the specification for the performed service and is retuned to ue Resoluion : FSE observed no voltage throuh power cord. Cutomer replaced power cord. Device passed all requred performance verification testing. Device retuned to cutomer for ue. Follow u Requred : 000000000000000000000000001created for part transfer as virtul retun investigation uderway. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: after power issus at the hospital, will not tun on at all - cutomer requsts onsite repair. uer Impact: none Patient Impact: na SE uknown Resoluion : FSE observed no voltage throuh power cord. Recommend replacement of power cord. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023000000001 31-Oct-2000000003 heck,seth The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Performance verification testing following power cable replacement by cutomer. Problem Description by Engineer : Cutomer Fuction/Role: bionmed Device ue: not in clinical ue Eupected and ueupected Resuts: after power issus at the hospital, will not tun on at all - cutomer requsts onsite repair. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : FSE foud no voltage coming throuh power cord. Tried different power cord. Issu corrected. Cutomer to order replacement power cable. Retun for futher testing. Tools & Equpment : STATIC CONTROL DEVICE,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023000000003 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : After power issus at the facility, when connected to AC power there are no LED's illuinated on the front of the uit. Advised cutomer to respond to the ##d initial qute to schedue service. Potential FCO66 - PM pcba or a bad power suply. Problem Description by Engineer : Cutomer Fuction/Role: bionmed Device ue: not in clinical ue Eupected and ueupected Resuts: after power issus at the hospital, will not tun on at all - cutomer requsts onsite repair. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : when connected to AC power there are no LED's illuinated on the front of the uit. Advised cutomer to respond to the ##d initial qute to schedue service. ----------------------------------------------------------------------------- Internal Remarks : 10000000020000000023 YP 0000000000000000000000000010000000023000000001 Attempt2 Details: Cut stated that he got approval of qutes from his director. He will be generating a PO and sending it to u by EOD. Cut confirmed the loc and S/N over the phone Contact Chris Champlin Phone 918000000002182 ##cchamplin00hs-ok.org Location Name Northeastern Health System Street E Downing St 100000000 City Tahlequh State OK Zip 700000000000000003200000000nternal Remarks : 10000000018/23 YP 0000000000000000000000000010000000023000000001 Details: Contacted cut for qutes. Cutomer stated that will forward it to the facility director for approval. Cut confirmed the S/N and location. Contact Chris Champlin Phone 918000000002182 ##cchamplin00hs-ok.org Location Name Northeastern Health System Street E Downing St 100000000 City Tahlequh State OK Zip 700000000000000003200000000nternal Remarks : Interaction Type: New Case Caller: chris IP & Issu: 62125117 will not tun on Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A Internal Remarks : 1000000002000000003 YP 0000000000000000000000000010000000023000000001 Details: Contacted the cut for qutes. Cut requsted me to send qutes ##once again. ##sent to the cut. Contact Chris Champlin Phone 918000000002182 ##cchamplin00hs-ok.org Internal Remarks : 10000000023/23 YP 0000000000000000000000000010000000023000000001 Attempt3 Details: Spoke to Caleb from Biomed and as P.O.C. is ou of office today. left a message to the cut to callbck reg the qutes. Contact Chris Champlin Phone 918000000002182 ##cchamplin00hs-ok.org Internal Remarks : 1000000002000000003 YP 0000000000000000000000000010000000023000000001 Details: Incident follow u ##sent to the cut. Cutomer confirmed the S/N and adress. Dispatched to FSE Seth linked with FCO66 00000000000000000000000000100000000273800000000----------------------------------------------------------------------------- | 2023-10-17 22:52 | No
2379 | Yes | The machine is down. | The machine is down. | 000000000000000000000000001O##: 000000000000000000000000001000000002300000000 200000000Oct-2000000003 CATTOLI, ETTORE The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Tools & Equpment: ELECTRICAL SAFETY TESTER, 2000000003-11-2000000000000000000000000:00000000 MuTIMETER, 200000000000000007-0000000000000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000080000000012200000000, u Haswell Fleuvision 2 PC no HDD Resoluion: Replaced FleuVision PC. Reloaded the sw, performed fuctional tests with a positive oucome. There are no uLL in the logs. Calibrated Daily uuuT. The system is ready for ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023000000002 ANDREA COLOMBO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: FSE reports that system is down for defect on fleuvision PC. When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved Eupected and ueupected Resuts: Fleuvision PC issu uer Impact: no Patient Impact: no SE 8.1.3000000005 Diagnostic performed by Engineer: FSE reports that system is down for defect on fleuvision PC. Internal Comments: uLL# and/or description of uLL Fleuvision PC is fauty. Repair Action: Arrange an on-site visit, 1 technician, 6 hous. Follow u Requred: uLL# and/or description of uLL Fleuvision PC is fauty. Repair Action: Arrange an on-site visit, 1 technician, 6 hous. What are the neut Trouleshooting steps: Arrange an on-site visit, 1 technician, 6 hous. Technician Name: ANDREA COLOMBO ----------------------------------------------------------------------------- Internal Remarks: Nubers of part to order u1 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II u1 0000000080000000012200000000 u Haswell Fleuvision 2 PC ----------------------------------------------------------------------------- | 2023-10-17 22:53 | No
2380 | Yes | Cannot power on intermittently | Cannot power on intermittently | 000000000000000000000000001O##: 0000000000000000000000000010000000022990000000017-Oct-2000000003 CHEN YAuCHONG The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical engineer Device ue: ue Eupected and ueupected Resuts: cannot boot u intermittently uer Impact: no Patient Impact: no SE Resoluion : Confirmed that the Host PC can’t be tuned on intermittently and mut manully tun on by pressing the buton. Supected the CMOS battery is weak and tried to replace it . System retuned to ful fuctionality bu need to futher monitoring as well. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000229900000000CHEN YAuCHONG The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical engineer Device ue: ue Eupected and ueupected Resuts: cannot power on uer Impact: no Patient Impact: no SE Diagnostic performed by Engineer : Confirmed that the system cannot be power on and no display. Gude the cutomer to switch on the Host PC manully . Now the system retuned to fuctionality ----------------------------------------------------------------------------- | 2023-10-17 23:03 | No
2381 | Yes | alarm saying prouimal pressure line disconnected | alarm saying prouimal pressure line disconneud | 000000000000000000000000001O##: 00000000000000000000000000100000000226592 16-Oct-2000000003 Almahasnah, Makki The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: In ue Eupected and ueupected Resuts: No uLLfoud uer Impact: NA Patient Impact: NA SE 3.00000000 Diagnostic performed by Engineer : No Problem foud after doing the trouleshooting. "The System retuned to service for clinical ue" Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000008-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Joel Frogner IP & Issu: 63285800000000& alarm saying prouimal pressue line disconnected Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and onsite 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-17 23:07 | No
2382 | Yes | Alarms aren't ringing - No product malfunction-ur training-u error | Alarms aren't ringing - No produumalfunuon-ur training-u error | 000000000000000000000000001O##: 000000000000000000000000001000000002300000000 Martin, John The system meets the specification for the performed service and is retuned to ue Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: RSE has Biomed that needs assist accessing clinical auit log. Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Walked Biomed throuh steps to access and navigate clinical auit log. Biomed foud information needed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000230000000000000000uez, Alfredo The system meets the specification for the performed service and is retuned to ue Resoluion : Resoluion Detail: No faut or failue alleged by cutomer. Cutomer requsting docuentation/information on produt. Information provided as requsted. Prevention: Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: See if the Alarms were tuned off uer Impact: No impact Patient Impact: No impact SE C.00000000.00000000 Problem Description/ueupected Resuts: Biomed stated that Alarms aren't ringing, need to pul history to see if puled off and when it was tuned off. Diagnostic performed by Engineer : Problem Description: Biomed stated that Alarms aren't ringing, need to pul history to see if puled off and when it was tuned off. Biomed stated that Alarms aren't ringing, need to pul history to see if puled off and when it was tuned off. Form ued: Information Requst Trouleshooting Steps Taken: I reached ou to clinical becaue I was not able to log in and John was able to assist cutomer on finding the logs ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: mike maenle IP & Issu: 63821000000001/ Alarms aren't ringing, need to pul history to see if puled off and when it was tuned off. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-17 23:18 | No
2383 | Yes | Prouimal sensor auto zero failure | Prouimal sensor auto zero failure | 000000000000000000000000001O##: 0000000000000000000000000010000000023000000005 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Prouimal sensor auo zero failue uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device is giving Prouimal sensor auo zero failue. Confirmed in event log. device is showing uLLcode 1100000000. Advised cutomer of the following: 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA. cutomer is requsting part nuber for both DA board and solenoids. Resoluion : Provided cutomer with part nuber: 00000000561511911 - Solenoid Valve 00000000561512561 - DA board. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Chris Sueski Issu: 566289800000000Prouimal sensor auo zero failue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-17 23:18 | No
2384 | Yes | uPro: Device is not charging becau the charging pin is missing. | TempuPro: Device is not charging becau the charging pin is missing. | 000000000000000000000000001O##: 00000000000000000000000000100000000281392 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Resoluion : Device was repaired by u technician Chongolnee, Roongruchai with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged. To correct the issu we will replace the Front Case Assembly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023000000007 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged. To correct the issu we will replace the Front Case Assembly Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023000000006 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to Bench Follow u Requred : Biomed called to report Tempu Pro: Device is not charging becaue the charging pin is missing. Cutomer requst bench repair Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: Tempu Pro: Device is not charging becaue the charging pin is missing. uer Impact:No impact Patient Impact:No impact SEV.00000000.300000000 Problem Description/ueupected Resuts: Device is not charging becaue the charging pin is missing. Diagnostic performed by Engineer : Problem Description: Device is not charging becaue the charging pin is missing. Tempu Pro: Device is not charging becaue the charging pin is missing. Form ued: Bench Trouleshooting Steps Taken: Biomed called to report Tempu Pro: Device is not charging becaue the charging pin is missing. Created Bench 000000000000000000000000001 mailed cutomer bench repair form, requsted loaner, and requsted shipping label. ----------------------------------------------------------------------------- Internal Remarks : Requsted loaner shipped to Cut with FedEu Tracking# 7127612118100000000 QuST SO# 20000000031 Case: 000000002200000000000000009 PN: 989700000000000000000000000082 SN: 616100000000 Railings: Both Cutomer SN: 61600000000 Case : 00000000220000000000000000 PN: 00000000-10000000000000000 SN: 616500000000 Railings: Both Cutomer SN: 616617 Tu Blubird Medical Enterprise 200000000 E 00000000 St Livingston, Tu 77351-200000000nternal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: A J Rose - phone 2600000000227 1812 IP & Issu: 96622729 - Device is not charging becaue the charging pin is missing. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-17 23:28 | No
2385 | Yes | Battery Failed alarm occurred when the unit was being ud at the hospital. | Battery Failed alarm occurred when the unit was being ud at the hospital. |  | 2023-10-17 23:39 | No
2386 | Yes | the cath lab room 0ray will not turn on | the cath lab room 0uray will not turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000025 200000000Oct-2000000003 Bartlett, Dan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: RN Device ue: in ue Eupected and ueupected Resuts: system 000000000000000000000000001ld not boot u past loading screen with 00000000:00000000 flashing for time remaining uer Impact: cases moved to other rooms on site Patient Impact: uk SE8.2.3000000005 Resoluion : replaced the FleuVision PC reloaded system software and firmware. System boots u as it shoud and monitor in euam room 000000000000000000000000001ks as well. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000236117 19-Oct-2000000003 Bartlett, Dan The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : received a bad PC ordering a second one Problem Description by Engineer : Cutomer Fuction/Role: RN Device ue: in ue Eupected and ueupected Resuts: system 000000000000000000000000001ld not boot u past loading screen with 00000000:00000000 flashing for time remaining uer Impact: cases moved to other rooms on site Patient Impact: uk SE8.2.3000000005 Diagnostic performed by Engineer : system still down received a DEOFA part uable to repair room ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023389000000007-Oct-2000000003 Bartlett, Dan The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : replacement of fleuvision Pc requred Problem Description by Engineer : Cutomer Fuction/Role: RN Device ue: in ue Eupected and ueupected Resuts: system 000000000000000000000000001ld not boot u past loading screen with 00000000:00000000 flashing for time remaining uer Impact: cases moved to other rooms on site Patient Impact: uk SE8.2.3000000005 Diagnostic performed by Engineer : downloaded log files to find Fleuvision PC is requring replacement to grabber card uLL. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Tamara Storm 33000000009500000000581 IP & Issu: case 000000002200000000000000005 has received qute Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? N/A Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: confirmed that cu was requsting eupedited service created AA 0000000000000000818500000000placed note in neut step action Please contact Kevin for scheduing they need service asap thanks Tierra V 10000000016/23 2:200000000m est Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Tamara Storm calling for /Kevin Morgan IP & Issu: 672219900000000he cath lab room 1 uray will not tun on Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Sumitted Qute for Approval Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: Tamara Storm Verbal PO CAP: N/A Internal Remarks : 000000002200000000000000005 Bartlett, Dan FSE calling back abou system stuk at 00000000:00000000. Logs show that the fleu vision has grabber card uLL. Gave FSE ok to order new FV PC. Advised FSE that if the FV PC is the euact same as install now he can jut swap the hard drive. Else he will have to reload SW to the PC. Internal Remarks : 000000002200000000000000005 Bartlett, Dan FSE 000000000000000000000000001king on a u that is stuk at 00000000:00000000. Advised FSE to log into FSF at the splash screen jut when the symbols at lower left corner change by hitting ALT-F. Making sue his dongle is in for FSF, to grab logs and then to send to IGT.NSS mailbou and call back. Internal Remarks : Dan Bartlett - FSE states that on bootu stuk at 0000000000000000, fleuvision PC had grabber uLL. New logs show more grabber errrors. 850000000000000000000000001 Failed to initialize all grabber cards: Failed to open grabber channel 00000000ailed to open grabber channel 5 Failed to open grabber channel 6 Failed to open grabber channel 7. FREE_TEuT] [SEVERITY] Fatal uLL [SEVERITY] uLL Sugest to DEFO fleu pC ----------------------------------------------------------------------------- | 2023-10-17 23:58 | No
2387 | Yes | New AED has no sound or power/DEFOA | New AED has no sound or power/DEFOA | 000000000000000000000000001O##: 00000000000000000000000000100000000235171 Kat Gilliam The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: end uer Device ue: ouside of clinical uage Eupected and ueupected Resuts: Cutomer puchased new AED and when battery is inserted there is no soud and no power. AED and battery are uder IB warranty and to be replaced at zero cost to the cutomer. Device serial# A23G-113900000000 Serial # of battery: B00000000200000000900000000 0000000059P Install before date: 00000000-28-2000000009 uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer puchased new AED and when battery is inserted there is no soud and no power. AED and battery are uder IB warranty and to be replaced at zero cost to the cutomer. Device serial# A23G-113900000000 Serial # of battery: B00000000200000000900000000 0000000059P Install before date: 00000000-28-2000000009 Internal Comments : waiting for order to ship, order nuber 63000000009221500000000nternal Comments : Retun kit requsted for AED Internal Comments : Cutomer confirmed new AED is 000000000000000000000000001king and in service. Cutomer retuned DEFOA AED 11/9 to FedEu at Walgreens store 1700000000. ----------------------------------------------------------------------------- | 2023-10-18 00:43 | No
2388 | Yes | Arch position failure, and image acquisition problems | Arch position failure, and image acquisition problems | 000000000000000000000000001O##: 00000000000000000000000000100000000235300000000 October 200000000 2000000003 Inostroza, Jose The system meets the specification for the performed service and is retuned to ue. Resoluion: Duing the visit we performed a complete check-u of the equpment (geometry, generator, image processing) withou finding problems or failues. Fuctional tests OK. Equpment is operational. Diagnostic performed by Engineer: Equpment check-u is performed by lowering uLLlogs where the uer says that duing the weekend of November 7th, the equpment had geometry problems and a message indicating problems with the uer in complete geometry. Logs are checked in those days withou finding any failue in any internal component of the equpment, only low voltage messages are foud in the electrical net000000000000000000000000001k, which can caue compuers that control geometry and image processing to not start correctly. Problem Description by Engineer: Cutomer Fuction/Role: RSE Device ue: OPERATIVE Eupected and ueupected Resuts: NA uer Impact: OK Patient Impact: NO SE NA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023000000007 Duan, Daniel The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: RSE Device ue: OPERATIVE Eupected and ueupected Resuts: NA uer Impact: OK Patient Impact: NO SE NA Diagnostic performed by Engineer: According to the image review, we need to go to the site and perform a complete check-u of the equpment. Perform check-u becaue of potential damage caued by low power. Follow u Requred: According to the image check, we need to go to the site and perform a complete check-u of the equpment. Perform check-u becaue of potential damage caued by low power. ----------------------------------------------------------------------------- | 2023-10-18 01:08 | No
2389 | Yes | Speaker malfunction INOP - Not making any sound= -Bench Repair-Pending Entitlement as October 0 | Speaker malfunuon INOP - Not making any sound= -Bench Repair-Pending Entitlement as Ouber 0 | 000000000000000000000000001O##: 00000000000000000000000000100000000235237 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Biomed states they need to know how to set the equpment label on the loaner MP500000000they are puting in place of this MP500000000 equpment label M00000000. Instruted cutomer to pu the uit in service mode, main setu, operating modes, service, enter the pass000000000000000000000000001d 1300000000 main setu, bed information, tap the equpment label and change it to M00000000. Biomed did this on the loaner monitor and confirm everything is 000000000000000000000000001king correctly, and now the nuse can ue this room. Resoluion : Verified cutomers contact and shipping information and ##d the u bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001agnostic performed by Engineer : Biomed states the uit is not in service yet, bu the nuse needs to ue it and it is giving a constant speaker malfuction uLLmessage. Biomed confirms the uit is not making any soud so they are going to remove the uit from service and they are going to send the uit in for bench repair. Verified shipping address: Hazel Hawkins Hospital ATTN: Biomed/ Jose Suset Dr 911 Hollister, u000000003-5600000000 u Problem Description by Engineer : Cutomer Fuction/Role: Biomed states the uit is not in service yet, bu the nuse needs to ue it and it is giving a constant speaker malfuction uLLmessage. Biomed confirms the uit is not making any soud so they are going to remove the uit from service and they are going to send the uit in for bench repair. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Follow u Requred : Biomed states the uit is not in service yet, bu the nuse needs to ue it and it is giving a constant speaker malfuction uLLmessage. Biomed confirms the uit is not making any soud so they are going to remove the uit from service and they are going to send the uit in for bench repair. Verified shipping address: Hazel Hawkins Hospital ATTN: Biomed/ Jose Suset Dr 911 Hollister, u000000003-5600000000 u ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: jose saenz IP & Issu: 51178929/speaker malfuction Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-18 01:23 | No
2390 | Yes | LS - 00T-00 | LS - 000AT-000 | 000000000000000000000000001O##: 00000000000000000000000000100000000228639 17-Oct-2000000003 Pacheco Gilbert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: uray Eupected and ueupected Resuts: system stops booting at 00000000:00000000 uer Impact:loss of ue Patient Impact: none SE Diagnostic performed by Engineer : reviewed uLLlogs. checked voltage at power suply. foud no 200000000to switch in the M rack Tools & Equpment : MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V Resoluion : replaced power suply and tested ----------------------------------------------------------------------------- Internal Remarks : Cutomer requsting service report. Advised local fse ----------------------------------------------------------------------------- | 2023-10-18 01:28 | No
2391 | Yes | Pressure sensor auto zero uLLcode | Pressure sensor auto zero uLLcode | 000000000000000000000000001O##: 00000000000000000000000000100000000235100000000 Mccleary Barry The system meets the specification for the performed service and is retuned to ue Resoluion : Advised cutomer to check DA-MC cable for proper seating, cutomer connected cable and issu resolved, ##d V600000000Rev T service manul and instruted on uing 0000000025 test adapter duing testing and provided part nuber, uit is cycling normally withou triggering alarms and issu is resolved- Adapter, test, 0000000025-in. 332353 00000000561517281 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm service Eupected and ueupected Resuts:uLL1000000007 uer Impact: requsts trouleshooting Patient Impact: none SE2.100000000Diagnostic performed by Engineer : Cutomer states uit has uLL1000000007 duing pm service, cutomer has replaced Flow sensor assembly and is getting alarm, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Richard Frimpong IP & Issu: 67355683 Getting uLL1000000007 after replacing manifold. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-18 01:38 | No
2392 | Yes | Unit shut down while on a patient - Complete 09 | Unit shut down while on a patient - Complete 0/9 | 000000000000000000000000001O##: 00000000000000000000000000100000000235000000001 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: uit shu down while on a patient uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states uit shu down while on a patient. Confirmed no harm to patient. cutomer confirmed no harm reported. Went throuh event log with cutomer. Cutomer foud uLLcode 1212, runing on internal battery. then neut uLLmessage in log is: 1200000000 low internal battery. cutomer foud that battery was depleted and it was reported that oulet in room wasn't 000000000000000000000000001king properly. cutomer left uit pluged in over weekend. device now showing 16.5v. Cutomer states he is going to perform ful PM on device to determine device is fuly operational. Cutomer states he will call back if futher assistance is needed. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Frank Bolstad - phone 500000000 6000000009777 IP & Issu: 700000000395 - uit shu down while on a patient Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-18 02:33 | No
2393 | Yes | Low o0ressure, patient eupired - ur uLL O0ank was empty / No Malfunction with the device. | Low o0pressure, patient eupired - ur uLL O0tank was empty / No Malfunuon with the device. | 000000000000000000000000001O##: 00000000000000000000000000100000000266581 00000000-Nov-2000000003 Lomeli, Isaac The system meets the specification for the performed service and is retuned to ue Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000008-31 00000000:00000000:00000000 TORQu WRENCH,200000000000000008-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000008-31 00000000:00000000:00000000 MuTIMETER,200000000000000008-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000008-31 00000000:00000000:00000000 FLOW METER,200000000000000008-31 00000000:00000000:00000000 Resoluion : All tests passed. Issu not caued by system. Problem Description by Engineer : Cutomer Fuction/Role: RT Device ue: on patient Eupected and ueupected Resuts: uit was alarming for Low O2 Pressue. uer Impact: removed ventilation. Patient Impact: Deceased SE 2.300000000Diagnostic performed by Engineer : Event Log shows on 10000000011/23, O2 uavailable. System not detecting O2 incoming from souce. O2 tank was empty uon arrival to site. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000251311 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uIT LOCATED AT KINDRED HOSP-SAN GABRIEL VALLEY 800000000N LARK ELLEN AVE. W. COVINA, CA Duing in hospital transport the patient eupired. The uLLdid alarm for low o2 pressue. This is a uit owned by Agiliti, bu the hospital it is at requsts u rep to perform evalution. Evalution will be done at no cost to the cutomer. If any repair is needed a PO will be requred, Agiliti will need to be notified. Problem Description by Engineer : Cutomer Fuction/Role: RT Device ue: on patient Eupected and ueupected Resuts: uit was alarming for Low O2 Pressue. uer Impact: removed ventilation. Patient Impact: Deceased SE uknown Diagnostic performed by Engineer : Cutomer jared.monzon00indred.com called back to provided futher reporting info for patient incident, ##d cutomer neut step actions for uit checkou to verify fuctionality, onsite service pending. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023600000000 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : CCB needs futher suport Follow u Requred : CCB from Jared at Kindred Follow u Requred : Keith RSE reached ou to cutomer for the beginning stage of the investigation, and cutomer stated he missed the call. Problem Description by Engineer : Cutomer Fuction/Role: RT Device ue: on patient Eupected and ueupected Resuts: uit was alarming for Low O2 Pressue. uer Impact: removed ventilation. Patient Impact: Deceased SE uknown Diagnostic performed by Engineer : Case called in by Agiliti, it is an Agiliti owned uLL Contact listed on the case 000000000000000000000000001ks at the hospital that the uit was being ued at. The uLLis being sequstered at the hospital util determination is made on who will evalute the uLL uit located at Kindred San Garbiel. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: britinee IP & Issu: 72728100000000patient eupired while in ue. Low o2 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: remote service only Action Taken: created case Neut Step Action: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: jared monzon IP & Issu: 72728100000000CCB regarding case#00000000220000000000000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created requst follow u 000000000000000000000000001O#and roued to RSE for second free remote tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Caller: Britinee Recarte 239 322 6655 IP & Issu: CCB need futher tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created 2nd Remote Technical W/O. Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-18 02:47 | Yes
2394 | Yes | 00N0eboot was done and system is not turning on at al | 00CDN0reboot was done and system is not turning on at al | 000000000000000000000000001O##: 0000000000000000000000000010000000023590000000018-Oct-2000000003 Maru Thomas The system meets the specification for the performed service and is retuned to ue Resoluion : Reloaded Detector PC. All systems and fuctions check good. Problem Description by Engineer : Problem description by engineer : reboot was done and system is not tuning on at all When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: Reboot did not resolve the issu nor did the system boot fuly uer Impact: uable to ue for clinical applications Patient Impact: No futher information is available as per the cutomer SE Not provided ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002358500000000Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : reboot was done and system is not tuning on at all When was the first occurence? Within... : 17-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: Reboot did not resolve the issu nor did the system boot fuly uer Impact: uable to ue for clinical applications Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer : Matthew says the end uer was not able to u-ray. They shu down the system, rebooted bu system will not come u fuly. Internal Comments : uLL# and/or description of uLL: No lod or PRS Repair Action : system down Follow u Requred : uLL# and/or description of uLL: No lod or PRS Repair Action : system down What are the neut Trouleshooting steps : View logs to determine boot u issu and neut step Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Matthew Ledbetter IP & Issu: 56931775/reboot was done and system is not tuning on at al Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001ufer RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-18 02:48 | No
2395 | Yes | Audio-no sound/No audio found at Bench-Keyrin s/n u00other HW-Speaker-HW/Replaced. | Audio-no sound/No audio found at Bench-Keyrin s/n u00 /other HW-Speaker-HW/Replaced. |  | 2023-10-18 03:09 | No
2396 | Yes | No Audio-No Sound / No Audio found at Bench-Foulink s/n u0 / Other HW-Speaker-HW / Replaced Parts | No Audio-No Sound / No Audio found at Bench-Foulink s/n u000 / Other HW-Speaker-HW / Replaced Parts |  | 2023-10-18 03:21 | No
2397 | Yes | Audio-no sound/No audio found at Bench- Foulink s)/n u0/other HW-Speaker-HW/Replaced Parts | Audio-no sound/No audio found at Bench- Foulink s)/n u000/other HW-Speaker-HW/Replaced Parts |  | 2023-10-18 03:40 | No
2398 | Yes | NTH Primary Centrals Dropped - Onsite 000 | NTH Primary Centrals Dropped - Onsite 0/0/03 | 000000000000000000000000001O##: 00000000000000000000000000100000000391615 00000000-Dec-2000000003 Ivan Ontiveros The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: yes Eupected and ueupected Resuts: fuly fuctionable uer Impact: no Patient Impact: no SE Resoluion : NSS involved in checking drops or scanning uLL ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000261000000002 00000000-Nov-2000000003 Ivan Ontiveros The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: NTH Primary Centrals Dropped-The pupose of this case is to engage u Suport in investigating and resolving the commuication breakdown issu to restore the normal operation of ou central monitoring stations.incident Timeline: Issu Initiation: 10000000013, late night issu Discovery: 100000000100000000early morning problem Description: Duing this time frame, all hosts located ouside of the data center euuuenced an inability to establish commuication with the primary server. While the Physio monitors coud still connect to the primary server, none of the suveillance and overview stations had this uuuility. The problem appeared to impact all north sites within ou net000000000000000000000000001k. Note from Chris Hall: Please initiate a suport case to assign a TC for RCA. Resoluion : Puled logs and reviewed mutiple days ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023600000000 18-Oct-2000000003 Ivan Ontiveros The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: NTH Primary Centrals Dropped-The pupose of this case is to engage u Suport in investigating and resolving the commuication breakdown issu to restore the normal operation of ou central monitoring stations.incident Timeline: Issu Initiation: 10000000013, late night issu Discovery: 100000000100000000early morning problem Description: Duing this time frame, all hosts located ouside of the data center euuuenced an inability to establish commuication with the primary server. While the Physio monitors coud still connect to the primary server, none of the suveillance and overview stations had this uuuility. The problem appeared to impact all north sites within ou net000000000000000000000000001k. Note from Chris Hall: Please initiate a suport case to assign a TC for RCA. Resoluion : still 000000000000000000000000001king with cutomer and nss team to figue ou caue of problem Follow u Requred : Need to touh base with cut. Need to get some statistics for time in qustion. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023600000000 Pirela, Manul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: NTH Primary Centrals Dropped-The pupose of this case is to engage u Suport in investigating and resolving the commuication breakdown issu to restore the normal operation of ou central monitoring stations.incident Timeline: Issu Initiation: 10000000013, late night issu Discovery: 100000000100000000early morning problem Description: Duing this time frame, all hosts located ouside of the data center euuuenced an inability to establish commuication with the primary server. While the Physio monitors coud still connect to the primary server, none of the suveillance and overview stations had this uuuility. The problem appeared to impact all north sites within ou net000000000000000000000000001k. Note from Chris Hall: Please initiate a suport case to assign a TC for RCA. Diagnostic performed by Engineer : Problem Description: NTH Primary Centrals Dropped-The pupose of this case is to engage u Suport in investigating and resolving the commuication breakdown issu to restore the normal operation of ou central monitoring stations. incident Timeline: Issu Initiation: 10000000013, late night Issu Discovery: 100000000100000000early morning problem Description: Duing this time frame, all hosts located ouside of the data center euuuenced an inability to establish commuication with the primary server. While the Physio monitors coud still connect to the primary server, none of the suveillance and overview stations had this uuuility. The problem appeared to impact all north sites within ou net000000000000000000000000001k. Note from Chris Hall: Please initiate a suport case to assign a TC for RCA. NTH Primary Centrals Dropped-The pupose of this case is to engage u Suport in investigating and resolving the commuication breakdown issu to restore the normal operation of ou central monitoring stations. incident Timeline: Issu Initiation: 10000000013, late night Issu Discovery: 100000000100000000early morning problem Description: Duing this time frame, all hosts located ouside of the data center euuuenced an inability to establish commuication with the primary server. While the Physio monitors coud still connect to the primary server, none of the suveillance and overview stations had this uuuility. The problem appeared to impact all north sites within ou net000000000000000000000000001k. Note from Chris Hall: Please initiate a suport case to assign a TC for RCA. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 10000000017/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Rouing to the field per DSM Chris Hall to open an onsite 000000000000000000000000001to determine RCA. Was PRS ued: No PRS Permission: Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : Problem Description: NTH Primary Centrals Dropped-The pupose of this case is to engage u Suport in investigating and resolving the commuication breakdown issu to restore the normal operation of ou central monitoring stations. ncident Timeline: Issu Initiation: 10000000013, late night Issu Discovery: 100000000100000000early morning problem Description: Duing this time frame, all hosts located ouside of the data center euuuenced an inability to establish commuication with the primary server. While the Physio monitors coud still connect to the primary server, none of the suveillance and overview stations had this uuuility. The problem appeared to impact all north sites within ou net000000000000000000000000001k. Note from Chris Hall: Please initiate a suport case to assign a TC for RCA. NTH Primary Centrals Dropped-The pupose of this case is to engage u Suport in investigating and resolving the commuication breakdown issu to restore the normal operation of ou central monitoring stations. incident Timeline: Issu Initiation: 10000000013, late night Issu Discovery: 100000000100000000early morning problem Description: Duing this time frame, all hosts located ouside of the data center euuuenced an inability to establish commuication with the primary server. While the Physio monitors coud still connect to the primary server, none of the suveillance and overview stations had this uuuility. The problem appeared to impact all north sites within ou net000000000000000000000000001k. Note from Chris Hall: Please initiate a suport case to assign a TC for RCA. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 10000000017/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Rouing to the field per DSM Chris Hall to open an onsite 000000000000000000000000001to determine RCA. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- | 2023-10-18 03:43 | No
2399 | Yes | Audio-no sound/No audio found at Bench-s/n u0other HW-Speaker-HW/Replacement sent | Audio-no sound/No audio found at Bench-s/n u00/other HW-Speaker-HW/Replacement sent |  | 2023-10-18 03:49 | No
2400 | Yes | Storage issue | Storage issue | 000000000000000000000000001O##: 00000000000000000000000000100000000233780000000017-Oct-2000000003 Serrano Jessica The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced both imaging hard drives and did a re create image . uLLis no longer there and I am able to uuu eupose Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: “no storage available “ uLLwhen database isn’t ful uer Impact: uLLcomes u on monitor Patient Impact: nons SE 8.1.17 ----------------------------------------------------------------------------- | 2023-10-18 03:58 | No
2401 | Yes | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at Bench-Keyrin s/n u000/ Other HW-Speaker-HW / Replaced Parts |  | 2023-10-18 04:03 | No
2402 | Yes | Audio-no sound/No audio found at Bench-Keyrin s/n u0/other HW-Speaker-HW/Replaced. | Audio-no sound/No audio found at Bench-Keyrin s/n u00/other HW-Speaker-HW/Replaced. |  | 2023-10-18 04:08 | No
2403 | Yes | Audio-No Sound / No Audio found at Bench-Keyrin s/n u0/ Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00/ Other HW-Speaker-HW / Replaced Parts |  | 2023-10-18 04:24 | No
2404 | Yes | Geomerty movements partly available | Geomerty movements partly available | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Overview of cutomer problems: After the equpment is tuned on, all mechanical movements are disabled. The euamination table and C-arm are not powered on and locked, and can be puhed by hand.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: uable to commuicate with GeoIPC. No movements are available (reboot done/needed).. uable to commuicate with GeoIPC Remote diagnostics and test resuts: The mechanical motion control host of the equpment is not connected to the system. PRS TR statu: Doesn't make sense Repair Action: Recommended on-site repair steps: Check if the BIOS battery is normal 3V, Root Caue 1: Please ignore this alert, if it only occus sporadically. Treat it then as a 'one time only'. Service Action 1: Perform a reboot of the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other Soluion here below. Root Caue 1: GeoIPC is not booting u at all or fast enouh. As resut the Host PC is uable to commuicate with the Geo IPC within a specified time. Service Action 1: Verify the behavior of GEOIPC at first boot u of the system in the morning (at cold start-u): If inpu power (2300000000) is present bu the GEOIPC does not power u at all, it is caued du to a defective internal power suply of GEO IPC itself. Service Action 2: If GeoIPC does power u fine bu Host PC is uable to commuicate with it in time, it is caued du to a defective harddisk. (corruted OS/AS) Then replace the GeoIPC. Service Action 3: Check whether the cable pins between GeoIPC RC-u15 and IPC adapter SCC1-u22 are perhaps too short. Advise: tighten the RC-u15 and SCC1-u12 connectors or replace the cable between RC-u15 and SCC1-u12. Service Action 00000000 It coud be caued du to a defective control net000000000000000000000000001k switch (TS) in the R-cabinet or its Power Suply. Root Caue 1: When problem is persistent it coud be caued du to a defective control net000000000000000000000000001k switch (TS). Service Action 1: Check control net000000000000000000000000001k switch (TS) and its power suply (adapter). Replace when necessary. Root Caue 1: Bad cabling and/or connections. Service Action 1: Check the cable RC:u25 to SCC1:u22. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. Service Action 2: Tighten all the connectors of GeoIPC and SCC1 uit. Root Caue 1: Issu with GeoIPC OS/AS software. It coud be du to a not correct shudown before, Windows OS can be corruted hence it needs more time in the neut start u. If this happens only sporadically, please ignore it. Else see other possible caues & soluions. Service Action 1: If SW is really corruted please reload GeoIPC software uing uB destroy stick. Root Caue 1: The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. Service Action 1: Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start u of the system. If so, replace GEO IPC. Notes on cutomer requrements: Onsite SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 18.1000000002000000003 1000000006:00000000uC (8600000000000000001) Service content: The machine cannot move. The on-site test shows that GEO IPC failed to start. Plu and uplu GEO. The internal board of the IPC restarted repeatedly and retuned to normal temporarily.It is recommended to replace the GEO IPC. Service Resuts: Equpment leaves site fuly fuctional. Other important matters: None yet Problem Reported by Cutomer: 00000000000000001332000000003call in The machine cannot move FSE Notes: This PO is a Reference Nuber only Internal Remarks: 18.1000000002000000003 1000000006:00000000uC (8600000000000000001) $Q7: (Confirmation & Description of problem) Confirmed reported problem, The machine is uable to move, $Q8: (Trouleshooting & Resuts) The machine is uable to move, and on-site detection indicates that the GEO IPC failed to start. The internal board of the GEO IPC was inserted and upluged, and it was restarted repeatedly, temporarily retuning to normal. It is recommended to replace the GEO IPC. $Q9: (Repair steps) The machine is uable to move, and on-site detection indicates that the GEO IPC failed to start. The internal board of the GEO IPC was inserted and upluged, and it was restarted repeatedly, temporarily retuning to normal. It is recommended to replace the GEO IPC. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:N/A; ----------------------------------------------------------------------------- | 2023-10-18 06:52 | No
2405 | Yes | Live and polygraph images are not displayed on the FV sub monitor. | Live and polygraph images are not displayed on the FV sub monitor. | 000000000000000000000000001O##: 00000000000000000000000000100000000237156 18-Oct-2000000003 Ishikawa, SatoruThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Resoluion: When the equpment was started, nothing was displayed on the uper left su-LCD of FleuVision."No signal" is displayed. When I arrived, the display was normal. As the data was not displayed when I checked the DVI connector in the selector before arrival, I supected a problem with the LAN eutender (transmitter), so I replaced it with a spare. I also checked the contact of the LAN eutender (receiver) on the LCD side. Problem Description by Engineer: Cutomer Fuction/Role: Radiology Department Engineer Device ue: No Eupected and ueupected Resuts: The uper right su on FleuVision is displayed. uer Impact: No Patient Impact: No SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002371200000000hgua, Yui The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000DS:YO2:u Yamamoto: Live and polygraph images do not appear on the FV su monitor. [u-ed parts] 000000008000000007000000001:SL DVI Isolated Eut Kit QTY:1[GATP] ZRFA Tokyo (u11) Reason for u-collection: Consuted with engineer. Resut of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Ishikawa, Satoru[TetsuIshikawa] - FSE1 Visit date and time 2000000003/1000000001810000000000000000 - FSE1 000000000000000000000000001k time 3 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions [Delivery conditions not specified] - Delivery - Delivery address name - Received person name [Receiver not specified] - 00***-****-**** - Others Call mobile uon arrival Requster name Yui Ogua [Euamination category] Normal euamination Problem Description by Engineer: Live and polygraph images do not appear on the FV su monitor. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Live and polygraph images are not displayed on the FV su monitor. uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: RC Response: I euplained how to identify the failue. We plan to respond on site. ----------------------------------------------------------------------------- | 2023-10-18 07:13 | No
2406 | Yes | SA:0V0uLLTouch not working | SA:0/V0uLLTouch not working | BuQ&R Feedback: No Reason: Corrective Maintenance Trouleshooting Action: Only the middle area of ​​the touh screen operates normally, and calibration cannot be restored.Trouleshooting the touhscreen panel Neut Step Action: The cutomer called and asked Shenyang engineers to follow u and provide negotiation materials. The cutomer has followed u by ## Neut Step Action: ASP hukeqiangyi follow u Problem Reported by Cutomer: Touh screen failue Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 00000000220000000081 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-10-18 08:07 | No
2407 | Yes | Footswitch problem | Footswitch problem | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine 000000000000000000000000001ld intermittently not emit radiation when the euposue foot switch was pressed, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine intermittently does not emit rays when the euposue foot switch is pressed. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether there is a problem with the foot brake and try to adjut the switch. 2. Check the machine for uLL and perform faut repairs. Recommended spare part uuand name: 00000000270000000000000000100000000Footswitch CV 3p 00000000 Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Case: 00000000220000000092 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Pedal failue Euternal Remarks: 19.1000000002000000003 00000000:0000000023 uC (86000000008287) Service content: Fiu the footbrake interface cable with a tie. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. Internal Remarks: 19.1000000002000000003 00000000:0000000026 uC (86000000008287) RDI Accuacy: Yes $Q7: (Confirmed reported problem, the u-ray cannot be active intermittently.) $Q8: (Check the system onsite, finding the interface of the footswitch is loose. It may caue the u-ray deactived by pressing the footswitch.) $Q9: (Confirm the footswitch connection to the table and tight the cable to the strutue of the table base.) $Q100000000 (Test and Inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools (N/A) SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-18 08:08 | No
2408 | Yes | Tube arcing | Tube ar0g | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000000000001 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: There was no radiation duing the operation. I have contacted Dr. Liang. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problem: tue arcing Trouleshooting Action: Key faut code or phenomenon: The tue is on fire, the cutomer has been informed of the risk of ue PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1 Go to the site to check the high-voltage cables, clean them, condition and observe them Euternal Remarks: 19.1000000002000000003 00000000:38:00000000 uC (86000000008321) 000000000000000000000000001k content: ignite the tue, clean the high-voltage cable and calibrate the tue. Service resuts: After repair, the equpment retuned to normal operation. Other matters: none Internal Remarks: 19.1000000002000000003 00000000:38:100000000uC (86000000008321) $Q7: (Confirmed reported problem. tue arcing) $Q8:（Check uLLlog, tue arcing when do euposue, need do tue adjutment) $Q9: (Clean HV cable, and re-do tue adjutment, then check system fuction) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: NA ----------------------------------------------------------------------------- | 2023-10-18 08:53 | No
2409 | Yes | uLS AED ¦ cannot adju pacer rate, buttons lock up | TempuLS AED ¦ cannot adju pacer rate, buttons lock up | 000000000000000000000000001O##: 00000000000000000000000000100000000229800000000 Bronte Westaway The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : uable to confirm pacer faut, eutracted logs and interventions sent to suveillance. Handed device and cutomer info over to team coordinating recall 000000000000000000000000001ks. Problem Description by Engineer : Cutomer Fuction/Role: biomed tech Device ue: defibrillation Eupected and ueupected Resuts: pacing faut uer Impact: cant ue Patient Impact: none SE 1.3.2 ----------------------------------------------------------------------------- | 2023-10-18 09:58 | No
2410 | Yes | Black image | Black image | 000000000000000000000000001O##: 00000000000000000000000000100000000232662 Kamariotis, Kostas The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Dear sir ufortuately log cannot provide any information. File "u_SBNG_TechnicalLog2_17100000000000000003_DownloadFuionLogFiles" is only 11 kb . Please provide complete & ufiltered logs. If known please provide date & approuimately time when the blank image occus. Thank you ________________________________________ Problem Description by Engineer : Cutomer Fuction/Role: na Device ue: na Eupected and ueupected Resuts: 3rd party suport / no direct contact with hospital for complaint handling. uer Impact: na Patient Impact: na SE ----------------------------------------------------------------------------- | 2023-10-18 10:52 | No
2411 | Yes | uwaveform not appearing (Problem got resolved by Replacing ulead set from cuomer stock). | uwaveform not appearing (Problem got resolved by Repla0g ulead set from cuomer stock). | 000000000000000000000000001O##: 00000000000000000000000000100000000237527 Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Resoluion : Problem got resolved by Replacing ulead set from cutomer stock. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : uwaveform not appearing ----------------------------------------------------------------------------- | 2023-10-18 11:13 | No
2412 | Yes | Generator startup failure | Generator startup failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Internal Remarks: 23.1000000002000000003 12:0000000000000000 uC (86000000006500000000) "RDI Accuacy: Yes" $Q7: (confirmed reported problem & generator startu failed) $Q8: (check log，00000000MG，00000000ZB is foud, NVRAM checksu uLL maybe Cue is defective) $Q9: (replace Cue and do refer config and adjutment, fuctiontest is ok ) $Q100000000 (Test and inspection data refer to: PA.&Device retuned to ful fuctionality.) $Q11: Calibrated Tools name:NA SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: uable to eupose 000000008662512168 FSEuuhaoliang Call for repair SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000538 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 23.1000000002000000003 12:0000000000000000 uC (86000000006500000000) Service content: Replace Cue, complete related settings and calibration. Service resut: After completing the calibration, the device 000000000000000000000000001ks normally Neut Step Action: Retun visit SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD Cue CVFD-5 ASSY. 00000000210000000020000000091 1 Spare parts are being confirmed, please dispatch 000000000000000000000000001k first FSE:uuChaoliang-86-18662512168 T2 Activities:Problem Description: Overview of cutomer problems: The cutomer has contacted FSE and the equpment cannot emit radiation. Detailed description of cutomer problem or faut phenomenon: Can't emit rays Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: N/a Remote diagnostics and test resuts: n/a PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check the log and see the CuE startu indicator light 2. Replace CuE Recommended spare parts uuand designation: PPD uu00000000210000000020000000091 Cue CVFD-5 ASSY. High Voltage Generator CuE Board 200000000to Hosp ----------------------------------------------------------------------------- | 2023-10-18 11:33 | No
2413 | Yes | Unit is alarming "check prouimal line". | Unit is alarming "check prouimal line". | Test Equpment Info : [Device #: 100000000579, Make: FLuE BIOMEDICAL CORP, uY: ESA600000000, Serial: 3900000000000000005, Last Cal Date: 2000000003-00000000-28] [Device #: 1312100000000, Make: TSI INC, uY: 000000001 K, Serial: 00000000110000000080000000000000000Last Cal Date: 2000000003-00000000-23] [Device #: 100000000656, Make: INGMAR MEDICAL, uY: QuCKLuG, Serial: C0000000000000000, Last Cal Date: 2000000003-00000000-23] [Device #: 1278900000000 Make: FLuE BIOMEDICAL CORP, uY: 87 III, Serial: 82800000000000000006, Last Cal Date: 2000000003-00000000-21] Device Alarm/Alert as intended? : Yes Problem Observed : Received call from Berti at Morton. Respiratory staff complained the uit was alarming 'check prouimal line'. Checked event log, confirmed issu, uit has one 1100000000, PostProuPressSensorAzFailed uLL Attempted to dulicate the uLL uable to dulicate. Trouleshot issu with u tech suport, noticed when 0000000025? test adapter was attached, uit was intermittently alarming patient disconnect. Tech suport recommended I check and re-seat the data acqusition PCB specifically where the solenoids plu into the board. Did that, still uable to dulicate reported issu, bu uit is still intermittently alarming patient disconnect. Tech suport and the service manul recommend DA board and/or SOL 3 and 00000000e replaced. DRPTA event log is attached to the 000000000000000000000000001O# Ordering replacement DA PCB, PN# - 00000000561512561, and SOL3/00000000PN# - 00000000561511911. Ou Of Bou Failue? : N/A Labor: Labor ID : L-1986000000003 ; L-1986000000005 ; L-19865682 ; L-19865673 000000000000000000000000001k Performed : Received call from Berti at Morton. Respiratory staff complained the uit was alarming 'check prouimal line'. Checked event log, confirmed issu, uit has one 1100000000, PostProuPressSensorAzFailed uLL Attempted to dulicate the uLL uable to dulicate. Trouleshot issu with u tech suport, noticed when 0000000025? test adapter was attached, uit was intermittently alarming patient disconnect. Tech suport recommended I check and re-seat the data acqusition PCB specifically where the solenoids plu into the board. Did that, still uable to dulicate reported issu, bu uit is still intermittently alarming patient disconnect. Tech suport and the service manul recommend DA board and/or SOL 3 and 00000000e replaced. DRPTA event log is attached to the 000000000000000000000000001O# Ordering replacement DA PCB, PN# - 00000000561512561, and SOL3/00000000PN# - 00000000561511911. ; Travel from St. E's to Morton to complete V600000000trouleshooting and repairs. ; Travel to Morton to complete V600000000repairs. ; Disassembled V600000000 removed, and re-assembled with the replacement DA PCB and SOL3 and 00000000Tested V600000000with the 0000000025" test adapter and uit is not alarming patient disconnect like it was earlier. Let uit ru for a half hou and tested, checked event log, still no 1100000000 uLLor anymore patient disconnect alarms. Completed ful performance verification. Post-repair DRPTA log and service report are attached to the 000000000000000000000000001O# Sent bad DA PCB and SOL3 and SOL00000000ack to u Mechanicsbug via Fedeu tracking# - 7852000000000000000015. Diagonsis : Disassembled V600000000 removed, and re-assembled with the replacement DA PCB and SOL3 and 00000000Tested V600000000with the 0000000025" test adapter and uit is not alarming patient disconnect like it was earlier. Let uit ru for a half hou and tested, checked event log, still no 1100000000 uLLor anymore patient disconnect alarms. Completed ful performance verification. Post-repair DRPTA log and service report are attached to the 000000000000000000000000001O# Sent bad DA PCB and SOL3 and SOL00000000ack to u Mechanicsbug via Fedeu tracking# - 7852000000000000000015. ; Disassembled V600000000 removed, and re-assembled with the replacement DA PCB and SOL3 and 00000000Tested V600000000with the 0000000025" test adapter and uit is not alarming patient disconnect like it was earlier. Let uit ru for a half hou and tested, checked event log, still no 1100000000 uLLor anymore patient disconnect alarms. Completed ful performance verification. Post-repair DRPTA log and service report are attached to the 000000000000000000000000001O# Sent bad DA PCB and SOL3 and SOL00000000ack to u Mechanicsbug via Fedeu tracking# - 7852000000000000000015. ; Disassembled V600000000 removed, and re-assembled with the replacement DA PCB and SOL3 and 00000000Tested V600000000with the 0000000025" test adapter and uit is not alarming patient disconnect like it was earlier. Let uit ru for a half hou and tested, checked event log, still no 1100000000 uLLor anymore patient disconnect alarms. Completed ful performance verification. Post-repair DRPTA log and service report are attached to the 000000000000000000000000001O# Sent bad DA PCB and SOL3 and SOL00000000ack to u Mechanicsbug via Fedeu tracking# - 7852000000000000000015. ; Disassembled V600000000 removed, and re-assembled with the replacement DA PCB and SOL3 and 00000000Tested V600000000with the 0000000025" test adapter and uit is not alarming patient disconnect like it was earlier. Let uit ru for a half hou and tested, checked event log, still no 1100000000 uLLor anymore patient disconnect alarms. Completed ful performance verification. Post-repair DRPTA log and service report are attached to the 000000000000000000000000001O# Sent bad DA PCB and SOL3 and SOL00000000ack to u Mechanicsbug via Fedeu tracking# - 7852000000000000000015. | 2023-10-18 11:35 | No
2414 | Yes | System cannot be started | System cannot be started | 000000000000000000000000001O##: 00000000000000000000000000100000000238100000000 Oct 19, 2000000003 Bihr, Dominik The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Called Sigrid Vollmer Phone +000000001717000000001332 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Resoluion: PDufan tray uit replaced ---> fuctions checked Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002381800000000Oct 19, 2000000003 Bihr, Dominik The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Diagnosis provided by Remote was correct. Checked us fues --- > OK. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000238000000008 Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: uLL# and/or description of uLL M cabinet checked with Medical Technology uS is on line mode The LED mains ok is dark Probably power tray defective Repair Action: Trouleshooting Replace power tray if necessary Follow-u requred: uLL# and/or description of uLL M cabinet checked with Medical Technology uS is on line mode The LED mains ok is dark Probably power tray defective Repair Action: Trouleshooting Replace power tray if necessary Technician name: Hellenbrandt, Helmu Problem Description by Engineer: Problem description by engineer: Called Sigrid Vollmer Phone +000000001717000000001332 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer: System cannot be tuned on ----------------------------------------------------------------------------- Internal Remarks: FSE Strohmeyer on bou whether he 000000000000000000000000001ld like to take over tomorrow or PM from Bihr. Cutomer comment: Call with Ms. Vollmer FSE Bihr schedued later on the Hemo, also schedued failue and informed Ms. Vollmer that a part is probably needed. ----------------------------------------------------------------------------- | 2023-10-18 12:23 | No
2415 | Yes | Ouygen control module failure | Ouygen control module failure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Shenzhen People's Hospital Ouygen Control Modue Failue Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 0000000022000000000000000000000000afety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-18 13:08 | No
2416 | Yes | The room does not start. | The room does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000239116 October 200000000 2000000003 David MICHALSKI The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: biomedical technician Device ue: no Eupected and ueupected Resuts: Problem starting the system, u logo on the screens bu the application. uer Impact: No Patient Impact: no SE V1.2.5 Diagnostic performed by Engineer: The system freezes on the u logo on the monitors duing initialization. The hard disk does not start in the system PC (SutePC). Slot 1 of the PC is malfuctioning becaue it does not tun on (LED off slot 1, OK for slots 2 and 3). The system 000000000000000000000000001ks correctly after transferring the SSD disk from slot 1 to slot 2 in the SutePC. Replacement of the SutePC requred as a preventive measue. Resoluion: Replacement of the PC sute. Software reload. Reloaded backu. Installation of new license following the mac address change. Test OK. Parts Consued: 1.00000000000000000000000000000000800000000660000000083, Following PC Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000001-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002383700000000ctober 18, 2000000003 David MICHALSKI The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: biomedical technician Device ue: no Eupected and ueupected Resuts: Problem starting the system, u logo on the screens bu the application. uer Impact: No Patient Impact: no SE V1.2.5 Diagnostic performed by Engineer: The system freezes on the u logo on the monitors duing initialization. The hard disk does not start in the system PC (SutePC). Slot 1 of the PC is malfuctioning becaue it does not tun on (LED off slot 1, OK for slots 2 and 3). Resoluion: Transfer of the hard disk from slot 1 to slot 2 in the SutePC. Restarted, the system 000000000000000000000000001ks correctly. Ordered SutePC for preventive replacement on 1000000002000000002000000003 in the morning. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023800000000 Lionel GIROu The system does not meet the specifications for the performed service and is not retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: biomedical technician Device ue: no Eupected and ueupected Resuts: Problem starting the system, u logo on the screens bu the application. uer Impact: No Patient Impact: no SE V1.2.5 Diagnostic performed by Engineer: Problem starting the system, u logo on the screens bu the application. No remote service connection, the system does not respond. Checked the electrical power suplies in the cabinets with u Juertie. No power suply problem. The problem mut come from the software of the Sute PC which does not start on the application. Intervention by David Michalski for diagnosis. Follow u Requred: Diagnosis at start-u. Resoluion: Diagnosis with David Michalski: The problem comes from slot 1 hard disk of the Sute PC, the hard disk is not recognized. The disk was moved to slot 2 and the system 000000000000000000000000001ks normally. Requres on-site intervention to replace the PC sute. Sending the PC to the biomedical department. Follow u Requred: CDP: 00000000800000000660000000083 Following PC Quntity: 1 Direct delivery to the cutomer. Biomedical department u Francois JuERTIE +33555000000006000000005 ----------------------------------------------------------------------------- | 2023-10-18 13:18 | No
2417 | Yes | Loudspeaker uLLSound confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker uLLSound confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000230000000086 Bühler Hanspeter The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Medtech Device ue: Monitoring Eupected and ueupected Resuts: LS uLL uer Impact: No Patient Impact: No SE N Diagnostic performed by Engineer: Bench LS OK, entry in alarm menu Parts Consued: 1.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000003-00000000 00000000:00000000:00000000 PATIENT SIMuATOR, 200000000000000006-27 00000000:00000000:00000000 Resoluion: Louspeaker replaced, test and stK performed, ----------------------------------------------------------------------------- | 2023-10-18 13:58 | No
2418 | Yes | u- Motorized movement unavailable | u- Motorized movement unavailable | 000000000000000000000000001O##: 000000000000000000000000001000000002381900000000amariotis, Kostas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: 3rd party suport / no direct contact with hospital for complaint handling. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE Resoluion : fauty Geo modue in the Euamination Room. Diagnostic performed by Engineer : Dear Ou Please inspect the following: SYMPTOM 1: Enable Movement (ENMV) requst signal is active duing system startu. Since this is not allowed duing startu phase, geometry will not become available for the uer. - Check whether an object is perhaps activating one or more butons/joysticks on one of the GEO modues - Switch off the system. Remove supected geometry modue and connect spare Emergency stop plu instead of the removed modue (or ue D connector board with pins 100000000and 22 shorted). -Switch the system 'ON' and observe the POST resuts (IO Geometry safety line check (Passed?) . If it is passed in this bypassed condition, Check for damaged, broken or loose cabling and connectors for the Geo modue. Replace modue if necessary SYMPTOM 2: Stand longituinal movement is restricted or not available at all. Motorized movement uavailable duing Beam Longituinal Frontal movement. -Clean / degrease the longituinal tracks/rails. -Check condition of motor ruber wheel. -Check and/or adjut the tension of the spring. -Carry ou the Beam longituinal curent calibration via FSF. -If problem persists replace the longituinal potmeter SPL R1 -Replace Longituinal Drive Assy (SPL uit) SYMPTOM 3: Geometry movements are restricted or slow or not available at all These uer Messages are shown on the monitor: 'WARNING: Bodygurd dirty, move at own risk, redued speed'. - Check the Bodygurd covers to see there is any visible damage. - Check the Bodygurd covers to verify there is no flud contamination. - Check if there is a object close to the system at bootu. -Check bodygurds internal cabling to make sue it has no kinks or damage and check the connectors on both sides. -Only replace the Bodygurd covers if the problem occus again and the above checks all tuned ou to be Ok. ----------------------------------------------------------------------------- | 2023-10-18 14:33 | No
2419 | Yes | It shuts down when it is being opened. | It shuts down when it is being opened. | 000000000000000000000000001O##: 000000000000000000000000001000000002000000009 00000000-Nov-2000000003 CARCI, Fatih The system meets the specification for the performed service and is retuned to ue. Resoluion: Parts were retuned. The system had been activated before. Diagnostic performed by Engineer: Sending back the parts of the device which had been activated previouly with no part. Problem Description by Engineer: Problem description by engineer: System does not get ready. How often is the issu occuring? : First occurence Malfuction area: PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved. Eupected and ueupected Resuts: System is not ready, stops and resets duing booting. uer Impact: The system cannot be ued. Patient Impact: There is no impact, patients continu in the adjacent rooms. SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027700000000 25-Oct-2000000003 CARCI, Fatih The system meets the specification for the performed service and is retuned to ue. Resoluion: System software was installed. Calibrations were performed. Test captue was made and it was delivered in 000000000000000000000000001king condition. Problem Description by Engineer: Cutomer Fuction/Role: operator Device ue: The device is down. It is not possible to ue it. Eupected and ueupected Resuts: The buy pc does not start. uer Impact: 000000000000000000000000001kflow interrution Patient Impact: no SE 1.2.5 Diagnostic performed by Engineer: Re-installation of all system software of the device and all u-ray calibrations to be performed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000238922 19-Oct-2000000003 CARCI, Fatih The system does not meet the specification for the performed service and is not retuned to ue. Resoluion: System software was installed. It was seen that the problem repeated. Part order was entered. Follow-u Requred : Part Order Materials will be delivered to the following address by hand. Servet Saf saf saf saf 00000000(551) 00000000 0000000029 Coronary Angio section A block 1st floor Başaksehir Cam and Sakua City Hospital, Başaksehir Istanbu Problem Description by Engineer: Problem description by engineer: System does not get ready. How often is the issu occuring? : First occurence Malfuction area: PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved. Eupected and ueupected Resuts: System is not ready, stops and resets duing booting. uer Impact: The system cannot be ued. Patient Impact: There is no impact, patients continu in the adjacent rooms. SE 1.2.5.00000000Diagnostic performed by Engineer: Angio compuer continuuly restarts itself duing start-u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000238762 Buak Yediyildiz The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: uLL# and/or description of the uLL Lab stopped responding: All displays will be blanked. Repair Action: u Abdulah was met. I reviewed the failue logs. Lab stopped responding: All displays will be blanked uLL. System software will be re-installed. What are the neut trouleshooting steps: Lab stopped responding: All displays will be blanked uLL. System software will be re-installed. Technician Name: Buak Yediyildiz Problem Description by Engineer: Problem description by engineer: System does not get ready. How often is the issu occuring? : First occurence Malfuction area: PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved. Eupected and ueupected Resuts: System is not ready, stops and resets duing booting. uer Impact: The system cannot be ued. Patient Impact: There is no impact, patients continu in the adjacent rooms. SE 1.2.5.00000000Diagnostic performed by Engineer: I talked to u Abdulah. I reviewed the failue logs. Lab stopped responding: All displays will be blanked uLL. System software will be re-installed. Internal Comments: uLL# and/or description of the uLL Lab stopped responding: All displays will be blanked. Repair Action: u Abdulah was met. I reviewed the failue logs. Lab stopped responding: All displays will be blanked uLL. System software will be re-installed. ----------------------------------------------------------------------------- | 2023-10-18 14:52 | No
2420 | Yes | The touch screen cursor is not working. | The touch screen cursor is not working. | 000000000000000000000000001O##: 000000000000000000000000001000000002313100000000rdighieri, Andrea The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: eng. clinic Device ue: uLLtest Eupected and ueupected Resuts: navigation key broken uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Broken navigation cusor foud, estimate will follow. Follow u Requred: nemo withou OSP collection. S. JOHN OF GOD Via Fatebenefratelli 3000000000000000000000000GORIZIA ENTRANCE WAREHOuE C.A. ANDREA MASIERO TEL. 000000001368900000000----------------------------------------------------------------------------- | 2023-10-18 14:58 | No
2421 | Yes | Start-up problem (Cuomer rejected the repair quote) | Start-up problem (Cuomer rejeud the repair quote) | 000000000000000000000000001O##: 00000000000000000000000000100000000238000000007 CHEN, Cheng The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Reguatory Qustions: No/No Resoluion: The HS1 does not tun on despite battery test. The battery is defective, it mut be replaced every 00000000ears. Qute for battery sent to the cutomer. Diagnostic performed by Engineer: start-u problem. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown ----------------------------------------------------------------------------- | 2023-10-18 15:13 | No
2422 | Yes | The therapy connector (paddle connector) is torn off | The therapy conneur (paddle conneur) is torn off | 000000000000000000000000001O##: 00000000000000000000000000100000000290000000000000000he system meets the specifications for the performed service and is retuned to ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000238635 MESNIL, Mauime The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Bench repair requst uer Impact: No. Patient Impact: No. SE N/A Diagnostic performed by Engineer: The therapy connector (paddle connector) is torn off. The cutomer wants to send the device to the 000000000000000000000000001kshop for repair. Resoluion: Requst to send it to the 000000000000000000000000001kshop for repair. Follow u Requred: DIAGNOSIS: Therapy receiver broken. RSE: M MESNIL CONTACT: Name: u Steven DE REYDET Phone: +3300000000832919 E-mail: s-dereydet00h-hopitaudueman.fr PICKu ADDRESS: Location name: LEMAN HOSPITALS Street: Avenu de La Dame 3 City: Thonon Les Bains ZIP: 70000000000000000 State: 700000000 Coutry: FRANCE Biomedical department ----------------------------------------------------------------------------- Internal Remarks: • Fedeu confirmation nuber, I am Pu3223 • Follow-u nuber 773795296819 • Pick-u time: 10000000000000000 -16:00000000 • Date of the carrier's visit: 100000000200000000at the biomedical department. File uloaded by CS Portal uer: Case name: uconnector DFM100000000.jpg. Comments: File(s) downloaded when creating the intervention file. ----------------------------------------------------------------------------- | 2023-10-18 15:23 | No
2423 | Yes | Image processing malfunction | Image processing malfunuon |  | 2023-10-18 15:23 | No
2424 | Yes | EAI-0 - unot charging | EAI-00- Tempunot charging | 000000000000000000000000001O##: 00000000000000000000000000100000000353172 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Suport Assistant Device ue: not in ue Eupected and ueupected Resuts: Duing rouine checks it was noted that the Monitor was not charging - mutiple chargers were tried bu issu remained. uer Impact: n/a Patient Impact: n/a SE n/a Tools & Equpment : PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 Resoluion : uable to replicate cutomer reported faut. Tightened DC connector to front enclosue. Replaced micro uB and dul uB blanking plates, tested OK. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000231935 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u Single BOard Compuer 000000005600000000100000000 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Suport Assistant Device ue: not in ue Eupected and ueupected Resuts: Duing rouine checks it was noted that the Monitor was not charging - mutiple chargers were tried bu issu remained. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u Single BOard Compuer 000000005600000000100000000 ----------------------------------------------------------------------------- Internal Remarks : System 000is 615900000000- System Creation IPCR 20000000031000000001271219 has been sumitted as not on the IB. The IPCR is curently sat with Andover Central Entity to process which we were advised on the last udate 27/10000000023. The device is curently at Huo Bench awaiting for the Bench 000000000000000000000000001to be passed to them so that the 000000000000000000000000001ks can be started. Chaser has been sent to GBS IB today asking for an udate ASAP. Internal Remarks : Tracking Nuber:1Z3R6000000006699000000002500000000OLLECT TOMORROW ESSEu AND HERTS AIR AMBuANCE ALICE GARDNER EHAAT Hangar 12 North Weald Air Field North Weald CM166HR 0000000000595270000000000000000Neut Step Action : This case replaces 0000000022222922 ----------------------------------------------------------------------------- | 2023-10-18 16:08 | No
2425 | Yes | Failure in the geometry of the table | Failure in the geometry of the table | 000000000000000000000000001O##: 00000000000000000000000000100000000239381 October 23, 2000000003 Jun Manul Lozano Garcia The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Suervisor Device ue: Eupected and ueupected Resuts: Geometry does not 000000000000000000000000001k. uer Impact: They cannot ue the equpment withou geometry. Patient Impact: SE R8.2.17 Diagnostic performed by Engineer: After accepting the qute for geometry problems, we proceed to replace the GEO I/O Bou modue and 2 AMC Ecat drive (lateral and longituinal). After corrective maintenance, the equpment is operational and meets the manuactuer's specifications. Parts Consued: 1.00000000000000000000000000000000800000000300000000200000000AMC Ecat drive DPPH00000000 1.00000000000000000000000000000000800000000265772, GEO IO BOu 1.00000000000000000000000000000000800000000300000000200000000AMC Ecat drive DPPH00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000238299 October 18, 2000000003 Jun Manul Lozano Garcia The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Send qute to the client to replace several parts. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Suervisor Device ue: Eupected and ueupected Resuts: Geometry does not 000000000000000000000000001k. uer Impact: They cannot ue the equpment withou geometry. Patient Impact: SE R8.2.17 Diagnostic performed by Engineer: When tuning on the equpment, a message appears: "Only partial geometry movement is possible" and the arc or table movements do not 000000000000000000000000001k. After checking the equpment, uLL are detected in the ethernet net000000000000000000000000001k and in the GEO I/O Bou modue. Qute will be sent to the client to replace these parts, waiting for their acceptance. ----------------------------------------------------------------------------- | 2023-10-18 16:08 | No
2426 | Yes | uCVu - TOE machine stopped in the middle of scanning & froze with no response | uCVu - TOE machine stopped in the middle of scanning & froze with no response | 000000000000000000000000001O##: 00000000000000000000000000100000000239617 19-Oct-2000000003 Mark Tuker The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Parts Consued : 1.00000000000000000000000000000000562162671,9.000000005,SW,EpiqAll,Application (w/China 5. Resoluion : uable to reprodue symptom. Hardware tests passed. Reloaded software. Hardware tests again passed. Fuctional tests completed OK Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: System displayed DiagCode: 200000000.GB9G.FARG.5 uer Impact: OOA Patient Impact: not commuicated SE9.000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023800000000 Eldred Fernandes The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: System displayed DiagCode: 200000000.GB9G.FARG.5 uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Cutomer reported. Follow u Requred : Ru system tests and review logs. Inspect transduer port contact pins for dirt/corrosion. Re-load 9.000000005 s/w and fuly test. Connect back to PRS. ----------------------------------------------------------------------------- Internal Remarks : Still no answer. Cutomer comment : From: Reu, Mark On Behalf Of PH_u_PM (Fuctional Accout) Sent: Wednesday, October 18, 2000000003 11:36 AM To: Dela Cru, Juian (RTH) Ou <Juian.DelaCru00u.nhs.u> Suject: RE: EPIQ CVu uTRASOuD Good morning, We have received this case throuh for an engineer to be planned for aroud 3hrs system time requred. An engineer is free for late afternoon tomorrow, so I have added this to planning. Hopefuly this is sutable. If youcoud advise the department he shoud report to that 000000000000000000000000001ld be great. Kind Regards, Mark Reu Resouce Planner Internal Remarks : Tried calling bu no answer from the department. Internal Remarks : Still no answer. Cutomer comment : Good Morning Mark, My apologies for the delayed reply. I was on leave yesterday. Here is the location and contact. uCARDIAC THEATRES, LEVEL 1, JOHN RADCLIFFE HOSPITAL CONTACT: Charlotte Huns – tel. no. 00000000865-857-800000000 I copied Charlotte for her awareness. Kind Regards, Juian Dela Cru Grou Administrator Clinical Engineering Level 5, John Radcliffe Hospital Tel. No. 00000000865-221-500000000 Ouford uiversity Hospitals NHS Foudation Trut John Radcliffe Hospital Headley Way Headington Ouford Ou3 9Du----------------------------------------------------------------------------- | 2023-10-18 16:18 | No
2427 | Yes | M0A - u00- Configuration parameter lost | u- Configuration parameter lost | 000000000000000000000000001O##: 00000000000000000000000000100000000232000000005 Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Sending the procedue to the cutomer to reload the configuation. Cutomer called back 00000000imes withou sucess. Contact by e-mail. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: There is a message saying that the configuation parameter is lost. ----------------------------------------------------------------------------- Internal Remarks: u Tenah called again, becaue he has not had any feedback since his call 2 hous ago... The cutomer is 000000000000000000000000001ried, becaue it concerns a defibrillator. ----------------------------------------------------------------------------- | 2023-10-18 16:18 | No
2428 | Yes | Therapy test failure (DFM0 ASSY Therapy is replaced) | Therapy test failure (DFM0 ASSY Therapy is replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000239185 Madhan, kuar The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : To replace 0000000056000000009000000001, DFM100000000 Therapy PCA ASSY Diagnostic performed by Engineer : To replace 0000000056000000009000000001, DFM100000000 Therapy PCA ASSY Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000239661 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy ----------------------------------------------------------------------------- | 2023-10-18 16:23 | No
2429 | Yes | CV u uLLFLUOSTR-IMAGEDS | CV u uLLFLUOSTR-IMAGEDS | 000000000000000000000000001O##: 00000000000000000000000000100000000232189 November 17, 2000000003 Malhas, Ali The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Angio Device ue: Angio Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: no impact SE 8.2.3000000005 Resoluion: The reported faut is confirmed. Trouleshooting: POST resuts checked and log analyzed. No problem detected with the image PC or hard drives. Soluion: In the couse of maintenance, defective images were foud in the database and deleted from the system. The uLLmessage becaue of ful hard drive has now eupired. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- | 2023-10-18 16:53 | No
2430 | Yes | Equipment power failure | Equipment power failure |  | 2023-10-18 17:13 | No
2431 | Yes | notification00 PICIu telemetry central station - loudspeaker problem | notification000- PI0 telemetry central station - loudspeaker problem | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000066 October 19, 2000000003 JEAN-FRANCOIS GEORGES The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Eupected and ueupected Resuts: uer Impact: None Patient Impact: SE Resoluion: Replacement of the louspeaker + test ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000Gu DE GREVE The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Eupected and ueupected Resuts: uer Impact: None Patient Impact: SE Resoluion: uder contract, do 000000000000000000000000001onsite for JFG for tomorrow at 1:00000000. Follow u Requred: uder contract, issu 000000000000000000000000001onsite for JFG for tomorrow at 1:00000000. ----------------------------------------------------------------------------- | 2023-10-18 17:33 | No
2432 | Yes | Machine is not switching getting on | Machine is not switching getting on | 000000000000000000000000001O##: 000000000000000000000000001000000002000000005000000008-Oct-2000000003 Sah Bikash Kuar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching on How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Not provided Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: Not provided SE 8.1.3000000005 Parts Consued : 1.00000000000000000000000000000000800000000200000000612,Nehalem Host PC Monoplane Rev_III NO_HDD Resoluion : Host Pc of the system replaced as per replacement manul. After replacing the host Pc checked complete 000000000000000000000000001king of the system and foud 000000000000000000000000001king satisfactory. System handed over to cutomer in good 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000239222 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: No uLL Repair Action : FSE to visit and replace Host PC and reload backu/Software Diagnostic performed by Engineer : Checked and foud Host PC is not getting On need to order the same. Problem Description by Engineer : Problem description by engineer : Machine is not switching on How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Not provided Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: Not provided SE 8.1.3000000005 Follow u Requred : uLL# and/or description of uLL: No uLL Repair Action : FSE to visit and replace Host PC and reload backu/Software Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-10-18 17:43 | No
2433 | Yes | Free space low delete euams | Free space low delete euams | 000000000000000000000000001O##: 000000000000000000000000001000000002000000003 Chhlam, Sy The system meets the specification for the performed service and is retuned to ue Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu Free disk space is too low.: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Repair Action : Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. We went and did a database inconsistency check and after it passed we did a recreate image store. System is now operational. 0000000080000000050000000022 - ClarityIQ_II IP PC no HDD 0000000080000000050000000000000000- Hard disk spare part Resoluion : Instruted caller what to do in FSF, system is now fuly operational. Provided part nubers for replacement parts shoud issu happen again. Agreed to keep case open util 10000000020000000023 and close if no call backs. Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 17-Oct-2000000003 17:18:23 RMT - IPu Free disk space is too low. 17-Oct-2000000003 13:00000000300000000 When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: cannot make new images, reboot does not resolve and had to move patient to a different room uer Impact: cannot make images Patient Impact: moved to different room SE 8.2.3000000005 Diagnostic performed by Engineer : Zac reported that they got a message, "free space low delete euam." Per caller, reboot does not resolve and uder the review tab it is completely empty. ----------------------------------------------------------------------------- | 2023-10-18 18:53 | No
2434 | Yes | Touchscreen malfunction | Touchscreen malfunuon | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000051 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order the parts to solve the touhscreen and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and they are having touhscreen issus. He attempted to calibrate the touhscreen, bu was uable to perform it. Part #000000005615390000000061, touhscreen Part #00000000561511821 end cap, bezel Part #0000000056150000000091, bezel rear, gray ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Chase 317600000000900000000 IP & Issu: SN 100000000100000000900000000/ Needing Parts ID for touhscreen Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-18 18:58 | No
2435 | Yes | Generator does not start sometimes | Generator does not start sometimes | 000000000000000000000000001O##: 00000000000000000000000000100000000233331 November 2, 2000000003 Matuina, Ingo The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MT Device ue: Coro Eupected and ueupected Resuts: Generator starts from time to time not too high. uer Impact: uknown Patient Impact: uknown SE 000Resoluion: The reported faut is confirmed. The following actions were performed for uLLcorrection: - MPD control uit replaced The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT ----------------------------------------------------------------------------- Internal Remarks: Arranged matuina appointment directly with u Gromnitza, 11/2 after 9:00000000 am## Internal Remarks: uu 00000000230000000016000000007 - MPD CONTROL uIT Cutomer ETA: OCT-31-2000000003 ----------------------------------------------------------------------------- | 2023-10-18 19:18 | No
2436 | Yes | System not booting up, connecting to instillation source 0% | System not booting up, conneung to instillation source 0% | 000000000000000000000000001O##: 00000000000000000000000000100000000233800000000 Coffman, Josh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - SWF: FleuViewing PC for FleuSpot is ureachable. 17-Oct-2000000003 00000000:28:23 When was the first occurence? Within... : 17-Oct-2000000003 11:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Fleuspot not booting u uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Resoluion : Replace fleuspot OS drive Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - SWF: FleuViewing PC for FleuSpot is ureachable.: 1. The Sute PC (TH) reports that the Device 'FleuViewing PC for FleuSpot' (TA) is ureachable on the control net000000000000000000000000001k. It coud indicate the application SW is not runing. Repair Action : Service actions for RMT - SWF: FleuViewing PC for FleuSpot is ureachable.: 1.1. Check whether the FleuViewing PC for FleuSpot (TA) is powered-On. Reboot the uand verify whether problem is resolved. If FleuViewing PC for FleuSpot still not powers-u replace FleuViewing PC (TA). 1.2. Check cabling to the FleuViewing PC for FleuSpot (TA), especially the net000000000000000000000000001k cable: TA:u100000000- TS:u11 Diagnostic performed by Engineer : Patrick said the system is not booting u all the way. The review monitor has gears on it with a message to abou connecting to installation souce. He tried to reboot the system bu no change to the issu. Also, in the euam room everything seems to be normal and 000000000000000000000000001king as eupected. I checked the log file and foud fleuspot commuication uLLand per the trouleshooting manul there is an issu with the fleuspot OS drive. I provided the following p/n for the drive, 0000000080000000050000000000000000- Hard disk spare part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000017 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : spoke to Patrick abou Fleuspot pc problem When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: no fleuspot screen Patient Impact: none SE 1.2.5.00000000Resoluion : cutomer is going to review docuent and perform the replacement of the hard drive in fleu viewing pc Internal Comments : uLL# and/or description of uLL: fleuspot uLL Repair Action : Pattern Name: RMT - SWF: FleuViewing PC for FleuSpot is ureachable.uY Name: 9b9786e9-100000000-cd21-1d5e-00000000d77a2f9b6c Root Caue 1: The Sute PC (TH) reports that the Device 'FleuViewing PC for FleuSpot' (TA) is ureachable on the control net000000000000000000000000001k. It coud indicate the application SW is not runing. Service Action 1: Check whether the FleuViewing PC for FleuSpot (TA) is powered-On. Reboot the uand verify whether problem is resolved. If FleuViewing PC for FleuSpot still not powers-u replace FleuViewing PC (TA). Service Action 2: Check cabling to the FleuViewing PC for FleuSpot (TA), especially the net000000000000000000000000001k cable: TA:u100000000- TS:u11 Diagnostic performed by Engineer : speaking with Patricj he is looking the repair docuent for the fleuing pc .provided Du12700000000revA verbally and as an ##d link . provided additional checks for system ,check drive loghts on fleu spot and check cabling and net000000000000000000000000001k connections to and from fleuspot pc , especially the net000000000000000000000000001k cable: TA:u100000000- TS:u11 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000235700000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-10-18 19:28 | No
2437 | Yes | Low battery | Low battery | 000000000000000000000000001O##: 000000000000000000000000001000000002553800000000100000000Nov-2000000003 DAMIANI, MAuIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Low battery uer Impact: no Patient Impact: no SE na Resoluion: Replaced battery; The system was left charging. Parts Consued: 1.000000000000000000000000989800000000626900000000 Assy Battery Li-Ion, 11ah, V600000000Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000003-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000233598 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Low battery uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Colleagu Mauizio Damiani asks to order battery in replacement of the dead one. Follow u Requred: When the qute is accepted, order the spare part and schedue a technician. ----------------------------------------------------------------------------- Internal Remarks: Nuber of parts to order u1 989800000000626900000000 Send c/o MVS BERGAMO ----------------------------------------------------------------------------- | 2023-10-18 19:57 | No
2438 | Yes | UPS not working system boot issues | UPS not working system boot issues | 000000000000000000000000001O##: 0000000000000000000000000010000000023368000000009-Oct-2000000003 Obuhowski, Christian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: uS blew fue like other rooms Eupected and ueupected Resuts: room 000000000000000000000000001ld t boot u uer Impact: no Patient Impact: no uuParts Consued : 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP Resoluion : Replaced uS and blown fue in PDu System fuctional now ----------------------------------------------------------------------------- | 2023-10-18 20:13 | No
2439 | Yes | Prouimal sensor auto zero uLL0 | Prouimal sensor auto zero uLL00 | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000900000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: check vent prouimal pressue auozero failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and they are getting a check vent prouimal pressue auozero failed. He verified there is a code 1100000000 in the significant log. He informed that they jut replaced the flow sensor assy. Recommended to check the pin alignment between the solenoids and the DA pcba. Cutomer stated he reinstalled the DA pcba back on to the flow sensor assy, and it did not solve the issu. He requsted to sent it to the bench for repair. Shipping boues were backordered, and once available the bench will ship it to the cutomer. Follow u Requred : Cutomer has a V600000000and they are getting a check vent prouimal pressue auozero failed. He verified there is a code 1100000000 in the significant log. He informed that they jut replaced the flow sensor assy. Recommended to check the pin alignment between the solenoids and the DA pcba. Cutomer stated he reinstalled the DA pcba back on to the flow sensor assy, and it did not solve the issu. He requsted to sent it to the bench for repair. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Roger Streidl - phone 937 700000000 000000005 IP & Issu: 56632100000000- Prouimal sensor auo zero uLL1258 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-18 20:32 | No
2440 | Yes | stuck at zero at stat up | stuck at zero at stat up | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000019 18-Oct-2000000003 Childers, Brauton The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : spoke to joe .staff reported system did not start u this moring When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: system down Patient Impact: none SE Not provided Resoluion : Rebooted system completely. Completed 5 power cycles with system powering u correctly each time. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : spoke to joe .staff reported system did not start u this moring When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: system down Patient Impact: none SE Not provided Diagnostic performed by Engineer : speaking with Joe .the system did not start u this morning , stuk at zero . talked Joe thrua shu down andpower back u .power shows boot of host pc in control and startu screen .check of cabinets show main geo and cabinets have power and pc drive lights .euam and control screens show stuk at 0000000000000000. Internal Comments : uLL# and/or description of uLL: 0000000000000000 Repair Action : no prs or rsw connection to review PD.SCOMP.DCPS_20000000012V_5V Material Nuber 00000000800000000387582 0000000080000000058600000000 FleuVision-2_Revised PC 1 TZ ¦ For u R8.1.16.u and R8.1.17.u and u R8.2.16 - 8.2.17.u 1600000000Follow u Requred : uLL# and/or description of uLL: 0000000000000000 Repair Action : no prs or rsw connection to review PD.SCOMP.DCPS_20000000012V_5V Material Nuber 00000000800000000387582 0000000080000000058600000000 FleuVision-2_Revised PC 1 TZ ¦ For u R8.1.16.u and R8.1.17.u and u R8.2.16 - 8.2.17.u 1600000000 What are the neut Trouleshooting steps : check logging and startu of system check net000000000000000000000000001k connections and boot of fleu pc and low voltage suply Technician Name : Sulivan Scott ----------------------------------------------------------------------------- | 2023-10-18 20:38 | No
2441 | Yes | 0 / u00/ The treatment implementation detection fails (split, ref. PR 00. | 00/ u000 / The treatment implementation deteuon fails (split, ref. PR 000). |  | 2023-10-18 20:40 | No
2442 | Yes | 00N0eometry movements partly available | 00CDN0Geometry movements partly available | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000053 George Owens The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Debriefed caller. He said they have rebooted nuerou times with no improvement. Caller saved LOD and a review shows system reporting "Application uLL: uable to commuicate with GeoIPC. SID is uknown and no movements are available." as well as "POST AD5Gu EuT Failed" and "Application uLL POST failed: TLT, digital inpus." Follow u Requred : GeoIPC is not booting u at all or fast enouh. As resut the Host PC is uable to commuicate with the Geo IPC within a specified time. Verify the behavior of GEOIPC at first boot u of the system in the morning (at cold start-u): If inpu power (2300000000) is present bu the GEOIPC does not power u at all, it is caued du to a defective internal power suply of GEO IPC itself. If GeoIPC does power u fine bu Host PC is uable to commuicate with it in time, it is caued du to a defective harddisk. (corruted OS/AS) Then replace the GeoIPC. If needed, 0000000021291500000000 - IPC WITH NEW BIOS INTGR + ROHS, COA Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown ----------------------------------------------------------------------------- Internal Remarks : Initials: JJ Action Taken: I sumitted AA#000000000000000086000000008 to cancel the Qute & Onsite 000000000000000000000000001since issu have been resolved per the cu ISAAC ARMENDARIZ Oucome: Waiting on the Qute Team to cancel the requst Last udated: 12:13pm EDT 11/1/23 ----------------------------------------------------------------------------- | 2023-10-18 20:47 | No
2443 | Yes | VT is registering when not connected - Final 05 | VT is registering when not conneud - Final 0/5 |  | 2023-10-18 20:50 | No
2444 | Yes | Disk uLL Impossible to do fluoroscopy | Disk uLL Impossible to do fluoroscopy | 000000000000000000000000001O##: 000000000000000000000000001000000002000000002 October 19, 2000000003 MARTIN LITT The system meets the specifications for the performed service and is retuned to ue. Resoluion: impossible to recreate image store (problem to acqure images), replacement of the image PC's hdd Problem Description by Engineer: Cutomer Fuction/Role: Infi Device ue: yes Eupected and ueupected Resuts: Impossible to print disk uLL uer Impact: / Patient Impact: / SE / Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000031 ANTOINE SANTANGELO The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: Martin onsite 19/1000000008-10000000000000000 Part: 2u 0000000080000000050000000000000000 Delivery: onsite Problem Description by Engineer: Cutomer Fuction/Role: Infi Device ue: yes Eupected and ueupected Resuts: Impossible to print disk uLL uer Impact: / Patient Impact: / SE / Diagnostic performed by Engineer: replace the 2 disk images. ----------------------------------------------------------------------------- | 2023-10-18 20:52 | No
2445 | Yes | Stopped working on patient. Several alarms present. Unable to turn off. Performed on 000ent 0 GFE 04 | Stopped working on patient. Several alarms present. Unable to turn off. Performed on 0000 Sent 0d GFE 0/4 | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : yes Ou Of Bou Failue? : No Labor: Labor ID : L-19856385 000000000000000000000000001k Performed : DISPATCHED FOR ONSITE CM / CASE# 000000002200000000200000000 Diagonsis : calling OEM FOR REPAIR | 2023-10-18 21:08 | No
2446 | Yes | Patient Disconnect resulting in Low Tidal volume alarm | Patient Disconneuresulting in Low Tidal volume alarm |  | 2023-10-18 21:38 | Unknown
2447 | Yes | The central monitoring station alarm and the patient scope alarm did not sound. GFE0sent 0ec0 | The central monitoring station alarm and the patient scope alarm did not sound. GFE0 sent 0Dec03 |  | 2023-10-18 21:42 | Yes
2448 | Yes | No Product Malfunction-ur training - u uLL- insufficient information | No ProduuMalfunuon-ur training - u uLL- insuffi0nt information | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000097 Miller, Dell The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Transferred cutomer to internal department.Sent information to Bufor any inpu on this issu. 1000000002000000003 001200000000 hrs: Buresponse: Suject: Re: u Case #00000000220000000087/New York Presby Hospital Hi Dell Thanks for following u....I always appreciate it...and you patience with u! I think this is going to requre logs….even if this issu is old, someone will be able to look at the logs and see of there is a time discrepancy. There are a few possible euplanations bu the one that comes to mind is that there is a time discrepancy….for euample if youlook at you watch and I look at my watch at the same time, there may be a slight difference in the euact time that we see on ou respective watches……normally there is a time server that the PC's all sync with bu I am not as well eduated on how this is done throuhou the PIC system and I think the ful configuation is more complicated than that… However, if that is the case, I 000000000000000000000000001ld eupect other alarms to show that as well...hmmmm This 000000000000000000000000001ld be something the technical team will need to look into....... Thanks, Jennifer Forwarded EM to CS. Will await response from CS if the cuwants to escalate this issu. 10000000030000000023: CS response from 10000000025/23 and no futher commuication from CS: From: Mcdonald, Kelley <kelley.mcdonalducom> Sent: Wednesday, October 25, 2000000003 000000006:00000000 PM To: Miller, Dell <dell.millerucom>; Miller, Jennifer <jennifer.miller1ucom>; Rogozinski, Lauen <Lauen.Rogozinskiucom> Suject: RE: u Case #00000000220000000087/New York Presby Hospital Hi Everyone, NYP is going to pul the PIC logs from that day and send them to me. The other qustion that they have is, is there a way to be able to see the time on the bedside monitor on the auit log? Or 000000000000000000000000001ld they have to pul logs from the monitor for that? Best Regards, Kelley Kelley McDonald, RN, BSN, MPH Senior Clinical Implementation Consutant Services and Soluions Delivery Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Other: u CS Device ue: In Clinical ue Eupected and ueupected Resuts: Shoud not be any delay. uer Impact: No impact Patient Impact: Denies any pt harm, safety concerns, et al. SE B.00000000 Problem Description/ueupected Resuts: Cuwants docuentation, if any available, is delay from alarm souding at PICIu suveillance/OVV. Diagnostic performed by Engineer : Problem Description: Cuwants docuentation, if any available, is delay from alarm souding at PICIu suveillance/OVV. CS reports, cu##d her w/qustions abou needing docuentation on how long, if any, delay for an alarm to soud at PICIu suveillance to the OVV. Per clinical auit trail, this occured in CMuon hosts: PH1_TELSuV00000000 (primary suveillance) souded at 10000000000000000/23, 00000000:00000000:11/MIL7GN_SuV00000000 (OVV) 00000000:00000000:16 for Rm #7GN00000000B (see screen shot attached to case). Form ued: Trouleshooting/Investigation Object: OVV Not Object: Other PICIu's in the hospital Software Version: B.00000000 Date and Time of Issu: 10000000018/2000000003 uer Actions: Pt monitoring. Where is the Deviation: CMu Trouleshooting Steps Taken: Euplained, there is no docuentation in IFuabou this issu.Will send to Buprodut specialist for any inpu on this issu. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- | 2023-10-18 21:42 | No
2449 | Yes | QACASE 000000After uLLis connected to ouygen, air leakage occurs inside the body-GFE 0/6 | QACASE 00000000After uLLis conneud to ouygen, air leakage occurs inside the body-GFE 00/6 |  | 2023-10-18 21:43 | No
2450 | Yes | THE CENTRAL MONITORING STATION ALARM AND THE PATIENT SCOPE ALARM DID NOT SOUND Duplicate to 00 | THE CENTRAL MONITORING STATION ALARM AND THE PATIENT SCOPE ALARM DID NOT SOUND Duplicate to 005 |  | 2023-10-18 21:43 | Yes
2451 | Yes | disc space is low | disc space is low | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000071 300000000Oct-2000000003 Rodriguz Antonio The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : disc space is low When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Few available cases to delete in real time. uer Impact: Difficut to continu with clinical applications Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion : Replaced IPPC and imaging hard drives Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000000000000019-Oct-2000000003 Rodriguz Antonio The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : disc space is low When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Few available cases to delete in real time. uer Impact: Difficut to continu with clinical applications Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion : Recreated image store. Tested system and it is operational. Will follow u to see if hard drive or imaging compuer shoud be replaced as well. Follow u Requred : Replace imaging compuer or hard drive ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000057 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : disc space is low When was the first occurence? Within... : 13-Oct-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Few available cases to delete in real time. uer Impact: Difficut to continu with clinical applications Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Frankie says uer message image disc low. He has limited nuber of cases to delete. Advised that uer message has been present for several days, Euposue not possible. Image disk ful. Internal Comments : uLL# and/or description of uLL: Pattern Name: RMT - IPu Euposue not possible. Image disk ful Root Caue 1: Euposue not possible. Image disk ful Service Action 1: Instrut cutomer to free u disk space. Service Action 2: Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Repair Action : * Follow u Requred : uLL# and/or description of uLL: Pattern Name: RMT - IPu Euposue not possible. Image disk ful Root Caue 1: Euposue not possible. Image disk ful Service Action 1: Instrut cutomer to free u disk space. Service Action 2: Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Repair Action : * What are the neut Trouleshooting steps : Determine condition of IPPC/disc. Replace disc if needed. Re-Create image store. Pattern Name: RMT - IPu Euposue not possible. Image disk ful Root Caue 1: Euposue not possible. Image disk ful Service Action 1: Instrut cutomer to free u disk space. Service Action 2: Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : Image disk space low after all patient procedues deleted. FSE had to go into FSF uing ALT+F to recreate image store. uLL show uLLuknown PCDiagResut InvalidHardwareIndeu [IMAGE_PROCESSING_ON_TIME] plu nuberou internal software uLL. Approved for replacement IPPC compuer. 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- | 2023-10-18 22:08 | No
2452 | Yes | uuunot working, GFE00OV0 Response received/IS complete | uuunot working, GFE0 0NOV00Response received/IS complete | 000000000000000000000000001O##: 0000000000000000000000000010000000023000000005 Vongchavana, Eakchai The system meets the specification for the performed service and is retuned to ue Resoluion : Replace front bezel Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in ue Eupected and ueupected Resuts: uuunot 000000000000000000000000001king uer Impact: no Patient Impact: no SE na Parts Consued : 1.00000000000000000000000000000000561530000000051,RP-Bezel, Front, w/NAV, Gray, CTD, V600000000----------------------------------------------------------------------------- | 2023-10-18 22:13 | No
2453 | Yes | M0A - NA - Shock not delivered | M00A - NA - Shock not delivered |  | 2023-10-18 22:14 | Yes
2454 | Yes | M0A - NA - Defibrillator showed ""Insert battery"" and the device malfunctioned. | M00A - NA - Defibrillator showed ""Insert battery"" and the device malfunuoned. |  | 2023-10-18 22:22 | No
2455 | Yes | QA00000_the machine did not alarm when the blood READY FOR SUMMARY==LETTER | QA00000000_the machine did not alarm when the blood READY FOR SUMMARY==LETTER |  | 2023-10-18 22:36 | No
2456 | Yes | Stuck at 00 | Stuck at 00 | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000079 25-Oct-2000000003 Kolbinger Patrick The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD Resoluion : Replaced Fleuvision PC Problem Description by Engineer : Problem description by engineer : system Stuk at 00000000:00000000. Inhoue looked at system and 000000000000000000000000001ld like it sent to the field When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u fuly uer Impact: can not ue system. will not boot Patient Impact: no patient impact SE no prs ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000087 Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: system stuk at 00000000:00000000 Repair Action : Cutomer did some troule shooting and 000000000000000000000000001ld like the FSE to be paged ou 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Diagnostic performed by Engineer : Spoke with Biomed. He asked that it be sent ou to the FSE Problem Description by Engineer : Problem description by engineer : system Stuk at 00000000:00000000. Inhoue looked at system and 000000000000000000000000001ld like it sent to the field When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u fuly uer Impact: can not ue system. will not boot Patient Impact: no patient impact SE no prs Follow u Requred : uLL# and/or description of uLL: system stuk at 00000000:00000000 Repair Action : Cutomer did some troule shooting and 000000000000000000000000001ld like the FSE to be paged ou 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Eckbold Glenn ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000600000000Armenia, Mauice The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : Advised to check DC power suply at M cabinet and verify TS switch is fuctional. If ok, ue FSF (pre boot u) and check if Fleuvision uLL are present to determine if PC is fauty. Diagnostic performed by Engineer : Cutomer is trained biomed with assist contract. He has tried shuting system off completely and at the breaker bu system still does not power u. Fleuvision PC shows it has power visully. Determined this is an 8.2 system. Problem Description by Engineer : Problem description by engineer : Cutomer is reporting the system is stuk at 0000000000000000 coutdown and not booting u all the way When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: not booting u uer Impact: system not booting u Patient Impact: none mentioned SE Not provided Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Advised to check DC power suply at M cabinet and verify TS switch is fuctional. If ok, ue FSF (pre boot u) and check if Fleuvision uLL are present to determine if PC is fauty. ----------------------------------------------------------------------------- | 2023-10-18 22:53 | No
2457 | Yes | fiber optic cable is intermittent | fiber optic cable is intermittent | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000069 Chhlam, Sy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : losing connection with flat detector When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Giorgio reported that he need a part nuber for the fiber cable as it loses signal intermittently. Internal Comments : uLL# and/or description of uLL: fiber optic cable is intermittent Repair Action : Provided part nuber 00000000230000000023200000000- FIBER OPTIC CABLE LAT. + FRONT. Resoluion : Provided part nuber for replacement. ----------------------------------------------------------------------------- | 2023-10-18 22:53 | No
2458 | Yes | M0A - QACASE for PCI imported product#00000 the defibrillation monitor did not pass the self-test | M00A - QACASE for P0imported produu00000000 the defibrillation monitor did not pass the self-test |  | 2023-10-18 22:53 | No
2459 | Yes | 00 CN00- 0 Processor supply out of range - Processor PCA replaced | u- 00 Processor supply out of range - Processor PCA replaced |  | 2023-10-18 22:58 | No
2460 | Yes | Alarms-Failure to Alarm / Analysis after Patient Death / No Trouble Found-Working to Spec / Provided Information | Alarms-Failure to Alarm / Analysis after Patient Death / No Trouble Found-Working to Spec / Provided Information |  | 2023-10-18 23:01 | Yes
2461 | Yes | Security issue_The cuomer officially complains that the spare parts have short life - GFE 0/0 | Security issue_The cuomer offi0lly complains that the spare parts have short life - GFE 00/0 |  | 2023-10-18 23:25 | No
2462 | Yes | Cuomer wants all logs pulled from the monitor patient Incident 0ug0. No Product Malfunction-ur. | Cuomer wants all logs pulled from the monitor patient In0ent 0Aug00 No ProduuMalfunuon-ur. | 000000000000000000000000001O##: 0000000000000000000000000010000000027000000003 25-Oct-2000000003 Johnny Surez The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Patient Monitoring Eupected and ueupected Resuts: None uer Impact: uknown Patient Impact: SE Resoluion : Biomed learned to pul logs ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Sanchez, Karbin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE TO PERFORM PMIR Biomed called to requst FSE onsite to pul the logs of the bedside monitor. There was a patient incident on 2/8/2000000002. The monitor has not been ued since the incident and it is located in the Qulity Department.located in the third floor of the hospital. Device ue: MP5 Eupected and ueupected Resuts: FSE TO PERFORM PMIR Biomed called to requst FSE onsite to pul the logs of the bedside monitor. There was a patient incident on 2/8/2000000002. The monitor has not been ued since the incident and it is located in the Qulity Department located on the third floor of the hospital. uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer : FSE TO PERFORM PMIR Biomed called to requst FSE onsite to pul the logs of the bedside monitor. There was a patient incident on 2/8/2000000002. The monitor has not been ued since the incident and it is located in the Qulity Department.located in the third floor of the hospital. Follow u Requred : FSE TO PERFORM PMIR Biomed called to requst FSE onsite to pul the logs of the bedside monitor. There was a patient incident on 2/8/2000000002. The monitor has not been ued since the incident and it is located in the Qulity Department located on the third floor of the hospital. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: u employee Aleu Perez - 80000000026000000000000000057 IP & Issu: 39323900000000 - Cutomer wants all logs puled from the monitor Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O# Adv cut that bc a patient was harmed I had to get him over to team so they can get additional info. Warm ufer cut to SIT rtac agent Karbin Sanchez Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-18 23:42 | Unknown
2463 |  | Security Issue_Cuomer Dissatisfcation_The cuomer officially complains that the spare parts have short life | Security Issue_Cuomer Dissatisfcation_The cuomer offi0lly complains that the spare parts have short life |  | 2023-10-18 23:43 | No
2464 | Yes | M0A - Shock is not delivered - | M00A - Shock is not delivered - |  | 2023-10-18 23:44 | Yes
2465 | Yes | Wired footswithch intermittent - and ur message button pued at start up | Wired footswithch intermittent - and ur message button pued at start up | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000075 19-Oct-2000000003 Palmer Deante The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Wired footswithch intermittent - and uer message buton puhed at start u When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: Reboot of system does not resolve issu uer Impact: uable to continuwith clinical applications Patient Impact: delayed, requred to move to alternate lab SE 000Parts Consued : 1.000000000000000000000000989600000000180000000015,HANDSWITCH 1.00000000000000000000000000000000800000000151715,GEO MODuE WP Kit Resoluion : Hand switch was stuk in the prep position and needed to be replaced, also the geometry modue was 000000000000000000000000001n and needed to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000062 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Wired footswithch intermittent - and uer message buton puhed at start u When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: Reboot of system does not resolve issu uer Impact: uable to continuwith clinical applications Patient Impact: delayed, requred to move to alternate lab SE 000Diagnostic performed by Engineer : Sean said that they were uable to fluro then rebooted the system. uon reboot he said they were getting the uer message, buton puhed on startu. Asked if they are able to cine at handswitch bu he was ucertain. Asked if they have the wireless footswitch bu it is on back order. Internal Comments : uLL# and/or description of uLL: No lod available. Repair Action : field Follow u Requred : uLL# and/or description of uLL: No lod available. Repair Action : field What are the neut Trouleshooting steps : View logs to determine if footswitch has failed and replace Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : Deante - buton active at startu duing boot. All movements 000000000000000000000000001k on GEO. GEO modue has a rattling soud, bu all movements do 000000000000000000000000001k. Wired FS did not 000000000000000000000000001k. FSE checked HS and foud to be depressed. Lifted eupose buton u and rebooted, system boots normally. Approved HS and GEO modue du to rattling ----------------------------------------------------------------------------- | 2023-10-19 00:08 | No
2466 | Yes | Ouygen Manifold is broken; air is going everywhere. | Ouygen Manifold is broken; air is going everywhere. | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000079 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:part id uer Impact: euternal O2 manifold defective Patient Impact: none SEun Resoluion : Provided part nuber and cutomer will order, issu resolved- Material Nuber 98980000000060000000000000000- KIT,O2 TRANSPORT,CART MTG,DISS-M Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states when O2 applied at back of uit the O2 does not flow throuh the euternal manifold and appears to be defective, no patient involvement, requsts part id to replace. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Brian Angelico ph: 77969632500000000 IP & Issu: Cutomer Requsting Suport - Parts ID. Ouygen Manifold is broken ; air is going everywhere; o2 manifold part needed; 73192135 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? NO No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created and roued free r000000000000000000000000001to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-19 00:33 | No
2467 | Yes | THE SCREEN SIGNAL IS FAULTY. THE LIVE IS DISPLAYED INTERMITTENTLY. | THE SCREEN SIGNAL IS FAULTY. THE LIVE IS DISPLAYED INTERMITTENTLY. | 000000000000000000000000001O##: 00000000000000000000000000100000000267351 200000000ov-2000000003 Vazquz, Clauio The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: TM Device ue: Yes Eupected and ueupected Resuts: Live signal interruted in euam room uer Impact: No Patient Impact: No SE 8.2 Diagnostic performed by Engineer: Replacement of spare parts Operating equpment Parts Consued: 1.00000000000000000000000000000000800000000000000003100000000Video Splitter DVI 1.00000000000000000000000000000000200000000000000002781,Single Link DVI eut. Transm. str_PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000575 200000000ov-2000000003 Vazquz, Clauio The system meets the specification for the performed service and is retuned to ue Resoluion: Live signal video chain replacement. Change of fleuvision monitor controller compuer Change of video converters, video splitter and u00000000commuication cable in fleuvision. The equpment will be left uder observation as it is an intermittent failue. Fuctional testing. Operative team. Problem Description by Engineer: Cutomer Fuction/Role: TM Device ue: YES Eupected and ueupected Resuts: FLASHING LIVE SIGNAL ON EuAM ROOM SCREEN uer Impact: NO Patient Impact: YES SE 8.2 Diagnostic performed by Engineer: Video signal from the live signal is intermittently lost on the fleuvision screen ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000288161 300000000Oct-2000000003 Francisca Gamboa The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Send to Hospital Dr. Carlos Van Buen. Hemodynamics uit. Contact: Abdiel Milla +56982168293. Problem Description by Engineer: Cutomer Fuction/Role: TM Device ue: YES Eupected and ueupected Resuts: FLASHING LIVE SIGNAL ON EuAM ROOM SCREEN uer Impact: NO Patient Impact: YES SE 8.2 Diagnostic performed by Engineer: Diagnosis continus, carrying ou tests with new splitter and converters withou being able to solve the problem. Additional tests are carried ou, which point to a failue in the video card of the compuer responsible for sending these signals to the euam room monitor. New spare parts will be requsted to continu with the diagnosis. Restricted team. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000083 23-Oct-2000000003 Francisca Gamboa The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: TM Device ue: Yes Eupected and ueupected Resuts: Live signal interruted in euam room uer Impact: No Patient Impact: No SE 8.2 Diagnostic performed by Engineer: Complete review of the "live" signal chain is performed, with the last signal affected by the failue. The correct connection of the different elements involved is corroborated. Converters and signal splitters are requsted to continu and corroborate diagnosis. Equpment remains in restricted state du to signal intermittency. Follow u Requred: Send to Dr Carlos Van Buen Hospital, Hemodynamics uit. Contact: Abdiel Milla +56982168293 ----------------------------------------------------------------------------- | 2023-10-19 00:58 | No
2468 | Yes | Alarm/Failure to alarm:Patient Heart Rate euceeded High Limit but no alarm: Unknow Cau: Provided Information | Alarm/Failure to alarm:Patient Heart Rate euceeded High Limit but no alarm: Unknow Cau: Provided Information |  | 2023-10-19 01:09 | No
2469 | Yes | LS - 00/ AT - 02 | LS - 000 / AT - 002 | 000000000000000000000000001O##: 00000000000000000000000000100000000235700000000200000000Oct-2000000003 Mandeville Nathan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside ue Eupected and ueupected Resuts: system is not tuning on all the way. Displays are black and cannot perform u-ray. uer Impact: cannot ue system Patient Impact: none SE000Parts Consued : 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD Resoluion : Cannot tun on the system. ----------------------------------------------------------------------------- Internal Remarks : Nathan Nanderville - FSE can't configue TSM after replacement of Fleu PC. Shows green in FSF download board SW, told FSE to reload one more time and check resuts. ----------------------------------------------------------------------------- | 2023-10-19 01:13 | No
2470 | Yes | Receiving uLLcode "low leak-co0rebreathing risk" - Sent 0 FE 06 | Receiving uLLcode "low leak-co0 rebreathing risk" - Sent 0d FE 0/6 | Internal Remarks : Interaction Type: Admin Caller: IP & Issu: IPCR IP: 7000000001738 issu: cu wants to do an udate... receiving uLLcode "low leak-co2, rebreathing risk" related to closed case 00000000223999500000000600000000uLLserial # 200000000000000003539 equpment owned and located at 700000000 Merton Minter St SAN ANTONIO, Tu 78229 (acct# 00000000900000000700000000) Event Type: Incident Service Requst: IPCR 3 Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Preferred Service Time: Entitlement Options: Reason for Entitlement Claim: Action Taken: Cutomer didn't provide suficient evidence for IB team to approve the requst so this case can be completed/closed du to rejection of OE sumitted on 1000000002000000003 Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Admin Caller: IP & Issu: IPCR IP: 7000000001738 issu: cu wants to do an udate... receiving uLLcode "low leak-co2, rebreathing risk" related to closed case 00000000223999500000000600000000uLLserial # 200000000000000003539 equpment owned and located at 700000000 Merton Minter St SAN ANTONIO, Tu 78229 (acct# 00000000900000000700000000) Event Type: Incident Service Requst: IPCR 3 Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Preferred Service Time: Entitlement Options: Reason for Entitlement Claim: Action Taken: Sent 3rd attempt for OE 10000000019/23 Please allow 200000000ous for cu to respond Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Admin Caller: IP & Issu: IPCR IP: 7000000001738 issu: cu wants to do an udate... receiving uLLcode "low leak-co2, rebreathing risk" related to closed case 00000000223999500000000600000000uLLserial # 200000000000000003539 equpment owned and located at 700000000 Merton Minter St SAN ANTONIO, Tu 78229 (acct# 00000000900000000700000000) Event Type: Incident Service Requst: IPCR 3 Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Preferred Service Time: Entitlement Options: Reason for Entitlement Claim: Action Taken: Sent cu OE in to IB team (10000000020000000023) for approval. The OE was provided to u on 10000000019 Neut Step Action: Awaiting Entitlement Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Admin Caller: IP & Issu: IPCR IP: 7000000001738 issu: cu wants to do an udate... receiving uLLcode "low leak-co2, rebreathing risk" related to closed case 00000000223999500000000600000000uLLserial # 200000000000000003539 equpment owned and located at 700000000 Merton Minter St SAN ANTONIO, Tu 78229 (acct# 00000000900000000700000000) Event Type: Incident Service Requst: IPCR 3 Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Preferred Service Time: Entitlement Options: Reason for Entitlement Claim: Action Taken: Sent another attempt for OE 10000000023/23 du to previou rej for OE on 3rd attempt Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Admin Caller: Mario Flores IP & Issu: "IPCR/ new objective evidence receive 100000000200000000-- sumitted IP Change Requst 20000000031000000001277837" Event Type: Incident Service Requst: IPCR 2 Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: sumitted IPCR new evidence to IB for review Internal Remarks : Interaction Type: New Case Caller: Mario 2100000000837-0000000000000000 IP & Issu: IPCR IP: 7000000001738 issu: cu wants to do an udate... receiving uLLcode "low leak-co2, rebreathing risk" related to closed case 00000000223999500000000600000000uLLserial # 200000000000000003539 equpment owned and located at 700000000 Merton Minter St SAN ANTONIO, Tu 78229 (acct# 00000000900000000700000000) Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: sent uknown ip ## Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Admin Caller: IP & Issu: IPCR IP: 7000000001738 issu: cu wants to do an udate... receiving uLLcode "low leak-co2, rebreathing risk" related to closed case 00000000223999500000000600000000uLLserial # 200000000000000003539 equpment owned and located at 700000000 Merton Minter St SAN ANTONIO, Tu 78229 (acct Event Type: Incident Service Requst: IPCR 2 Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Preferred Service Time: Entitlement Options: Reason for Entitlement Claim: Action Taken: Sent 2nd attempt ##for OE on 10000000018/23 Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-19 01:18 | No
2471 | Yes | ETO u0 probe not recognized/ France due 00 CFM-GFE0FollowUp WED 0EC0 - Requesting uReturn info | ETO u00 probe not recognized/ France due 0-0/ CFM-GFE0 FollowUp WED 0DEC00- Requesting uReturn info |  | 2023-10-19 01:23 | No
2472 | Yes | when they press button to move crm side to side it doesnt lock so when they release the button the crm... | when they press button to move crm side to side it doesnt lock so when they release the button the crm... | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000031 200000000ct-2000000003 Sean Kiernan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - u: POST failue of the local uer Interface on the Stand (release brakes butons in yellow handgrip). 13-Oct-2000000003 10000000025:19 When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Carm not to move/carm is still moving when the buton is released uer Impact: Carm does not lock into place Patient Impact: No impact to patient or procedue SE 8.1.17.00000000Resoluion : Arrived onsite and foud that the longituinal rail the motor drives on was dirty. Cleaned rail as well as wheel that drives C-arm longituinally and C-arm was then moving properly. System retuned to ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000098 Vilkaukas Viktorija The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - u: POST failue of the local uer Interface on the Stand (release brakes butons in yellow handgrip). 13-Oct-2000000003 10000000025:19 When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Carm not to move/carm is still moving when the buton is released uer Impact: Carm does not lock into place Patient Impact: No impact to patient or procedue SE 8.1.17.00000000Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - u: POST failue of the local uer Interface on the Stand (release brakes butons in yellow handgrip).: 1. In most cases this POST failue is caued du to one of the uer Interface (brake release) butons on the yellow hand grip of the Stand is stuk or pressed duing startu. Repair Action : Service actions for RMT - u: POST failue of the local uer Interface on the Stand (release brakes butons in yellow handgrip).: 1.1. Replace buton assy if problem persists. (SAM HANDGREEP L-ARM) 1.2. Check butons in yellow handgrip and release it again. 1.3. Sometimes the butons gets physically stuk. Clean and re-align the buton with respect to its houing (cover) so that the u butons will not be stuk anymore. 1.00000000Check the internal cabling of the uer Interface butons inside the Yellow handgrip. The cabling can be damaged and making a short against groud. Technician Name : Vilkaukas Viktorija Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - u: POST failue of the local uer Interface on the Stand (release brakes butons in yellow handgrip).: 1. In most cases this POST failue is caued du to one of the uer Interface (brake release) butons on the yellow hand grip of the Stand is stuk or pressed duing startu. Repair Action : Service actions for RMT - u: POST failue of the local uer Interface on the Stand (release brakes butons in yellow handgrip).: 1.1. Replace buton assy if problem persists. (SAM HANDGREEP L-ARM) 1.2. Check butons in yellow handgrip and release it again. 1.3. Sometimes the butons gets physically stuk. Clean and re-align the buton with respect to its houing (cover) so that the u butons will not be stuk anymore. 1.00000000Check the internal cabling of the uer Interface butons inside the Yellow handgrip. The cabling can be damaged and making a short against groud. Diagnostic performed by Engineer : ui is reporting that when they press buton to move crm side to side it doesnt lock so when they release the buton the crm still moves. he was not at the system and was not able to create a logfile. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Yuiy 71883000000001000000002 (POC David) IP & Issu: 5000000008800000000 when they press buton to move crm side to side it doesnt lock so when they release the buton the crm still moves Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created remote 0000000000000000000000000010000000020000000098 dispatched to HC-NA-Tier2-CV-Cut,Cut-Tech-CV Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Yuiy 71883000000001000000002 (POC David) IP & Issu: 5000000008800000000 when they press buton to move crm side to side it doesnt lock so when they release the buton the crm still moves Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: created remote 0000000000000000000000000010000000020000000098 dispatched to HC-NA-Tier2-CV-Cut,Cut-Tech-CV Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-19 02:18 | No
2473 | Yes | replacing Fleuvision PC and system stuck at 00, how to get into FSF | repla0g Fleusion PC and system stuck at 00, how to get into FSF |  | 2023-10-19 02:33 | No
2474 | Yes | HS0Cuomer called stating device will not turn on. (Battery Calls) | HS0 Cuomer called stating device will not turn on. (Battery Calls) | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000029 Simpson,Kreshia The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED green light wasn’t flashing and the AED was powering u uer Impact: NONE Patient Impact: NONE SEE.00000000.000000003 Cutomer reported the AED green light wasn’t flashing and the AED was powering u Diagnostic performed by Engineer : Cutomer reported the AED green light wasn’t flashing and the AED was powering u, cutomer stated that the AED was chirping since Auut. Cutomer stated that the school was closed so no one was able to inspect the AED. Instruted cutomer to remove and reinsert the battery and the AED didn’t power back u. Will replace battery that is uder warranty AED Serial# A18B-0000000000000000 Battery Install Before# 9/3000000002000000007 Lot#00000000768P Pads Eup Date: 00000000/20000000000000000 Pads Lot # Y100000000721-00000000Internal Comments : SCM Reguar Order 6300000000925800000000 Internal Comments : Tracking# 711191976000000001 SCHEDuED DELIVERY DATE Friday 10000000027/23 by end of day Internal Comments : DELIVERED Monday 10000000030000000023 at 11:36 AM Signed for by: JMCBuNEY. ##cutomer to confirm AED is 000000000000000000000000001king as it shoud Resoluion : Closing case cutomer has confirmed that AED is 000000000000000000000000001king as it shoud. ----------------------------------------------------------------------------- | 2023-10-19 03:08 | No
2475 | Yes | System is down(powers up by itself, without hitting the power button - Sent 0 GFE 06 | System is down(powers up by itself, without hitting the power button - Sent 0d GFE 0/6 | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000900000000ccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to disconnect battery and leave uit off and pluged into AC power overnight, if problem not dulicated replace battery to correct issu, if problem reoccus replace u pcb to correct issu, cutomer acknowledged and will continu trouleshooting, provided part nuber- Material Nuber 00000000561528871 - RP-PCBA,uer Interface,u 2nd GEN,V600000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: powers on by itself uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states the uit powers on by itself while not in ue and has troule powering off, PM pcb was recently replaced per FCO86600000000000000006 and cutomer replaced power switch overlay bu problem persists, verified uit powers on by itself with jut battery connected and withou AC, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Michael Romero IP & Issu: "5700000000833 System is down(powers u by itself, withou hitting the power buton)." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-19 03:17 | No
2476 | Yes | This device is DOWN becau it's FAILING to DEFIBRILLATE | This device is DOWN becau it's FAILING to DEFIBRILLATE | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000062 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Tempu LS failed to defibrillate. Diagnostic performed by Engineer : Problem Description: Tempu LS failed to defibrillate. Tempu LS failed to defibrillate. Form ued: Trouleshooting/Investigation Object: Tempu LS Not Object: Software Version: NA Date and Time of Issu: 10000000018/2000000003 uer Actions: Where is the Deviation: Tempu LS Trouleshooting Steps Taken: NA Was PRS ued: No PRS Permission: Resoluion : Is Caue Known: No Caue: Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.Ordered replacement Tempu LS 000000005600000000181. Follow u Requred : Tu Victoria Fire Department 1700000000 E Airline Rd. Victoria, Tu 77900000000-00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : ##d cutomer - requst for patient information and log/rescu files. Internal Remarks : Interaction Type: New Case Caller: Mark MARTYN IP & Issu: " 93683817 This device is DOWN becaue it's FAILING to DEFIBRILLATE" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: PCCI Action Taken: I created a case, CSP, R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-19 04:02 | No
2477 | Yes | [Q] - one of the cover of hose has fell off. | [Q] - one of the cover of hose has fell off. | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000033 31-Oct-2000000003 Pitts, Jared The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : one of the cover of hose has fell off. When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Resoluion : Replaced hose hanger bracket Parts Consued : 1.00000000000000000000000000000000800000000123121,Hose carrier ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000017 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : one of the cover of hose has fell off. When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Dennis says local fse is aware of problem and has sent pictue of item reported. Internal Comments : uLL# and/or description of uLL: mechanical issu Repair Action : field Follow u Requred : uLL# and/or description of uLL: mechanical issu Repair Action : field What are the neut Trouleshooting steps : order and replace item as described by biomed Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Dennis Ruz IP & Issu: " 7813200000000one of the cover of hose has fell off." Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-19 04:43 | No
2478 | Yes | WFS footswitch is not working. | WFS footswitch is not working. | 000000000000000000000000001O##: 00000000000000000000000000100000000237878 19-Oct-2000000003 Michael SMITH The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: rad Device ue: wireless footswitch Eupected and ueupected Resuts: not 000000000000000000000000001king uer Impact: none Patient Impact: none SE R8.2.16 Resoluion : Tested operation of the of the wireless foot switch. Checked base station. Coud not repeat reported faut ----------------------------------------------------------------------------- | 2023-10-19 06:08 | No
2479 | Yes | Speaker malfunction-Sound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Speaker malfunuon-Sound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 000000000000000000000000001000000002000000006000000003-Oct-2000000003 Bugannavar, Praveen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Engineer Device ue: Ouside of ue Eupected and ueupected Resuts: Speaker malfuction occuring. uer Impact: Low Patient Impact: Low SE Not requred Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-19 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000005-22 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000056000000008000000001,SS CBL Speaker Assy FS Diagnostic performed by Engineer : CONFIRMED FAuT. REPLACED THE SPEAKER ASSY AND FuL PM PASSED. RETuNED TO SERVICE. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000037 Edward Carson The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Engineer Device ue: Ouside of ue Eupected and ueupected Resuts: Speaker malfuction occuring. uer Impact: Low Patient Impact: Low SE Not requred Diagnostic performed by Engineer : Speaker malfuction occuring when tuning on machine. Occus everyttime Resoluion : Speaker assembly requres replacing Follow u Requred : FSE to site to replace speaker Assembly ----------------------------------------------------------------------------- | 2023-10-19 06:13 | No
2480 | Yes | 0/u FD0Geometry movements partly available | 0R/u FD0Geometry movements partly available | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 21.1000000002000000003 13:10000000036 uC (8600000000000000007) Service content: The equpment often fails to move the rack and bed, and reports the uLLGeometry movement not available. On-site inspection logs foud that the system and IPC lost commuication. Sometimes it was jut started, sometimes duing ue, GEO IPC commuication time ou. I supect there is a problem with the IPC motherboard and order an IPC replacement. Service resuts: Replace the IPC, download the software, and retun to normal after debuging. The equpment is fuly fuctioning when leaving the site. Other important matters: None yet. SAP Transfer: Case Transferred to SAP Neut Step Action: The frame is locked when the machine is tuned on and the bed cannot be moved. SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: 00000000R/u uuuFD200000000Internal Remarks: 21.1000000002000000003 13:10000000039 uC (8600000000000000007) RDI Accuacy: Yes $Q7: (Confirmation & Description of problem) Confirmed reported problem, system cannot ready for move always, uLLGeometry movement not available $Q8: (Trouleshooting & Resuts) On site check the event log find GEO IPC commuication time ou IPC problem $Q9: (Repair steps) Disassemble the houing and replace IPC, download IPC's OS and APP, observe and ue a period of time $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:N/A; Safety Qustions: Case: 000000002200000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: After the equpment is tuned on, all mechanical movements are disabled. The euamination bed and C-arm are not powered on and locked, and can be puhed by hand.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: uable to commuicate with GeoIPC. No movements are available (reboot done/needed).. uable to commuicate with GeoIPC Remote diagnostics and test resuts: The mechanical motion control host of the equpment is not connected to the system. PRS TR statu: Doesn't make sense Repair Action: Recommended on-site repair steps: Check if the BIOS battery is normal 3V, Root Caue 1: Please ignore this alert, if it only occus sporadically. Treat it then as a 'one time only'. Service Action 1: Perform a reboot of the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other Soluion here below. Root Caue 1: GeoIPC is not booting u at all or fast enouh. As resut the Host PC is uable to commuicate with the Geo IPC within a specified time. Service Action 1: Verify the behavior of GEOIPC at first boot u of the system in the morning (at cold start-u): If inpu power (2300000000) is present bu the GEOIPC does not power u at all, it is caued du to a defective internal power suply of GEO IPC itself. Service Action 2: If GeoIPC does power u fine bu Host PC is uable to commuicate with it in time, it is caued du to a defective harddisk. (corruted OS/AS) Then replace the GeoIPC. Service Action 3: Check whether the cable pins between GeoIPC RC-u15 and IPC adapter SCC1-u22 are perhaps too short. Advise: tighten the RC-u15 and SCC1-u12 connectors or replace the cable between RC-u15 and SCC1-u12. Service Action 00000000 It coud be caued du to a defective control net000000000000000000000000001k switch (TS) in the R-cabinet or its Power Suply. Root Caue 1: When problem is persistent it coud be caued du to a defective control net000000000000000000000000001k switch (TS). Service Action 1: Check control net000000000000000000000000001k switch (TS) and its power suply (adapter). Replace when necessary. Root Caue 1: Bad cabling and/or connections. Service Action 1: Check the cable RC:u25 to SCC1:u22. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. Service Action 2: Tighten all the connectors of GeoIPC and SCC1 uit. Root Caue 1: Issu with GeoIPC OS/AS software. It coud be du to a not correct shudown before, Windows OS can be corruted hence it needs more time in the neut start u. If this happens only sporadically, please ignore it. Else see other possible caues & soluions. Service Action 1: If SW is really corruted please reload GeoIPC software uing uB destroy stick. Root Caue 1: The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. Service Action 1: Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start u of the system. If so, replace GEO IPC. Notes on cutomer requrements: Onsite ----------------------------------------------------------------------------- | 2023-10-19 06:13 | No
2481 | Yes | 0/u FD0startup failture | 0R/u FD0startup failture | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 18.1000000002000000003 00000000:15:200000000uC (86000000006568)"SuT. Ready for Clinical ue local time" : " 2000000003-10000000018 11:00000000 SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/u uuuFD200000000Black screen at startu SAP Transfer: Dispatched for Onsite Suport Neut Step Action: Send 000000000000000000000000001kers to zhang cheng T2 Activities:Problem Description: Overview of cutomer problems: /startu failue Detailed description of cutomer problem or faut phenomenon: Host PC is not started Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: When I tun on the compuer, only the review display shows u Shield. Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: 1. Check whether the HOST PC indicator light is on. If it is not on, check the 2200000000 power suply behind the host PC. If not, check the fue or replace the power backplane 0000000080000000000000000321 2. If the host PC has power, it is recommended to check the power indicator light of the hard disk bou. 3. It is recommended to manully switch the host pc on and off. If it starts normally, it is recommended to replace the bois battery. Recommended spare parts uuand designation: uu00000000800000000200000000612 Nehalem Host PC Monoplane Rev_III NO_HDD or 00000000800000000122391 u Haswell Monoplane Host PC no HDD Notes on cutomer requrements: Please contact the cutomer for on-site repair Safety Qustions: Case: 000000002200000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 18.1000000002000000003 00000000:17:37 uC (86000000006568) Faut description: uable to boot Service content: Measue the low battery voltage of the HOST PC motherboard, and boot it normally after replacing the battery Service resuts: After repair, the equpment retuned to normal operation. Internal Remarks: 18.1000000002000000003 00000000:37:00000000 uC (86000000006568) RDI Accuacy:YES $Q7: (Confirm reported problem & cannot boot u) $Q8: (Measue the motherboard battery voltage too low) $Q9: (Replace the motherboard battery, the system retuned to normal operation) $Q100000000 (Test and inspection data refer to PA, Device retun to ful fuctionality) $Q11:Calibrated Tools name (Mutimeter Flue 287) S/N (1000000000000000017) neut Calibration date (00000000/11/20000000000000000 For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue ue:N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-19 06:13 | No
2482 | Yes | 0_0MICU_DFM0_ufailure CN06 | 0N_00MICU_DFM0_ufailure CN006 | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes : This PO is a Reference Nuber only Safety Qustions : Case: 000000002200000000315 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sabrina Lin Last udated by: Sabrina Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer : 00000000N_162_MICuDFM100000000_ufailue CN32676226 SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-19 08:43 | No
2483 | Yes | 0uDFM0 Defibrillator/Monitor/The operation check fails | 0/Effi0 DFM0 Defibrillator/Monitor/The operation check fails | BuQ&R Feedback: No Reason: Cutomer Relations Safety Qustions: Case: 000000002200000000327 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Majiang Couty People's Hospital of GuzhouProvince Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-19 08:48 | No
2484 | Yes | Audio & Sounds/Loudspeaker uLLSpeaker | Audio & Sounds/Loudspeaker uLLSpeaker | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000039 Problem Description by Engineer: HP faut Diagnostic performed by Engineer: Hp fauty Resoluion: Louspeaker Changed Consued parts: 1.000000000000000000000000000000005600000000000000001,IV2-STAT MECHASY Louspeaker ----------------------------------------------------------------------------- | 2023-10-19 08:58 | No
2485 | Yes | uL+ operational check fail therapy test disabled (Device is EoL and no work performed by uYy) | uL+ operational check fail therapy test disabled (Device is EoL and no work performed by uYy) | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Produt is EOL, Informed cutomer. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Operational check fail/Du therapy test disabled ----------------------------------------------------------------------------- | 2023-10-19 10:13 | No
2486 | Yes | Mu0 show speaker abnormal. Onsite service pending | Mu0 show speaker abnormal. Onsite service pending | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 00000000R_1600000000OR_mu500000000_show speaker abnormal SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for Onsite Suport SAP Transfer:Case Transferred to SAP FSE Notes:This PO is a Reference Nuber only Safety Qustions:Case: 000000002200000000600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Wang Emily Last udated by: Wang Emily 1. was the device in clinical uage when the problem occured? No 2. was the patient or uer hut? No 3. for devices with alarm/warning featues, did the device send ou alarm/warning when the problem occured in time? No 00000000was the problem open-bou failue? No ----------------------------------------------------------------------------- | 2023-10-19 10:48 | No
2487 | Yes | System does not boot even after shut down | System does not boot even after shut down | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000013 November 7, 2000000003 Hofbeck, Ralf The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Called 00000000000000001199132700000000 problem confirmed Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Parts Consued: 1.0000000000000000000000000000000080000000058600000000, FleuVision 2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.00000000000000000000000000000000800000000387582,PD. SCOMP. DCPS_20000000012V_5V Resoluion: HDD replaced, software installed, system handed over ready for operation ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000300000000ölck, Jan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: Called 00000000000000001199132700000000 problem confirmed Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: System boots and hangs at 00000000:00000000, Touh pad is also dark Cold start withou sucess Internal Comments: uLL# and/or description of uLL Fleuvison boots disk read uLL Repair Action: Te 3 hous to replace the Fleuvision PC or possibly 200000000PDM Power Suply Follow-u requred: uLL# and/or description of uLL Fleuvison boots disk read uLL Repair Action: Te 3 hous to replace the Fleuvision PC or possibly 200000000PDM Power Suply Technician name: Völck, Jan ----------------------------------------------------------------------------- Cutomer comment: ## Service with Ms. Theile for 200000000100000000 from approu. 1:00000000 pm. ----------------------------------------------------------------------------- | 2023-10-19 11:33 | No
2488 | Yes | 0S:SS: The equipment shut down on its own. | 0CS:SS: The equipment shut down on its own. | 000000000000000000000000001O##: 00000000000000000000000000100000000261595 31-Oct-2000000003 Takebayashi, TakanobuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The equpment shu down on its own. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The equpment shu down on its own. uer Impact: No effect Patient Impact: No effect SE 8.2.3000000005 Resoluion: It is the PARTS processing 000000000000000000000000001 so I will close this one. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000700000000200000000Oct-2000000003 Takebayashi, TakanobuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The equpment shu down on its own. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The equpment shu down on its own. uer Impact: No effect Patient Impact: No effect SE 8.2.3000000005 Resoluion: I will close this becaue it is the 000000000000000000000000001for parts processing. Parts Consued: 1.00000000000000000000000000000000800000000387582,PD. SCOMP. DCPS_20000000012V_5V ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000009 19-Oct-2000000003 Takebayashi, TakanobuThe system meets the specification for the performed service and is retuned to ue Resoluion: I replaced the 200000000power suply uit in the R cabinet and reset the circut breaker. When I finished, I checked the operation. Parts ued Part nuber: 00000000800000000387582 Part name: PD. SCOMP. DCPS_20000000012v_5_v: 1 Problem Description by Engineer: The equpment shu down on its own. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The equpment shu down on its own. uer Impact: No effect Patient Impact: No effect SE 8.2.3000000005 Reguatory Qustions: N/A Follow u Requred: N/A Resoluion: Regarding additional 000000000000000000000000001k, I replaced the 00for the remote connection. The DICOM Writer does not display on large monitors. The converter is fauty, so I have connected the DIVI from the PC to the inpu bou. Follow u Requred: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000aka, Shuji The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: The equpment shu down on its own. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The equpment shu down on its own. uer Impact: No effect Patient Impact: No effect SE 8.2.3000000005 Diagnostic performed by Engineer: RC Response: After it shu down on its own, when the power was tuned on the arm bed 000000000000000000000000001ld not move. It did not 000000000000000000000000001k after a warm restart, so he tuned the power off/on and he was in the middle of restarting the system. I checked to see if the power off buton had been left on, and jut to be sue I had him pul the switch back on and wait for it to start, and then it 000000000000000000000000001ked normally. I checked the log. 1000000009:200000000Power Down event received from MPD There was a power off command. At the 1st time of the lauch, there was an uLLwith both the arm bed interface bou and side flat panel commuications and front side collimeter.(R cabi power is common) Visit and investigate. Internal Comments: [Suject] 00000000S:SS:u Kuimama: The equpment shu down on its own. [Parts from the equpment] 00000000800000000000000005100000000Service part REVIEW MODuE QTY:1 ZRFA BackOrder 00000000230000000025312:CABLE GEO_REVIEW MODuE QTY:1 ZRFA Osaka (u15) 00000000800000000387582:PD. SCOMP. DCPS_20000000012V_5V QTY:1 ZRFA/Traceable Osaka (u15) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [TakanobuTakebayashi] Takebayashi, Takanobu - FSE1 Visit date and time 2000000003/1000000001917:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Osaka Office - Received person name TakanobuTakebayashi - 0000000000000000002180000000019300000000 - Others Overseas order 00000000800000000000000005100000000Service part REVIEW MODuE QTY:1 BackOrder Requster name Toshihito Saka ----------------- shipping instrutions #2----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name TakanobuTakebayashi - 0000000000000000002180000000019300000000 - Others Call mobile uon arrival 00000000230000000025312:CABLE GEO_REVIEW MODuE QTY:1 Osaka (u15) 00000000800000000387582:PD. SCOMP. DCPS_20000000012V_5V QTY:1 Osaka (u15) Requster name Toshihito Saka [Euamination category] Normal euamination Follow u Requred: ⋆ ----------------------------------------------------------------------------- | 2023-10-19 11:58 | No
2489 | Yes | Equipment alarm | Equipment alarm | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000N/HS1 (W B) Defib Simp. ChineseEuchg Safety Qustions: Case: 000000002200000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No Internal Remarks: 200000000000000002000000003 00000000:52:33 uC (860000000000000000100000000 $ASP engineer:[honglejia wudehong] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.Alarm euception $Q8: The battery is low $Q9: Cutomers are advised to puchase replacement batteries $Q100000000 Test and inspection data refer to: Basic performance check: Passed. Device retuned to ful fuctionality. $Q11:Calibrated Tools:N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 200000000000000002000000003 00000000:52:31 uC (860000000000000000100000000Department information: [Department: Logistics Department Contact: Chen Gong Tel: 18565000000005563 ]* $ASP engineer:[honglejia wudehong] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.Alarm euception $Q8: The battery is low $Q9: Cutomers are advised to puchase replacement batteries $Q100000000 Test and inspection data refer to: Basic performance check: Passed. Device retuned to ful fuctionality. $Q11:Calibrated Tools:N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Cutomers reported that the device kept alarming The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Replacement of electrode bou and battery bu ineffective Repair Action: T2 Recommended on-site repair steps: It is recommended to check the host Recommended spare part uufor T2 with description: #0000000056000000002711 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA ----------------------------------------------------------------------------- | 2023-10-19 12:07 | No
2490 | Yes | UPS defect u00Centrul Medical Transilvania Cluj | UPS defeuu0B0Centrul Medical Transilvania Cluj | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000016 26-Oct-2000000003 From, Titian Nicolae The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: operator Device ue: startu Eupected and ueupected Resuts: no boot uer Impact: none Patient Impact: none SE 2.2 L3 Diagnosis performed by Engineer: The defective uS is replaced, fuctionality tests are performed. The angiograph 000000000000000000000000001ks within the appropriate parameters. Parts Consued: 1.00000000000000000000000000000000800000000117161, uS 1Phase Data Integrity Controller SP Tools & Equpment : MuTIMETER, 2000000003-10000000028 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000032 19-Oct-2000000003 From, Titian Nicolae The system meets the specification for the performed service and is retuned to ue Follow u Requred: uS Controller replacement Problem Description by Engineer: Cutomer Fuction/Role: operator Device ue: startu Eupected and ueupected Resuts: no boot uer Impact: none Patient Impact: none SE 2.2 L3 Diagnosis performed by Engineer: The uS Controller is foud to be defective. The fuctional angiograph is left uder suervision. It is necessary to replace the defective part. Tools & Equpment : MuTIMETER, 2000000003-10000000028 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-19 12:13 | No
2491 | Yes | 0/u FD0Geometry movements problem | 0R/u FD0Geometry movements problem | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Overview of cutomer problems: After the equpment is tuned on, all mechanical movements are disabled. The euamination bed and C-arm are not powered on and locked, and can be puhed by hand.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: uable to commuicate with GeoIPC. No movements are available (reboot done/needed).. uable to commuicate with GeoIPC Remote diagnostics and test resuts: The mechanical motion control host of the equpment is not connected to the system. PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Check if the BIOS battery is normal 3V, Root Caue 1: Please ignore this alert, if it only occus sporadically. Treat it then as a 'one time only'. Service Action 1: Perform a reboot of the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other Soluion here below. Root Caue 1: GeoIPC is not booting u at all or fast enouh. As resut the Host PC is uable to commuicate with the Geo IPC within a specified time. Service Action 1: Verify the behavior of GEOIPC at first boot u of the system in the morning (at cold start-u): If inpu power (2300000000) is present bu the GEOIPC does not power u at all, it is caued du to a defective internal power suply of GEO IPC itself. Service Action 2: If GeoIPC does power u fine bu Host PC is uable to commuicate with it in time, it is caued du to a defective harddisk. (corruted OS/AS) Then replace the GeoIPC. Service Action 3: Check whether the cable pins between GeoIPC RC-u15 and IPC adapter SCC1-u22 are perhaps too short. Advise: tighten the RC-u15 and SCC1-u12 connectors or replace the cable between RC-u15 and SCC1-u12. Service Action 00000000 It coud be caued du to a defective control net000000000000000000000000001k switch (TS) in the R-cabinet or its Power Suply. Root Caue 1: When problem is persistent it coud be caued du to a defective control net000000000000000000000000001k switch (TS). Service Action 1: Check control net000000000000000000000000001k switch (TS) and its power suply (adapter). Replace when necessary. Root Caue 1: Bad cabling and/or connections. Service Action 1: Check the cable RC:u25 to SCC1:u22. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. Service Action 2: Tighten all the connectors of GeoIPC and SCC1 uit. Root Caue 1: Issu with GeoIPC OS/AS software. It coud be du to a not correct shudown before, Windows OS can be corruted hence it needs more time in the neut start u. If this happens only sporadically, please ignore it. Else see other possible caues & soluions. Service Action 1: If SW is really corruted please reload GeoIPC software uing uB destroy stick. Root Caue 1: The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. Service Action 1: Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start u of the system. If so, replace GEO IPC. Notes on cutomer requrements: Onsite Euternal Remarks: 23.1000000002000000003 00000000:21:17 uC (8600000000000000009) Service content: The device power suply fails and needs to be reset. Soluion/Repair: After resetting the power suply, complete the relevant corrections and retun to normal. Problem Reported by Cutomer: Operation panel failue Internal Remarks: 23.1000000002000000003 00000000:21:200000000uC (8600000000000000009) RDI Accuacy: Yes $Q7: (Confirmation & Description of problem)Confirmed reported problem, the problem is cannot boot u $Q8: (Trouleshooting & Resuts)Checked the system statu is not ok, the problem is cannot boot u then check the system statu foud the system need to reset. $Q9: (Repair steps)Complete reset the system, system statu ok now. $Q100000000 (Test and Inspection data & Device statu)Test and inspection data refer to PA. Device retuned to ful fuctionality $Q11: $Calibrated tools name（NA）S/N (NA) Neut calibration date (NA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:NA ^Curent Tue uage: NA Safety Qustions: Case: 000000002200000000000000001 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-19 12:23 | No
2492 | Yes | Speaker malfunction -euchange - other HW Speaker/ Replacement sent | Speaker malfunuon -euchange - other HW Speaker/ Replacement sent | 000000000000000000000000001O##: 000000000000000000000000001000000002638100000000 Parts Consued : 1.0000000000000000000000000000000056000000002531,TELE Mu00000000 2.00000000hz, uonly, euchange ----------------------------------------------------------------------------- | 2023-10-19 12:23 | No
2493 | Yes | System stuck on "0:0" | System stuck on "0:0" | 000000000000000000000000001O##: 00000000000000000000000000100000000256213 October 2000000002000000003 MARTIN LITT The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: u Eupected and ueupected Resuts: issu to select fleu vision inpus. uer Impact: no Patient Impact: na SE Resoluion: Replacement of the new uuubecaue video signals are not displayed normally. The diagnosis of the PC with the diagnostic tool shows a problem with the grabber cards. Replaced the PC and installed the software. Test OK Parts Consued: 1.0000000000000000000000000000000080000000058600000000, FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000253879 October 200000000 2000000003 MARTIN LITT The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: grabber card issu, uuuneeds to be replaced twice. Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: u Eupected and ueupected Resuts: issu to select fleu vision inpus. uer Impact: no Patient Impact: na SE Diagnostic performed by Engineer: uuuneeds to be replaced becaue the grabber cards are defective. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002538800000000ctober 19, 2000000003 PATRICK TERRAS The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer: After replacing the fleuvision PC (yesterday), the cutomer reports that they are not able to have the Hemo Live video on the Fleuvision monitor. No remote soluion possible, an on-site visit has been schedued. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000300000000ctober 19, 2000000003 NICOLAS MEDINANE The system meets the specifications for the performed service and is retuned to ue. Resoluion: Install a new Fleuvision PC. Reinstall, the OS, the firm and the application. After test, everything 000000000000000000000000001ks properly. Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: start-u Eupected and ueupected Resuts: System stuk at 00000000:00000000 uer Impact: / Patient Impact: / SE / Diagnostic performed by Engineer: Test the fleuvision PC with the PC Diagnostic tool. It comes ou that there is a problem with one of the Grabber cards. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000058600000000, FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000200000000ANTOINE SANTANGELO The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: Nicolas onsite 12-16:00000000 19/100000000 Part: 0000000080000000058600000000 + 0000000080000000050000000000000000 Delivery: TAuI ASAP Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: start-u Eupected and ueupected Resuts: System stuk at 00000000:00000000 uer Impact: / Patient Impact: / SE / Diagnostic performed by Engineer: Fleuvision PC probably down ----------------------------------------------------------------------------- | 2023-10-19 12:33 | No
2494 | Yes | uuutest discharges incorrect values - | u test discharges incorreulues - | 000000000000000000000000001O##: 0000000000000000000000000010000000023800000000 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: test discharges incorrect valus uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Foud valus that are not always correct in the discharge test, intervention will follow. Follow u Requred: as per talk with technician Milani, he carried ou the test several times withou detecting anomalies, in the discharge test, the impoedence valu is abou 00000000ohm instead of 500000000 proceeded with order of condenser for replacement The defibrillator was taken to the laboratory of Clinical Engineering, floor -1 San Bassiano Hospital (Bassano del Grappa (VII)) ----------------------------------------------------------------------------- | 2023-10-19 13:18 | No
2495 | Yes | Triple Chirp | Triple Chirp | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : uLLmessage SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 000000002200000000539 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-19 13:58 | No
2496 | Yes | It shows a failure in the auto-test. | It shows a failure in the auto-test. | 000000000000000000000000001O##: 00000000000000000000000000100000000311978 Ruz, Ignacio The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: Defibrillator. Eupected and ueupected Resuts: uer Impact: uknown. Patient Impact: uknown. SE N.00000000.00000000 Diagnostic performed by Engineer: El equpo tiene u uLLal pasar los test del "Operational Check". Se observa en el log u uLL qu confirma qu el condenador está gastado. Se pruba con uo de culificación y se compruba el correcto fucionamiento del equpo. Se recomienda sutiturlo para soluionar el problema. Follow u Requred: El equpo tiene u uLLal pasar los test del "Operational Check". Se observa en el log u uLL qu confirma qu el condenador está gastado. Se pruba con uo de culificación y se compruba el correcto fucionamiento del equpo. Se recomienda sutiturlo para soluionar el problema. ----------------------------------------------------------------------------- | 2023-10-19 14:23 | No
2497 | Yes | M0A / u00 HeartStart uL Defibrillator/Monitor / Discharge amount is greater than the set energy. | M00A / u000/ HeartStart uL Defibrillator/Monitor / Discharge amount is greater than the set energy. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Zhenjiang First People's Hospital The discharge amout is greater than the set energy Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Internal Remarks: pending timeou: DM metting Safety Qustions: Case: 0000000022000000006200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-19 14:23 | No
2498 | Yes | M0A - u00 Energy delivery meterage failed | u Energy delivery meterage failed | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: The discharge amout is greater than the set energy Internal Remarks: pending timeou: DM metting Safety Qustions: Case: 000000002200000000623 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-19 14:23 | No
2499 | Yes | Defective uuu GFE00OV0 No response/IS complete | Defeu uuu GFE0 0NOV00No response/IS complete | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000072 October 26, 2000000003 BASHIR, MouamadouSeeer The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: uuuDEFECTIVE uer Impact: No. Patient Impact: No. SE uknown Resoluion: The front cover was replaced. Consued parts: 1.0000000000000000000000000000000056153300000000, RP-Bezel, Front, w/NAV, Wht, CTD, V600000000Tools & Equpment: ELECTRICAL SAFETY TESTER, 2000000003-12-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000238981 Raffenne, Virginie The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: uuuDEFECTIVE uer Impact: No. Patient Impact: No. SE uknown Follow u Requred: SERVICE QuTE WITHOu TRAVEL COSTS. 2H M1 The cutomer will call MouamadouBashir as soon as the part is received at the biomedical department and he will come at the same time as for the FCO. Thank youResoluion: Sending the service qute with only the labor costs and the part. 0000000056153300000000 Bezel, front, Non-English (white/coated) for SN >= 100000000111553 & 200000000000000007800000000 Diagnostic performed by Engineer: Futher to the preventive maintenance, it was observed that the uuuis defective. Requsts a qute for the intervention withou travel. The technician will come to replace the part at the same time as the replacement of the card schedued in the frame000000000000000000000000001k of the 2000000001-CC-HRC-000000003. ----------------------------------------------------------------------------- | 2023-10-19 14:57 | No
2500 | Yes | The device cannot complete the operation test | The device cannot complete the operation test | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000075 Hueyin TasselluThe system does not meet the ful specification for the performed service and is retuned to ue with limitation as accepted by the cutomer, no follow-u requsted. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Ouside of ue Eupected and ueupected Resuts: The device cannot complete the operation test. uer Impact: No Patient Impact: No SE Resoluion: The cutomer has agreed that the device shoud be sent to ou shop for detailed analysis. Diagnostic performed by Engineer: The cutomer has agreed that the device shoud be sent to ou shop for detailed analysis. ----------------------------------------------------------------------------- | 2023-10-19 15:28 | No
2501 | Yes | CV*System down* u-ray not possible | CV*System down* u-ray not possible | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000012 22-Oct-2000000003 Mohamed Maher The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : CV*System down* u-ray not possible Information to suport the complaint handling process Cutomer Fuction/Role: Technician Device ue: No patient involved Eupected and ueupected Resuts: CV*System down* u-ray not possible, Call Service , eupected: both fluroscopy and euposue 000000000000000000000000001king uer Impact: no impact Patient Impact: no impact SE 8.2.3000000005 Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell IP PC no HDD Diagnostic performed by Engineer : After checking the system, it was foud that the IP-PC is failed in power suply and t was replaced by new one and reinstall the software on it, system u and 000000000000000000000000001king fine. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000Khaled Megud Fawzy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : CV*System down* u-ray not possible Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: CV*System down* u-ray not possible uer Impact: no impact Patient Impact: no impact SE 8.2.3000000005 Diagnostic performed by Engineer : CV*System down* u-ray not possible Internal Comments : uLL# and/or description of uLL: CV*System down* u-ray not possible Repair Action : CV*System down* u-ray not possible onsite visit needed to replace Haswell IPPC. Resoluion : CV*System down* u-ray not possible onsite visit needed to replace Haswell IPPC. Follow u Requred : uLL# and/or description of uLL: CV*System down* u-ray not possible Repair Action : CV*System down* u-ray not possible onsite visit needed to replace Haswell IPPC. Technician Name : Khaled Megud Fawzy ----------------------------------------------------------------------------- | 2023-10-19 15:38 | No
2502 | Yes | uLS AED ¦ screen is freezing SPLIT (cancelled) | TempuLS AED ¦ screen is freezing SPLIT (cancelled) |  | 2023-10-19 16:20 | No
2503 | Yes | uPro ¦ physical damage on display caung intermittent touch issues in certain areas of the touch screen SPLIT | TempuPro ¦ physical damage on display caung intermittent touch issues in certain areas of the touch screen SPLIT |  | 2023-10-19 16:25 | No
2504 | Yes | Misalignment in the touch position was confirmed | Misalignment in the touch position was confirmed |  | 2023-10-19 17:23 | No
2505 | Yes | navigation ring failure | navigation ring failure |  | 2023-10-19 17:36 | No
2506 | Yes | Backup alarm failed (EC 00 | Backup alarm failed (EC 00) |  | 2023-10-19 17:46 | No
2507 | Yes | C arc blocked | C arc blocked | 000000000000000000000000001O##: 00000000000000000000000000100000000339619 22-Nov-2000000003 Bernasconi Nicola The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: trm Device ue: yes Eupected and ueupected Resuts: Check Switch and Order Part uer Impact: none Patient Impact: none SE8.1 Resoluion: Material consued ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000030000000027-Oct-2000000003 Bernasconi Nicola The system meets the specification for the performed service and is retuned to ue Follow u Requred: Replacement of bow handle C Resoluion: Replacement schedued for 11/100000000roblem Description by Engineer: Cutomer Fuction/Role: trm Device ue: yes Eupected and ueupected Resuts: Check Switch and Order Part uer Impact: none Patient Impact: none SE8.1 Parts Consued: 1.00000000000000000000000000000000230000000012551,SAM HANDGREEP L-ARM Y. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 19-Oct-2000000003 Bernasconi Nicola The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: trm Device ue: yes Eupected and ueupected Resuts: Check Switch and Order Part uer Impact: none Patient Impact: none SE8.1 Resoluion: C bow blocked Geo Modue fuctioning checked E-stop checked Bow brake butons checked: sometimes it gets stuk Handle tidy Follow u Requred: Replace handle and brake butons for bow movement ----------------------------------------------------------------------------- Internal Remarks: Package not today, we are 000000000000000000000000001king with warehoue shipped to know the reason behind the same, will get back with the inpus post having the feedback with warehoue Internal Remarks: order come after cuoff will be shipped today. AWB# 1Z39R500000000000000003000000009000000002 ----------------------------------------------------------------------------- | 2023-10-19 17:58 | No
2508 | Yes | [Q] u ray unavailable error | [Q] u ray unavailable error | 000000000000000000000000001O##: 00000000000000000000000000100000000255715 27-Oct-2000000003 Karthik, R The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.000000000000000000000000000000002100000000125700000000oco Velara Replacement Kit Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000001-17 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000001-17 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000001-17 00000000:00000000:00000000 Resoluion : Replaced RoCo system 000000000000000000000000001king ok Problem Description by Engineer : Problem description by engineer : u ray uavailable uLL When was the first occurence? Within... : 19-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: u ray uavailable uLL uer Impact: No uer impact Patient Impact: No patient impact SE 7.2.16.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000025 200000000Oct-2000000003 Karthik, R The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace Cue Problem Description by Engineer : Problem description by engineer : u ray uavailable uLL When was the first occurence? Within... : 19-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: u ray uavailable uLL uer Impact: No uer impact Patient Impact: No patient impact SE 7.2.16.00000000Diagnostic performed by Engineer : Replaced u tue , foud Cue and KVmA malfuctioning ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000016 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: u ray uavailable uLL Repair Action : FSE visit requred for futher checks Resoluion : FSE visit requred for futher checks Follow u Requred : uLL# and/or description of uLL: u ray uavailable uLL Repair Action : FSE visit requred for futher checks Technician Name : Vijayakuar, K.S Problem Description by Engineer : Problem description by engineer : u ray uavailable uLL When was the first occurence? Within... : 19-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: u ray uavailable uLL uer Impact: No uer impact Patient Impact: No patient impact SE 7.2.16.00000000Diagnostic performed by Engineer : FSE visit requred for futher checks ----------------------------------------------------------------------------- | 2023-10-19 18:03 | No
2509 | Yes | PSA_ System geometry works not as designed while emergency patient was in treatment | PSA_ System geometry works not as designed while emergency patient was in treatment |  | 2023-10-19 18:28 | No
2510 | Yes | Backup Alarm Failed | Backup Alarm Failed |  | 2023-10-19 18:50 | No
2511 | Yes | Machine is not switching getting on | Machine is not switching getting on | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000027 200000000Oct-2000000003 Puohit, Vatsal The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching getting on When was the first occurence? Within... : 19-Oct-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching getting on uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.17.00000000Resoluion : Done system power recycle and reset of Host PC connections. System now 000000000000000000000000001king fine. Tools & Equpment : MuTIMETER,2000000003-12-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000097 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Machine is not switching getting on Repair Action : FSE visit requred for futher checks on GPDu Resoluion : FSE visit requred for futher checks on GPDu Follow u Requred : uLL# and/or description of uLL: Machine is not switching getting on Repair Action : FSE visit requred for futher checks on GPDu Technician Name : Vijayakuar, K.S Problem Description by Engineer : Problem description by engineer : Machine is not switching getting on When was the first occurence? Within... : 19-Oct-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching getting on uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.17.00000000Diagnostic performed by Engineer : FSE visit requred for futher checks on GPDu ----------------------------------------------------------------------------- | 2023-10-19 19:03 | No
2512 | Yes | System stuck at 0:0 | System stuck at 0:0 | 000000000000000000000000001O##: 0000000000000000000000000010000000025100000000 200000000Oct-2000000003 Salazar William The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Resoluion : FleuVision PC replaced sucessfuly and system 000000000000000000000000001king again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000000000000019-Oct-2000000003 Salazar William The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Fleuvision PC needs replacement Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Fleuvision PC needs replacement ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000085 George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ue alt + f to start fsf, and review POST resuts, download board software, and event log to determine caue of problem. Verify fleuvision PC is powered and commuicating. If needed, 00000000800000000119123 - FleuVision 2 PC No HDD, COA Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Debriefed caller. He said uers reported the system fuly bootred first thing this morning eucept Fleuvision did not come u. When caller arrived, he rebooted system and it is know stuk at 00000000:00000000 each time system reboots. Recommended he remove front cover of B cabinet and verify fleuvision PC is powered and if not, ue switch on PC to try and tun on. Also recommended he ue alt + f to start fsf. Coud not resolve the issu and requres onsite suport ----------------------------------------------------------------------------- Internal Remarks : Approved for replacement 00000000800000000119123 - FleuVision 2 PC No HDD, COA. PC Diag Tool showing grabber cards are not itializing, replace pc. ----------------------------------------------------------------------------- | 2023-10-19 19:03 | No
2513 | Yes | Primary Alarm Failed (EC 00- Final 07 | Primary Alarm Failed (EC 00) - Final 0/7 | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: uLLcode 1100000000 uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states while performing PM foud uLLcode 1100000000 on device. 1100000000: Primary alarm failue. Troule shot with cutomer: 1. Listen for auible soud from the speakers. 2. Verify that the speakers are connected. 3. Verify that the braided wires are not touhing the speaker houing. 00000000Verify that the resistance of each speaker is 6.8 to 9.2 Ω (Figue 6-2). 5. Replace speaker #1 (MC PCBA). 6. Replace speaker #2 (PM PCBA). 7. Replace the CPuPCBA. Cutomer states he isn't curently in front of device. he will troule shot with device and callback if futher assistance is needed. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Travis Dickey IP & Issu: 9200000000723 - The V600000000is alarming power speakers failed; primary alarm failue. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case, R000000000000000000000000001and roue to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-19 19:08 | No
2514 | Yes | Request for spare parts | Request for spare parts | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000027 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: patient monitoring Eupected and ueupected Resuts: The system does not detect the utrace either with the cable or throuh the plates uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u. Angelica from ELCAMM reports that technician u Michelangelo Totaro foud that the defibrillator does not detect the utrace either with the cable or throuh the plates and asks for the delivery of spare parts. Follow u Requred: Make NEMO withou collection at the same time. Tu Elcamm S.r.L. Via Roma, 139 (ASI Area) 70000000026 Moduno (BA) Ac. u Michelangelo Totaro Tel.0000000000000000337726 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027800000000 Parts Consued: 1.0000000000000000000000000000000056000000001751,Intrepid Processor PCA ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000263732 Parts Consued: 1.0000000000000000000000000000000056000000001700000000Intrepid Therapy PCA ASSY 1.0000000000000000000000000000000056000000001811, Intrepid uSpO2+NBP+CO2 Frame ASSY ----------------------------------------------------------------------------- Internal Remarks: Good morning, We ask for the suply of the following spare parts uder warranty: 0000000056000000001700000000 0000000056000000001811 The defibrillator is u property, delivered to u on loan and located at ou offices. Awaiting feedback. The requst is ugent. Kind regards Elcamm Srl Via Roma 139 70000000026 Moduno (BA) 00 00000000000000005337726 Fau: 00000000000000005330000000059 E-mail: elcamm00lcamm.it Pec: elcammsrl00ec.it Internal Remarks: Good morning, We ask if the spare parts requsted with the case in qustion were sent. Thank you Kind regards Elcamm Srl Via Roma 139 70000000026 Moduno (BA) 00 00000000000000005337726 Fau: 00000000000000005330000000059 E-mail: elcamm00lcamm.it Pec: elcammsrl00ec.it Internal Remarks: Hi, Opened IPCR 20000000031000000001272681 for the location ----------------------------------------------------------------------------- | 2023-10-19 19:18 | No
2515 | Yes | 0CDN0uck at 00 does not start | 00CDN0tuck at 00 does not start | 000000000000000000000000001O##: 00000000000000000000000000100000000252395 21-Oct-2000000003 Coyne Stephen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : spoke to Rus ,he system did not start u this morning When was the first occurence? Within... : 19-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system du not boot as eupected uer Impact: system down Patient Impact: none SE 8.1.3000000005 Resoluion : replaced fleuvision pc with original HDD downloaded board software additional problem with table geometry uLL foud connection issu with Geometry PC performed safety and tested operation packed and retuned parts Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000027 200000000Oct-2000000003 Coyne Stephen The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : parts ordered Problem Description by Engineer : Problem description by engineer : spoke to Rus ,he system did not start u this morning When was the first occurence? Within... : 19-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system du not boot as eupected uer Impact: system down Patient Impact: none SE 8.1.3000000005 Diagnostic performed by Engineer : checked logs - fleuvision uLL fleuvision not connected - diag test no display on fleuvision monitor - checked HDD connection Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000051 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: stuk at 0000000000000000 Repair Action : from install base Produt Name FleuVision-2_Revised PC Produt Modality Produt Code 0000000080000000058600000000 to 0000000080000000058600000000 FleuVision-2_Revised PC no HDD 1 For u R8.u.16.u and onwards; withou OS Disk; with COA label; Mandatory for u uuuR8.u.25.u systems Diagnostic performed by Engineer : speaking to Rus,he had already shu down system after seeing the 0000000000000000 .it took longer than he supected We tried one more startu of system same resut . he can hear power to stand bu fleu vision screen is black. Problem Description by Engineer : Problem description by engineer : spoke to Rus ,he system did not start u this morning When was the first occurence? Within... : 19-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system du not boot as eupected uer Impact: system down Patient Impact: none SE 8.1.3000000005 Follow u Requred : uLL# and/or description of uLL: stuk at 0000000000000000 Repair Action : from install base Produt Name FleuVision-2_Revised PC Produt Modality Produt Code 0000000080000000058600000000 to 0000000080000000058600000000 FleuVision-2_Revised PC no HDD 1 For u R8.u.16.u and onwards; withou OS Disk; with COA label; Mandatory for u uuuR8.u.25.u systems What are the neut Trouleshooting steps : check start u of fleuvision pc and power to pc Technician Name : Sulivan Scott ----------------------------------------------------------------------------- Internal Remarks : 0000000022000000007200000000Coyne Stephen FSE calling in again on u 8.1.3000000005 where he needs part nubers for Fleu PC and hard drive. FSE is frutrated becaue he can not seem to get IC2 parts to 000000000000000000000000001k. I shared screen with him and showed him how I maneuered in IC2 parts and foud the part nubers he needed. Internal Remarks : Steve Coin - fleu monitor and acqusition monitor are black. RSE sugested to replace fleu. I concu, along with HDD. RSE suplied the uus Internal Remarks : Steve Coyne - connectors on new fleu pC did not match old one. FSE was looking at the wrong PC (GEO). Instruted on where to find euisting FLEu PC Internal Remarks : 0000000022000000007200000000Coyne Stephen FSE 000000000000000000000000001king on u FD200000000where system is stuk at 00000000:00000000. Advised FSE to try Alt-F at splash screen when lower left corner symbols jut change from initial symbol. Shoud ask for his dongle pass000000000000000000000000001d and get into FSF to grab logs. ----------------------------------------------------------------------------- | 2023-10-19 19:23 | No
2516 | Yes | The units touch screen is not working correctly-split | The units touch screen is not working correuy-split |  | 2023-10-19 19:33 | No
2517 | Yes | Remote Alert: RMT - IPU: Fluoro storage not possible due to an Image disk problem (>Rel0 | Remote Alert: RMT - IPU: Fluoro storage not possible due to an Image disk problem (>Rel00 | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000093 19-Oct-2000000003 Bouans Ashley The system meets the specification for the performed service and is retuned to ue Resoluion : Puled logfiles, checked over logfiles and spoke with cutomer. Low image storage message was on the screen when I arrived on-site. After looking thrulogfiles I didn’t see any issus ouside of the HDD being ful. Completed the (re)create image and rebooted system. Message disappeared and system seemed to be fuly fuctional afterwards. Completed test shot and fluro jut to be safe. Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: in ue Eupected and ueupected Resuts: image storage low uLLappearing in screen uer Impact: none Patient Impact: none SE none ----------------------------------------------------------------------------- Internal Remarks : Called FSE Ashley Bouans they advised dispatching today Oct 18 at 00000000 ----------------------------------------------------------------------------- | 2023-10-19 19:33 | No
2518 | Yes | No image, loud u-ray tube | No image, loud u-ray tube |  | 2023-10-19 19:43 | No
2519 | Yes | Smoke coming form generator cabinet | Smoke coming form generator cabinet | 000000000000000000000000001O##: 0000000000000000000000000010000000025000000006000000003-Oct-2000000003 Jose Diaz, The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Smoke coming form generator cabinet When was the first occurence? Within... : 19-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Smoke coming form generator cabinet uer Impact: System is down Patient Impact: None/ no procedue SE 1.2.5.00000000Diagnostic performed by Engineer : FSE replaced frontal generator cooling uit. All fuctional tests passed. FSE also had lab tech ue system to verify fuctionality. Testing and verification complete and OTR report completed. Parts Consued : 1.0000000000000000000000000000000022000000000000000052,Cu3100000000 for Certeray (service) Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 MuTIMETER,2000000005-00000000-28 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000037 19-Oct-2000000003 Swanson Jeffrey The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Smoke coming form generator cabinet When was the first occurence? Within... : 19-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Smoke coming form generator cabinet uer Impact: System is down Patient Impact: None/ no procedue SE 1.2.5.00000000Diagnostic performed by Engineer : On site 10000000019 foud system off. Inspected and foud coolant flud on floor and inside generator cabinet, level in cooler at bottom. Testing foud that braided hose leaking. Requst to order cooler uit. Follow u Requred : Follow u to replace cooling uit in generator frontal plane ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000009 Whitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Smoke coming form generator cabinet When was the first occurence? Within... : 19-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Smoke coming form generator cabinet uer Impact: System is down Patient Impact: None/ no procedue SE 1.2.5.00000000Diagnostic performed by Engineer : Steven reports that on first boot of the morning smoke was coming ou of the top of one of the generator cabinets. The smoke alarm went off, he shu the room down and it is not smoking at this time. Steven reports that it smells like coolant Internal Comments : uLL# and/or description of uLL: None reported Repair Action : Verify were smoke was coming from, check coolant hoses and TCufor leaks, repair or replace damaged parts and test 0000000022000000000000000052 - Cu3100000000 for Certeray (service) 000000002100000000173 - OIL CAN WITH OIL 2,5L Follow u Requred : uLL# and/or description of uLL: None reported Repair Action : Verify were smoke was coming from, check coolant hoses and TCufor leaks, repair or replace damaged parts and test 0000000022000000000000000052 - Cu3100000000 for Certeray (service) 000000002100000000173 - OIL CAN WITH OIL 2,5L Technician Name : Whitaker, Steven ----------------------------------------------------------------------------- Internal Remarks : Planner Notes Planner Initials: TB Action Taken: The planner contacted the site's secondary FSE Jeffery Swanson to schedue 0000000000000000000000000010000000020000000037 and discus u remote engineer diagnosis via the follow-u activity AN0000000088582, and the engineer accepted 0000000000000000000000000010000000020000000037 schedued for today 10000000019 with onsite arrival between 11-11:300000000AM PST. The planner contacted the listed cutomer: Steven McGauh to provide an ETA of 0000000000000000000000000010000000020000000037, bu the call jut rang util I received a "call times ou uLL Oucome: DSPCH Neut Steps: FSE accepted Last udated: 10000000019/23 11:19AM EST ----------------------------------------------------------------------------- | 2023-10-19 20:13 | No
2520 | Yes | System did not start up | System did not start up | 000000000000000000000000001O##: 000000000000000000000000001000000002000000007 Oct 19, 2000000003 Tutuowskyj Erhard The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: Reading and analysis of the log file. Log file forwarded to Helpdesk, as no clear uLLcan be foud in the log file. System restarted several times. uLLdid not occu. Problem Description by Engineer: Cutomer Fuction/Role: RT Device ue: yes Eupected and ueupected Resuts: System did not shu down uer Impact: yes Patient Impact: no SE1.2 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000003 Tutuowskyj Erhard The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: System did not start u in the morning. Cutomer had to perform several restarts. Reading ou of the log files requred Problem Description by Engineer: Cutomer Fuction/Role: RT Device ue: yes Eupected and ueupected Resuts: System did not start uer Impact: yes Patient Impact: no SE1.2 ----------------------------------------------------------------------------- | 2023-10-19 20:23 | No
2521 | Yes | During the intervention procedure, sudden interruption of ray supply and shutdown of the room monitors | During the intervention procedure, sudden interruption of ray supply and shutdown of the room monitors | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000058 23-Oct-2000000003 ALBANO, GIOELE The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: Duing the intervention procedue, suden interrution of ray suply and shudown of the room monitors When was the first occurence? Within... : 19-Oct-2000000003 How often is issu occuring? : uknown Malfuctioning area: PC's Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: duing the last euamination in the morning video signals disappeared on the monitors in the room, no u-rays, very slow restart and uLLmessage on image space uer Impact: no Patient Impact: no SE 8.1.3000000005 Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000006-22 00000000:00000000:00000000 MuTIMETER, 200000000000000006-29 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part Resoluion: Replaced data disks of the system database for uLL duing the image acqusition phase. The fuction tests and checks were carried ou with the reguar oucome. The equpment was left 000000000000000000000000001king and ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000700000000NDREA COLOMBO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: Duing the intervention procedue, suden interrution of ray suply and shudown of the room monitors When was the first occurence? Within... : 19-Oct-2000000003 How often is issu occuring? : uknown Malfuctioning area: PC's Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: duing the last euamination in the morning video signals disappeared on the monitors in the room, no u-rays, very slow restart and uLLmessage on image space uer Impact: no Patient Impact: no SE 8.1.3000000005 Diagnostic performed by Engineer: reported by the cutomer. "Duing the intervention procedue, suden interrution of rays suply and shudown of the room monitors" Internal Comments: uLL# and/or description of uLL The system is not reachable remotely, RSN is connected, bu the doors to be able to access are inactive. Tried to contact the cutomer, bu the landline nuber of the room does not respond 0000000000000000397100000000, on cell u Davino is not reachable. Some information is provided by Gioele: "Duing the last euamination in the morning video signals disappeared on the monitors in the room, no u-rays, restart is very slow and uLLmessage on image space". In the history there are no files to check to see the release (from rsw portal it seems 8.1.3000000005) and what type of ippc (from smau it is a nehalem) sugests to intervene with 2u hdd for diagnosis. Repair Action: Schedue onsite intervention for diagnosis with parts (2u hdd) 1 technician 1 day Follow u Requred: uLL# and/or description of uLL The system is not reachable remotely, RSN is connected, bu the doors to be able to access are inactive. I tried to contact the cutomer, bu the landline nuber of the room does not respond 0000000000000000397100000000, on u Davino is not reachable. Some information is provided by Gioele: "Duing the last euamination in the morning video signals disappeared on the monitors in the room, no u-rays, restart is very slow and uLLmessage on image space". In the history there are no files to check to see the release (from rsw portal it seems 8.1.3000000005) and what type of ippc (from smau it is a nehalem) sugests to intervene with 2u hdd for diagnosis. Repair Action: Schedue onsite intervention for diagnosis with parts (2u hdd) 1 technician 1 day What are the neut Trouleshooting steps: Schedue onsite intervention for diagnosis with parts (2u hdd) 1 technician 1 day Technician Name: ANDREA COLOMBO ----------------------------------------------------------------------------- | 2023-10-19 20:28 | No
2522 | Yes | c arm | c arm | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000053 25-Oct-2000000003 Sheets Mike The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: in ue Eupected and ueupected Resuts: System to perform to spec's uer Impact: had to ue a different room Patient Impact: none SEuknown Tools & Equpment : MuTIMETER,200000000000000002-29 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000800000000276513,Control Modue Tilt ER Resoluion : Geo modue not 000000000000000000000000001king on carm prop and roll positions. Reloaded software and it 000000000000000000000000001ked. Rebooted the system several time to verify fuctionality. Geo modue did not 000000000000000000000000001k and 000000000000000000000000001ld no longer take a sw load. FSE replaced geo modue. ----------------------------------------------------------------------------- | 2023-10-19 20:48 | No
2523 | Yes | GS failure, no fluoro or cine | GS failure, no fluoro or 0e | 000000000000000000000000001O##: 00000000000000000000000000100000000266812 200000000ct-2000000003 Nichols,Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: None Eupected and ueupected Resuts: Grid switch uLL uer Impact: intermittent Patient Impact: None SE 7.2 Resoluion : Ordered and replaced u u-ray tue and performed calibrations. Tools & Equpment : MuTIMETER,200000000000000009-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000250000000060000000000000000ct-2000000003 Lantz Verl The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: Diagnostic Eupected and ueupected Resuts: Eupect to not have Gridswitch uLL duing fluro uer Impact: Intermittently can’t fluro Patient Impact: None SE 7.2.16 Resoluion : Performed the Gridswitch test prior to staring the tue replacement and it failed with 115.5 microgy. Replaced the u u-Ray Tue. Performed the following calibrations: Tue Conditioning, Tue Adaptation, Tue Yield, Grid switch passed at 100000000nanogy, Grid Switch Suervisor, Tue Yield and u-ray Field Limitations. Performed all the requred Test and Verifications and all passed, see attached Suvey all resuts. Parts Consued : 1.000000000000000000000000989000000000000000085100000000,u 200000000 000000007 ROT-GS 10000000000000000ools & Equpment : DOSE-KVP-LIGHT METER,200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000000000000031 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000000000000031 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000011 19-Oct-2000000003 Lantz Verl The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: Diagnostic Eupected and ueupected Resuts: Eupect to not have Gridswitch uLL duing fluro uer Impact: Intermittently can’t fluro Patient Impact: None SE 7.2.16 Diagnostic performed by Engineer : Arrived on site and the system had the uLLmessage "Grid Switch uavailable", Tested the system and I was able to make fluro and cine acqusitions. I inspected the log file and there were nuerou uLL abou the Grid switch not available and u-ray not possible over the past 3 days. Foud that the GSS Bou in the generator had a red uLLLED lit showing that it detected a grid switch failue. Shudown the system and powered back on the GSS bou and the system booted u withou any issus or uLL. Performed the test low and high kV tests and had no failues. Performed the Grid Switch test and it failed it was >500000000ngy, measued 25.3uy. Measued the grid switch cable and it tested good. Also checked the small and large filaments with the LCR meter and they both were good at 0000000000000000mH. I coud not reprodue the problem and tried for 300000000minues. Techs came back from luch and said that it 000000000000000000000000001ld show u at steep angles. They 000000000000000000000000001ld move ou of the steep angle and it 000000000000000000000000001ld clear then move back and it 000000000000000000000000001ld show u again. I then tried the steep angles and it did show u bu it 000000000000000000000000001ld not clear for me. Discused my findings with NSS Patrick Helfer and he said to order the tue. Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000000000000031 00000000:00000000:00000000 Follow u Requred : Need follow u 000000000000000000000000001to retun and complete the repair ----------------------------------------------------------------------------- Internal Remarks : Assited FSE Verl Lantz with gridswitch uLL. Approved APM tue, bu no TPR tues are available according to Parts Department. DSM Richard Patke approved the new u Tue ----------------------------------------------------------------------------- | 2023-10-19 20:58 | No
2524 | Yes | Speaker test alarm (EC 0 | Speaker test alarm (EC 00 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000782 November 17, 2000000003 Alejandro Rodrigo Mauier Córdoba The system meets the specification for the performed service and is retuned to ue. Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000000000000023 00000000:00000000:00000000 Resoluion: Client reports internal speaker failue. The uLLcannot be retrieved. The system is left uder testing for three weeks withou observing any anomalies. Louspeaker is preventively sused. It is left uder testing for one week. Fuctional parameters are checked. It is sent to the client. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: normal Eupected and ueupected Resuts: An alarm appears with uLL5000000001 uer Impact: Another V600000000Patient Impact is ued: None SE 3.200000000Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000336565 November 13, 2000000003 Alejandro Rodrigo Mauier Córdoba The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: normal Eupected and ueupected Resuts: An alarm appears with uLL5000000001 uer Impact: Another V600000000Patient Impact is ued: None SE 3.200000000Resoluion: Client reports uLLmessage 5000000001. The client says that the uLLmessage has not occured again. uLLcode 5000000001 makes reference to Speaker. We removed the fan from the service. The equpment is left in tests withou observing a repetition of the uLLmessage. Both louspeakers are measued and they are within parameters. Power Manager louspeaker is replaced. Proper operation tests are performed. The equpment is left uder testing. Follow u Requred: Schedue Tools & Equpment: FLOW METER, 2000000003-12-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000286700000000 00000000-Nov-2000000003 Alejandro Rodrigo Mauier Córdoba The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: normal Eupected and ueupected Resuts: An alarm appears with uLL5000000001 uer Impact: Another V600000000Patient Impact is ued: None SE 3.200000000Resoluion: Client reports failue in operation of V600000000fan. Louspeaker failue message. The equpment is left uder testing. No uLLmessage is observed. Follow u Requred: Schedue Tusday 7th Parts Consued: 1.00000000000000000000000000000000561511921,RP-Speaker Assy, V600000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000285226 October 26, 2000000003 Moreno Álvarez, Pedro The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: normal Eupected and ueupected Resuts: An alarm appears with uLL5000000001 uer Impact: Another V600000000Patient Impact is ued: None SE 3.200000000Diagnostic performed by Engineer: The equpment has an intermittent uLL(5000000001), which refers to the main speakers. When starting the system once in a while the SPEAKER POWER is not recognized. Material is requsted (the speaker) to try to solve the problem. Follow u Requred: Order material to resolve the failue. The equpment is at the bench right now. ----------------------------------------------------------------------------- | 2023-10-19 20:58 | No
2525 | Yes | System has too high value during release | System has too high value during release | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: System has too high valu duing release Patient Impact: - SE - Diagnostic performed by Engineer: Referred to specialty dealer by ##----------------------------------------------------------------------------- Internal Remarks: Referred cutomer to retailers, bu he 000000000000000000000000001ld also like to receive technical information from the manuactuer. Cutomer comment: By ##from >>> d.weidenmaier00endergrupe.de (udelivered): 55000000005.00000000uable to relay (see attachment) ----------------------------------------------------------------------------- | 2023-10-19 21:08 | No
2526 | Yes | Display is dim - Pending parts for repair 04 | Display is dim - Pending parts for repair 0/4 | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000062 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: need parts id ugrade kit /display uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states display is dim, ordered a replacement LCD and foud that other parts are needed to ugrade for 2nd gen LCD to 000000000000000000000000001k with device. Provided cutomer with part nubers and price for parts needed to ugrade LCD. Resoluion : Provided cutomer with part nuber and price: 0000000056153600000000 [RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000 $138.00000000 00000000561528911 [RP-Cable, u PCBA to LCD, V600000000 $66.00000000 00000000561528871 [RP-PCBA, uer Interface, u 2nd GEN,V600000000 $332.00000000 00000000561531651 [RP-Tray,LCD,2nd GEN V600000000 $229.00000000 00000000561528900000000 [RP-Screw,M2u3 Socket hd, PKG/100000000V600000000 $90000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: brett piles IP & Issu: 566276900000000need parts id ugrade kit /display Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-19 21:23 | No
2527 | Yes | Display issue, No GFE required/IS complete | Display issue, No GFE required/IS complete | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000031 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: Thin line across LCD screen - needs parts ID uer Impact: none Patient Impact: na SE 2.00000000Diagnostic performed by Engineer : Discused ugrading from 1st gen LCD to 2nd gen to inclue ugrading software to 2.1 or higher. provided: u assembly, withou nav-ring, V600000000(gray) 1100000000000000000000000056150000000091 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ramon Mejia IP & Issu: 5663000000002 - The LCD part needs replacing. Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-19 21:28 | No
2528 | Yes | Speaker Issue/Malfunction-No Speaker sound for alarm-Device-other-HW/Replacement sent | Speaker Issue/Malfunuon-No Speaker sound for alarm-Device-other-HW/Replacement sent | 000000000000000000000000001O##: 00000000000000000000000000100000000255591 00000000-Nov-2000000003 Gregory Desire The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed/ clinical engineering Device ue: in ue Eupected and ueupected Resuts: speaker malfuction displayed on the device uer Impact: no impact, device is sent ou for euchange Patient Impact: no impact SE N.00000000 Parts Consued : 1.0000000000000000000000000000000056000000005331,TELE PWM,800000000.11a/b/g,uSp00000000,Eu,u only Resoluion : Receive euchanged part. Performed incoming inspection verify device is 000000000000000000000000001king properly. Performed manuactue recommendation T&V all test passed. Transfer licenses and configuation Euchange has been retuned to BIOMED in order for device to be back in rotation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000098 19-Oct-2000000003 Gregory Desire The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A Follow u Requred : need follow u activity in order to order a replacement part in order to complete this repair. Resoluion : Follow u activity created. replacement device ordered Problem Description by Engineer : Follow u activity created. replacement device ordered ----------------------------------------------------------------------------- | 2023-10-19 21:33 | No
2529 | Yes | Check vent backup alarm failed | Check vent backup alarm failed | 000000000000000000000000001O##: 000000000000000000000000001000000002000000005 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Check vent backu alarm failed uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states uit is giving backu alarm failue uLLmessage. confirmed in event log, cutomer foud 1100000000 backu alarm failue. Informed cutomer manuactuer recommendation is to replace the following parts in this order: 1. Replace the MC PCBA. 2. Replace the CPuPCBA. 3. Replace the PM PCBA. Resoluion : Provided cutomer with part nuber and price: 00000000561535700000000- CPu2.100000000SW 00000000561537531 - MC board Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tyler 00000000-500000000-6300000000 IP & Issu: SN 10000000010000000062 / Check vent backu alarm failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-19 21:38 | No
2530 | Yes | THE EQUIPMENT DOES NOT MAKE AN u-ray | THE EQUIPMENT DOES NOT MAKE AN u-ray | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000600000000ctober 200000000 2000000003 Acua, Jose The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: An intervention is performed in the equpment becaue of a rotation uLLin the tue ado. Checked the cooling uit level. Measuements of uV-W coils are performed of the rotation of the ado and measuement of resistances of the same. When trying to fire, the equpment makes a lou noise inside the tue, which indicates that the ado is mechanically defective. It is necessary to replace the u-ray tue. Follow u Requred: u-ray tue replacement Tools & Equpment: MuTIMETER, 200000000000000005-23 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-19 22:03 | No
2531 | Yes | The cover of the arm of the lead screen fell during an operation. | The cover of the arm of the lead screen fell during an operation. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000700000000October 2000000002000000003 MARC RICAu The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Resoluion: A tightening collar was added on all the arm covers in order to avoid their fall. Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: yes Eupected and ueupected Resuts: A plastic cover of the arm of the lead protection fell. uer Impact: No Patient Impact: no SE u ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Lionel GIROu The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: yes Eupected and ueupected Resuts: A plastic cover of the arm of the lead protection fell. uer Impact: No Patient Impact: no SE u Diagnostic performed by Engineer: A plastic cover of the arm of the lead protection fell. I did not manage to get throuh to u. Diana LENORMAND on the phone, I left a message with my contact details. I do not know if the cover fell on the groud or on someone. I asked Marc Ricau to go on site to pu the covers back in place and to fiu them. Intervention on Monday 10000000023/2000000003 Follow u Requred: Fastening plastic covers of the arm of the lead protection. ----------------------------------------------------------------------------- | 2023-10-19 22:18 | No
2532 | Yes | Not powering up | Not powering up | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000076 Hall Patrick The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Is Caue Known: No Caue: Resoluion Detail: Ordering cutomer replacement device uder contractQty (1) 000000005600000000181San Auutine EmsATTN:Bethany Buler Bethany Buler700000000 S Liberty St San Auutine Tu, 75972-2517 Follow u Requred : Replacement device requsted uder contract Qty (1) 000000005600000000181 San Auutine Ems ATTN:Bethany Buler Bethany Buler 700000000 S Liberty St San Auutine Tu, 75972-2517 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Monitoring Vitals uer Impact: inoperable Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Device will not power u Diagnostic performed by Engineer : Problem Description: Device will not power u Not powering u Form ued: Trouleshooting/Investigation Object: Tempu LS Not Object: N/A Software Version: uu Date and Time of Issu: 10000000019/2000000003 uer Actions: General Where is the Deviation: N/A Trouleshooting Steps Taken: Device will not power u, it is totally inoperable. No logs available. The device was not on a patient. Cutomer jut wants a replacement ASAP uder contract. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000011 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-10-19 22:22 | No
2533 | Yes | Audio-No Sound / Speaker No Sound-Keyrin s/n u0 / Device-Other-HW / Replacement Sent | Audio-No Sound / Speaker No Sound-Keyrin s/n u00 / Device-Other-HW / Replacement Sent |  | 2023-10-19 22:39 | No
2534 | Yes | UNITS NEED uYy BATTERY INSTALLED IN THE V0 | UNITS NEED uYy BATTERY INSTALLED IN THE V0 |  | 2023-10-19 22:43 | No
2535 | Yes | Capnography module is malfunctioning, intermittent curve | Capnography module is malfunuoning, intermittent curve | 000000000000000000000000001O##: 000000000000000000000000001000000002000000001 31-Oct-2000000003 Nuo, Vilas Boas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: SIE Device ue: in ue Eupected and ueupected Resuts: fail co3 reading uer Impact: no impact Patient Impact: no impact SE n\a Diagnostic performed by Engineer: CO2 sensor calibration Operational equpment and complies with manuactuer's specifications. Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000003-11-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-19 23:08 | No
2536 | Yes | uTrunk Cable worn-out - | uTrunk Cable worn-out - | 000000000000000000000000001O##: 000000000000000000000000001000000002533500000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000018 RENZO FORESTIERI The system meets the specification for the performed service and is retuned to ue. Follow u Requred: make nemo withou collection at the same time to: VITO FAZZI HOSP. MuATORE SQuRE 73100000000 LECCE LE DEPARTMENT OF CARDIOLOGY AC. COORDINATOR FRANCESCO GERMINIAL 3280000000086283 The system does not recognize the code. 989800000000160000000000000000 please order throuh the commercial channel Problem Description by Engineer: Cutomer Fuction/Role: Caterina Nicolardi Device ue: defibrillator Eupected and ueupected Resuts: requst for material uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u Germinal requsts spo2 cable and ucable becaue they are 000000000000000000000000001n and not 000000000000000000000000001king, the parts will be sent to the address indicated ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Parts Consued: 1.00000000000000000000000098980000000010000000071, Reuable Adut SpO2 Sensor ----------------------------------------------------------------------------- | 2023-10-19 23:08 | No
2537 | Yes | High flow and machine stopped working properly | High flow and machine stopped working properly | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000088 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient providing HFT Eupected and ueupected Resuts: RT reports that the high flow stopped. they powered the uit off then back no and there still was no flow. uer Impact: swapped vents Patient Impact: vent was swapped with no patient issu SE uknown Diagnostic performed by Engineer : cutomer has ran the uit in High Flow uing same settings as RT, flow of 500000000FIO2 of 700000000, for 1.5 days with no issus. Reviewed the even log and there were no alarms or ueupected issus logged. Advised to perform a ful sucessfu PVT before retuning uit to service. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Sam Hamilton +1970000000098200000000 IP & Issu: 566281800000000igh flow and machine stopped 000000000000000000000000001king properly Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: R000000000000000000000000001created and roued call to RSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-19 23:13 | No
2538 | Yes | Tidal volume suddenly reached 0 | Tidal volume suddenly reached 0 |  | 2023-10-19 23:44 | Unknown
2539 | Yes | Need RTECH / Battery is depleting, Splitted to 00 | Need RTECH / Battery is depleting, Splitted to 000 | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000056 Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient Eupected and ueupected Resuts: RT reports that on 1000000001000000002000000003 this uit shu down on pateint and they coud not silence the alarm becaue the screen was black/blank. uer Impact: swapped vent Patient Impact: vent was swapped with no patient issu SE uknown Diagnostic performed by Engineer : advised that what RT described was a vent inop condition. See DRPTA attached to this 000000000000000000000000001 Log shows uit was operated for 8 hous on 10000000000000000/2000000003 while on battery power and uit produed a "low internal battery" 1200000000alarm. on 1000000001000000002000000003 the uit was powered on at 00000000 am and logged a "runing on battery power" alarm right away then a "low internal battery" alarm duing set u at 00000000. A 2000000002 shudown was logged at00000000 then powered right back on and alarmed for "runing on internal battery" at 00000000. THe log shows setting changes being made util 000000002 then "Low Tidal Volue" alarms util the last entry at 5:00000000. Cutomer verified the oulet that was in ue did NOT have any AC power when tested. Cutomer has connected the vent to AC power and allowed it to fuly charge as indicated by a steady batter LED when off and connected to AC power, and batter voltage is over 16vdc. Advised cutomer to perform PVT battery testing. Advised to dulicate the way the uit was operated by RT and allow it to ru util the low battery alarm, then connect to AC power and verify the vent does transition to AC power with a low battery and allow the battery to charge fuly before placing back into clinical ue. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Arthu Lacy 619-700000000000000006 IP & Issu: 200000000000000003336 / Need RTECH / Battery is depleting Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-20 00:03 | No
2540 | Yes | It locks and when pressing the pedal to help, the u-ray does not work. | It locks and when pressing the pedal to help, the u-ray does not work. | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Equpment Problem It locks and when pressing the pedal to project the u-rays it doesn't 000000000000000000000000001k Hemodynamic room sector Search for Pauo 00 5599991361320000000000 E-mail centralhemodinamica00mail.com Priority stopped uit HOSPITAL SÃO RAFAEL Address Avenida Dorgival Pinheiro de Soua 1178 Imperatriz-MA 65900000000-2700000000BR Repair Action: FSE visit Schedued with client: yes Schedued date/time: 11/00000000 from 12pm Date schedued at the client's requst? no SAP/1EMS priority changed? no Estimated duation: 00000000Is the machine necessary? equpment stopped Smart schedue carried ou? Yes Name of FSE: Leonardo Reis FSE trained according to qulification sheet? Yes T2 Activities:Problem Description: Problem description by engineer : Ru crashes intermittently Trouleshooting Action: What steps are taken duing the system diagnosis : Checked cooling uit, pressue and oil level OK. I ued force 00000000ith a difference of 2 hous and the voltage did not change between the phases: L1-L2: 00000000Vac, L1-L3: 00000000Vac and L2-L3: 00000000Vac. Intermittent defect, MIuwill be requsted for testing, as directed by RADAR uLL# and/or description of uLL: Generator Buy OneEMS internal remarks: Problem description by engineer: Ru locks intermittently What steps are taken duing the system diagnosis: Checked cooling uit, pressue and oil level OK. I ued force 00000000ith a difference of 2 hous and the voltage did not change between the phases: L1-L2: 00000000Vac, L1-L3: 00000000Vac and L2-L3: 00000000Vac. Intermittent defect, MIuwill be requsted for testing, as per RADAR gudance uLL# and/or description of uLL: Generator Buy Repair Action "Internal" : Send PART+FSE+ 00000000ous of machine downtime. System u, treat case as P00000000ecommendations to the FSE ► Measue inpu voltage of the phases at the MIuinpu connector; ► If OK, change MIuand leave uder observation for at least 25 days. Follow the logs before applying the part Parts requred: 1 Part : 1, uu: 00000000210000000019971, Description : Mains Interface uit, Quntity : 1, Priority : 1 Eutra Internal comment field : RSE Marcel Ogassawara SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 12:20000000036 POLARIS (POLARIS) ======================== Leonardo Reis - 000000000000000057500000000000000000000000020000000000000000000000005 ===================== == When arriving at the equpment, it was in normal operating condition. Electrical net000000000000000000000000001k analyzer was installed in the cutomer's electrical panel. The MIuwas replaced sucessfuly. Acqusition tests were sucessfu. Equpment was released for testing on patients. Duing euamination, the equpment failed, even with the new MIu Duing a moment of failue, the electrical net000000000000000000000000001k analyzer measued drops in the phases. It was also identified that the TRAFO 3800000000-00000000V was making a very lou noise. Analysis carried ou by the cutomer's electrician identified that there was a problem with the equpment's main power switch (euternal to the bulding). The cutomer will repair the problem identified in their electrical net000000000000000000000000001k. There is no evidence of failue in the equpment components. MIuwas sucessfuly destroyed. Service was completed in accordance with the u Qulity Policy. Tests and Calibrated Tools in PA Tool. Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Gislaine Silva Last udated by: Gislaine Silva 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Was the equpment defective or failed at the time of installation? uknown SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-20 00:08 | No
2541 | Yes | Dim Display | Dim Display | 000000000000000000000000001O##: 000000000000000000000000001000000002000000000000000000000000-Nov-2000000003 Mack, Isaac The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : I powered on the v600000000and confirmed the uit display was black with no visibility of display screen. Problem Description by Engineer : Cutomer Fuction/Role: Bioengineering/Will Auon & Isaac Mack Device ue: v600000000 Eupected and ueupected Resuts: Replaced PCBA-power mgmt with fco66 and then brightness was depleted and became black uer Impact: None Patient Impact: None SE3.200000000Resoluion : I.M COMPLETED 11/1/2 REPLACED THE RP-PCBA , uer Interface, u 2nd GEN,V600000000AND RP-Cable, uer Intf To Pwr Mgmt, V600000000.THE SYSTEM MEETS THE SPECIFICATION FOR THE PERFORMED SERVICE.uIT SuCESSFuLY PASSED PERFORMANCE AND VERIFACATION.uIT WAS RETuNED TO CuTOMER FOR uE. Parts Consued : 1.00000000000000000000000000000000561528871,RP-PCBA,uer Interface,u 2nd GEN,V600000000 1.00000000000000000000000000000000561511661,RP-Cable, uer Intf To Pwr Mgmt, V600000000Tools & Equpment : TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 STATIC CONTROL DEVICE,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : 10000000019/23 YP 0000000000000000000000000010000000020000000000000000 Details: Dispatched to FSE Isaac. (Parent case) Contact Will Auon Phone (5700000000 570000000000000000 ##william.auon00swhealth.com Location Name Baylor uiversity Medical Center Street Gaston Ave 3500000000 City Dallas State Tu Zip 752000000002000000008 ----------------------------------------------------------------------------- | 2023-10-20 00:43 | No
2542 | Yes | Alarm or Messages/No Sound: Alarm not heard at central station: No Trouble Found/Cau Unknown: Reboot | Alarm or Messages/No Sound: Alarm not heard at central station: No Trouble Found/Cau Unknown: Reboot | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000079 Doulas, Norman The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Problem Description: Suv (GMCPEDSSS00000000 00172.29.33.100000000is displaying a VTACH Alarm bu not produing any Alarm Tone at the Suv Speaker Suv (GMCPEDSSS00000000 00172.29.33.100000000is displaying a VTACH Alarm bu not produing any Alarm Tone at the Suv Speaker Form ued: Trouleshooting/Investigation Object: GMCPEDSSS00000000 Not Object: All other Suvs Software Version: C.00000000.00000000 Date and Time of Issu: 10000000019/2000000003 uer Actions: Checking Suv Where is the Deviation: GMCPEDS Trouleshooting Steps Taken: Cureported that Patient is on a MP300000000s and in Room 296 Cuwas able to find speak for the Suv hidden behind the desk and it appears to be connected Rebooted Suv = and we coud hear the Suv boot u to monitoring and cureported that he can hear the Alarms as well PRS Accout: https://pww.portal.rs.healthcare.u.com/PrsPortal/Defaut.aspu?PrsID=256168 Location Name Geisinger Medical Center Street N Academy Ave 100000000 City Danville State PA Zip 17822-9800000000 Checked logs and foud no uLL that indicated a speaker/alarm issu Cuwill monitor System performance and will call back as needed Was PRS ued: Yes PRS Permission: Ron Boyer https://pww.portal.rs.healthcare.u.com/PrsPortal/Defaut.aspu?PrsID=256168 Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:In Clinical ue Eupected and ueupected Resuts:Monitor Vitals uer Impact:No impact Patient Impact:No impact SE C.00000000.00000000 Problem Description/ueupected Resuts: Suv (GMCPEDSSS00000000 00172.29.33.100000000is displaying a VTACH Alarm bu not produing any Alarm Tone at the Suv Speaker Resoluion : Is Caue Known: Yes Caue: Possible Speaker Connection issu Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Resolved Remotely Resoluion : From: Boyer, Ronald S. <rsboyer100sssoluions.com> Sent: Friday, October 200000000 2000000003 5:00000000 AM To: Doulas, Norman <norman.doulasucom> Suject: RE: [Euternal] Case Follow u Cauion: This e-mail originated from ouside of u, be carefu for phishing. This message was sent secuely uing Ziu® Yes, it has been corrected. Thank you From: Norman Doulas <norman.doulasucom> Sent: Thusday, October 19, 2000000003 7:57 PM To: Boyer, Ronald S. <rsboyer100sssoluions.com> Suject: [Euternal] Case Follow u Euternal: Be carefu with links and files. If supiciou, alert the Information Secuity Office. Qustions? Call 5700000000271-8000000002 Dear Ron Boyer Thank youfor being a valud u services cutomer. I am writing to follow u on the statu of you open case. Case Nuber: 000000002200000000163 Accout Name: Geisinger Health System IP/Site ID: 0000000056000000009000000001 000u62627Z9 Produt Name: PIIC iu ST rp58100000000Dul Disk PC Case Suject Line: Suv (GMCPEDSSS00000000 00172.29.33.100000000is displaying a VTACH Alarm bu not produing any Alarm Tone at the Suv Speaker udates: Please reply to let u know if the issu is corrected. If it is not corrected we want to assist futher. If youprefer to call, youcan reach u at 800000000 722 9377 and provide you case nuber. Sincerely, Norman Doulas CONFIDENTIAL ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002500000000200000000iller, Dell The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: No Caue: Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Resolved Remotely.Issu resolved w/PICIu reboot. Diagnostic performed by Engineer : Problem Description: Cannot hear alarm tones on PICIu. Cureports uable to hear alarms on PICIu. Form ued: Trouleshooting/Investigation Object: Speaker Not Object: Other fuctions Software Version: uu Date and Time of Issu: 10000000019/2000000003 uer Actions: Pt monitoring. Where is the Deviation: Single PICIu. Trouleshooting Steps Taken: No speaker malfuction msgs viewed on PICIu main screen/speaker is pluged in.RSE instruted cuto reboot the PICIu/now hear alarm tones. Was PRS ued: No PRS Permission: Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Shoud hear the alarm tones. uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Cannot hear alarm tones on PICIu. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ron Boyur IP & Issu: 69877728/ alarm not heard at central station Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-10-20 01:03 | No
2543 | Yes | Prouimal pressure sensor autozero failed | Prouimal pressure sensor autozero failed | 000000000000000000000000001O##: 00000000000000000000000000100000000250000000063 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: prouimal pressue censor auo zero failed uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states he is getting prouimal pressue sensor uLL cutomer confirmed in event log. Cutomer foud uLLcode 1100000000: The prouimal pressue is not measued and pressue-related alarms are compromised. Informed cutomer of the following: Manuactuer recommendation is to replace in the following order: 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA. Cutomer states he verified da board placement, replaced Sol 3 and 00000000nd still getting same uLL Informed cutomer to attempt to replace DA board. Cutomer is requsting part nuber and price for replacement DA board. Resoluion : Provided cutomer with part nuber and price: 00000000561512561 - RP-PCBA, Data Acqusition, V600000000- 1286.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Follow u Requred : N/A ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Callback Caller: Mark Hanna IP & Issu: prouimal pressue censor auo zero failed/89962000000007 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created Follow u 000000000000000000000000001& Roued to RTAC Internal Remarks : Interaction Type: New Case Caller: Mark Hanna IP & Issu: prouimal pressue censor auo zero failed/89962000000007 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create case and R000000000000000000000000001----------------------------------------------------------------------------- | 2023-10-20 02:07 | No
2544 | Yes | 00- Speaker issue/malfunction-Other HW/ Replacement sent | 0000- Speaker issue/malfunuon-Other HW/ Replacement sent | 000000000000000000000000001O##: 00000000000000000000000000100000000250000000097 Lefker, Harry The system meets the specification for the performed service and is retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Follow u Requred : VA Medical Center West Haven Street Campbell Ave 9500000000 City West Haven State CT Zip 00000000516-2700000000 Attn Biomed Joe Cianciolo RM 2132 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Problem Description/ueupected Resuts: Remote Speaker is not 000000000000000000000000001king Diagnostic performed by Engineer : Problem Description: Remote Speaker is not 000000000000000000000000001king Remote Speaker is not 000000000000000000000000001king Form ued: Other Trouleshooting Steps Taken: Biomed replaced bad speaker with a spare and is now 000000000000000000000000001king. Biomed needs a replacement ordered uit uder warranty.Order part per biomed Qty 1 0000000056000000000000000081PCACY 25V Remote Speaker w/Volue Control ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002500000000100000000Parts Consued : 1.0000000000000000000000000000000056000000000000000081,PCACY 25V Remote Speaker w/Volue Contro ----------------------------------------------------------------------------- Internal Remarks : Caller: Joe 200000000 927 000000006 IP & Issu: 99136800000000- Speaker does not 000000000000000000000000001k Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A ----------------------------------------------------------------------------- | 2023-10-20 02:33 | No
2545 | Yes | AED initially recognized the replacement pads but then gave ““install pads cartridge cover” uLLmessage | AED initially recognized the replacement pads but then gave ““install pads cartridge cover” uLLmessage | 000000000000000000000000001O##: 00000000000000000000000000100000000251000000005 Apiruvaraku, Ida The system meets the specification for the performed service and is retuned to ue Internal Comments : 2nd attempt - called cut bu uable to leave vm du to cut not having set u a vm bou. ##d cut. Resoluion : Cut responded that the AED initially recognized the replacement pads bu then after 18 hrs later it started talking again with the same uLLmsg. Advised cut to remove device from service and will replace device uder warranty. See new case 000000002250000000000000000 for device replacement Internal Comments : FedEu tracking 7853300000000729300000000 DELIVERED Tusday 1000000002000000003 at 11:00000000 AM Called John as a follow u and he said he will test it later and will give u a cb Problem Description by Engineer : Cutomer Fuction/Role: Owner Device ue: no Eupected and ueupected Resuts: uLLmsg uer Impact: none Patient Impact: none SE E.00000000.000000003 See ref Case 000000002200000000800000000 Cut called stating that the device 000000000000000000000000001ld intermittently be voice activated with instrutions as if it is in an emergency mode. Diagnostic performed by Engineer : Had cut install battery to ru BIT and it started to beep. Had cut hit the blu i-buton and it emitted an uLLmsg “install pads cartridge cover”. Had cut remove and replace pads cover bu it did not resolve the issu. Cut stated that they bouht this device early this year and he takes the device with him in a vehicle as he 000000000000000000000000001ks with St Johns Ambuance. He said the issu started to happen over a week ago. Pads are the original set that were already installed with the AED when puchased. Will replace pads uder warranty to see if this resolves the issu. uable to get pads back for evalution du to its location in Canada. Pads lot# C100000000722-00000000 Eup date: 2000000005-0000000017 Internal Comments : Adut pads Order 6300000000927571 ----------------------------------------------------------------------------- | 2023-10-20 03:13 | No
2546 | Yes | Touchscreen not responsive in areas | Touchscreen not responsive in areas | 000000000000000000000000001O##: 0000000000000000000000000010000000025100000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the touhscreen is not operating good. Calibration did not solve the issu Part #00000000561539000000001, touhscreen, list price $990000000000000000Resoluion : Cutomer will order the touhscreen and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : Caller: Chris 700000000637 3512 IP & Issu: 56629000000005-Issus with touh screen Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-20 03:22 | No
2547 | Yes | HS0AED is emitting the triple chirp (Replacing Device) | HS0 AED is emitting the triple chirp (Repla0g Device) | 000000000000000000000000001O##: 00000000000000000000000000100000000251278 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Tracking#785295650000000016, DELIVERED Friday 10000000020000000023 at 100000000100000000AM, DELIVERED Friday 10000000020000000023 at 100000000100000000AM, ##cutomer to confirmed the AED is 000000000000000000000000001king as it shoud. Cutomer has confirmed that AED is 000000000000000000000000001king as it shoud Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below Duham Fire Department 51 College Road Duham, NH 000000008200000000 Attn: Dave Blatchford Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping and will not pass the self-test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported the AED kept chirping and will not pass the self-test Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and will not pass the self-test. Cutomer removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Remove and reinsert battery to test. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Advise cutomer to remove device ou of service. util she received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# A17F-00000000200000000 Battery Install Before# 2000000009/6/300000000 Lot#00000000000000005P Pads Eup Date: 7/10000000020000000000000000 Pads Lot #Y000000001000000002-00000000 Warranty 6/26/2000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000251300000000 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-10-20 03:43 | No
2548 | Yes | HS0Cuomer called stating device will not turn on. | HS0 Cuomer called stating device will not turn on. | 000000000000000000000000001O##: 00000000000000000000000000100000000251300000000Simpson,Kreshia The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that the AED will not tun on uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that the AED will not tun on Diagnostic performed by Engineer : Cutomer reported that the AED will not tun on. Cutomer stated that the AED is not ued heavy and uaware if the AED was chirping. Requst the cutomer to insert the battery in the AED, did not power u. Replacing battery that is uder warranty. AED Serial# A200000000-00000000832 Battery Install Before# 2000000006/00000000 Lot#95998P Pads Eup Date: 5/17/2000000003 Pads Lot # Y111720000000000000000Resoluion : Closing case, called ##cutomer to confirm AED is 000000000000000000000000001king, cutomer never called or response back Internal Comments : Tracking #53900000000368393 SCHEDuED DELIVERY DATE Monday 10000000030000000023 by end of day Estimated between 1:15 PM - 5:15 PM Internal Comments : DELIVERED Monday 10000000030000000023 at 1:300000000PM Signed for by: LLAMEL ##d cutomer to confirm the AED is 000000000000000000000000001king as it shoud. 2nd attempt Internal Comments : SCM Reguar Order 63000000009258700000000---------------------------------------------------------------------------- | 2023-10-20 03:53 | No
2549 | Yes | PM correction host cold boot issue | PM correuon host cold boot issue | 000000000000000000000000001O##: 000000000000000000000000001000000002000000002000000006-Oct-2000000003 Byron Kirk The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: fse Device ue: ou Eupected and ueupected Resuts: u uer Impact: none Patient Impact: none SE Resoluion : replaced host pc Parts Consued : 1.00000000000000000000000000000000800000000200000000612,Nehalem Host PC Monoplane Rev_III NO_HDD ----------------------------------------------------------------------------- Internal Remarks : 00000000220000000000000000 Byron Kirk FSE calling in on u, FSE says that the Host PC drive bay does not auo start, led does not light u. FSE tried to reseat and no luk. Host will not start u on boot u. Gave FSE ok to order new host PC. Internal Remarks : Planner Initials: GM Action Taken: FSE Kirk Byron requsted call open and dispatched for 100000000200000000 Oucome: dispatched ----------------------------------------------------------------------------- | 2023-10-20 04:13 | No
2550 | Yes | EQUIPMENT HAS TO BE TURNED ON SEVERAL TIMES SINCE ONLY HALF OF THE MONITOR GOES UP | EQUIPMENT HAS TO BE TURNED ON SEVERAL TIMES SINCE ONLY HALF OF THE MONITOR GOES UP | 000000000000000000000000001O##: 000000000000000000000000001000000002966100000000ovember 100000000 2000000003 Muoz, Aleuander The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: EQuPMENT HAS TO BE TuNED ON SEVERAL TIMES SINCE ONLY HALF OF THE MONITOR GOES u When was the first occurence? Within... : October 18th-2000000003 18:15:00000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: EQuPMENT HAS TO BE TuNED ON SEVERAL TIMES SINCE ONLY HALF OF THE MONITOR GOES u. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. uu00000000Resoluion: Check monitor connectors and observe a false contact. The adapters are disconnected. They are cleaned and then connected again. The equpment tuns on again and we observe that the monitor tuns on on both sides. The equpment is fuctional. We recommend to follow u the temperatue and huidity in the room and technical room. Resoluion: Auo-verification of the system ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000100000000November 100000000 2000000003 Castillo, Moises The system meets the specification for the performed service and is retuned to ue. Resoluion: We provide suport to the field service eng. we proceed to check the monitor connectors. We observe a false contact. The adapters are disconnected. Then they are cleaned and then connected again. The equpment tuns on again and we observe that the monitor tuns on on its 2 sides. Then we proceed to install field eutensions u7. We check that they have been installed correctly. Observed that they appear in the application, proceeded to perform tests, they show positive resuts, left the equpment fuctional. Problem Description by Engineer: Cutomer Fuction/Role: BIOMEDICINE Device ue: NO Eupected and ueupected Resuts: N/A uer Impact: NO Patient Impact: NO SE 2.2 L3 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Koc, Migul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: EQuPMENT HAS TO BE TuNED ON SEVERAL TIMES SINCE ONLY HALF OF THE MONITOR GOES u When was the first occurence? Within... : October 18th-2000000003 18:15:00000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: EQuPMENT HAS TO BE TuNED ON SEVERAL TIMES SINCE ONLY HALF OF THE MONITOR GOES u. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. uu00000000Diagnostic performed by Engineer: They have to restart the equpment continully to finish tuning it on, becaue half of the screen tuns on, checking the uLLlog file, shows a FleuVision PC failue. A field service engineer is needed to clean internally, ru HD and PC tests, replace battery if necessary and test operation, if the faut persists, format disks, replace gravers, reinstall software or evalute replacing HD or PCs. Internal Comments: uLL# and/or description of uLL EQuPMENT HAS TO BE TuNED ON SEVERAL TIMES SINCE ONLY HALF OF THE MONITOR GOES u Repair Action: They have to restart the equpment continully to finish tuning it on, becaue half of the screen tuns on, checking the uLLlog file, it shows a FleuVision PC failue. A field service engineer is needed to clean internally, ru HD and PC tests, replace battery if necessary and test operation, if the faut persists, format disks, replace gravers, reinstall software or evalute replacing HD or PCs. Follow u Requred: uLL# and/or description of uLL EQuPMENT HAS TO BE TuNED ON SEVERAL TIMES SINCE ONLY HALF OF THE MONITOR GOES u Repair Action: They have to restart the equpment continully to finish tuning it on, becaue half of the screen tuns on, checking the uLLlog file, it shows a FleuVision PC failue. A field service engineer is needed to clean internally, ru HD and PC tests, replace battery if necessary and test operation, if the faut persists, format disks, replace gravers, reinstall software or evalute replacing HD or PCs. What are the neut Trouleshooting steps: A field service engineer is needed to clean internally, ru HD and PC tests, replace battery if necessary and test operation, if the faut persists, format disks, replace gravers, reinstall software or evalute replacing HD or PCs. Technician Name: Koc, Migul ----------------------------------------------------------------------------- | 2023-10-20 05:17 | No
2551 | Yes | 0/UNIQ FD0Space low | 0R/UNIQ FD0Space low | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 2000000001000000002000000003 11:20000000012 uC (86000000000000000091) Service content (requred): There is not enouh space for the image.Replace IPPC. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None Euternal Remarks: 27.1000000002000000003 00000000:50000000011 uC (86000000000000000091) Service content (requred): There is not enouh space for the image.Replace the IPPC and the 3 hard drives inside. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None Problem Reported by Cutomer: 00000000R/uIQ FD200000000 Neut Step Action: PPD0000000080000000050000000092*1, no need to order, please send 000000000000000000000000001k FSE: Dai Chaojiang-86-1881500000000257 Neut Step Action: NTD: The cutomer wants engineers to be present. Refue to transfer. The disk is ful and uLLis reported. YNN Internal Remarks: 2000000001000000002000000003 11:200000000100000000TC (86000000000000000091) RDI Accuacy: YES $Q7: comfirmed the reported problem, space is low, system can't euposue. $Q8: check the system find there is only one patient record, system still report space is low. $Q9: replace the IPPC uu 0000000080000000012200000000 $Q100000000 (Test and Inspection data refer to:PA.Device retuned to ful fuctionality.) $Q11: Calibrated Tools refer to:PA SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 27.1000000002000000003 00000000:500000000100000000TC (86000000000000000091)RDI Accuacy: YES $Q7: comfirmed the reported problem, space is low, system can't euposue. $Q8: check the system find there is only one patient record, system still report space is low. $Q9: replace the IPPC and OS disk, image disk uu 0000000080000000012200000000, uu0000000080000000050000000000000000 $Q100000000 (Test and Inspection data refer to:PA.Device retuned to ful fuctionality.) $Q11: Calibrated Tools refer to:PA T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the image storage space is low and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Free disk space is too low. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Euecue Database Consistency Test and Repair, delete Repaired patients, and see if the space is restored. 2. If there is no effect, euecue Recreate Image Store. 3. If there is no effect, try Bypass hard disk, and then reinstall the IPPC software. 00000000If there is no effect, ue PC. Diagnostic Tool performs diagnostic tests and replaces corresponding spare parts 5. Check machine uLLreports and perform faut repairs Recommended spare part uuand name: " 0000000080000000050000000092 IP PC Revised_II no HDD &0000000080000000050000000000000000Hard disk spare part " Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Case: 000000002200000000931 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Repair Action: 0000000080000000012200000000 u Haswell IP PC no HDD ----------------------------------------------------------------------------- | 2023-10-20 07:38 | No
2552 | Yes | 0/u00Poweron failed | 0K/uPoweron failed | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine coud not be tuned on and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine cannot be tuned on Remote diagnosis and test resuts: Cabinet M is not powered on PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the FanTray and DCPS boards are normal. If there is any problem, replace it. 2. Check whether the power-on control board is normal. 3. Check whether the insuance is normal. 00000000Check the machine uLLand perform faut repairs. Recommended spare parts uuand names: 0000000080000000000000000900000000 PDuFanTray and DCPS &uu 0000000080000000000000000900000000, Description: PDuDC Modue&uu 0000000080000000000000000900000000, Description: PDuFues Oupus&uu 0000000080000000000000000900000000, Description: PDuMains Interface Modue Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000955 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 2000000001000000002000000003 13:35:500000000uC (86000000000000000013)"SuT. Ready for Clinical ue local time" : " 2000000003-10000000020000000010000000000000000 Service content: uable to power on, PDuFan Tray and DCPS damaged, replacement parts Service resuts: After repair, the equpment retuned to normal operation SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000K/u7 M200000000 Neut Step Action: PPD spare parts 0000000080000000000000000900000000*1 Arrange 200000000to the hospital, Shangzheng-86-13669196600000000Internal Remarks: 2000000001000000002000000003 13:36:00000000 uC (86000000000000000013) $Q7:FSE confirmed reported problem Poweron failed $Q8:FSE checked the PDuFan Tray and DCPS of power suply; foud it is 2300000000 AC inpu, bu there is no oupu; juged it was broken $Q9: Replacement the PDuFan Tray and DCPS $Q100000000System tun onis ok. Test and inspection data refer to PA, device retuned to ful fuctionality $Q11: Calibrated Tools name: N/A SAP Transfer: Case Transferred to SAP Neut Step Action: 00000000152310000000067 kouteacher’s landline phone call ----------------------------------------------------------------------------- | 2023-10-20 07:48 | No
2553 | Yes | M0A / u0 / uALS monitor / Alarm: Paddle malfunction. | M00A / u00 / HeartStart u ALS monitor / Alarm: Paddle malfunuon. | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ##d ASP humei for follow u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: Huntong 15600000000317691 Problem Reported by Cutomer: ECR:： 00000000/HeartStart uuuALS monitor/ Safety Qustions: Case: 000000002200000000979 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-20 07:53 | No
2554 | Yes | 0 - u0- Intermittent blacksreen (Repair quote is rejected by cuomer) | 00- u00 - Intermittent blacksreen (Repair quote is rejeud by cuomer) | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022000000009800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000/HeartStart uL+ Defibrillator/Monitor/ ----------------------------------------------------------------------------- | 2023-10-20 07:53 | No
2555 | Yes | One side with 0onitors is not showing live flouro | One side with 0monitors is not showing live flouro | Internal Remarks : Interaction Type: New Case Caller: Garrett IP & Issu: 800000000761700000000ne side with 2 monitors is not showing live flouo Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to quu. Adv cutomer of callback neut buiness day Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-20 08:23 | No
2556 | Yes | 0/u FD0 Euposure not possible. Image disk full | 0R/u FD0 Euposure not possible. Image disk full | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 25.1000000002000000003 100000000:00000000uC (86000000006572)on job training Euternal Remarks: 23.1000000002000000003 0000000012:25 uC (86000000006572) On job training Euternal Remarks: 29.1000000002000000003 00000000:10000000000000000TC (86000000000000000091) Service content (requred): There is not enouh space for the image.Replacing IPPC Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None T2 Activities:Problem Description: Overview of cutomer problems: Euposue not possible. Image disk ful Detailed description of cutomer problem or faut phenomenon: Small image space Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The cutomer requres the engineer to be present Image Processor 18-Oct-2000000003 00000000:20000000031.0000000000000000000000000000000000000000000000003 POST "Frontal IP-PC" Not Done [TEST_RESuT] FALSE [TEST_NAME] Frontal IP-PC POST [TEST_REPORT] Image Processor 18-Oct-2000000003 00000000:20000000031.0000000000000000000000000000000000000000000000003 POST "Frontal IP-PC OS disk 1" Not Done [TEST_RESuT] FALSE [TEST_NAME] Frontal IP-PC OS disk 1 POST [TEST_REPORT] Image Processor 18-Oct-2000000003 00000000:20000000031.0000000000000000000000000000000000000000000000003 POST "Frontal IP-PC Image disk 1" Not Done [TEST_RESuT] FALSE [TEST_NAME] Frontal IP-PC Image disk 1 POST [TEST_REPORT] Image Processor 18-Oct-2000000003 00000000:20000000032.0000000000000000000000000000000000000000000000003 POST "Frontal IP-PC Image disk 2" Not Done [TEST_RESuT] FALSE [TEST_NAME] Frontal IP-PC Image disk 2 POST [TEST_REPORT] Image Processor 18-Oct-2000000003 00000000:20000000032.0000000000000000000000000000000000000000000000003 POST "Frontal IP-PC FRuFan 1" Not Done [TEST_RESuT] FALSE [TEST_NAME] Frontal IP-PC FRuFan 1 POST [TEST_REPORT] Image Processor 18-Oct-2000000003 00000000:20000000032.0000000000000000000000000000000000000000000000003 POST "Frontal IP-PC FRuFan 2" Not Done [TEST_RESuT] FALSE [TEST_NAME] Frontal IP-PC FRuFan 2 POST [TEST_REPORT] Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: 1. It is recommended to replace the IPPC bios battery 2. bypass hard drive enclosue 3. Reinstall IPPC software 00000000ecreate Notes on cutomer requrements: Please contact the cutomer for on-site repair Internal Remarks: 29.1000000002000000003 00000000:1000000007 uC (86000000000000000091)RDI Accuacy: YES $Q7: confirmed the reported problem, space is low. $Q8: check the system find there is only a few patients, system still report space is low. $Q9: replace the IPPC uu0000000080000000050000000092 $Q100000000 (Test and Inspection data refer to:PA.Device retuned to ful fuctionality.) $Q11: Calibrated Tools refer to:PA Neut Step Action: FSE Chen Chuu 1881727300000000 Call and add OJT Internal Remarks: 23.1000000002000000003 0000000012:27 uC (86000000006572) On job training Safety Qustions: Case: 0000000022000000000000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 25.1000000002000000003 100000000:00000000uC (86000000006572)on job training SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/u uuuFD200000000 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-20 08:28 | No
2557 | Yes | Loudspeaker malfunction - no sound - | Loudspeaker malfunuon - no sound - | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000076 Problem Description by Engineer: No soud Diagnostic performed by Engineer: the loudspeaker is breakdown Resoluion: the speaker changed, ----------------------------------------------------------------------------- | 2023-10-20 09:03 | No
2558 | Yes | QA 000000 Device cannot be ud normally during u | QA 00000000 - Device cannot be ud normally during u |  | 2023-10-20 09:19 | No
2559 | Yes | Low Leak-CO0ebreathing Risk, GFE00OV0/IS complete*** | Low Leak-CO0Rebreathing Risk, GFE0 0NOV00IS complete*** | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Sensor is broken The First Affiliated Hospital of GungzhouMedical uiversity FSE Notes: This PO is a Reference Nuber only SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 00000000:57:00000000 uC (86000000000000000000000000 *Department information [Department: The First Affiliated Hospital of GungzhouMedical uiversity Contact person: Pang Gong Tel: 19925759315]* $ASP engineer:[weigung zhangweihu] $Q1:NO. $Q2:NO. $Q3:YES. $Q00000000ES. $Q7: Confirmed reported problem. Low Leak-CO2 Rebreathing Risk. $Q8:,Enter the diagnostic mode, the air delivery flow accuacy test verifies the accuacy of the air flow sensor and fuction of the blower. $Q9: Replace Flow Sensors. $Q100000000 Test Pressue .Test and inspection data refer to : PA Checklist $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000000000007) Neut Calibration Date (00000000/22/20000000000000000 TSI SN (000000008100000000000000009) Neut Calibration date (00000000/22/20000000000000000Internal Remarks: 00000000.11.2000000003 00000000:57:00000000 uC (86000000000000000000000000 $ASP engineer:[weigung zhangweihu] $Q1:NO. $Q2:NO. $Q3:YES. $Q00000000ES. $Q7: Confirmed reported problem. Low Leak-CO2 Rebreathing Risk. $Q8:,Enter the diagnostic mode, the air delivery flow accuacy test verifies the accuacy of the air flow sensor and fuction of the blower. $Q9: Replace Flow Sensors. $Q100000000 Test Pressue .Test and inspection data refer to : PA Checklist $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000000000007) Neut Calibration Date (00000000/22/20000000000000000 TSI SN (000000008100000000000000009) Neut Calibration date (00000000/22/20000000000000000Safety Qustions: Case: 000000002200000000283 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Sensor failue The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: It is recommended to check the sensor Recommended spare part uufor T2 with description: #989800000000617100000000 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA ----------------------------------------------------------------------------- | 2023-10-20 09:23 | No
2560 | No | Appear alarm CO0ervice Require with OP check failed. (Parts are replaced). | Appear alarm CO0Service Require with OP check failed. (Parts are replaced). |  | 2023-10-20 09:35 | No
2561 | Yes | SA：0u FD0System cannot boot up | SA：0/u FD0System cannot boot up | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. FSE Notes: This PO is a Reference Nuber only SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Power failue T2 Activities:Problem Description: Overview of cutomer problems: When I tuned off the power switch of the 3D display, sparks appeared and the device ueupectedly lost power. When I restarted it, the device stayed on the coutdown interface and coud not enter the system.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: When I tuned off the power switch of the 3D display, sparks appeared and the device ueupectedly lost power. When I restarted it, the device stayed on the coutdown interface and coud not enter the system.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: When I tuned off the power switch of the 3D display, sparks appeared and the device ueupectedly lost power. When I restarted it, the device stayed on the coutdown interface and coud not enter the system.Engineers are requred to be present Notes on cutomer requrements: Onsite Safety Qustions: Case: 000000002200000000333 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 23.1000000002000000003 10000000023:00000000 uC (8600000000800000000) Service Content: The equpment coud not be tuned on. When I went to the site for inspection, I foud that the fue was buned ou. I replaced the fue and the equpment retuned to normal. Service resuts: After repair, the equpment is normal. Others: none Neut Step Action: SS Zhang Yan has been ##d to follow u, waiting for SA leader’s reply whether to go throuh the green channel Internal Remarks: 23.1000000002000000003 10000000023:11 uC (8600000000800000000) $Q7: (Confirmed reported problem, cannotbootu ) $Q8: (check the uLLlog, no uLLfue defected) $Q9: (replace fue .) $Q100000000 (Test and Inspection data refer to PA. Device retuned to ful fuctional) $Q11: Calibrated Tools ( NA ) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: na ^Curent Tue uage: na SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-20 09:43 | No
2562 | Yes | Can’t enable u-ray | Can’t enable u-ray | 000000000000000000000000001O##: 000000000000000000000000001000000002000000003000000001-Oct-2000000003 KITTI SuAN The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Not provide Device ue: CV Eupected and ueupected Resuts: Can’t enable u-ray. uer Impact: Not provide Patient Impact: Not provide SE 8.1.17 Parts Consued : 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD Resoluion : Replaced IPPC. Performed download board software IPPC lateral. ----------------------------------------------------------------------------- | 2023-10-20 12:12 | No
2563 | Yes | The touchscreen was malfunctioning. | The touchscreen was malfunuoning. | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000862 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Gao Lui Last udated by: Gao Lui 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Bu PARTS Neut Step Action: ASP.Qiangyi.Shi Changsen13521675291 call ----------------------------------------------------------------------------- | 2023-10-20 12:13 | No
2564 | Yes | System does not start up | System does not start up | 000000000000000000000000001O##: 0000000000000000000000000010000000025000000007 Oct 300000000 2000000003 RONALD BOK The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: system down Eupected and ueupected Resuts: Does not start u, ending at 00000000:00000000 uer Impact: n/a Patient Impact: n/a SE R8.2.300000000Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Resoluion: new PC, however, had to re-install software becaue of Windows 7 disk encrepition ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000252900000000 Oct 200000000 2000000003 RONALD BOK The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: system down Eupected and ueupected Resuts: Does not start u, ending at 00000000:00000000 uer Impact: n/a Patient Impact: n/a SE R8.2.300000000Diagnostic performed by Engineer: hard disk connected directly to mother board received auhorization from Ton Sanders. System is 000000000000000000000000001king again. PC does need to be replaced Follow u Requred: fleuvision PC needs to be replaced 0000000080000000058600000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000252865 ROB VAN GAMEREN The system does not meet the specifications for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: system down Eupected and ueupected Resuts: Does not start u, ending at 00000000:00000000 uer Impact: n/a Patient Impact: n/a SE R8.2.300000000Diagnostic performed by Engineer: Ronald Bok will do a diagnosis on site, no log file available. Follow u Requred: System does not start u, ending at 00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-20 12:18 | No
2565 | Yes | Does not turn on | Does not turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000253000000003 October 200000000 2000000003 Christophe LAuENT The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Problem at start-u, the screens tun on, only the u logo is displayed. uer Impact: No Patient Impact: no SE u Diagnostic performed by Engineer: Assistance for corrective action. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000253369 October 200000000 2000000003 uavier NEYRET The system meets the specifications for the performed service and is retuned to ue. Resoluion: Replacement of the internal battery on PC HOST. Reparametred the BIOS. Several stop/000000000000000000000000001k tests of the u. The room is fuly operational again. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: u Eupected and ueupected Resuts: power on system uer Impact: No. Patient Impact: No. SE n/a Diagnostic performed by Engineer: In the CORO room, impossible to start the room, the screens remain black. When forcing the pc host to start, we can see that it is no longer at the correct date. The internal battery of the PC is ou of service. Tools and Equpment: MuTIMETER, 2000000005-00000000-16 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000252700000000Lionel GIROu The system does not meet the specifications for the performed service and is not retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Problem at start-u, the screens tun on, only the u logo is displayed. uer Impact: No Patient Impact: no SE u Diagnostic performed by Engineer: Problem at start-u, the screens tun on, only the u logo is displayed. No remote service feedback. The cutomer shu off the power suply and restarted it, still the problem. Requres on-site intervention for start-u diagnosis. Follow u Requred: Diagnosis at start-u. ----------------------------------------------------------------------------- | 2023-10-20 12:33 | No
2566 | Yes | Fluoro storage not possible due to an Image disk problem | Fluoro storage not possible due to an Image disk problem | 000000000000000000000000001O##: 00000000000000000000000000100000000252900000000300000000Oct-2000000003 Ashley Goodwin The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: rad Device ue: in ue Eupected and ueupected Resuts: fauty image disk uer Impact: na Patient Impact: na SE000Resoluion : Please see case 000000002200000000311 for faut reporting. A follow u activity will be raised against case 000000002200000000311. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000252933 Ahmed Marzou The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: rad Device ue: in ue Eupected and ueupected Resuts: fauty image disk uer Impact: na Patient Impact: na SE000Diagnostic performed by Engineer : Logfiles checked , i foud image disk uLL, FSE to site to replace it . Follow u Requred : Logfiles checked , i foud image disk uLL, FSE to site to replace it . ----------------------------------------------------------------------------- Cutomer comment : tried calling no answer. TSS advised to plan 000000002200000000311 alongside this case and only order hard drives on one case Internal Remarks : ##d SDM, as on this system we are waiting for TSS advice before planning 000000002200000000311. Checking if this case needs planning asap or will be linked with 000000002200000000311 ----------------------------------------------------------------------------- | 2023-10-20 12:38 | No
2567 | Yes | Speaker malfunction - (GFE0SENT Dec0 F/U= Dec0 | Speaker malfunuon - (GFE0 SENT Dec0) F/U= Dec0) |  | 2023-10-20 14:23 | No
2568 | Yes | It doesn't work. | It doesn't work. | 000000000000000000000000001O##: 0000000000000000000000000010000000028600000000 00000000-Nov-2000000003 Magdy Maarou Khalil The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Doctor Device ue: No Eupected and ueupected Resuts: It does not start. uer Impact: No Patient Impact: No SE Resoluion: I have changed the PC, installed the SOFTWARE, Now everything 000000000000000000000000001ks well. Parts Consued: 1.00000000000000000000000000000000800000000122391, u Haswell Monoplane Host PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000025000000003 October 26, 2000000003 Magdy Maarou Khalil The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Install new HOST PC and SW. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Doctor Device ue: No Eupected and ueupected Resuts: It does not start. uer Impact: No Patient Impact: No SE Diagnostic performed by Engineer: The new PC did not 000000000000000000000000001k, I left the PC 000000000000000000000000001king temporarily, the client accepted the temporary soluion util we received the new PC. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000253612 October 200000000 2000000003 Magdy Maarou Khalil The system does not meet the specification for the performed service and is not retuned to ue. Diagnostic performed by Engineer: Host PC does not start, I have ordered it and it will arrive when possible. Problem Description by Engineer: Problem description by engineer: The equpment does not do the start-u sequnce, nor does it show anything on the screen, althouh the euternal video signals can be seen on the FleuVision monitor. The Host PC seems to be failing, need to go and check it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment cannot be ued. Patient Impact: No futher information is available as per the cutomer. SE 8.1.3000000005 Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000253262 Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not do the start-u sequnce, nor does it show anything on the screen, althouh the euternal video signals can be seen on the FleuVision monitor. The Host PC seems to be failing, need to go and check it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment cannot be ued. Patient Impact: No futher information is available as per the cutomer. SE 8.1.3000000005 Diagnostic performed by Engineer: The equpment does not do the start-u sequnce, nor does it show anything on the screen, althouh the euternal video signals can be seen on the FleuVision monitor. The Host PC seems to be failing, need to go and check it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The equpment does not do the start-u sequnce, nor does it show anything on the screen, althouh the euternal video signals can be seen on the FleuVision monitor. The Host PC seems to be failing, need to go and check it. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: The equpment does not do the start-u sequnce, nor does it show anything on the screen, althouh the euternal video signals can be seen on the FleuVision monitor. The Host PC seems to be failing, need to go and check it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-10-20 14:38 | No
2569 | Yes | 0/uDFM0 Defibrillator/Monitor/ufailure (Replaced 0eadset, Snap, IEC, ICU) | 0N/Effi0 DFM0 Defibrillator/Monitor/ufailure (Replaced 0Leadset, Snap, IEC, ICU) | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022000000005900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: The newly installed lead wire operation check failed Problem Reported by Cutomer: Yanling Town Health Center, Danyang City There is a problem with the ulead wires T2 Activities:Repair Action: Service content: If the uoperation check fails, replace the lead wire throuh spare parts eupress service. SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-20 14:43 | No
2570 | Yes | The device does not start | The device does not start | 000000000000000000000000001O##: 00000000000000000000000000100000000253958 00000000-Nov-2000000003 000000000000000000000000001ciech Filipowicz The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: diagnostics Eupected and ueupected Resuts: a malfuction of Sute PC uer Impact: no Patient Impact: no uu Resoluion: List of replacement parts ued for repair. Parts Consued: 1.00000000000000000000000000000000800000000660000000083,Sute PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000000000002-Nov-2000000003 000000000000000000000000001ciech Filipowicz The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: diagnostics Eupected and ueupected Resuts: a damage of the compuer Sute PC uer Impact: no Patient Impact: no uu Resoluion: Replacement of the Sute PC compuer. A reinstallation of the software. Repair tests. The defect was eliminated. Parts Consued: 1.00000000000000000000000000000000800000000660000000083,Sute PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000253112 200000000Oct-2000000003 Kamil Tarachanowicz The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Electrophysiology Eupected and ueupected Resuts: The system does not switch on. uer Impact: no Patient Impact: no SE 1.2.5. Diagnostic performed by Engineer: Diagnostics of the system was performed as a resut of a faut consisting in the system not being able to start. A damage of the PC Sute compuer was foud. The compuer needs to be replaced. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000021 19-Oct-2000000003 000000000000000000000000001ciech Filipowicz The system does not meet the specification for the performed service and is not retuned to ue. Resoluion: The system was started in the changed hardware configuation. Changes were made only in order to verify the operation of other components. Du to a high risk of a futher damage of the Sute PC compuer, at an impossible to predict moment, the system is deemed fauty. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: diagnostics Eupected and ueupected Resuts: The system does not start uer Impact: It is not possible to perform Patient Impact euaminations: no uu Diagnostic performed by Engineer: Duing diagnostics it was confirmed that starting the system is not possible. An attempt to fuly reinstall the system with the application was performed with a negative resut. Futher diagnostics revealed that the HDD is not 000000000000000000000000001king correctly du to a damage within the Sute PC compuer. ----------------------------------------------------------------------------- | 2023-10-20 14:48 | No
2571 | Yes | Hardware Failure (DPM | Hardware Failure (DPM | 000000000000000000000000001O##: 0000000000000000000000000010000000025300000000 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : End uer to be provided with service euchange device 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Lead Nuse Device ue: not in ue Eupected and ueupected Resuts: Hardward (DPM Failue) INOP displayed on LS uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : End uer to be provided with service euchange device 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000253812 Parts Consued : 1.000000000000000000000000000000005600000000000000001,Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- | 2023-10-20 15:12 | No
2572 | Yes | Display/Unable to operate: Screen freezes when u0s connected to Mu0: Insufficient Info/Unknown: Provided Information | Display/Unable to operate: Screen freezes when u0is conneud to Mu00 Insuffi0nt Info/Unknown: Provided Information | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000032 James Gillham The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In ue Eupected and ueupected Resuts: Intermittent issu uer Impact: None Patient Impact: None SE NA Diagnostic performed by Engineer : Intermittent issu with an u3 attached to the Mu00000000 monitor. When the u00000000s attached to the Mu00000000 the screen will freeze and they have to restart the monitor. The u3 is on software rev P00000000.00000000- 2 and the Mu00000000 is on rev M00000000.00000000. Requsted the cutomer to ugrade the Mu00000000 to the most latest version on Rev M to see if that fiues the issu. Not had a reply from the cutomer so not sue if this resolved the issu. Going to close the case and if the problem persists another case will be raised. ----------------------------------------------------------------------------- | 2023-10-20 15:18 | No
2573 | Yes | The wireless footswitch is still not operational - Medimover Storage, Burrfields Road, P00 | The wireless footswitch is still not operational - Medimover Storage, Burrfields Road, P0 0N | 000000000000000000000000001O##: 00000000000000000000000000100000000253863 David Couens The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Follow u Requred : FSE to site to replace WFS and base station and pair them Diagnostic performed by Engineer : This footswitch has had the FCO udate for the WFS and the Base station. There has been issus with this FCO so I will order a new WFS and BS Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: The wireless footswitch is still not operational - Medimover Storage, Burfields Road, P00000000 5NN uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- Internal Remarks : Hi Lauence, Apologies for the late reply, youcan cancel the parts order if youhaven’t already. uit is leaving u soon. Best Regards, Mindi ----------------------------------------------------------------------------- | 2023-10-20 15:28 | No
2574 | Yes | Hardware Failure (DPM) | Hardware Failure (DPM) | 000000000000000000000000001O##: 0000000000000000000000000010000000025300000000Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Resoluion : End uer to be provided with service euchange device 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN Problem Description by Engineer : Cutomer Fuction/Role: Lead Nuse Device ue: not in ue Eupected and ueupected Resuts: Hardward (DPM Failue) INOP displayed on LS uer Impact: n/a Patient Impact: n/a SE n/a Follow u Requred : End uer to be provided with service euchange device 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN Reguatory Qustions : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000253981 Parts Consued : 1.000000000000000000000000000000005600000000000000001,Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- | 2023-10-20 15:38 | No
2575 | Yes | Device damaged following a shock. Bench Repair Pending as of 00 | Device damaged following a shock. Bench Repair Pending as of 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002600000000he system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Raffenne, virginia The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: total repair qute requst uer Impact: No Patient Impact: No SEuknown Follow u Requred: The device mut have fallen. The rear houing is damaged. Some screws have been lost. The power suply connector have been damaged as well. Need for a complete diagnosis and qute for repair and tests. CSR Virginie Raffenne PICK u GENERAL STORE 8 a.m. 12 p.m. Resoluion: Requst for qute for 000000000000000000000000001kshop repair. Sending the label for complete repair diagnosis and testing. Diagnostic performed by Engineer: The device has been dropped and is very damaged. The rear hul is damaged, dislocated, the screws have broken. The sector base is also damaged. ----------------------------------------------------------------------------- Internal Remarks: • FEDEu confirmation nuber: LYSA3000000005 • Tracking nuber 7738000000007395800000000 Picku time: 8:00000000 - 1:00000000 p.m. • Carrier delivery date: 10000000023 at general store ----------------------------------------------------------------------------- | 2023-10-20 15:53 | No
2576 | Yes | 0 / u0/ Device defective: Battery is not recognized. | 00/ u000/ Device defeu: Battery is not recognized. |  | 2023-10-20 16:08 | No
2577 | Yes | uLL database space is full. | uLL database space is full. | 000000000000000000000000001O##: 00000000000000000000000000100000000510000000000000000ESTOLIN, DARIO The system meets the specification for the performed service and is retuned to ue Internal Comments: uLL# and/or description of uLL: . Repair Action: . Resoluion: Tu 28-Nov-23 - 10000000000000000 AM - contacted FSE Schillacci and susequntly also uRandisi - they confirm that they have no details of the report, probably after a reboot the system resued reguar operation. Curently the system is 000000000000000000000000001king normally and the case can be closed Problem Description by Engineer: Problem description by engineer: No problem highlighted Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer: Mon 27-Nov-23 - 15:00000000PM - Attempted 00contact with the cutomer several times at the nubers indicated - bu no response. 00contact made with u Randisi - he has no information on the matter. The call was opened on 19 October, bu du to the absence of the contract, it was blocked awaiting renewal. As agreed with the cutomer, we proceed by monitoring the system util tomorrow ----------------------------------------------------------------------------- Cutomer comment: Other Technical intervention is requsted. as per agreements with Ing.Romani, a qute is accepted pending the signing of the contract. ----------------------------------------------------------------------------- | 2023-10-20 16:23 | No
2578 | Yes | 0 - u0- urhythm is not established. | 00- u00 - urhythm is not established. | File uloaded by CS Portal uer : 파일명: 당시상황사진.png. 댓글: File(s) uloaded duing case creation Cutomer comment : 000000003000000000000000000000000LQ26ZQAT ----------------------------------------------------------------------------- | 2023-10-20 16:27 | No
2579 | Yes | Loudspeaker uLLSound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker uLLSound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000253276 Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: Louspeaker uLL Patient Impact: - SE - Follow-u requred: NEMO 0000000056000000003831 Mu_100000000/u3 Speaker Assembly Diagnostic performed by Engineer: NEMO ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000250000000000000000 Parts Consued: 1.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- File uloaded by CS Portal uer: File Name: Laus.docu. Comments: File(s) uloaded duing process creation ----------------------------------------------------------------------------- | 2023-10-20 16:28 | No
2580 | Yes | ulead malfunction (Replaced DFM0 Measurement Module u | ulead malfunuon (Replaced DFM0 Measurement Module u | 000000000000000000000000001O##: 0000000000000000000000000010000000025300000000 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Follow u Requred : To replace 0000000056000000009131, Measuement Modue uonly Diagnostic performed by Engineer : To replace 0000000056000000009131, Measuement Modue uonly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000025000000001 Parts Consued : 1.0000000000000000000000000000000056000000009131,DFM100000000 Measuement Modue u----------------------------------------------------------------------------- | 2023-10-20 16:38 | No
2581 | Yes | Menu button broken. | Menu button broken. | 000000000000000000000000001O##: 000000000000000000000000001000000002500000000 November 13, 2000000003 Ruén Álvarez Villa The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Diagnostic performed by Engineer: The menubuton is broken. Since there is no availability of spare parts becaue the equpment is old, it cannot be repaired, its ue is not recommended and an obsolescence letter is sent. Problem Description by Engineer: Cutomer Fuction/Role: We talk to Mario Balarezo Espín 6000000007700000000 from Electromedicine. Device ue: Not in clinical ue. Eupected and ueupected Resuts: Menubuton is broken. uer Impact: None Patient Impact: None SE Not apply. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000025300000000 Rau Cordero Calderón The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We talk to Mario Balarezo Espín 6000000007700000000 from Electromedicine. Device ue: Not in clinical ue. Eupected and ueupected Resuts: Menubuton is broken. uer Impact: None Patient Impact: None SE Not apply. Diagnostic performed by Engineer: We talk to Mario Balarezo Espín 6000000007700000000 electromedicine and he says that the menubuton is broken. 8612900000000HeartStart uL+ Defibrillator/Monitor obsolete since December 31, 2000000002 Assign fSE. There is no material 0000000056000000006161 - uL+ Kit-Front Case Assy uGeneral Emergency Room, Critical Room B Follow u Requred: We talk to Mario Balarezo Espín 6000000007700000000 electromedicine and he says that the menubuton is broken. 8612900000000HeartStart uL+ Defibrillator/Monitor obsolete since December 31, 2000000002 Assign fSE. There is no material 0000000056000000006161 - uL+ Kit-Front Case Assy uGeneral Emergency Room, Critical Room B ----------------------------------------------------------------------------- | 2023-10-20 17:23 | No
2582 | Yes | Legal Issue_Death // Due to a seriouevent, please urgently pull a log file | Legal Issue_Death // Due to a seriouevent, please urgently pull a log file | 000000000000000000000000001O##: 00000000000000000000000000100000000255000000005 Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: ONSITE PLEASE WITHOu PARTS PARTS: uTelemetry HTG 1 TODO: Pul logs and send to RCE Also send logs of the monitors TIME: 1h AP: Andreas Rupel +000000005315952300000000Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: Du to a seriou event, please ugently pul a log file Patient Impact: - SE - Diagnostic performed by Engineer: Onsite requred to pul logs Remote not reachable ----------------------------------------------------------------------------- Internal Remarks: u Roth is consuting with Remote and will contact u again. Internal Remarks: Gruning, Steffen by ## A ticket was already requsted by the cutomer in standby! Case#: 000000002200000000362, maybe the logs have already been saved. Internal Remarks: 00lanner: Cutomer has contract, is LSP cutomer Actully claim process 000000000000000000000000001ld still have to be driven, was now omitted du to time constraints Cutomer comment: Andreas Rupel by ##to u Schardin, Coud youcall me again regarding this at +00000000315952300000000 Thanks in advance. Internal Remarks: Thorsten Gerdes was informed abou the procedue by ##Cutomer comment: By ##to u Schardin, Please find the filled ou qustionnaire as far as I can answer the qustions. ufortuately, I don't have any futher data. Andreas Rupel Tel.: +00000000315952300000000 ## a.rupel00kbs.de Internal Remarks: After consutation with u Roth, this shoud actully be done by remote. Internal Remarks: Via u Roth: Remote puls the logs and will clarify and give u info on Monday whether an onsite service is still requred. Cutomer comment: By ##u Rupel Good day u Schardin, Here is the SN of the uilized telemetry transmitter Mu00000000 the SN is: u156J18700000000software SN: 3Z000000002BNM-W. ----------------------------------------------------------------------------- | 2023-10-20 18:02 | Yes
2583 | Yes | Touch screen is defective - touch screen broken | Touch screen is defeu - touch screen broken | 000000000000000000000000001O##: 0000000000000000000000000010000000025000000006 Consued parts: 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000----------------------------------------------------------------------------- Internal Remarks: FSE George Kamvosouis called to log an Onsite 000000000000000000000000001for a screen bu Cu is not covered for onsite. FSE adv'd to do Material only 000000000000000000000000001 I called POD, spoke to Matthew Tran, he created the Material 000000000000000000000000001 Pinged FSE George asking to sned The Part #and entitlement. ----------------------------------------------------------------------------- | 2023-10-20 18:23 | No
2584 | Yes | The screen sometimes moves significantly from side to side. | The screen sometimes moves significantly from side to side. |  | 2023-10-20 18:41 | No
2585 | Yes | Device will not power on; main host pc has no power supply ; device in clinical u at time of incident | Device will not power on; main host pc has no power supply ; device in clinical u at time of in0ent | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000026 200000000Oct-2000000003 Clement Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: boot u Eupected and ueupected Resuts: system 000000000000000000000000001ldnt boot properly uer Impact: 000000000000000000000000001kflow stop Patient Impact: none SE2.2.3 Diagnostic performed by Engineer : foud power to Sute PC available bu Sute pc was not booting u. Resoluion : Replaced Sute PC. Reloaded and restored system software to the Sute PC. System is fuctioning properly. Parts Consued : 1.0000000000000000000000000000000080000000051700000000,SuTE PC mDP Tools & Equpment : MuTIMETER,2000000003-11-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Joby Hyman ph: 000000002700000000500000000 IP & Issu: Device will not power on; main host pc has and power suply ; device in clinical ue at time of incident; 9366900000000 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: asap Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created Qute Neut Step Action: Awaiting Qute Verbal PO Nuber: Not Provided Verbal PO Approver: Not Provided Verbal PO CAP: N/A Internal Remarks : Michael Clement - FSE states that the usute PC will not tun. FSE has uu told to remove all power and reboot. if no resuts, approved. Internal Remarks : 0000000022000000009100000000 Clement Michael FSE replacing sute pc on u. FSE needs license swap. FSE sending to the IGT mailbou. Swapped and informed FSE to only load the PBL and not the Coronary Tools License. ----------------------------------------------------------------------------- | 2023-10-20 19:08 | No
2586 | Yes | The electrical board of the spare part is damaged and cannot be started | The eleuical board of the spare part is damaged and cannot be started |  | 2023-10-20 19:14 | No
2587 | Yes | Unspecified alarm, device inoperable, EVAL ADDED, IS complete | Unspe0ied alarm, device inoperable, EVAL ADDED, IS complete |  | 2023-10-20 20:05 | No
2588 | Yes | Wireless footswitch does not work, they want to connect another one that they have, question how this works | Wireless footswitch does not work, they want to conneuanother one that they have, question how this works | 000000000000000000000000001O##: 00000000000000000000000000100000000255563 MARK DAAMEN The system meets the specification for the performed service and is retuned to ue Follow u Requred: cutomer replaces footswitch receiver himself Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: y Eupected and ueupected Resuts: footswitch defect uer Impact: low Patient Impact: low SE see IB Diagnostic performed by Engineer: different footswitch no effect, receiver vv ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000255726 Parts Consued: 1.00000000000000000000000000000000800000000257862,WFS Base Station ----------------------------------------------------------------------------- | 2023-10-20 20:08 | No
2589 | Yes | Questioning the management of supraventricular tachycardias by the SMART ANALYSE software | Questioning the management of supraventricular tachycardias by the SMART ANALYSE software |  | 2023-10-20 20:47 | Yes
2590 | Yes | No power, will not boot up, GFE00OV0 No response/IS complete | No power, will not boot up, GFE0 0NOV00No response/IS complete | 000000000000000000000000001O##: 000000000000000000000000001000000002557100000000ewcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : w/AC power connected there are no LED's illuinated on the front of the vent. Foud no power at the power cord. Cutomer replaced the power cord, uit now powers u. Battery LED is flashing indicating it is charging, battery volts show 100000000c. Advised to let the battery fuly charge before retuning the vent to service. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: uit w/not power on uer Impact: none Patient Impact: na SE uknown ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: dong kim IP & Issu: 5950000000073/no power will not boot Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-20 21:03 | No
2591 | Yes | DEFOA-0-00_Speaker malfunction-Sound not confirmed. Waiting product return | DEFOA-00000_Speaker malfunuon-Sound not confirmed. Waiting produureturn |  | 2023-10-20 21:22 | No
2592 | Yes | the therapy PCA needs replacement | the therapy PCA needs replacement |  | 2023-10-20 21:37 | No
2593 | Yes | 00N0-ray working intermittently | 00CDN0u-ray working intermittently | 000000000000000000000000001O##: 00000000000000000000000000100000000256572 23-Oct-2000000003 Kolbinger Patrick The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : fluro pedal 000000000000000000000000001k intermittently When was the first occurence? Within... : 200000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: u-ray 000000000000000000000000001king intermittently cannot reboot at this time Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : replaced footswitch Parts Consued : 1.000000000000000000000000000000008000000000000000060000000000000000iplane Footswitch (000000002) 00000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000025600000000 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : fluro pedal 000000000000000000000000001k intermittently When was the first occurence? Within... : 200000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: u-ray 000000000000000000000000001king intermittently cannot reboot at this time Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Jaime reported that the u-ray 000000000000000000000000001king intermittently. Cutomer is uable to reboot the system as directed by local fse to keep system on. Log indicates that uray was enable and then disabled right away with message, "System not ready for flurscopy u-ray acqusition." Caller requsting onsite suport. Internal Comments : uLL# and/or description of uLL: u-ray 000000000000000000000000001king intermittently Repair Action : Instruted caller to clean TSM and inspect for cracked glass. FSE inspect footswitch and cable and replace defective component if any. 00000000800000000000000006000000003 - Biplane Footswitch (000000002) 00000000Follow u Requred : uLL# and/or description of uLL: u-ray 000000000000000000000000001king intermittently Repair Action : Instruted caller to clean TSM and inspect for cracked glass. FSE inspect footswitch and cable and replace defective component if any. 00000000800000000000000006000000003 - Biplane Footswitch (000000002) 00000000 Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- | 2023-10-20 22:13 | No
2594 | Yes | uLS Manual ¦ failed to pace "DPM Hardware failure" | TempuLS Manual ¦ failed to pace "DPM Hardware failure" | 000000000000000000000000001O##: 00000000000000000000000000100000000257182 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Clinical Device ue:Not in Clinical ue Eupected and ueupected Resuts: Failed to pace "DPM Hardware failue" uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Failed to pace "DPM Hardware failue" Diagnostic performed by Engineer : Problem Description: Failed to pace "DPM Hardware failue" Tempu LS: Failed to pace "DPM Hardware failue" Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 1000000002000000002000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: u Clinical Manager called to report Tempu LS failed to pace. "DPM Hardware failue" The following information was requsted from cutomer: *Log and Rescu Files: Requsted *Date of event? *Brief description of the event? *Operator of the produt? PARAMEDIC *Did the uit display any uLLmessages or warnings before the incident? "DPM Hardware failue" *Was the Patient or uer injued or caued any harm? uu *Patient Oucome? *Age of Patient? *Patient Gender? Was PRS ued: No PRS Permission: Follow u Requred : Please Tu East Pierce Fire and Rescu Attention: Jeff Berry 27723 Suner Bukley Hwy E Bukley WA 98321-9567 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000257288 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-10-20 22:57 | No
2595 | Yes | EC 0A - DA PCBA ADC reference failed | EC 0A - DA PCBA ADC reference failed | 000000000000000000000000001O##: 00000000000000000000000000100000000256900000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1100000000 uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and he stated the note on the device said broken. He verified there was a code 100000000A. Before we spoke, he noticed a few cables that were not secuely connected. He reconnected the DA pcba to MC pcba cable properly. The uit has been runing for a coule hous and no issus. Resoluion : Cutomer plans to do the PVT. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Darrel Sanders IP & Issu: 5910000000000000000/ uit listed as broken Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-21 00:13 | No
2596 | Yes | movements unavailable error | movements unavailable error | Internal Remarks : Interaction Type: New Case Caller: "Barry Fisher 500000000-72000000001100000000 IP & Issu: 5876000000005 movements uavailable uLL Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: Qute Action Taken: Created Qute Neut Step Action: N/A Verbal PO Nuber: N/A ----------------------------------------------------------------------------- | 2023-10-21 00:23 | No
2597 | Yes | Low Tidal Volume, Splitted from 00 | Low Tidal Volume, Splitted from 000 |  | 2023-10-21 00:36 | No
2598 | Yes | Part Id back light on the screen, screws at the bottom of unit to hold it up | Part Id back light on the screen, screws at the bottom of unit to hold it up | 000000000000000000000000001O##: 00000000000000000000000000100000000255957 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Part Id back light on the screen, screws at the bottom of uit to hold it u uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 00000000561528891 [RP-LCD, Assy,2nd GEN, V600000000 $1,387.00000000 0000000056153600000000 [RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000 $138.00000000 00000000561528911 [RP-Cable, u PCBA to LCD, V600000000 $66.00000000 00000000561528871 [RP-PCBA, uer Interface, u 2nd GEN,V600000000 $332.00000000 00000000561531651 [RP-Tray,LCD,2nd GEN V600000000 $229.00000000 00000000561528900000000 [RP-Screw,M2u3 Socket hd, PKG/100000000V600000000 $90000000000000000 00000000561500000000131 – Foot kit (package of 00000000- $32.00000000 989800000000630000000061 - RP-Thubwheel ,V600000000 uStand – $61.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states that the backlight has gone ou on LCD and is requsting a replacement backlight. informed cutomer that youneed to replace the whole LCD, there isn't jut a backlight that can be puchased. Foud that cutomer has 1st gen LCD. Informed cutomer of difference between 1st gen and 2nd gen. Cutomer requsting ##with part nuber and price for everything needed to ugrade. also requsting part nuber for replacement feet and thubwheel screws. ----------------------------------------------------------------------------- | 2023-10-21 00:48 | No
2599 | Yes | Alarms or Messages/ No sound : Speaker Malfunction : Unknown if Sound Worked: Speaker/HW: Replaced Part (000 | Alarms or Messages/ No sound : Speaker Malfunuon : Unknown if Sound Worked: Speaker/HW: Replaced Part (00000 |  | 2023-10-21 01:02 | No
2600 | Yes | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 / Other HW-Speaker-HW / Replaced Parts |  | 2023-10-21 01:03 | No
2601 | Yes | will not allow storage of images | will not allow storage of images | 000000000000000000000000001O##: 00000000000000000000000000100000000265200000000 27-Oct-2000000003 Brue Galen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 200000000Oct-2000000003 00000000:0000000058 RMT - IPu Free disk space is too low. 200000000Oct-2000000003 00000000:00000000:500000000 RMT - IPu IPPC failed to boot u - Frontal 200000000Oct-2000000003 00000000:22:00000000 When was the first occurence? Within... : 200000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System to save fluroscopy rus/System is not saving any rus uer Impact: System is not able to save fluroscopy rus Patient Impact: Delay of abou 10000000015 minues SE 8.2.3000000005 Resoluion : Install of IPpc test of system and log checks. Memory issus on post logs. Move of memory banks util no uLL. Diagnostic performed by Engineer : Install of IPpc test of system and log checks. Memory issus on post logs. Move of memory banks util no uLL. Repair of consistency and recheck Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000258000000002 Ketcherside Jeff The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 200000000Oct-2000000003 00000000:0000000058 When was the first occurence? Within... : 200000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: will not allow storage of images uer Impact: uable to do cine Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : cutomer states the system will not allow for cine euposues, ued remote console to do re-create image store Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Repair Action : Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Hard disk spare part 0000000080000000050000000000000000Resoluion : ued remote desktop to perform re-create image store, system now fuctional, cutomer wants the case to stay open till monday and needs a call back monday ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000025700000000 Vilkaukas Viktorija The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 200000000Oct-2000000003 00000000:0000000058 RMT - IPu Free disk space is too low. 200000000Oct-2000000003 00000000:00000000:500000000 RMT - IPu IPPC failed to boot u - Frontal 200000000Oct-2000000003 00000000:22:00000000 When was the first occurence? Within... : 200000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System to save fluroscopy rus/System is not saving any rus uer Impact: System is not able to save fluroscopy rus Patient Impact: Delay of abou 10000000015 minues SE 8.2.3000000005 Diagnostic performed by Engineer : Blake is reporting that the system is not allowing fluroscopy and keeps saying to delete euaminations. They have deleted everything eucept a teaching case and the message remains. The patient is on the table and the doctor has started the procedue bu withou fluroscopy they can not continu at this point. Remoted into the system and performed a (re)create image store. The system started to 000000000000000000000000001k and the procedue coud be continud. Blake created a logfile and it is showing that IPPC was slow to boot u. Euplained to Blake that the database might be corrut on the system and recreating it has a possibility of fiuing the issu. Called and informed FSE Thomas Hammer of the issu and what was done. We will monitor the system for a week and I will check in with Blake on Tusday 100000000200000000000000003. EuT 51000000005 Phone 31000000001000000007311 Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu Free disk space is too low.: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Root Caue for RMT - IPu IPPC failed to boot u - Frontal: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power or the OS/AS did not boot u in-time. Repair Action : Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu IPPC failed to boot u - Frontal: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu on InCenter in the produt manul of the PC. 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IP-PC and perform Cold reboot of the system. 1.3. If LED statu is not Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IP-PC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct! 1.5. Measue/replace the CMOS battery inside the IPPC becaue a bad/weak battery can caue the PC not to start the first time. 1.6. This pattern can also be caued if the BIOS is set wrong. please check: Advanced->Boot Featue->Restore On AC Power Loss set to 'Last State' or 'Power Off', defaut setting mut be 'Power On'. Diagnostic performed by Engineer : System euhibited the same issu again. My co-000000000000000000000000001ked performed another re create image store. Checked RSW and foud that the IPPC bay is degraded and at this point IPPC needs to be replaced. 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD Hard disk spare part 0000000080000000050000000000000000QTY 3 Follow u Requred : System euhibited the same issu again. My co-000000000000000000000000001ked performed another re create image store. Checked RSW and foud that the IPPC bay is degraded and at this point IPPC needs to be replaced. 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD Hard disk spare part 0000000080000000050000000000000000QTY 3 ----------------------------------------------------------------------------- Internal Remarks : 0000000022000000000000000000000000SE Brue reporting that newly installed IPPC has a memory 1 uLLin the logfile; however, 000000000000000000000000001ks fine. Recommend FSE swap memory sticks from 1 to 2...1 to 3 or jut wherever...if the issu follows the memory stick, then replace the PC, if the uLLgoes away, then it simply needed to be reseated. Internal Remarks : Interaction Type: New Case Caller: Kristen Galler IP & Issu: will not allow storage of images Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? N/A (iuR) Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created 000000000000000000000000001 roued to tech suport Internal Remarks : FSE Galen called in asking what to check for an intermittent issu with the image storage. I was able to retrieve the logging and noticed the IPPC and host are not commuicating. no control net000000000000000000000000001k connection IPPC - Host Which indicates the PC is not booting u. Sugest replacement of IPPC and drives if necessary. 0000000080000000050000000022 ClarityIQ_II IP PC non HDD 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II (u3 only if needed.) ----------------------------------------------------------------------------- | 2023-10-21 01:08 | No
2602 | Yes | No Sound / No Audible Tones-Keyrin s/nu0/ No Trouble-Found /HW- Replaced Parts | No Sound / No Audible Tones-Keyrin s/nu000/ No Trouble-Found /HW- Replaced Parts | 000000000000000000000000001O##: 00000000000000000000000000100000000257259 Thomas, Shea The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Cutomer ureachable, uable to determine Eupected and ueupected Resuts: Bench Repair-Mu00000000LCD ureadable and Speaker Malf (Withou Auio) uer Impact: Yes, Bench Repair-Mu00000000LCD ureadable and Speaker Malf (Withou Auio) Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Bench Repair-Mu00000000LCD ureadable and Speaker Malf (Withou Auio) Diagnostic performed by Engineer : Problem Description: Bench Repair-Mu00000000LCD ureadable and Speaker Malf (Withou Auio) Bench Repair-Mu00000000LCD ureadable and Speaker Malf (Withou Auio) Form ued: Bench Trouleshooting Steps Taken: Bench Repair-Mu00000000LCD ureadable and Speaker Malf (Withou Auio) Resoluion : Sent to Bench Follow u Requred : Sent to Bench ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: brittany IP & Issu: 62725167 speaker and screen not 000000000000000000000000001king Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: Created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-10-21 01:33 | No
2603 | Yes | Intermittently saying fluoroscopy not available reselect application | Intermittently saying fluoroscopy not available reseleuapplication | 000000000000000000000000001O##: 00000000000000000000000000100000000258000000000000000000000000ct-2000000003 Fedorka Timothy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: bio med Device ue: yes Eupected and ueupected Resuts: uable to Fluro FDC uLL uer Impact: no Patient Impact: no SE8.1.16 Resoluion : Replaced flat detector controller. Loaded software and configued. Tested system Parts Consued : 1.000000000000000000000000989710000000000000000573,Flashlite High End KIT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000256923 200000000Oct-2000000003 Fedorka Timothy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: bio med Device ue: yes Eupected and ueupected Resuts: uable to Fluro FDC uLL uer Impact: no Patient Impact: no SE8.1.16 Diagnostic performed by Engineer : uable to Fluro. Foud FDC uLL Follow u Requred : Replace FDC controller Tools & Equpment : MuTIMETER,200000000000000008-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Tom Petrozza Issu: 5582000000007 Intermittently saying fluroscopy not available reselect application Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: No contract or warranty Action Taken: Created case an requsted qute Neut Step Action: Awaiting Qute Verbal PO Nuber: Tom Petrozza Verbal PO Approver: 100000000000000009000000002 Verbal PO CAP: 0000000000000000 ----------------------------------------------------------------------------- | 2023-10-21 01:58 | No
2604 | Yes | Audio-No Sound / No Audio found at Bench-Foulink s/n u0/ Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at Bench-Foulink s/n u000/ Other HW-Speaker-HW / Replaced Parts |  | 2023-10-21 02:01 | No
2605 | Yes | The count down reaching zero/its not booting up/cuomer requested tech support | The count down reaching zero/its not booting up/cuomer requested tech support | 000000000000000000000000001O##: 00000000000000000000000000100000000258132 200000000ct-2000000003 Stone, Matthew The system meets the specification for the performed service and is retuned to ue Resoluion : Installed fleu vision, PC into uit calibrated downloaded board software performed accepting Settings tested had cutomer test retuned ful service Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: cath Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE Diagnostic performed by Engineer : No image Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000025000000003 200000000Oct-2000000003 Stone, Matthew The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: cath Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE Diagnostic performed by Engineer : No display on fleuvision bad grabber card Follow u Requred : Install fleuvision ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Samul Medina IP & Issu: 65566796/The cout down reaching zero/its not booting u/cutomer requsted tech suport Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created O000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : assisted the FSE with trouleshooting the no boot problem. Told him how to get into FSF at the boot u screen dots. He is downloading the uLLlogs for evalution. foud that the grabber cards in the fleuvision pc are not initializing cauing the system to not boot u. TPR approval coming ----------------------------------------------------------------------------- | 2023-10-21 02:07 | No
2606 | Yes | Unit will not power on | Unit will not power on | Internal Remarks : Caller: Mike 731 00000000 2693 IP & Issu: CCB - Awaiting service, and have not been notified of anything. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: Requsted Qute Action Taken: Created an AA to Qutation Neut Step Action: Internal Remarks : Caller: Mike Walker 731 00000000 2693 IP & Issu: 62719278 - uit will not power on Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: Requsted Qute Action Taken: Awaiting Qute- Approved eupedite fee $16500000000 Neut Step Action: Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mike Walker 7315000000008900000000 IP & Issu: Cutomer already sent hard copy back for eupedited service, PO - 953677. They will re-send back PO and qute to the Servicequtesucom ## needing this to be entitled ASAP Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Created Case and Page FSE Neut Step Action: Awaiting Entitlement Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Caller: Adam Mauwell 731 00000000 3113 IP & Issu: CCB to see if PO was recievd Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: Requsted Qute Action Taken: Awaiting Qute - Informed PO has been received attached to case Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-21 02:13 | No
2607 | Yes | Primary alarm failure | Primary alarm failure | 000000000000000000000000001O##: 00000000000000000000000000100000000258300000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to replace speaker to correct issu, cutomer acknowledged and will order part, no follow u needed, issu resolved- Material Nuber 00000000561511921 - RP-Speaker Assy, V600000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states duing uit set u it alarms Primary alarm failed message at start u, no patient involvement, complaint dulicated in biomed and noted uLLcode 1100000000/5000000001 PWR speaker failed in event log, requsts part to replace. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:Primary alarm failed message uer Impact: not uable needs repair Patient Impact: none SEun ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Denny Budah Issu: 59395537 Primary alarm failue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-21 02:53 | No
2608 | Yes | OP check is fault | OP check is fault | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Contacted ASP honglejia for buiness follow-u Neut Step Action: LiuWeiming, 18916700000000559 Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: SA:00000000/uDFM100000000 Defibrillator/Monitor Neut Step Action: The equpment is at the Yuang Campu of the First Affiliated Hospital of Zhejiang uiversity School of Medicine Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service ----------------------------------------------------------------------------- | 2023-10-21 06:03 | No
2609 | Yes | Speaker is malfunctioning | Speaker is malfunuoning | Internal Remarks : Interaction Type: New Case Caller: Budy IP & Issu: 95322153- Speaker is malfuctioning Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-21 06:53 | No
2610 | Yes | 0/u00Geometry movements partly available | 0N/uGeometry movements partly available | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 25.1000000002000000003 12:59:00000000 uC (8600000000000000007)approve all of the docuentation in this s000000000000000000000000001and confirm this repiar in a safe and complete manner Internal Remarks: 26.1000000002000000003 100000000000000008 uC (86000000000000000000000000RDI Accuacy: YES $Q6: (The geometry can't boot u.) $Q7:(Comfirmed reported problem.The geometry can't boot u.) $Q8:(Check and find it is the L-arm AMC problem, need to replace the AMC.) $Q9:(Replace the L-arm AMC and do the calibration.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: N/A. Problem Reported by Cutomer: 00000000N/u7 M200000000uLLreported when uing Euternal Remarks: 26.1000000002000000003 100000000000000005 uC (86000000000000000000000000 Service content: The mechanical movement part cannot be powered on. The inspection foud that it was the AMC problem behind the L arm. Replace the AMC. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Dispatched for Onsite Suport Neut Step Action: PPD AMC Triple Servo Drive HP-LP-LP 00000000800000000990000000051 1 200000000ou shipment has been arranged, please send 000000000000000000000000001k FSE: Liuuuu-86-18813361000000005 T2 Activities:Repair Action: PPD/ 00000000800000000990000000051 AMC Triple Servo Drive HP-LP-LP Safety Qustions: Case: 000000002200000000667 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No T2 Activities:Problem Description: Overview of cutomer problems: After the device is tuned on, all mechanical movements fail, even if it is restarted.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: Withou PRS and logs, the caue of the failue cannot be determined.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Engineers arrive on site to perform repairs based on uLLlogs and indicator light statu. Notes on cutomer requrements: Onsite SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 25.1000000002000000003 12:59:12 uC (8600000000000000007) Approve all of the docuentation in this s000000000000000000000000001and confirm this repiar in a safe and complete manner SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-21 06:58 | No
2611 | Yes | Speaker malfunction. Onsite repair pending | Speaker malfunuon. Onsite repair pending |  | 2023-10-21 07:47 | No
2612 | Yes | The machine touchscreen is malfunctioning. | The machine touchscreen is malfunuoning. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Internal Remarks: 200000000000000002000000003 00000000:56:36 uC (86000000001000000005) $ASP engineer:[Qiangyi Gao uinyu $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7:（Confirmed the reported problem & The machine touhscreen is malfuctioning） $Q8:（Restart the device and check the device. The problem persists After reconnecting the touhscreen) $Q9:（Replace the TouhScreen） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11:Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration Date (00000000/11/20000000000000000SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 200000000000000002000000003 00000000:56:33 uC (86000000001000000005) *Department information: [Department: Equpment Department Contact person: Teacher Zhao Contact nuber: 000000002-2000000001556 ]* $ASP engineer:[Qiangyi Gao uinyu $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7:（Confirmed the reported problem & The machine touhscreen is malfuctioning） $Q8:（Restart the device and check the device. The problem persists After reconnecting the touhscreen) $Q9:（Replace the TouhScreen） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11:Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration Date (00000000/11/20000000000000000SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA Problem Reported by Cutomer: 00000000/V600000000uLLbu parts Safety Qustions: Case: 0000000022000000007900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhang Caiqin Last udated by: Zhang Caiqin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-21 07:48 | No
2613 | Yes | Unit will not power on | Unit will not power on | 000000000000000000000000001O##: 00000000000000000000000000100000000258219 21-Oct-2000000003 Bolton, Brian The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: OR manager Device ue: uk Eupected and ueupected Resuts: system will not boot u du to fleuvision pc uer Impact: system will not boot Patient Impact: none SE 8.1.17 Diagnostic performed by Engineer : System stuk in boot cycle showing time to system start 00000000:00000000. Software level is 8.1.17 requring an IST puh to retrieve uLLlog. Logs sent to NSS and confirmed that FleuVision PC is in need of replacement to resolve issu. Qute generated and sent to cutomer for parts and labor, pending approval. Follow u Requred : Onsite service needed pending approval of additional time and material ----------------------------------------------------------------------------- Internal Remarks : Fatal: Failed to initialize all grabber cards. Details: Failed to open grabber channel 00000000Failed to open grabber channel 5, Failed to open grabber channel 6, Failed to open grabber channel 7 Approved for replacement 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD Internal Remarks : Instruted FSE to shu down system since stuk at zeros. Power on system and ue ALT + F key to login FSF and view logging for possible grabber cards failed to initialize. FSE will call back in. ----------------------------------------------------------------------------- | 2023-10-21 08:13 | No
2614 | Yes | 0u FD0Carm move fault | 0/u FD0Carm move fault | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: 00000000/u uuuFD200000000Neut Step Action: C-arm movement problem, display screen problem, 0000000016200000000700000000 Call for repair, Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service Safety Qustions: Case: 000000002200000000900000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Contacted ASP zhutianzhufor buiness follow-u Euternal Remarks: 00000000.11.2000000003 00000000:16:00000000 uC (86000000008287) Service content: Check equpment problems and reinstall rack cable snakeskin tues. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. Internal Remarks: 00000000.11.2000000003 00000000:16:00000000 uC (86000000008287) RDI Accuacy: Yes $Q7: (Confirmed reported problem, the cutomer report several problems and ask FSE onsite to check.) $Q8: (Check the system onsite, finding the stand longituinal drive's ruber wheel is broken; the table rail clamps is stuked; the floating buton on GEO modue is dropped; the table side u-ray protection shield is ou of shape; the table base connection board is corroded; the cables condut of stand is dropping down.) $Q9: (Fiu the cables condut of stand and order the broken parts to replaced. Another S0000000000000000000000000017500000000100000000s created to report the parts.) $Q100000000 (Test and Inspection data refer to PA. Device retuned to service.) $Q11: Calibrated Tools (N/A) ----------------------------------------------------------------------------- | 2023-10-21 08:48 | No
2615 | Yes | The equipment does not pass a pre-specification inspection. | The equipment does not pass a pre-spe0ication inspeuon. | Internal Remarks: Failue content 2 The pre-specification inspection did not pass. There is no problem with new pads and batteries. ----------------------------------------------------------------------------- | 2023-10-21 10:03 | No
2616 | Yes | AED emits triple beep - log analysis | AED emits triple beep - log analysis | 000000000000000000000000001O##: 00000000000000000000000000100000000252576 Suui, Kosue The system meets the specification for the performed service and is retuned to ue Resoluion: When the AED log was analyzed at the u manuactuing plant, there was a record that the diagnostics failed becaue the pad impedance was higher than the permissible valu. This alarm occured becaue either the pad case was not properly closed, or a load was applied to the cable, suh as a pad wire getting cauht in the cover. There was no record of any other problem, and there was no record of a problem with the AED. From the above, there is no faut with the AED and the pad, so I am retuning the AED. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred.November 9, 2000000003 Fuua Denshi Tokyochu Hanbai I will send the 000000000000000000000000001k report along with the equpment to the person in charge. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit a 000000000000000000000000001k report on paper? ＜YES＞ Final check: I checked the operation following the 000000000000000000000000001k based on the set of orders. Yes Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of Sales at Fuda Electronics Device ue: Not ued Eupected and ueupected Resuts: The AED beeped three times. uer Impact: None Patient Impact: None SE None Diagnostic performed by Engineer: When I received the AED, I performed a battery test and confirmed that it passed. I did not confirm the reproduibility of the abnormal alarm even with standby monitoring. Tools & Equpment : MuTIMETER, 200000000000000003-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 A PAD uLLoccus and when it's new the uLLdisappears LOT Adut pad 220000000016-0000000051 ----------------------------------------------------------------------------- | 2023-10-21 10:52 | No
2617 | Yes | Self-Test Failed | Self-Test Failed | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 000000002200000000362 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : CM - SELF TEST FAILED ----------------------------------------------------------------------------- | 2023-10-21 11:28 | No
2618 | Yes | CM - SELF TEST FAILED (split) | CM - SELF TEST FAILED (split) | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : CM - SELF TEST FAILED+Battrey Issu SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 000000002200000000368 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable ----------------------------------------------------------------------------- | 2023-10-21 11:28 | No
2619 | Yes | SELF TEST FAILED | SELF TEST FAILED | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 0000000022000000003700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : CM - SELF TEST FAILED SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-21 11:28 | No
2620 | Yes | Malfunctioning of the power supply unit. | Malfunuoning of the power supply unit. | 000000000000000000000000001O##: 00000000000000000000000000100000000252700000000Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: nemo withou collection ELCAMM SRL 139 ROMA STREET 70000000026 ASI MODuNO (BA) AREA AC. MICHELANGELO TOTARO 0000000000000000337726 Problem Description by Engineer: Cutomer Fuction/Role: dealer Device ue: defibrillator test Eupected and ueupected Resuts: power equpment misalucion uer Impact: no impact Patient Impact: no impact SE na Contacted elcamm technician Totaro Michelangelo, the device always reports power equpment malfuction and from the uLL there seem to be problems on power suply, proceeded with order of part. Diagnostic performed by Engineer: message: power equpment malfuction ----------------------------------------------------------------------------- | 2023-10-21 11:48 | No
2621 | Yes | Speaker malfunction-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Speaker malfunuon-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000252918 Seidel, Bernhard The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: Dear Sir or Madam, the following answers to you list of qustions: Did the speaker no longer have any soud or was there a slight, possibly distorted soud or a beep? - No soud Were any INOP uLLmessages displayed? - no A. If an INOP uLLwas displayed, was there auio present duing that message? - Did the device 000000000000000000000000001k as a standalone device or was it connected to a host monitor? - On host monitor, uLLdisplay even when ued as a standalone device When was the problem with the speaker malfuction discovered? Duing setu, maintenance or duing patient ue? - Duing patient operation What software was installed on the device when the speaker malfuction occured? - M.0000000000000000100000000Even a cold start of the device coud not resolve the uLLDiagnostic performed by Engineer: +++NEMO KVA+++ Device, spare part is defective and mut be replaced. The system is ou of ue util the replacement part is installed. Parts: 0000000056000000003831 Follow u Requred: 0000000056000000003831 Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Device/component is reported as defective. uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer: Dear Sir or Madam, as requsted, I will send youthe following data: uuuProcessing nuber: 000000002200000000598System name: IntelliVu Muti-Measuement Modue You reference: 000000000000000086321u Du to internal reguatory requrements, qustions regarding ue and possible behavior in the event of an alarm mut be clarified before delivery of certain spare parts. This particuarly applies to speakers. The main point here is that if the alarm is lost, this may have to be reported to the responsible auhority. This resuts in the following set of qustions: · Did the louspeaker no longer have any soud or was there a slight, possibly distorted soud or a beep? · · Were any INOP uLLmessages displayed? A. If an INOP uLLwas displayed, was there auio present duing that message? · · Did the device fuction as a standalone device or was it connected to a host monitor? · · When was the speaker malfuction problem discovered? Duing setu, maintenance or duing patient ue? · · What software was installed on the device when the speaker malfuction occured? ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 Parts Consued: 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- Cutomer comment: Cutomer not reached by phone Cutomer comment: u Moewes reports by ## Dear Sir or Madam, the following answers to you list of qustions: Did the louspeaker no longer have any soud or was there a slight, possibly distorted soud or a beep? No soud Were any INOP uLLmessages displayed? no A. If an INOP uLLwas displayed, was there auio present duing that message? . Did the device fuction as a standalone device or was it connected to a host monitor? On host monitor, uLLdisplay even when ued as a standalone device · When was the problem with the speaker malfuction discovered? Duing setu, maintenance or duing patient ue? Duing Patient Operation What software was installed on the device when the speaker malfuction occured? M.0000000000000000100000000Even a cold start of the device coud not resolve the uLL Cutomer comment: KVA sent by ## Cutomer comment: Dear u Schuz, attached is the approved KV. Kind regards Maic Regner Maic Regner Head of Medical Technology Cutomer comment: Cutomer not reached by phone. ----------------------------------------------------------------------------- | 2023-10-21 12:22 | No
2622 | Yes | SPLIT - 00- Paddle cable issue | SPLIT - 0000- Paddle cable issue | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 1600000000000000002 - Paddle cable & ucable issu Safety Qustions : Case: 000000002200000000811 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Karen Lee Last udated by: Karen Lee 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-21 12:58 | No
2623 | Yes | 0V0uLLTouch screen malfunction | 0/V0uLLTouch screen malfunuon | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/V600000000uLLTouh screen malfuction Safety Qustions: Case: 000000002200000000833 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhang Caiqin Last udated by: Zhang Caiqin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-21 13:08 | No
2624 | Yes | uLS AED ¦ Pacemaker function does not work, have to press the “discharge” button several times to activate it'. | TempuLS AED ¦ Pacemaker funuon does not work, have to press the “discharge” button several times to aute it'. |  | 2023-10-21 13:12 | No
2625 | Yes | 0/AED/Voice message:Unavaible | 0N/AED/Voice message:Unavaible | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000N/AED/ SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: A200000000-00000000753 The cutomer informed that the actul location is at Moleu Connector ChengduCo., Ltd. Safety Qustions: Case: 000000002200000000936 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . SAME DAY HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on of day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLLVoice message: Device not available ue, reinstall the battery and detect that it cannot be repaired. Repair Action: ---------RA-----REPAIR ACTION--------------- IF PART REQuRED - 12 NC RECOMMENDED : 0000000056000000002711 HS1 (W B) Defib Simp. Chinese Euchg RECOMMENDED REPAIR ACTION: Trouleshooting HS1 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-21 13:23 | No
2626 | Yes | SC: The Live monitor in the euamination room didn't display anything. | SC: The Live monitor in the euamination room didn't display anything. | 000000000000000000000000001O##: 00000000000000000000000000100000000253323 25-Oct-2000000003 Takebayashi, TakanobuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiology Device ue: Duing euamination Eupected and ueupected Resuts: None uer Impact: None Patient Impact: None SE2、2、3 Resoluion: I replaced the DVI signal distribuor in the R cabinet in the machine room. When I finished, I checked the operation. Parts ued Part nuber: 000000008000000000000000000000000 Part name: DVI Splitter Modue uits: 1 Parts Consued: 1.000000000000000000000000000000008000000000000000000000000,DVI Splitter modue ----------------------------------------------------------------------------- | 2023-10-21 13:33 | No
2627 | Yes | 0S:KN: "System starting 0" is being ud and the system does not start. | 0DS:KN: "System starting 0" is being ud and the system does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000259336 21-Oct-2000000003 Matsuaki, MasakazuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: System starting 00000000 is being ued and it 000000000000000000000000001't start Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: System starting 00000000 is displayed and the system does not start. uer Impact: No effect Patient Impact: No effect SE 8.2.3000000005 Resoluion: The equpment did not start, bu when he pressed the power buton on the large monitor control PC (FleuVision PC), it retuned to normal. Based on the log from when the problem was occurence, the caue was the above FleuVision PC startu failue. I repeated startu tests, bu the problem was not reproduibility, so I think it was a temporary startu problem. Jut to be sue, I changed the power suply souce to the PC and checked the PCs uing diagnostic tools. I confirmed normal operation in the final operation test. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000259276 Noda, Kotaro The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: System starting 00000000 is being ued and it 000000000000000000000000001't start Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: System starting 00000000 is displayed and the system does not start. uer Impact: No effect Patient Impact: No effect SE 8.2.3000000005 Diagnostic performed by Engineer: When he started the system, the coutdown was 00000000 and the system 000000000000000000000000001ld not start.Restarting did not improve the sitution.The inside of the FV is dark and the indicator lamp is orange. I had him tun on the FleuVision PC in cabinet B directly, and then it started, and the FleuVision also started and the equpment started.The equpment log also shows that the FleuVision uit power did not come on. An FSE will check the part. Internal Comments: [Suject] 00000000DS:KN:u Takata: "System starting 00000000" is being ued and it 000000000000000000000000001't start [Internal Comments] FleuVisionControlService:Observer000000000000000000000000001ker PING(FLEuVISION_PC_CONTROL) is not suceeded after 100000000 seconds. Please check if FLEuVISION_PC_CONTROL is powered on and all (net000000000000000000000000001k) cables are correctly connected. [u-ed parts] 0000000080000000012200000000:u Haswell Fleuvision 2 PC no HDD QTY:1 ZREP Osaka (u15) Reason for u-collection: Consuted with engineer. Resut of 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Osaka (u15) Reason for u-collection: Consuted with engineer. Resut of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name --None-- - FSE1 Visit date and time 2000000003/10000000021--None-- - FSE1 000000000000000000000000001k time 2 [Euamination category] Normal euamination Follow u Requred: * ----------------------------------------------------------------------------- | 2023-10-21 13:43 | No
2628 | Yes | CR.NO PASSES THE TEST, IT RESTARTS BY MISTAKE | CR.NO PASSES THE TEST, IT RESTARTS BY MISTAKE | 000000000000000000000000001O##: 00000000000000000000000000100000000278131 Vega Zevallos, Juissa Kriss The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred:  The device displays the fiued signal u on the RFuindicator with a beep and restarts constantly. - This failue is uully related to the processor, when placing a test one it is observed that it does not completely solve the faut, cauing failue in the blades. In the uLLlog, a critical failue with code 7:.29 is observed, which corresponds to the therapy plate.  A test therapy plate is placed. The failue of the ignition and blades can be solved. - The equpment passes all the tests and the operational check. - Therapy PCA mut be replaced. - The equpment has a damaged and broken front casing. The front case Assy, Therapy Select knob and Label set mut be replaced. - The DFM100000000 has a broken printer, the roller is missing. The Assy printer mut be replaced. Equpment received with paddles, test plu, battery and software version 2.00000000.21, the software will be udated when replacing the therapy plate. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: Cutomer indicates failed test and restarts. uer Impact: uknown Patient Impact: uknown SE N/A Diagnostic performed by Engineer:  The equpment displays the fiued signal u on the RFuindicator with a beep and restarts constantly. - This failue is uully related to the processor, when placing a test one it is observed that it does not completely solve the faut, cauing failue in the blades. In the uLLlog, a critical failue with code 7:.29 is observed, which corresponds to the therapy plate.  A test therapy plate is placed. The failue of the ignition and blades can be solved. - The equpment passes all the tests and the operational check. - Therapy PCA mut be replaced. - The equpment has a damaged and broken front casing. The front case Assy, Therapy Select knob and Label set mut be replaced. - The DFM100000000 has a broken printer, the roller is missing. The Assy printer mut be replaced. Equpment received with paddles, test plu, battery and software version 2.00000000.21, the software will be udated when replacing the therapy plate. ----------------------------------------------------------------------------- | 2023-10-21 14:08 | No
2629 | Yes | SA:0/u FD0can't eupose | SA:0R/u FD0can't eupose | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 23.1000000002000000003 00000000:10000000000000000uC (8600000000800000000) 23.1000000002000000003 00000000:58:21 uC (8600000000800000000) Service content: The device cannot be euposed, on-site inspection ippc system is lost, reinstall the system. Service resut: The equpment is ued normally. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 21.1000000002000000003 13:19:00000000 uC (86000000009567) Only on job training Neut Step Action: No rays Problem Reported by Cutomer: SA:00000000R/u uuuFD200000000Neut Step Action: CHAI YANFEI Internal Remarks: 23.1000000002000000003 00000000:100000000500000000uC (8600000000800000000)$Q7: (Confirmed reported problem, can't eupouse ) $Q8: (Check the logfile no report uLLippc systerm cannot boot u) $Q9: (install IPPC OS and APP) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: not ue tools ^Curent Tue installed date: ^Curent Tue uage: SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: can't eupose Trouleshooting Action: Key faut codes or phenomena: The ippc hard drive is directly connected PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Euecue database consistency test, database consistency repair & re-create image storage Recommended spare parts uuand designation: 0000000080000000012200000000 ippc; 00000000800000000122391 host pc; 0000000080000000050000000000000000hard disk Safety Qustions: Case: 000000002200000000100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 21.1000000002000000003 13:19:00000000 uC (86000000009567) only on job training ----------------------------------------------------------------------------- | 2023-10-21 16:23 | No
2630 | Yes | Loudspeaker uLL(Sound not confirmed) Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part | Loudspeaker uLL(Sound not confirmed) Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part | 000000000000000000000000001O##: 0000000000000000000000000010000000025000000003 Diagnostic performed by Engineer: The uLLcannot be reprodued uder 000000000000000000000000001kshop conditions Problem Description by Engineer: shows sporadic "Speaker failue" Parts Consued: 1.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly Resoluion: Speaker changed ----------------------------------------------------------------------------- | 2023-10-21 16:43 | No
2631 | Yes | The u-ray were blocked during the euaminations twice | The u-ray were blocked during the euaminations twice | 000000000000000000000000001O##: 00000000000000000000000000100000000255200000000ANDREA COLOMBO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: The u-ray e rays were blocked duing the euaminations twice. When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : Intermittent Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: no u-ray available, 2 times uer Impact: no Patient Impact: no SE 000Diagnostic performed by Engineer: reported 2 episodes of no ray emission. downloaded and analyzed log files of the system. Internal Comments: uLL# and/or description of uLL Reported 2 episodes of no ray emission. The log files do not indicate a faut on the biplane. Repair Action: Reported 2 episodes of no ray emission. downloaded and analyzed log files of the system. Resoluion: +39 0000000082 500000000737 asked the technical u. Manula to tun on the biplane to be able to download the data before October 200000000 informed cutomer u Campi that the report is in monitoring. on October 300000000 since no other elements emerge, in agreement with u Campi closes this report. at the moment the system is 000000000000000000000000001king. ----------------------------------------------------------------------------- | 2023-10-21 18:33 | No
2632 | Yes | System down | System down | 000000000000000000000000001O##: 00000000000000000000000000100000000259200000000 21-Oct-2000000003 TRONG HIEuNGuEN The system does not meet the specification for the performed service and is not retuned to ue Tools & Equpment : MuTIMETER,200000000000000008-21 00000000:00000000:00000000 Diagnostic performed by Engineer : System down becaue DPumain interface is defected and waiting for PO Problem Description by Engineer : Cutomer Fuction/Role: technician Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE2.u ----------------------------------------------------------------------------- | 2023-10-21 19:43 | No
2633 | Yes | system wont boot in | system wont boot in | 000000000000000000000000001O##: 00000000000000000000000000100000000325000000006 00000000-Nov-2000000003 WEINGART, SASHA The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: Boot u of system Eupected and ueupected Resuts: System sticks on boot u screen and will not boot all the way u uer Impact: uable to ue system Patient Impact: None. SE8.1.17 Resoluion : Had in-houe engineer retun these t000000000000000000000000001uued parts. Proper parts consuption - coutesy time no charge. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032500000000 00000000-Nov-2000000003 Caldwell, Pau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: cardiac Eupected and ueupected Resuts: no video uer Impact: down Patient Impact: SE Resoluion : Assisted with fleuvision replacement and software load ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000026760000000000000000-Nov-2000000003 WEINGART, SASHA The system meets the specification for the performed service and is retuned to ue Follow u Requred : replace parts again, system down Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: Boot u of system Eupected and ueupected Resuts: System sticks on boot u screen and will not boot all the way u uer Impact: uable to ue system Patient Impact: None. SE8.1.17 Resoluion : Replaced the Fleuvision PC, took several attempts to get the SW to load correctly bu eventully got it 000000000000000000000000001king properly. Tested, verified normal operation; ready for ue. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000256698 23-Oct-2000000003 Hargrave Brue The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace Fleu PC and Hard drive in Fleu PC Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: Boot u of system Eupected and ueupected Resuts: System sticks on boot u screen and will not boot all the way u uer Impact: uable to ue system Patient Impact: None. SE8.1.17 Diagnostic performed by Engineer : Troule shoot system uing Docuent DEP110000000023 Rev C from u InCenter. Foud that the Fleu PC power suply has issu on boot u. Ordering new Fleu PC on follow u call. Puled uLLlogs also. Tools & Equpment : MuTIMETER,2000000005-0000000030000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action Taken: Follow u activity for FSE Brue Hargrave 1000000002500PM part order was canceled Oucome: dispatched Internal Remarks : Interaction Type: New Case Caller: Robert Victor IP & Issu: 555293-system 000000000000000000000000001t boot in Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Qute Reason for Entitlement Claim: Qute Action Taken: created case and qute Neut Step Action: Awaiting Qute Internal Remarks : Planner Initials: GM Action Taken: Called Cu no answer left vm no response. Called FSE Brue Hargrave he accepted call for 1000000002300AM he will confirm with the cutomer Oucome: dispatched ----------------------------------------------------------------------------- | 2023-10-21 19:58 | No
2634 | Yes | Problem on the u in particular in the hardware error | Problem on the u in particular in the hardware error | 000000000000000000000000001O##: 00000000000000000000000000100000000256182 DI FOGGIA, FRANCESCO The system does not meet the ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: yes Eupected and ueupected Resuts: hw uLLu uer Impact: no Patient Impact: no SE na HW uLLu Diagnostic performed by Engineer: Heard from: Domenico Romano in charge of 118 Brindisi and Eng Domenico Giordano at Iredeem. At the company they installed some defibrillators that occasionally duing the auo test give the uLLof fauty u The check of the uLLhistory reports the following uLL"ucommon leakage curent". In agreement with ING Giordano, in first analysis I send complete ucable with the management to verify the conditions of the cables, the connections and any presence of disposable electrodes connected To the derivation withou patient. If there are futher problems, the cutomer will open a new report. Follow u Requred: Nemo withou collection. Iredeem SPA Piazza dei Martiri 1900000000- 1900000000 1, 0000000021 Bologna BO Domenico Giordano,+390000000010000000035879 ----------------------------------------------------------------------------- | 2023-10-21 20:23 | No
2635 | Yes | Problem on the u in particular in the hardware error | Problem on the u in particular in the hardware error | 000000000000000000000000001O##: 00000000000000000000000000100000000256181 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: yes Eupected and ueupected Resuts: hw uLLu uer Impact: no Patient Impact: no SE na HW uLLu Diagnostic performed by Engineer: Heard from: Domenico Romano in charge of 118 Brindisi and Eng Domenico Giordano at Iredeem. At the company they installed some defibrillators that occasionally duing the auo test give the uLLof fauty u The check of the uLLhistory reports the following uLL"ucommon leakage curent". In agreement with ING Giordano, in first analysis I send complete ucable with the management to verify the conditions of the cables, the connections and any presence of disposable electrodes connected To the derivation withou patient. If there are futher problems, the cutomer will open a new report. Follow u Requred: Iredeem spa Piazza dei Martiri 1900000000- 1900000000 1, 0000000021 Bologna BO Domenico Giordano,+390000000010000000035879 ----------------------------------------------------------------------------- | 2023-10-21 20:23 | No
2636 | Yes | During the self-test give the uLLof a faulty ECG | During the self-test give the uLLof a faulty ECG | 000000000000000000000000001O##: 0000000000000000000000000010000000025600000000 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: yes Eupected and ueupected Resuts: hw uLLu uer Impact: no Patient Impact: no SE na HW uLLu Diagnostic performed by Engineer: Heard from: Domenico Romano in charge of 118 Brindisi and Eng Domenico Giordano at Iredeem. At the company they installed some defibrillators that occasionally duing the auo test give the uLLof ugusto.Il problem comes back when doing the fuctional test And then it reompt occasionally duing the auo test. The check of the uLLhistory reports the following uLL"ucommon leakage curent". In agreement with ING Giordano, in first analysis I send complete ucable with the management to verify the conditions of the cables, the connections and any presence of disposable electrodes connected To the derivation withou patient. If there are futher problems, the cutomer will open a new report. Follow u Requred: Iredeem spa Piazza dei Martiri 1900000000- 1900000000 1, 0000000021 Bologna BO Domenico Giordano,+390000000010000000035879 ----------------------------------------------------------------------------- | 2023-10-21 20:58 | No
2637 | Yes | Request for replacement under warranty of ucable Intrepid | Request for replacement under warranty of ucable Intrepid | 000000000000000000000000001O##: 0000000000000000000000000010000000025600000000 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: yes Eupected and ueupected Resuts: hw uLLu uer Impact: no Patient Impact: no SE na HW uLLu Diagnostic performed by Engineer: Heard from: Domenico Romano in charge of 118 Brindisi and Eng Domenico Giordano at Iredeem. At the company they installed some defibrillators that occasionally duing the auo test give the uLLof ugusto.Il problem comes back when doing the fuctional test And then it reompt occasionally duing the auo test. The check of the uLLhistory reports the following uLL"ucommon leakage curent". In agreement with INg Giordano, in first analysis I send complete ucable with the management to verify the conditions of the cables, the connections and any presence of electrodes. Disposable connected to the derivation withou patient. If there are futher problems, the cutomer will open a new report. Follow u Requred: Iredeem spa Piazza dei Martiri 1900000000- 1900000000 1, 0000000021 Bologna BO Domenico Giordano,+390000000010000000035879 ----------------------------------------------------------------------------- | 2023-10-21 21:02 | No
2638 | Yes | M0A - u00 uLLmessage | u uLLmessage | BuQ&A Feedback: No Reason: Corrective Maintenance Trouleshooting Action: Qute sent SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.12.2000000003 00000000:15:51 POLARIS (POLARIS) ================= André Cifone Peiuoto - 000000000000000057600000000600000000000000002000000000000000000000000300000000=================== === Application of part of invoice 0000000078. After replacing the uuuitor, the equpment was sujected to tests and passed all of them. Tests: Visul, Power on, Operational Check and Load triggering in Jauím etro (Simuator) Neut Step Action: Analysis carried ou and directed to sending a buget Repair Action: OPP-198376 Qute: 00000000113100000000 Shipping ## Engenhariaclinica00abara.com.br SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 15.11.2000000003 00000000:00000000:00000000POLARIS (POLARIS) ================= André Cifone Peiuoto - 000000000000000057600000000600000000000000001000000000000000000000000300000000=================== === When carrying ou the operational test, the equpment fails the defibrillation item and when firing with a 200000000J load in the simuator, the resut is 1700000000. Both tests show the inefficiency of the uuuitor, therefore, for perfect operation of the equpment, it will be necessary to replace this defective part. SAP Transfer: Case Transferred to SAP Trouleshooting Action: schedued for 11/100000000 FSE Cifone SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000368 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT Internal Remarks: 15.11.2000000003 00000000:00000000:00000000POLARIS (POLARIS) ========== André Cifone Peiuoto - 000000000000000057600000000600000000000000001000000000000000000000000300000000=================== === When carrying ou the operational test, the equpment fails the defibrillation item and when firing with a 200000000J load in the simuator, the resut is 1700000000. Both tests show the inefficiency of the uuuitor, therefore, for perfect operation of the equpment, it will be necessary to replace this defective part. GG00000000: Test resuts - NOK - GG13: Calibratable tools - N/A Problem Reported by Cutomer: Equpment problem: Equpment uLLmessage Priority: STOPPED uit: FuDAção JOSE LuZ EGYDIO SETuAL Address: Avenida Angeliu87 Sao Pauo-SP 00000000227-200000000 Search for: Igor Silva Clinical engineering 000000001 3155 28900000000ngenhariaclinica00abara.com.br SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.12.2000000003 00000000:15:51 POLARIS (POLARIS) ======================== André Cifone Peiuoto - 000000000000000057600000000600000000000000002000000000000000000000000300000000=================== === Application of part of invoice 0000000078. After replacing the uuuitor, the equpment was sujected to tests and passed all of them. Tests: Visul, Power on, Operational Check and Load Triggering in the Jauím etro (Simuator) GG00000000: Test resuts - OK - GG13: Calibrable tools - N/A SAP Transfer: Case Transferred to SAP Repair Action: Delivery made and service schedued for FSE Cifone on 00000000-12 ----------------------------------------------------------------------------- | 2023-10-21 21:07 | No
2639 | Yes | In u/No impact/uLLmessage on the equipment | In u/No impauuLLmessage on the equipment | BuQ&A Feedback: No Reason: Corrective Maintenance Trouleshooting Action: Qute sent Trouleshooting Action: OPP-19800000000 Qute: 00000000113200000000 ## Engenhariaclinica00abara.com.br Safety Qustions: Case: 000000002200000000397 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT Problem Reported by Cutomer: Equpment problem: Equpment uLLmessage Priority: STOPPED uit: FuDAção JOSE LuZ EGYDIO SETuAL Address: Avenida Angeliu87 Sao Pauo-SP 00000000227-200000000 Search for: Igor Silva Clinical engineering 000000001 3155 28900000000ngenhariaclinica00abara.com.br Euternal Remarks: 00000000.12.2000000003 00000000:15:36 POLARIS (POLARIS) ================= André Cifone Peiuoto - 000000000000000057600000000600000000000000002000000000000000000000000300000000=================== === Application of part of invoice 0000000039. After replacing the uuuitor, the equpment was sujected to tests and passed all of them. Tests: Visul, Power on, Operational Check and Load triggering in Jauím etro (Simuator) Euternal Remarks: 15.11.2000000003 00000000:00000000:29 POLARIS (POLARIS) ================= André Cifone Peiuoto - 000000000000000057600000000600000000000000001000000000000000000000000300000000=================== === When carrying ou the operational test, the equpment fails the defibrillation item and when firing with a load of 200000000J in the simuator, the resut is 171.5J. Both tests show the inefficiency of the uuuitor, therefore, for perfect operation of the equpment It will be necessary to replace this defective part. Internal Remarks: 00000000.12.2000000003 00000000:15:36 POLARIS (POLARIS) ======================== André Cifone Peiuoto - 000000000000000057600000000600000000000000002000000000000000000000000300000000=================== === Application of part of invoice 0000000039. After replacing the uuuitor, the equpment was sujected to tests and passed all of them. Tests: Visul, Power on, Operational Check and Load Triggering in the Jauím etro (Simuator) GG00000000: Test resuts - OK - GG13: Calibrable tools - N/A Internal Remarks: 15.11.2000000003 00000000:00000000:29 POLARIS (POLARIS) ========== André Cifone Peiuoto - 000000000000000057600000000600000000000000001000000000000000000000000300000000=================== === When carrying ou the operational test, the equpment fails the defibrillation item and when firing with a load of 200000000J in the simuator, the resut is 171.5J. Both tests show the inefficiency of the uuuitor, therefore, for perfect operation of the equpment It will be necessary to replace this defective part. GG00000000: Test resuts - NOK - GG13: Calibratable tools - N/A SAP Transfer: Case Transferred to SAP Trouleshooting Action: Schedued for FSE Cifone on 11/100000000AP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to sending a buget SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Repair Action: Delivery made and service schedued for FSE Cifone on 12/00000000 ----------------------------------------------------------------------------- | 2023-10-21 21:07 | No
2640 | Yes | Battery failed. (EC 0) - Final 00 | Battery failed. (EC 00) - Final 0/0 |  | 2023-10-21 21:12 | No
2641 | Yes | system down emergency cuomer on the table. | system down emergency cuomer on the table. | 000000000000000000000000001O##: 00000000000000000000000000100000000259651 21-Oct-2000000003 Fedorka Timothy The system meets the specification for the performed service and is retuned to ue Tools & Equpment : DOSE-KVP-LIGHT METER,2000000005-00000000-29 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: yes Eupected and ueupected Resuts: tue curent issu uer Impact: no Patient Impact: no SE 1.2.5 Resoluion : Performed tue adaption and tue yield also EDL calibration. ----------------------------------------------------------------------------- Internal Remarks : 0000000022000000006900000000 Fedorka Timothy 000000000000000000000000001king with FSE over weekend. FSE calling in on u7 M200000000where cutomer complaint is uGEN not available. FSE sent logs to the IGT mailbou. Logs show From Generator device: uSC: Tue Curent ou of range. uLL Asked FSE to try to readapt and calibrate tue. If not to check HiVo high tension connections for signs of arcing and possible at tue for signs of arcing. FSE was able to complete the tue adaptation with sucess. FSE is now testing system. Internal Remarks : **FSE TIM CALLED FOR IGT NSS, uFERED TO ON CALL (NEW PVC uDATE SENT CALL AuOMATICALLY WHEN SELECTED...WIERD) Internal Remarks : Interaction Type: Cutomer Call Back Caller: Nicole Jacabs IP & Issu: 000000002200000000600000000priorty one patient on table emergency Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Priority one/ASAP Entitlement Options: Entitled Reason for Entitlement Claim: P1 heart attack on the table Action Taken: set order u sent to RP Neut Step Action: N/A Verbal PO Nuber: cutomer provided a PO jut in case as well 10000000000000000000000002 if we wanted it Verbal PO Approver: na Verbal PO CAP: na ----------------------------------------------------------------------------- | 2023-10-21 22:03 | No
2642 | Yes | Device not charging, running on internal battery | Device not charging, running on internal battery | 000000000000000000000000001O##: 00000000000000000000000000100000000296816 00000000-Nov-2000000003 Reu, Timothy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Ouside of ue Eupected and ueupected Resuts: Not charging/operating on AC power. uLLcode 1100000000 uer Impact: None Patient Impact: None SE Diagnostic performed by Engineer : Replaced Power Management PCBA to correct the problem of device not charging/operating on AC power. Properly seated speaker cable on Motor Controller PCBA to correct the 1100000000 uLLcode. All tests passed, device is ready for patient ue. Parts Consued : 1.000000000000000000000000000000005615000000001,RP-PCBA,PWR MGMT,GEN00000000.300000000PIC, (FCO) Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 MuTIMETER,200000000000000001-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-21 23:18 | No
2643 | Yes | ISSSTE Contract - Damaged paddle connector | ISSSTE Contrau- Damaged paddle conneur | 000000000000000000000000001O##: 00000000000000000000000000100000000317197 Arrieta, Jorge The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: The connector of the paddles of the equpment is damaged, its lansue does not 000000000000000000000000001k properly. It needs to be replaced. Follow u Requred: Replacement of therapy cable (paddles) Problem Description by Engineer: Cutomer Fuction/Role: Maintenance Manager Device ue: in ue Eupected and ueupected Resuts: The paddle cable connector lay is damaged. uer Impact: None Patient Impact: None SE N/A ----------------------------------------------------------------------------- | 2023-10-22 03:02 | No
2644 | Yes | 0V0uLLAlarm：Prouimal pressure monitoring tube fall off | 0/V0uLLAlarm：Prouimal pressure monitoring tube fall off | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: Connection/Cable/Interconnect Repair Activity Code: Instruted Cutomer Problem Reported by Cutomer: 00000000/V600000000uLL SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Duing Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Alarm: Prouimal pressue monitoring tue fall off WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Checked and foud that the prouimal pressue monitoring tue was disconnected from patient mask, Iinstrut the cutomer to reconnect it, and then the alarm disappeared and device retun to ful fuctionality Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: NA Test/Inspection: Power on test and visul test are passed, device retun to ful fuctionality Test Equpment: NA Safety Qustions: Case: 000000002200000000000000008 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-22 09:41 | No
2645 | Yes | Fluoroscopy is not functioning | Fluoroscopy is not funuoning | 000000000000000000000000001O##: 0000000000000000000000000010000000026000000001 Vijayakuar, K.S The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Fluroscopy is not fuctioning When was the first occurence? Within... : 22-Oct-2000000003 18:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Fluroscopy is not fuctioning uer Impact: No uer impact Patient Impact: No patient impact SE 000Diagnostic performed by Engineer : Generator power was tripped. The same was released and check the system - 000000000000000000000000001king fine. Internal Comments : uLL# and/or description of uLL: Fluroscopy is not fuctioning Repair Action : Generator power was tripped. The same was released and check the system - 000000000000000000000000001king fine. Resoluion : Generator power was tripped. The same was released and check the system - 000000000000000000000000001king fine. ----------------------------------------------------------------------------- | 2023-10-22 18:26 | No
2646 | Yes | In an euamination ung the breast coil, there was a crack in the ribs - CFB | In an euamination ung the breast coil, there was a crack in the ribs - CFB | 000000000000000000000000001O##: 00000000000000000000000000100000000261330000000023-Oct-2000000003 Michishita, Yoshitaka The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: In an euamination uing the breath coil, there was a spine in the ribs Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: In an euamination uing the breath coil, there was a spine in the ribs uer Impact: No effect Patient Impact: No effect SE uknow Resoluion: 000000000000000000000000001k details: I conduted a hearing abou the problem that occured this time. I will report the details that were received to the head office. I checked the operation and there was no problem with the equpment or the breath coil. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000261313 Akashi, Kazua The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: In an euamination uing the breath coil, there was a spine in the ribs Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: In an euamination uing the breath coil, there was a spine in the ribs uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: The power was from FSE Michishita. Case creation Follow u Requred: * ----------------------------------------------------------------------------- | 2023-10-23 06:11 | Yes
2647 | Yes | Not starting up | Not starting up | 000000000000000000000000001O##: 00000000000000000000000000100000000261398 23-Oct-2000000003 Sturt RuSELL The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: rad Nichole Device ue: normal startu post weekend Eupected and ueupected Resuts: System cout down to 00000000.00000000 and will not boot any futher. uer Impact: uable to do patients, system down Patient Impact: no patients possible SE 000Resoluion : Fitted new FleuVision PC after swapping the HardDrive over, and tested. All aspects of the system were tested from Images and layous to uray to Geometry. another restart and another test carried ou. System 000000000000000000000000001king perfectly and handed back to staff for ue. Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Resoluion : Ordered new FleuVision PC, initial indications were that part was coming from Sydney. Advised cutomer, bu then 000000000000000000000000001d got back that there was the part in Brisbane. logistics arranged for same day couier to get part to site which arrived aroud 3PM. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002611300000000RSE_Sahdia Husain The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : uLL# and/or description of uLL: NA Repair Action : FSE to attend site: 1. Please ue the PC diagnostic tool to see if any issus noted with the PC. 2. Replace the fleuvision PC if requred Technician Name : RSE_Sahdia Husain Diagnostic performed by Engineer : Discused and confirmed the issu with the cutomer. The cutomer stated that when they started u the system this morning, it 000000000000000000000000001ld not start u, the monitor in the euam room is completely blank with the green LED lit. The cutomer stated when they started u the system, it went throuh the coutdown and was stuk on start u at 0000000000000000. The cutomer then carried ou a hard shu down via the circut breaker and when the system started u again, it got stuk again at 0000000000000000. We are uable to access log files as the system is down, however, I supect the issu is with the Fleuvision PC and we will requre an engineer to attend to site to carry ou futher trouleshooting. Problem Description by Engineer : Problem description by engineer : Not starting u When was the first occurence? Within... : 23-Oct-2000000003 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Internal Comments : uLL# and/or description of uLL: NA Repair Action : FSE to attend site: 1. Please ue the PC diagnostic tool to see if any issus noted with the PC. 2. Replace the fleuvision PC if requred ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000026110000000023-Oct-2000000003 Sturt RuSELL The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Not starting u When was the first occurence? Within... : 23-Oct-2000000003 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Logs show the FleuVision PC video cards are failing and not starting the PC. New PC being ordered and delivered to site tomorrow. Follow u Requred : To replace the FV PC, load SW and test ----------------------------------------------------------------------------- | 2023-10-23 09:01 | No
2648 | Yes | RR uray system cannot turn on. | RR uray system cannot turn on. | 000000000000000000000000001O##: 00000000000000000000000000100000000262355 SILA TuRuA The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : uray system cannot tun on. Engineer need to check this problem on site. Problem Description by Engineer : Cutomer Fuction/Role: RadTech Device ue: Ou of ue Eupected and ueupected Resuts: uray system cannot tun on. uer Impact: uray system cannot tun on. Patient Impact: not impact. SE R7 ----------------------------------------------------------------------------- | 2023-10-23 12:36 | No
2649 | Yes | uLLmessages during the drive up | uLLmessages during the drive up | 000000000000000000000000001O##: 0000000000000000000000000010000000027000000009 Oct 25, 2000000003 Maris, Ezra The system meets the specification for the performed service and is retuned to ue Tools & Equpment : DOSE-KVP-LIGHT METER, 2000000005-00000000-26 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-26 00000000:00000000:00000000 Problem Description by Engineer: Problem description by engineer: RR Auel Kurau Phone 0000000000000000121900000000600000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion: Host PC replaced and software installed ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000262376 Oct 2000000002000000003 Maris, Ezra The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: RR Auel Kurau Phone 0000000000000000121900000000600000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer: Host PC replaced and software reinstalled Follow-u requred: Finish repair Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000262100000000Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: RR Auel Kurau Phone 0000000000000000121900000000600000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer: System hangs duing start-u Internal Comments: uLL# and/or description of uLL According to Bios, hard drive is not foud With u Kurau hard drive was installed and installed. Now the Bios rus throuh - now an uLLoccus when loading software Curently installed is a: Host PC Monoplane Revised no HDD 00000000800000000251781 It is curently not available. When changing to Revised II compuer, the reinstallation of the complete system is requred. Repair Action: Check start-u sequnce Replace compuer or hard drive if necessary Follow-u requred: uLL# and/or description of uLL According to Bios, hard drive is not foud With u Kurau hard drive was installed and installed. Now the Bios rus throuh - now an uLLoccus when loading software Curently installed is a: Host PC Monoplane Revised no HDD 00000000800000000251781 It is curently not available. When changing to Revised II compuer, the reinstallation of the complete system is requred. Repair Action: Check start-u sequnce Replace compuer or hard drive if necessary Technician name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Internal Remarks: Curently installed is a: Host PC Monoplane Revised no HDD 00000000800000000251781 It is curently not available. When changing to Revised II compuer, the reinstallation of the complete system is requred. Please clarify again with FSE Cutomer comment: ## Appt. arranged with u Kuraufor 20000000000000000 from approu. 9:00000000/10000000000000000 am. ----------------------------------------------------------------------------- | 2023-10-23 12:41 | No
2650 | Yes | System is not starting up successfully | System is not starting up successfully | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000358 31-Oct-2000000003 DAPRATI, DAVIDE The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Confirmed the reported problem. Log file check. Post uLL were foud on IPC image processor. Disassembled and replaced IPC. Complete software loading. Fuctional tests were carried ou with a positive oucome. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000262581 31-Oct-2000000003 MARCO GIuSANI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: At the first start-u the system does not start correctly. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.1.3000000005 Resoluion: Foud reported problem. Carried ou faut search. Replaced IP-PC and reloaded sw. Carried ou tests and fuctional checks with PASS oucome. Parts Consued: 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD Tools & Equpment: MuTIMETER, 20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002622300000000SAVERIO DIVITTORIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: At the first start-u the system does not start correctly. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer: Analyzed the log files, foud that the IMAGE MANAGEMENT PC is not on at the first start-u of the system. u. Giuia Muarò was talked to. I advised her not to tun the angiograph off util ou intervention. Internal Comments: uLL# and/or description of uLL POST "Frontal IP-PC" Not Done POST "Frontal IP-PC OS disk 1" Not Done POST "Frontal IP-PC Image disk 1" Not Done POST "Frontal IP-PC Image disk 2" Not Done POST "Frontal IP-PC FRuFan 1" Not Done POST "Frontal IP-PC FRuFan 2" Not Done Repair Action: . Resoluion: An intervention will follow to repair the faut. Follow u Requred: uLL# and/or description of uLL POST "Frontal IP-PC" Not Done POST "Frontal IP-PC OS disk 1" Not Done POST "Frontal IP-PC Image disk 1" Not Done POST "Frontal IP-PC Image disk 2" Not Done POST "Frontal IP-PC FRuFan 1" Not Done POST "Frontal IP-PC FRuFan 2" Not Done Repair Action: . What are the neut Trouleshooting steps: The logs show that at the first start-u the IP-PC does not start. It is advised to replace the backu battery of BIOS (CR2000000002), however I order the PC in case the defect is the power suply, evalute whether to replace it or not. On November 6 pm is schedued (00000000000000000000000000100000000835182), evalute with the cutomer the possibility to do both interventions together. System is not down, I sugested to the cutomer to leave it on util ou intervention. Technician Name: SAVERIO DIVITTORIO ----------------------------------------------------------------------------- | 2023-10-23 13:11 | No
2651 | Yes | System did not go on | System did not go on | 000000000000000000000000001O##: 00000000000000000000000000100000000262575 Oct 2000000002000000003 Pugs, Volker The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The problem was confirmed. Consutation held with u Steigmüller at +0000000033170000000076900000000uber. How often is issu occuring? : First occurence Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion: uS and battery pack replaced, fuction test of the system Parts Consued: 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP Tools & Equpment: MuTIMETER, 200000000000000006-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000262221 Malhas, Ali The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: uLL# and/or description of uLL . Repair Action: . Problem Description by Engineer: Problem description by engineer: The problem was confirmed. Consutation held with u Steigmüller at +0000000033170000000076900000000uber. How often is issu occuring? : First occurence Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: M cabinet checked. uLLmessage can be seen on the 1-phase uS and is lit red = 1-phase no longer available. After several system restart attempts, the system coud boot, bu not completely. Probably 2000000002.5V power suply defective. Onsite check requred. Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: 2 services! If possible, please schedue an FSE today withou a part with 2 hous of 000000000000000000000000001k time as a bypass 1-phase uS and check on site. Check of 2000000002.5V power suply in M cabinet also requred - > temporary soluion util the 2u spare parts arrive for the 2nd service with 2 hous of 000000000000000000000000001k. Technician name: Malhas, Ali ----------------------------------------------------------------------------- Cutomer comment: ## Service with u Steigmüller for 100000000200000000fter 10000000000000000 am. According to Remote, the service will take place in the engineering room, the euamination room can be ued elsewhere util final inspection. Informed u Pugs on bou. Internal Remarks: 00CC: Appointment consutation please with u Steigmüller Medical Technology 000000000000000033170000000076900000000utomer comment: u Steigmüller asks for info whether the technicians have to go into the room. Requst for contact 000000000000000033170000000076900000000nternal Remarks: u Steigmüller Medical Technology is not at the system. Asked for a 9:00000000 hrs. to be requsted. ----------------------------------------------------------------------------- | 2023-10-23 13:11 | No
2652 | Yes | Message at start-up. | Message at start-up. | 000000000000000000000000001O##: 00000000000000000000000000100000000262600000000October 2000000002000000003 Ajeeth MARIATHAS The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Engineer Device ue: no Eupected and ueupected Resuts: System start-u problem, the screens tun on bu the software does not start. uer Impact: system not uable. Patient Impact: no SE V1.2.5 Diagnostic performed by Engineer: The Pc Sute does not start anymore, with the hard disk not detected. Resoluion: Replaced the hard disk and reinstalled the software. Parts Consued: 1.00000000000000000000000000000000800000000660000000083, Following PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000262100000000 Lionel GIROu The system does not meet the specifications for the performed service and is not retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Engineer Device ue: no Eupected and ueupected Resuts: System start-u problem, the screens tun on bu the software does not start. uer Impact: system not uable. Patient Impact: no SE V1.2.5 Diagnostic performed by Engineer: System start-u problem, the screens tun on bu the software does not start. Diagnosis with the help of u Bergerat. The problem comes from the PC sute, the hard disk is no longer recognized on the disk rack. Ordered pc sute and intervention by Ajeeth Mariathas on 100000000200000000000000003. Follow u Requred: CDP: 00000000800000000660000000083 Following PC Quntity: 1 Direct delivery to the cutomer. Biomedical department u Bergerat +3300000000618956 ----------------------------------------------------------------------------- | 2023-10-23 13:16 | No
2653 | Yes | SA:0/u FD0start failure | SA:0R/u FD0start failure | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 23.1000000002000000003 12:30000000013 uC (8600000000600000000) $Q7: (Confirmation & Description of problem/start failue) $Q8: (cutomer report device start failue, check and foud fleuvision PC not startu) $Q9: (replace PC BIOS battery and system startu ok, test system fuction ok then system OK) $Q100000000 (Test system ru ok and inspection data refer to: PA. Device retuned to ful fuctionality. ) $Q11: N/A SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: u uuuFD200000000T2 Activities:Problem Description: Overview of cutomer problems: Start failue Trouleshooting Action: Key faut codes or phenomena: Boot stuk at 00000000 PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC startu is normal, clean the gold finger and start the compuer to observe Recommended spare parts uuand designation: 0000000080000000012200000000 SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000512 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 23.1000000002000000003 12:300000000100000000uC (8600000000600000000) Service content (requred): Equpment startu failue, check Fleuvision The PC did not start. After replacing the motherboard BIOS battery, the test was normal and the machine retuned to normal ue. Service resuts (requred): [After repair, the equpment retuned to normal operation.] Other important matters: [none] ----------------------------------------------------------------------------- | 2023-10-23 13:41 | No
2654 | Yes | CV*system down* SW issue | CV*system down* SW issue | 000000000000000000000000001O##: 00000000000000000000000000100000000262733 200000000ct-2000000003 Saif, Islam The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer : system is not starting u How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer: Sute pc not runing Follow u Requred: Spare part will be ordered Diagnostic performed by Engineer: Sute pc will be ordered ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000272389 25-Nov-2000000003 Saif, Islam The system meets the specification for the service performed and is retuned to ue Diagnostic performed by Engineer: this 000000000000000000000000001assigned to me only to retun part to local warehoue as uued, no interaction done with the system Problem Description by Engineer: Problem description by engineer : system is not starting u How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000262596 Ahmed Kotb The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer : system is not starting u How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer: system is not starting u - it seems that the system software is corruted. Onsite visit is needed. Internal Comments: uLL# and/or description of uLL: system is not starting u Repair Action : system is not starting u - it seems that the system software is corruted. Onsite visit is needed. Resoluion: The system is not starting u - it seems that the system software is corruted. Onsite visit is needed. Follow u Requred: uLL# and/or description of uLL: system is not starting u Repair Action : system is not starting u - it seems that the system software is corruted. Onsite visit is needed. What are the neut Trouleshooting steps: system is not starting u - it seems that the system software is corruted. Onsite visit is needed. check and reload the sute PC sw... do a check on power oupus. Technician Name: Ahmed Kotb ----------------------------------------------------------------------------- | 2023-10-23 13:41 | No
2655 | Yes | Fm: Geometry movement only possible to a limited eutent | Fm: Geometry movement only possible to a limited eutent | 000000000000000000000000001O##: 000000000000000000000000001000000002628300000000ct 2000000002000000003 Biebl, Josef The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: RR Marku Schreiber Phone 000000000000000071000000007500000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 000Resoluion: Geo PC replaced. Software installation. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000931363,GEO IPC Ivybridge with ZMP CAN and IO ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000262291 Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: RR Marku Schreiber Phone 000000000000000071000000007500000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: Geomology failed Internal Comments: uLL# and/or description of uLL According to log file: Application uLL uable to commuicate with GeoIPC. SID is uknown and no movements are available. u Schreiber restarted the Geo compuer manully - withou sucess According to his statement, the switch is OK Repair Action: Trouleshooting on site If necessary, replace Geo compuer Follow-u requred: uLL# and/or description of uLL After compuer replacement on Friday, the system is curently not reachable remotely - RSN admin informed u Schreiber looked at the R cabinet. Switch is runing Compuer has power on the SCC1, the LED of post H1 lights u. B On supicion SCC! also ordered - not saved Repair Action: Trouleshooting on site Technician name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Internal Remarks: u Biebl is to be contacted Everything remains as schedued Internal Remarks: 00emote: Please call u Biebl. Cutomer comment: ## Service with u Schreiber and Nuse Kathrin for 100000000200000000fter 9:00000000 am. ----------------------------------------------------------------------------- | 2023-10-23 14:01 | No
2656 | Yes | IT DOES NOT LOAD THE SYSTEM. | IT DOES NOT LOAD THE SYSTEM. | 000000000000000000000000001O##: 0000000000000000000000000010000000026000000007 200000000ct-2000000003 MIGLIACCIO, NICOLA The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: technician Device ue: normal Eupected and ueupected Resuts: it doesn't boot uer Impact: system down Patient Impact: nothing SE 8.2.17 Resoluion: Host PC replacement, license reinstallation, System configuation, fuctional checks sucessfuly completed. Tools & Equpment: MuTIMETER, 2000000005-00000000-00000000 00000000:00000000:00000000 Diagnostic performed by Engineer: Intermittent boot problem on Host PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000263121 200000000ct-2000000003 MARCO GIuSANI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: Software loading is not complete. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: start u failed uer Impact: no Patient Impact: no SE 8.2.17.00000000Resoluion: Problem: It does not load the sw. Performed faut search, replaced Host PC and reconfigued parameters. Carried ou tests and tests with a PASS oucome. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000, Host PC Monoplane Revised_II no HDD Tools & Equpment: MuTIMETER, 20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000262831 ANDREA COLOMBO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: Software loading is not complete. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: start u failed uer Impact: no Patient Impact: no SE 8.2.17.00000000Diagnostic performed by Engineer: confirmed by the client u Pirandelli, head nuse, the angiograph does not complete the start-u. System is down, intervention will follow. Internal Comments: uLL# and/or description of uLL The cutomer tried to restart the system, duing the 00discusion, asked for another restart. the monitors in the control room tun on bu the screen is always black, then no signal appears and then they go to st.by. while the live monitor in the euam room shows the first bios activity screen, then after a moment of black screen, the u message appears, bu shortly after the signal disappears and the screen remains black. The geometry does not 000000000000000000000000001k. remotely the system is offline. Repair Action: Sugested onsite intervention for diagnosis (attempt to restore) with parts (1uhdd host PC), 1 technician 1 day Follow u Requred: uLL# and/or description of uLL The cutomer tried to restart the system, duing the 00discusion, asked for another restart. the monitors in the control room tun on bu the screen is always black, then no signal appears and then they go to st.by. while the live monitor in the euam room shows the first bios activity screen, then after a moment of black screen, the u message appears, bu shortly after the signal disappears and the screen remains black. The geometry does not 000000000000000000000000001k. remotely the system is offline. Repair Action: Sugested onsite intervention for diagnosis (attempt to restore) with parts (1uhdd host PC), 1 technician 1 day Technician Name: ANDREA COLOMBO ----------------------------------------------------------------------------- Cutomer comment: 000000003000000000000000000000000CrrQtQAJ ----------------------------------------------------------------------------- | 2023-10-23 14:36 | No
2657 | Yes | two patients report pressure sore from the mattress (patient 0 | two patients report pressure sore from the mattress (patient 0 | 000000000000000000000000001O##: 00000000000000000000000000100000000239797 Barry White The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: t000000000000000000000000001patients report pressue sore from the mattress uer Impact: NA Patient Impact: NA SEu Diagnostic performed by Engineer : Mattress not damaged and confirmed no thicker alternate available. Staff to monitor ----------------------------------------------------------------------------- Internal Remarks : Called , no answer . Cutomer has not yet called my mobile nuber Internal Remarks : Will call back 19/100000000and speak to Rachel Hathaway , 0000000017 00000000721 Internal Remarks : Called again , still no answer. Internal Remarks : Rachel to call back on my mobile nuber ----------------------------------------------------------------------------- | 2023-10-23 14:46 | Yes
2658 | Yes | [Continuation of case 000Critical uLLis displayed (No issue) | [Continuation of case 00000 Critical uLLis displayed (No issue) | 000000000000000000000000001O##: 00000000000000000000000000100000000263200000000Hübler Christina The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: no Eupected and ueupected Resuts: Critical uLLis displayed uer Impact: no Patient Impact: no SE Resoluion: duing the inspection of the device, the uer is displayed: "critical uLL, the medical technician cannot fiu this display and will send the device to ou 000000000000000000000000001kshop, problems are detected when loading the new software, software logs are euamined with the help of the Bu continu reading in the primary case 0000000022172161/escalation with investigation in the Bu ------------------- A software bu confirmed, the root caue investigation is thrive. An RFS ticket has been sumitted to requst problem resolving. Item ID: 0000000000000000581 Short Description: DFM100000000 2.00000000.32 sw bu - "ugradeAssitant" malfuction ---------------------------- Device was sent back to the cutomer after software udate and check in bench: see 0000000022172161 ----------------------------------------------------------------------------- | 2023-10-23 14:51 | No
2659 | Yes | It does not start. | It does not start. | 000000000000000000000000001O##: 000000000000000000000000001000000002626900000000drián Sancho López The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The uers say that the equpment did not start, bu when trying with them on the phone it starts normally. The event log of the system is checked and a normal start-u is checked. Agreed to close the notice and they will call back if they notice any anomalies in the operation. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment did not start. uer Impact: No impact Patient Impact: No impact SE 2.2.6.00000000Resoluion: Agreed to close the notice and that they will call back if they notice any anomalies in the operation. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The uers say that the equpment did not start, bu when trying with them on the phone it starts normally. The event log of the system is checked and a normal start-u is checked. Agreed to close the notice and they will call back if they notice any anomalies in the operation. Diagnostic performed by Engineer: The uers say that the equpment did not start, bu when trying with them on the phone it starts normally. The event log of the system is checked and a normal start-u is checked. ----------------------------------------------------------------------------- | 2023-10-23 15:01 | No
2660 | Yes | It does not start | It does not start | 000000000000000000000000001O##: 00000000000000000000000000100000000263861 15-Nov-2000000003 FAuTO FERRARO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: The sw does not start at the first start-u. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.1.3000000005 Resoluion: Carried ou replacement of the spare part previouly ordered and reinstalled software and all the necessary configuations; completed diagnostic and fuctional tests with a positive oucome. Parts Consued: 1.00000000000000000000000000000000800000000122391, u Haswell Monoplane Host PC no HDD 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part Tools & Equpment: ELECTRICAL SAFETY TESTER, 2000000000000000000000000000000000000000000000000:00000000 MuTIMETER, 200000000000000003-28 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000263812 SAVERIO DIVITTORIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Internal Comments: uLL# and/or description of uLL . Repair Action: . Resoluion: Intervention will follow for replacement of the PC. Problem Description by Engineer: Problem description by engineer: The sw does not start at the first start-u. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer: Ou technician Ferraro is already present for another activity on the system, he reports an occasional problem with tuning on the PC that manages the angiograph. Follow u Requred: uLL# and/or description of uLL . Repair Action: . What are the neut Trouleshooting steps: Fauto Ferraro is on site for another intervention. He foud that the Host PC sometimes does not tun on when the system is tuned on. It is suficient to tun it on from the front buton of the PC to solve the problem. Most likely the batt. backu is discharged, bu we are forced to replace all the PC that I order. It is important to specify in the internal notes of the CSR that we are changing the PC for the backu battery that is flat. The system is not down. Technician Name: SAVERIO DIVITTORIO ----------------------------------------------------------------------------- | 2023-10-23 16:41 | No
2661 | Yes | Patient euuuenced desat and bradycardia but alarms didn’t sound. No Trouble Found-Evaluation Confirmed Device Working. | Patient euuuenced desat and bradycardia but alarms didn’t sound. No Trouble Found-Evaluation Confirmed Device Working. | 000000000000000000000000001O##: 00000000000000000000000000100000000263883 Marco Vianello The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer: Allegation of missing Desat (SpO2) alarm: There are several "SpO2 sensor off" INOPs aroud the time of the alleged event (19-OCT-2000000003 aroud 13:300000000: 13:25:00000000 SpO2 Sensor off ended 13:25:00000000 SpO2 Sensor off generated at 13:25:00000000 13:25:25 SpO2 sensor off ended 13:26:33 SpO2 sensor off generated at 13:26:25 13:26:35 SpO2 sensor off ended 13:26:51 SpO2 sensor of generated at 13:26:0000000013:26:57 SpO2 sensor off ended 13:28:31 acknowledge 13:32:55 SpO2 sensor off generated 13;32:00000000Generally, if an INOP interruts monitoring and alarm detection, the measuement nueric will be replaced by a "?". When an "SpO2 sensor off" INOP is indicated, an SpO2 measuement is not possible, the SpO2 nueric is replaced by a "?" and an INOP tone is issud. The presence of an "SpO2 sensor off" INOP euplains, why there were no "DESAT" alarms aroud 13:300000000on 19-OCT-2000000003: An SpO2 measuement was not possible at this time. Futhermore, it shoud be noted that a DESAT alarm is only issud after the DESAT alarm delay has eupired. The Desat alarm delay defines the amout of time (in the cutomers case 200000000sec, see bed5.cfg) that the averaged SpO2 valu needs to be below the DESAT alarm limit (in this case 800000000 see bed5.cfg) before a DESAT alarm is activated. One single valu above the DESAT limit (in this case 800000000 see bed5.cfg) 000000000000000000000000001ld trigger the DESAT alarm delay to start again and inhibit the DESAT alarm. In the translated log information, there is no indication that the SpO2 alarm was switched off at the time of the event. However, this might have happened before or after the event. The Alarm Review Window of the IntelliVu Patient Monitors has only a limited nuber of entries and is overwritten with new data once the limit is reached. Therefore, the entries for the alleged time were no longer present in the Alarm Review Window when it was checked at a later time. Allegation of missing Brady (u alarm: There are mutiple ualarms listed aroud the time of the alleged event (19-OCT-2000000003 aroud 13:300000000, patient alarms incluing "eutreme brady" as well as the INOP "Cannot analyze u". 13:25:00000000 * HR 00000000500000000generated at 13:2000000002 13:26:15 HR 00000000500000000ended 13:26:15 ***eutreme brady 38<00000000generated at 13:26:00000000 13:26:15 red alarm soud annouced 13:26:31 ***eutreme brady 300000000 00000000ended 13:26:31 ***asystole generated at 13:26:23 13:26:58 Cannot analyze ugenerated at 13:26:50000000013:27:12 ***asystole ended 13:27:12 ***Vent Fib / Tach generated at 13:27:0000000013:28:00000000 Cannot analyze uended 13:28:11 Vent Fib / Tach ended 13:28:11 non-sutain VT generated at 13:28:00000000 13:28:11 yellow alarm soud annouced 13:28:13 non-sutain VT ended 13:28:13 *** VTach generated at 13:28:0000000013:28:13 red alarm soud annouced 13:28:13 31 acknowledge 13:28:32 ***VTach ended 13:28:32 yellow alarm soud annouced 13:28:33 ***VTach generated at 13:28:25 13:28:33 red alarm soud annouced 13:28:39 ***VTach ended 13:28:39 yellow alarm soud annouced 13:28:00000000*HR 127>1200000000generated 13:28:500000000 HR 1000000001200000000ended 13:28:500000000**eutreme tachy 100000000100000000generated at 13:28:00000000The allegation of failue to alarm for "brady" (ualarms) is contradicted by the alarms logged aroud the time of the event. There is one "eutreme brady" alarm at 13:26 and afterwards several tachycardia alarms. Allegation of missing alarm souds It was stated in the reply to Action Response 2: "Allegation of failue to alarm was made by ICudoctor who was passing by the monitor. She said she missed auible alarms (SpO2 desat and asystole) at the monitor. When I asked abou visul alarms she replayed "no visul alarms too". Then I asked abou alarms at PICiu. She replayed that she didn't hear alarms at PICiu." As far as I uderstood the statement, the ICudoctor was neut to the monitor, when she made the allegation of missing auible alarms. It cannot be determined retrospectively what the alarm volue at the monitor was at the time of the event and if the respective alarm volue was lou enouh for the environment. The monitor configuation shows that the minimu allowed alarm volue is 00000000see bed5.cfg). Futhermore, it can be seen in the configuation, that for some profile/monitor settings both, "visul latching" and "auible latching" are set to "Off". This means, if the alarms had been acknowledged, the alarm lamps and the alarm tone 000000000000000000000000001ld be switched off, even if the alarm condition was still present. At 13:28 there is an acknowledgment logged in the PICiu log: 13:28:31 acknowledge Also, in the PICiu logs there is evidence that the central station issud auible alarms: 13:26:15 red alarm soud annouced 13:28:11 yellow alarm soud annouced 13:28:31 acknowledge 13:28:33 red alarm soud annouced 13:28:39 yellow alarm soud annouced Concluion All in all, the analysis of the data does not reveal any produt malfuction. Please provide the cutomer with the investigation resuts in an appropriate manner. Please let me know if there are any futher qustions regarding this event. If I do not hear from youby 28 November 2000000003, I will close the escalation action requst. ----------------------------------------------------------------------------- | 2023-10-23 16:46 | No
2662 | Yes | The equipment does not complete start-up. | The equipment does not complete start-up. | 000000000000000000000000001O##: 00000000000000000000000000100000000263926 October 25, 2000000003 Alberto Mateos Chacón The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce, so it is not possible to access remotely to check uLL; We need to go and check it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not complete the start-u sequnce. uer Impact: The equpment does not start. Patient Impact: No futher information is available as per the cutomer. SE 8.1.3000000005 Resoluion: I know ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000263771 Adrián Sancho López The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce, so it is not possible to access remotely to check uLL; We need to go and check it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not complete the start-u sequnce. uer Impact: The equpment does not start. Patient Impact: No futher information is available as per the cutomer. SE 8.1.3000000005 Diagnostic performed by Engineer: The equpment does not finish the start-u sequnce, so it cannot be accessed remotely to check uLL; We need to go and check it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The equpment does not finish the start-u sequnce, so it is not possible to access remotely to check uLL; We need to go and check it. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: The equpment does not finish the start-u sequnce, so it is not possible to access remotely to check uLL; We need to go and check it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-10-23 16:51 | No
2663 | Yes | It is giving uEquipment Failure. | It is giving uEquipment Failure. | 000000000000000000000000001O##: 0000000000000000000000000010000000026390000000000000000-Nov-2000000003 Buak Tezel The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Not ue. Eupected and ueupected Resuts: uEqupment Failue. uer Impact: No. Patient Impact: No. SE 2.00000000.32 Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000000000000000000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000003-16 00000000:00000000:00000000 MuTIMETER, 200000000000000002-16 00000000:00000000:00000000 Parts: *1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor 1.0000000000000000000000000000000056000000009131,DFM100000000 Measuement Modue uResoluion: The defective parts of the device were replaced with new ones. All tests and checks were performed in according to the procedue listed in the service docuent. No problem was seen. The device was delivered to the uer in 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000263889 Buak Tezel The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: The checks showed that the Processor PCA and uboard of the device were defective. According to the failue detection, 0000000056000000009000000001,0000000056000000009131 - It shoud be ordered. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Not ue. Eupected and ueupected Resuts: uEqupment Failue. uer Impact: No. Patient Impact: No. SE 2.00000000.32 Resoluion: Duing the checks, it was detected that the Processor PCA and uboard of the device were defective. According to the failue detection, 0000000056000000009000000001,0000000056000000009131 - It shoud be ordered. ----------------------------------------------------------------------------- | 2023-10-23 17:01 | No
2664 | Yes | System movements not working | System movements not working | 000000000000000000000000001O##: 00000000000000000000000000100000000263358 TSCHuTZ NIKLAS The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow-u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: na Device ue: Eupected and ueupected Resuts: uer Impact: Na Patient Impact: na SE Follow-u Requred: Order new GEO PC while this issu has occured 2 times in the last 2 weeks. Diagnostic performed by Engineer: Checked logs and GEO not available Resoluion: Restarted the GEO PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000263937 Parts Consued: 1.00000000000000000000000000000000800000000931373,GEO IPC Ivybridge no ZMP with CAN and IO ----------------------------------------------------------------------------- | 2023-10-23 17:06 | No
2665 | Yes | It is not possible to move the arc. | It is not possible to move the arc. | 000000000000000000000000001O##: 0000000000000000000000000010000000026000000000000000000000000ct-2000000003 MARIANO GENNARI The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.00000000000000000000000000000000800000000262813, CLEA2 Force Sensor Resoluion: Replaced and calibrated sensor force of the movement of the sensor. Fuctional tests were normal. The system is 000000000000000000000000001king. Problem Description by Engineer: Problem description by engineer: Arc movement blocked. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: collision detected ids, rot, prop, long uer Impact: no Patient Impact: no SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000263325 ANDREA COLOMBO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: Arc movement blocked. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: collision detected ids, rot, prop, long uer Impact: no Patient Impact: no SE 1.2.5.00000000Diagnostic performed by Engineer: confirmed by the cutomer 338900000000600000000Sandro Fantini, arc movements are not available. Internal Comments: uLL# and/or description of uLL The arc movements do not 000000000000000000000000001k, eucept for the rotation of the "L" in vascuar position, also the table does not respond to the u/down movement, while the release/block of the horizontal brakes 000000000000000000000000001ks. the system is however in ue, euamination uderway. The cutomer tried to tun the system on many times. From today's log files, there are variou collision alarms of almost all the movements (sid, rot, prop, long), in RSW in the history there are other similar episodes in the previou three months. Yesterday's log file shows an ad7 collision force sensor alarm, which does not appear today. The cutomer will call me back at the end of the euamination in order to try the bodygurd calibration (it was sucessfu), bu however the movements of the arc are blocked. sugested intervention with parts for diagnosis (ad7 height potmente, force sensor ad7 and clea2 force sensor) Repair Action: Schedue 1 technician 1 day system "almost down" Follow u Requred: uLL# and/or description of uLL The arc movements do not 000000000000000000000000001k, eucept for the rotation of the "L" in vascuar position, also the table does not respond to the u/down movement, while the release/block of the horizontal brakes 000000000000000000000000001ks. the system is however in ue, euamination uderway. The cutomer tried to tun the system on many times. From today's log files, there are variou collision alarms of almost all the movements (sid, rot, prop, long), in RSW in the history there are other similar episodes in the previou three months. Yesterday's log file shows an ad7 collision force sensor alarm, which does not appear today. The cutomer will call me back at the end of the euamination in order to try the bodygurd calibration (it was sucessfu), bu however the movements of the arc are blocked. sugested intervention with parts for diagnosis (ad7 height potmente, force sensor ad7 and clea2 force sensor) Repair Action: Schedue 1 technician 1 day system "almost down" What are the neut Trouleshooting steps: Schedue 1 technician 1 day system "almost down" Technician Name: ANDREA COLOMBO ----------------------------------------------------------------------------- | 2023-10-23 17:41 | No
2666 | Yes | Audio & Sounds/Loudspeaker problem Wants to receive a quote./Speaker | Audio & Sounds/Loudspeaker problem Wants to receive a quote./Speaker | 000000000000000000000000001O##: 0000000000000000000000000010000000026000000007 MESNIL, Mauime The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Qute requst uer Impact: No. Patient Impact: No. SE N/A Diagnostic performed by Engineer: The louspeaker is defective, the cutomer wants to receive a qute to order a new louspeaker. Resoluion: Qute requst sent to the qutation department. REF: 00000000261000000009000000001 Speaker Assembly QTY: 1 Follow u Requred: Qute to be issud. QR on 10000000023 at 10000000000000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000026000000007 Parts Consued: 1.00000000000000000000000000000000261000000009000000001,IV-MP5 CBL Speaker Assembly ----------------------------------------------------------------------------- | 2023-10-23 18:11 | No
2667 | Yes | The application remains uninitialized at 0:0 - SJU Brasov | The application remains uninitialized at 0:0 - SJU Brasov | 000000000000000000000000001O##: 000000000000000000000000001000000002638700000000Neagu Cristian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: operator Device ue: no Eupected and ueupected Resuts: app failed duing initialization uer impact: no Patient Impact: no SE 8.2.17 Diagnostic performed by Engineer: following the remote checks with the operator, it is foud that the FleuVision uit is not initialized. The compuer is restarted manully, the device retuning to normal parameters. It is recommended to replace the bios battery. Resoluion: The system is in 000000000000000000000000001king condition. If the uLLreappears, restart the FleuVison calcuator manully. The BIOS battery will be replaced. ----------------------------------------------------------------------------- | 2023-10-23 18:31 | No
2668 | Yes | Alarm went off, checked the indicator and it showed a flashing red u. | Alarm went off, checked the indicator and it showed a flashing red u. |  | 2023-10-23 18:53 | No
2669 | Yes | FRu: Cuomer called stating the device uLLplug in connector | FRu: Cuomer called stating the device uLLplug in conneur | 000000000000000000000000001O##: 00000000000000000000000000100000000265153 Simpson,Kreshia The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED chirping with uLLplu in connector uer Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Cutomer reported the AED chirping with uLLplu in connector Diagnostic performed by Engineer : Cutomer reported that the AED keeps saying device uLLplu in connector. Instruted cutomer to uplu pad remove battery, then we pluged the pads back in and reinsert the battery and same uLLcode. Cutomer stated that they tried different pads bu same uLLmessage. Replacing AED. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. AED Serial# B16A-00000000392 Battery Install Before# 6/3000000002000000009 Lot#00000000116P Pads Eup Date: 11/23/2000000005 Pads Lot # 230000000023-0000000098 Resoluion : Closing case cutomer has confirmed that AED is 000000000000000000000000001king as it shoud. Internal Comments : Tracking # 78500000000726000000002, DELIVERED Tusday 1000000002000000003 at 9:00000000AM Signed for by: S.STEVE ##cutomer to confirm that AED is 000000000000000000000000001king as it shoud Internal Comments : 2nd attempt ##cutomer to confirm that AED is 000000000000000000000000001king as it shoud. Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer Rowan uiversity 200000000 Mulica Hill Rd Glassboro, NJ 00000000000000008-1700000000 Attn: Vernon Howery ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002652800000000 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-10-23 19:41 | No
2670 | Yes | CS00000It does not start | CS0000000It does not start | 000000000000000000000000001O##: 0000000000000000000000000010000000026000000002 October 26, 2000000003 Luas Santolino Simón The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce. The HostPC coud be failing, L. Santolino is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment does not start. Patient Impact: No impact SE 8.2.17.00000000Resoluion: Host PC and hard disk are replaced. Software is reloaded. Adjutments are performed and correct operation of the system is confirmed. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000, FleuVision-2_Revised PC not HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000026000000006 October 26, 2000000003 Luas Santolino Simón The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: maintenance Device ue: Vascuar Eupected and ueupected Resuts: System down. uer Impact: No activity Patient Impact: SE 8.2.17 Diagnostic performed by Engineer: Problem located in the host PC. New host PC is requsted. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000262757 Adrián Sancho López The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce. The HostPC coud be failing, L. Santolino is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment does not start. Patient Impact: No impact SE 8.2.17.00000000Diagnostic performed by Engineer: The equpment does not complete the start-u sequnce. The HostPC coud be failing, L. Santolino is in charge of checking it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The equpment does not finish the start-u sequnce. The HostPC coud be failing, L. Santolino is in charge of checking it. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: The equpment does not finish the start-u sequnce. The HostPC coud be failing, L. Santolino is in charge of checking it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-10-23 19:41 | No
2671 | Yes | 00T0rouble reviewing images, patient on table | 00CDT0rouble reviewing images, patient on table | 000000000000000000000000001O##: 00000000000000000000000000100000000266627 200000000ct-2000000003 Saypack John The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : system 000000000000000000000000001’t store images RMT - uer gudance: Image storage is not possible becaue of an image disk problem. 23-Oct-2000000003 00000000:00000000:00000000 When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : First occurence constantly Malfuction area : u-ray pc u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: troule reviewing images, patient on table uer Impact: uable to make uray Patient Impact: case stopped/other room in ue at the moment SE 1.2.5.00000000Resoluion : Replace u-ray PC and reload SW ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000026000000002 23-Oct-2000000003 Saypack John The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : system 000000000000000000000000001’t store images RMT - uer gudance: Image storage is not possible becaue of an image disk problem. 23-Oct-2000000003 00000000:00000000:00000000 When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : First occurence constantly Malfuction area : u-ray pc u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: troule reviewing images, patient on table uer Impact: uable to make uray Patient Impact: case stopped/other room in ue at the moment SE 1.2.5.00000000Diagnostic performed by Engineer : Attempt sw loading of the u-ray pc Follow u Requred : Replace u-ray pc ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002600000000 Ketcherside Jeff The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - uer gudance: Image storage is not possible becaue of an image disk problem. 23-Oct-2000000003 00000000:00000000:00000000 When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: troule reviewing images, patient on table uer Impact: uable to make uray Patient Impact: case stopped/other room in ue at the moment SE 1.2.5.00000000Diagnostic performed by Engineer : cutomer states troule reviewing images, patient on table. reviewed logfile, foud image disk issus, cutomer deleted patients and rebooted the system mutiple times withou sucess. cutomer did not have a c-arm to continu the case, only other lab was in ue. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - uer gudance: Image storage is not possible becaue of an image disk problem.: 1. The Image Storage is not possible becaue of an image disk problem. Repair Action : Service actions for RMT - uer gudance: Image storage is not possible becaue of an image disk problem.: 1.1. The Image disk of the u-Ray PC (TP) is defective and needs to be replaced. u-ray PC Monoplane 00000000800000000660000000093 SSD Data Haswell 00000000800000000660000000032 Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - uer gudance: Image storage is not possible becaue of an image disk problem.: 1. The Image Storage is not possible becaue of an image disk problem. Repair Action : Service actions for RMT - uer gudance: Image storage is not possible becaue of an image disk problem.: 1.1. The Image disk of the u-Ray PC (TP) is defective and needs to be replaced. u-ray PC Monoplane 00000000800000000660000000093 SSD Data Haswell 00000000800000000660000000032 Technician Name : Ketcherside Jeff ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: orlando IP & Issu: 72800000000622 troule reviewing images, patient on table Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A Internal Remarks : 000000002200000000300000000 FSE Saypack is reporting that the uRAY PC OS drive does not appear to be tuning on...likely a hard drive bay issu. FSE can try runing the drive euternal of the PC...if this 000000000000000000000000001ks, replace PC. Approve: 00000000800000000660000000093 - u-ray PC Monoplane PC will deliver with new drives...informed FSE not to swap any drives between PCs and to make sue system backu is available to restore after the swap. ----------------------------------------------------------------------------- | 2023-10-23 19:41 | No
2672 | Yes | [Q] Replaced board spc0running detector does not work | [Q] Replaced board spc0 running deteur does not work | 000000000000000000000000001O##: 000000000000000000000000001000000002653300000000ct 300000000 2000000003 Frank, Nieuenhujse The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Certain motorized movement party available SPC6 is defective Diagnostic performed by Engineer: board replaced. calibrations performed. Parts Consued: 1.0000000000000000000000000000000023000000002300000000,MVR LOW POWER BOARD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000265265 RONALD BOK The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Engineer needs to replace low powerboard spc7 and do calibrations again Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Certain motorized movement party available SPC6 is defective Diagnostic performed by Engineer: spc 6 changed with spc7 so system 000000000000000000000000001ks again eucept the rotation of the detector. Engineer needs to do new low powerboard in spc7 and perform calibrations ----------------------------------------------------------------------------- | 2023-10-23 19:51 | No
2673 | Yes | uuuucable worn out | u ucable worn out | 000000000000000000000000001O##: 000000000000000000000000001000000002700000000000000008-Nov-2000000003 Gabriel Jimenez Strzelecki The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.000000000000000000000000989800000000125891,5 Leadset, Grabber, Chest, IEC, ICu1.00000000000000000000000098980000000010000000091,100000000Lead uTruk AAMI/IEC 2m 1.000000000000000000000000989800000000125851,5 Leadset, Grabber, IEC, ICuTools & Equpment : PATIENT SIMuATOR, 200000000000000009-0000000000000000:00000000:00000000 Resoluion: Signal failue is detected in the ucable with lot nuber: M1663A 7A M1971A 00000000/200000000M1978A 00000000/200000000 Defective cables are replaced. The necessary verifications have been carried ou to certify the restoration to the operating conditions prior to the breakdown. To ensue proper operation of the equpment, we recommend that youue only accessories and consuables approved by u. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000026500000000 200000000ct-2000000003 Gabriel Jimenez Strzelecki The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Follow u Requred: Change defective ucables. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Diagnostic performed by Engineer: Duing the review of the defibrillator it is observed that the ucables are very 000000000000000000000000001n. New ucables are ordered. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000263368 Ghazwan Hueidi Deaf The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Reguatory Qustions: N/A Resoluion: HeartStart uuumonitor/defib u000000003169600000000roblem in uuuCPR Requst replacement assign Gabriel Jimenez Order: 1 u 989800000000158661 Pads/CPR Cable Neut gen QCPR Jun José García Castillo # +3000000001180000000089 # OT 722100000000ollow u Requred: HeartStart uuumonitor/defib u000000003169600000000roblem in uuuCPR Requst replacement assign Gabriel Jimenez Order: 1 u 989800000000158661 Pads/CPR Cable Neut gen QCPR Jun José García Castillo # +3000000001180000000089 # OT 722100000000---------------------------------------------------------------------------- | 2023-10-23 20:01 | No
2674 | Yes | 00N0or the last couple of the months the wireless foot switch is not working | 00CDN0for the last couple of the months the wireless foot switch is not working | 000000000000000000000000001O##: 0000000000000000000000000010000000026500000000 Ketcherside Jeff The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : for the last coule of the months the wireless foot switch is not 000000000000000000000000001king When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: for the last coule of the months the wireless foot switch is not 000000000000000000000000001king uer Impact: wireless footswitch is not 000000000000000000000000001king Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Follow u Requred : uLL# and/or description of uLL: wireless footswitch is not 000000000000000000000000001king Repair Action : created case as per cutomer instrutions Technician Name : Ketcherside Jeff Internal Comments : uLL# and/or description of uLL: wireless footswitch is not 000000000000000000000000001king Repair Action : created case as per cutomer instrutions Diagnostic performed by Engineer : cutomer states for the last coule of the months the wireless foot switch is not 000000000000000000000000001king. cutomer realizes there is a recall on the wireless footswitch bu wanted to pu a case in . uable to resolve the issu remotely, created onsite 000000000000000000000000001--------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ronald Ragbir IP & Issu: 65872966/for the last coule of the months the wireless foot switch is not 000000000000000000000000001king Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-23 20:16 | No
2675 | Yes | [Q]PSE_QACASE for PCI imported product#000000 the device suddenly fails, can not eupose normally, can | [Q]PSE_QACASE for P0imported produu00000000 the device suddenly fails, can not eupose normally, can |  | 2023-10-23 20:19 | No
2676 | Yes | [Q] Having uLLcodes and had an emergency yesterday with patient on the table . Previoucase# 000 | [Q] Having uLLcodes and had an emergency yesterday with patient on the table . Previoucase# 00000 | 000000000000000000000000001O##: 00000000000000000000000000100000000283116 26-Oct-2000000003 Fedorka Timothy The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.000000000000000000000000989000000000000000086700000000u 200000000+ 000000007 ROT-GS 10000000000000000esoluion : Pick u part and stands to ship back Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: yes Eupected and ueupected Resuts: tue curent ou of spec uer Impact: no Patient Impact: no uu----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000275399 26-Oct-2000000003 Jozwiak David The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End uer Device ue: Tue Eupected and ueupected Resuts: Replace tue uer Impact: N/A Patient Impact: N/A SE N/A Diagnostic performed by Engineer : Tue replacement ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000026700000000 25-Oct-2000000003 Fedorka Timothy The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced u-ray tue and calibrated system. Tested system. Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: yes Eupected and ueupected Resuts: tue curent ou of spec uer Impact: no Patient Impact: no uuFollow u Requred : Monitor system operations Tools & Equpment : TORQu WRENCH,200000000000000008-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-29 00000000:00000000:00000000 MuTIMETER,200000000000000008-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000272100000000200000000ct-2000000003 Colley, Jeremy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: u-Ray Eupected and ueupected Resuts: u-Ray uer Impact: Down system Patient Impact: N/A SE Diagnostic performed by Engineer : Parts did not arrive. Reschedued for following day. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000260000000099 23-Oct-2000000003 Fedorka Timothy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: yes Eupected and ueupected Resuts: tue curent ou of spec uer Impact: no Patient Impact: no uuDiagnostic performed by Engineer : Troule shoot with NSS determined u-ray tue is bad Follow u Requred : Order and replaced u-ray tue Tools & Equpment : LCR METER,200000000000000008-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : FSE Tim called in after getting "From Generator device: uSC: Tue Curent ou of range. uLL over the weekend and today. I reviewed the logging and it indicates nine uLL. FSE was in over the weekend and completed a tue adaptation that appears to resolve for a day. I asked the FSE to inspect all the HV cabling and clean. (HiVo side was checked on previou visit). After this is checked the FSE will complete 2u conditioning and the rest of the generator calibrations and call back in to see if the issu persists. The following was also foud in the logs: "Tap feedback state does not match the configued state logical tap: 8 tap name: CRCB_GrouB uu Slot 5 DSM SW1 (11u1n) [Cont] HwTap/PropIo: 8" I asked the FSE to check the connections on the MPDuto ensue they are pluged into the correct locations. Internal Remarks : New Case Caller: Natalie Beckinger Issu: 70000000055600000000aving uLLcodes and had an emergency yesterday with patient on the table . Previou case# 0000000022000000006900000000 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case an onsite 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Case: 000000002200000000823 FSE: Fedorka Timothy Tue curent ou of range. FSE states getting constant Tue Arcing after performing Tue Conditioning. Requsted logs for review. Foud "u-Ray tue curent (Ie) is ou of range. This is an uLL" Instruted FSE to place part order for 989000000000000000086700000000- u 200000000+ 000000007 ROT-GS 10000000000000000----------------------------------------------------------------------------- | 2023-10-23 20:21 | No
2677 | Yes | Duplicate of PR00 | Duplicate of PR000 |  | 2023-10-23 20:25 | Yes
2678 | Yes | 0 - u00- QACASE for PCI imported product#0000000 Daily self-test fails, report uLL | 00- u0000- QACASE for P0imported produu00000000 Daily self-test fails, report uLL |  | 2023-10-23 20:35 | No
2679 | Yes | Unable to playback on cine | Unable to playback on 0e | 000000000000000000000000001O##: 00000000000000000000000000100000000266893 00000000-Nov-2000000003 Jepsen Larry The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000717991,u-ray PC Monoplane W7 + W100000000Resoluion : Replaced u-ray PC loaded software, restored backu and tested operation. System 000000000000000000000000001king properly at this time. Problem Description by Engineer : Problem description by engineer : RMT - uer gudance: Image storage is not possible becaue of an image disk problem. 23-Oct-2000000003 00000000:52:32 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: To be able to playback cine/uable to playback cine uer Impact: uable to playback cine Patient Impact: No impact to patient or procedue case is continud SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000265757 23-Oct-2000000003 Jepsen Larry The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - uer gudance: Image storage is not possible becaue of an image disk problem. 23-Oct-2000000003 00000000:52:32 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: To be able to playback cine/uable to playback cine uer Impact: uable to playback cine Patient Impact: No impact to patient or procedue case is continud SE 1.2.5.00000000Resoluion : Tested system reviewed logs, image disk issu, Tried to reload sute pc software, failed with uLL. Ordering new pc with drives. Follow u Requred : install received parts ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000265100000000Vilkaukas Viktorija The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - uer gudance: Image storage is not possible becaue of an image disk problem.: 1. The Image Storage is not possible becaue of an image disk problem. Repair Action : Service actions for RMT - uer gudance: Image storage is not possible becaue of an image disk problem.: 1.1. The Image disk of the u-Ray PC (TP) is defective and needs to be replaced. 00000000800000000660000000093 - u-ray PC Monoplane 00000000800000000660000000032 - SSD Data Haswell Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - uer gudance: Image storage is not possible becaue of an image disk problem.: 1. The Image Storage is not possible becaue of an image disk problem. Repair Action : Service actions for RMT - uer gudance: Image storage is not possible becaue of an image disk problem.: 1.1. The Image disk of the u-Ray PC (TP) is defective and needs to be replaced. 00000000800000000660000000093 - u-ray PC Monoplane 00000000800000000660000000032 - SSD Data Haswell Technician Name : Vilkaukas Viktorija Problem Description by Engineer : Problem description by engineer : RMT - uer gudance: Image storage is not possible becaue of an image disk problem. 23-Oct-2000000003 00000000:52:32 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: To be able to playback cine/uable to playback cine uer Impact: uable to playback cine Patient Impact: No impact to patient or procedue case is continud SE 1.2.5.00000000Diagnostic performed by Engineer : Bob is reporting that the uers are reporting that they are not able to playback cine. He said they are getting a buch of messages on the system bu he did not know what those messages. Created a logfile and foud that the image storage is having an issu. uers are getting a message "uer gudance: Image storage is not possible becaue of an image disk problem". ----------------------------------------------------------------------------- Internal Remarks : Case: 000000002200000000728 FSE: Jepsen Larry System uable save images. Reviewed logs. Instruted FSE to perform recommended soluions: Caue The Image Storage is not possible becaue of an image disk problem. Soluion The Image disk of the u-Ray PC (TP) is defective and needs to be replaced. Parts TB Gude: https://u.mizecu.com/knowledgeDocuent.html?5fa38fae38abbe000000002df81eaaa1870000000076dcf5ac1262bbad000000006000000005251a000000005 u-Ray PC SW Load: Repair - M-cabinet OS Disk of u-ray PC u- InCenter (mizecu.com) Internal Remarks : uray pc failed software reload. Replace uray pc. ----------------------------------------------------------------------------- | 2023-10-23 20:36 | No
2680 | Yes | IT DOES NOT START | IT DOES NOT START | 000000000000000000000000001O##: 000000000000000000000000001000000002657200000000ctober 2000000002000000003 Ortega, Roberto The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: IT DOES NOT START When was the first occurence? Within... : October 23rd-2000000003 00000000:30000000000000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: IT DOES NOT START uer Impact: No Patient Impact: No SE 000Resoluion: We received an operational equpment that had problems when starting the application. When restarting the system, the failue was foud, we proceeded to check it completely and foud fauts in the battery of the uuu Corrective maintenance of cleaning and battery replacement is performed and the failue is corrected. Operation tests are sucessfu and delivery is sucessfu. Equpment is operational. Tools & Equpment: MuTIMETER, 200000000000000005-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000260000000000000000oc, Migul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: IT DOES NOT START When was the first occurence? Within... : October 23rd-2000000003 00000000:30000000000000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: IT DOES NOT START uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 000Diagnostic performed by Engineer: The equpment does not finish tuning on, the couter stays at zero, the field service engineer says that apparently it is uuu monitor suport monitor does not give a signal, contact restarts only uuuand the equpment finishes starting, recommended to the client not to tun off the equpment util the field service engineer visits. A field service engineer is needed to clean and reconnect the cards and cables internally, replace the battery if necessary and test operation, if the faut persists, reinstall the software or do a disk scan and repair, otherwise perform HD and PC tests, evalute its replacement. Internal Comments: uLL# and/or description of uLL IT DOES NOT START Repair Action: The equpment does not complete start-u, the couter stays at zero, the FSE says that apparently it is uuu monitor suport monitor does not give a signal, contact restarts only uuuand the equpment finishes starting u, recommended to the client not to tun off the equpment util the visit of the FSE. A field service engineer is needed to clean and reconnect the cards and cables internally, replace the battery if necessary and test operation, if the faut persists, reinstall the software or do a disk scan and repair, otherwise perform HD and PC tests, evalute its replacement. Requred part: (u1), 00000000800000000119123 - FleuVision 2 PC No HDD, COA - For u R8.1 - 8.1.17.u and u R8.2 - 8.2.17.u; withou OS disk; Follow u Requred: uLL# and/or description of uLL IT DOES NOT START Repair Action: The equpment does not complete start-u, the couter stays at zero, the FSE says that apparently it is uuu monitor suport monitor does not give a signal, contact restarts only uuuand the equpment finishes starting u, recommended to the client not to tun off the equpment util the visit of the FSE. A field service engineer is needed to clean and reconnect the cards and cables internally, replace the battery if necessary and test operation, if the faut persists, reinstall the software or do a disk scan and repair, otherwise perform HD and PC tests, evalute its replacement. Requred part: (u1), 00000000800000000119123 - FleuVision 2 PC No HDD, COA - For u R8.1 - 8.1.17.u and u R8.2 - 8.2.17.u; withou OS disk; What are the neut Trouleshooting steps: A field service engineer is needed to clean and reconnect the cards and cables internally, replace the battery if necessary and test operation, if the faut persists, reinstall the software or do a disk scan and repair, otherwise perform HD and PC tests, evalute its replacement. Technician Name: Koc, Migul ----------------------------------------------------------------------------- | 2023-10-23 20:41 | No
2681 | Yes | uPSA_PSE_No uon Monitors on room 0nd room 0mergency Department/Maintenance | uPSA_PSE_No uon Monitors on room 0and room 0Emergency Department/Maintenance |  | 2023-10-23 20:42 | Yes
2682 | Yes | 00N0nable to take cine runs and flouro snapshots, "Image disk full, delete euams". | 00CDN0Unable to take 0e runs and flouro snapshots, "Image disk full, delete euams". | 000000000000000000000000001O##: 00000000000000000000000000100000000267293 26-Oct-2000000003 Peterson Daniel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : IPPC is not booting u properly. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : Troule shot system and foud mutiple occurences of image disk ful messages. Reinitializing the curent drives did not resolve. Ordered replacement drives and formatted. TEsted and ran diagnostics. all pass Tools & Equpment : MuTIMETER,200000000000000002-29 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000265719 23-Oct-2000000003 Peterson Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : IPPC is not booting u properly. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Checked uLLlog and foud image disk issus. Performed reinitialize format and tested. Drives passed and lab is fuctional. Ran image pc and host pc diagnostics. Both pass. Check lab service history and foud same issu last month with image disk ful. Will order replacement hard drives for Monday Oct 300000000h at 1pm Follow u Requred : replace image hard drives. Ship parts overnight to fnloc 16000000003. 363600000000Couar place, Murieta, u563 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000265632 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : IPPC is not booting u properly. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Cutomer states that they are uable to take cine rus and flouo snapshots, "Image disk ful, delete euams". Follow u Requred : uLL# and/or description of uLL: Logs indicate that the IPPC did not boot u properly. Repair Action : ***Check to see if the IPPC has a red backplane and if that is the case replace the IPPC regardless*** IPPC Failed to Boot u - Frontal CAuE 1: Frontal IPPC coud be defective, not powered on or net000000000000000000000000001k cabling incorrectly attached SOLuION 1: The caue of this problem is either du to lack of net000000000000000000000000001k connection or IP-PC itself is not booted u: -Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs here: http://pww.uraywiki.u.com:80000000000000000uray-wiki/indeu.php/Post_code_LED_boot_statu_signals_of_PCs - If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IP-PC and perform Cold reboot of the system - If LED statu is not Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IP-PC if problem persists/repeats. Futher checking: 1. Before changing IPPC, check if the Power for the PC is swiched off when the system is OFF, if Power is ON, the IPPC will not boot correct. 2. Change CMOS battery, a bad battery can caue the PC not to start first time. 3. This pattern can also be caued if the BIOS is set wrong, like Advanced->Boot Featue->Restore On AC Power Loss set to 'Last State' or 'Power Off', defaut setting mut be 'Power On', if not green power LED at IPPC does not switch ON at boot u and PC needs to be started with Front Buton, in this case issu is not intermittant bu always present. Technician Name : Johansson Anders Resoluion : Created on-site 000000000000000000000000001ternal Comments : uLL# and/or description of uLL: Logs indicate that the IPPC did not boot u properly. Repair Action : ***Check to see if the IPPC has a red backplane and if that is the case replace the IPPC regardless*** IPPC Failed to Boot u - Frontal CAuE 1: Frontal IPPC coud be defective, not powered on or net000000000000000000000000001k cabling incorrectly attached SOLuION 1: The caue of this problem is either du to lack of net000000000000000000000000001k connection or IP-PC itself is not booted u: -Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs here: http://pww.uraywiki.u.com:80000000000000000uray-wiki/indeu.php/Post_code_LED_boot_statu_signals_of_PCs - If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IP-PC and perform Cold reboot of the system - If LED statu is not Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IP-PC if problem persists/repeats. Futher checking: 1. Before changing IPPC, check if the Power for the PC is swiched off when the system is OFF, if Power is ON, the IPPC will not boot correct. 2. Change CMOS battery, a bad battery can caue the PC not to start first time. 3. This pattern can also be caued if the BIOS is set wrong, like Advanced->Boot Featue->Restore On AC Power Loss set to 'Last State' or 'Power Off', defaut setting mut be 'Power On', if not green power LED at IPPC does not switch ON at boot u and PC needs to be started with Front Buton, in this case issu is not intermittant bu always present. ----------------------------------------------------------------------------- Internal Remarks : Cutomer called back with issu with system. FSE schedued for 100000000300000000bu needs sooner. Reached ou to FSE who will stop by today 10000000025 at aroud 2pm local time to look at system. Internal Remarks : Planner Initials: GM Action Taken: Called Cu spoke to John Lauenbuger he stated system was down and not in ue. Called FSE Dan Peterson schedued for 10000000023@11AM and called cu back and confirmed Oucome: dispatched Internal Remarks : Planner Initials: GM Action Taken: Follow u activity for FSE Dan Peterson for10000000030000000000000000 Oucome: dispatched Internal Remarks : Interaction Type: Cutomer Call Back Caller: Lee IP & Issu: " 52000000001983 cutomer called to get statu udate on the case" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Gave cutomer schedued time and date for 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-23 20:51 | No
2683 | Yes | Heating incident with patient under sedation during an MR euamination of the spine | Heating in0ent with patient under sedation during an ueuamination of the spine | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 218132 2000000003-10000000018 18:23:32 POLARIS (POLARIS) ================= Luz Carlos Evangelista Bezerra - 000000000000000057510000000000000000200000000000000000000000000000000 =================== ==== Problem Description: I 000000000000000000000000001ld like to log patient 10000000018 euam. Description of the service performed: Collection of data to investigate the immatuity in a patient that occured in 181000000003, collection of logs and completion of the "heating qustionnaire" case treated by the Bu Post-service equpment statu: 000000000000000000000000001king withou restriction. Safety Qustions: Case: 00000000220000000067 Safety Qustion Statu: Sumitted PSA: false PSE: tru PSF Statu: Sumitted Created By: Fernando Michelan Last udated by: Fernando Michelan 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? Yes 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A 1. Date of occurence: (mm/dd/yyyy) 10000000018/2000000003 2. Date PH employee became aware of the incident: (mm/dd/yyyy) 10000000023/2000000003 3. Potential safety issu related to: Injuy 00000000Was there a patient or uer involved? Yes a. Patient/uer age: 28 Years b. Patient / uer weight: 100000000 Kilos c. Patient/uer gender: Female d. Describe patient/ uer Oucome (requred if 00000000s answered yes): The patient has been taken to the hospital emergency for evalution and patient care e. If medical intervention occured, provide details (requred if 00000000s answered yes): The patient has been taken to the hospital emergency for evalution and patient care f. Pre-euisting conditions (requred if 00000000s answered yes): - 5. Procedue oucome: The female patient was sedated, duing a Total Spine Euam acqusition. Cutomer said that the euamination took 1 hou long and the patient´s arm did touh the magnet cover, cauing the skin bun. 6. Device Info (6 a - g requred for all): - a. Curent device statu: Action taken & In ue b. Is this device available for evalution? Yes c. What was device being ued for? Diagnosis d. How was the device being ued: Spine Euamination - Diagnostic e. Is the problem intermittent? No f.Problem symptoms and frequncy: na g.Software Version (if Applicable): na 7. Has the issu been reported to any reguatory body(FDA,ISO,Other CA, etc.)? uknown 8. Has action been taken? Yes a. Action taken: Evaluted system log file and sent to the cutomer. B. Tests performed and resuts: System has been ued since the event, withou futher problems c. Additional observations: na d. Parts replaced or ordered: na e. Parts retuning for evalution: na 9. Additional Information: - a. Are there monitoring strips, log files or other information available? Yes i. Type of additional information (requred if 9a is answered yes): System Logfile available. B. If additional information is needed, who shoud be contacted? - i. Name Gulherme Ferreira ii. Phone 220000000037 2000000003-10000000018 iii. Title Eng. c. Risk Manager/Contact: - d. Has anyone from u legal been contacted? at the SAP Transfer: Case Transferred to SAP Internal Remarks: 215671 2000000003-10000000018 18:23:32 POLARIS (POLARIS) ================= Luz Carlos Evangelista Bezerra - 000000000000000057510000000000000000200000000000000000000000000000000 =================== ==== Alignment with CSR: No. Action Plan: NOk, following the diagnostic and testing steps sugested by the manuactuer. SAP Transfer: Case Transferred to SAP Trouleshooting Action: Schedued with the cutomer:- Schedued date/time: 10000000011 AT 00000000H T2 Activities:Problem Description: Problem description by engineer : Observed report of bun reported by the cutomer in the region of the patient's arm. When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Where in the end-uer 000000000000000000000000001kflow is the issu occuring? : Duing clinical ue (Acqusition/Scan) Has the cutomer created a specific system event report of the issu? : No Trouleshooting Action: What steps are taken duing the system diagnosis : Contact was made with cutomer Leonardo Dias who reported having received a complaint of buns reported by the patient in an euamination carried ou on 10000000018. Euamination of the spine region with a Spine coil lasting aroud 600000000minues according to the client (between 11:00000000 am and 12:00000000 pm). Client mentions female patient (28 years old, 100000000kg) uder sedation duing the euam, who 000000000000000000000000001ld have touhed her arm against the cover on the magnet, cauing the bun. The patient was reportedly taken to the hospital emergency room for treatment. SAR valus ​​between the reported time, eutracted from the equpment log: 212185 2000000003-10000000018 10000000051:33.700000000 uerMessage oc-scan-engine 12200000000MPuR: PO: SAR: limiting type whole body, 0000000000000000/kg 21000000001 2000000003-10000000018 10000000052:18.8500000000uerMessage oc-scan-engine 12200000000MPuPR : PO: SAR: limiting type whole body, 0000000000000000/kg 00000000MPuR : PO: SAR: limiting type whole body, 1.86W/kg 00000000PuR :PO:SAR: limiting type whole body, 1.97W/kg MPuR :PO:SAR: limiting type whole body, 0000000000000000/kg 00000000MPuR : PO: SAR: limiting type whole body, 0000000000000000/kg 2200000000MPuR : PO: SAR: limiting type whole body, 0000000000000000/kg 12200000000MPuR : PO: SAR: limiting type whole body, 0000000000000000/kg 00000000MPuR : PO: SAR: limiting type whole body, 1.00000000/kg 200000000MPuR : PO: SAR: limiting type whole body, 1.00000000/kg 200000000MPuR : PO: SAR: limiting type whole body, 2.56W/kg 12200000000MPuR : PO: SAR: limiting type whole body, 00000000800000000kg 2200000000MPuR : PO: SAR: limiting type whole body, 0000000000000000/kg 2200000000MPuR : PO: SAR: limiting type whole body, 2.28W/kg 12200000000MPuR :PO:SAR: limiting type whole body, 1.00000000/kg 2200000000MPuR : PO: SAR: limiting type whole body, 2.300000000/kg uLL# and/or description of uLL: Patient skin bun Malfuction area : No parts OneEMS internal remarks: Problem description by engineer : Observed report of bun reported by the cutomer in the region of the patient's arm. When was the first occurence? Within... : 18-Oct-2000000003 How often is issu occuring? : First occurence Where in the end-uer 000000000000000000000000001kflow is the issu occuring? : Duing clinical ue (Acqusition/Scan) Has the cutomer created a specific system event report of the issu? : No What steps are taken duing the system diagnosis : Contact was made with cutomer Leonardo Dias who reported having received a complaint of buns reported by the patient in an euamination carried ou on 10000000018. Euamination of the spine region with a Spine coil lasting aroud 600000000minues according to the client (between 11:00000000 am and 12:00000000 pm). Client mentions female patient (28 years old, 100000000kg) uder sedation duing the euam, who 000000000000000000000000001ld have touhed her arm against the cover on the magnet, cauing the bun. The patient was reportedly taken to the hospital emergency room for treatment. SAR valus ​​between the reported time, eutracted from the equpment log: 212185 2000000003-10000000018 10000000051:33.700000000 uerMessage oc-scan-engine 12200000000MPuR: PO: SAR: limiting type whole body, 0000000000000000/kg 21000000001 2000000003-10000000018 10000000052:18.8500000000uerMessage oc-scan-engine 12200000000MPuPR : PO: SAR: limiting type whole body, 0000000000000000/kg 00000000MPuR : PO: SAR: limiting type whole body, 1.86W/kg 00000000PuR :PO:SAR: limiting type whole body, 1.97W/kg MPuR :PO:SAR: limiting type whole body, 0000000000000000/kg 00000000MPuR : PO: SAR: limiting type whole body, 0000000000000000/kg 2200000000MPuR : PO: SAR: limiting type whole body, 0000000000000000/kg 12200000000MPuR : PO: SAR: limiting type whole body, 0000000000000000/kg 00000000MPuR : PO: SAR: limiting type whole body, 1.00000000/kg 200000000MPuR : PO: SAR: limiting type whole body, 1.00000000/kg 200000000MPuR : PO: SAR: limiting type whole body, 2.56W/kg 12200000000MPuR : PO: SAR: limiting type whole body, 00000000800000000kg 2200000000MPuR : PO: SAR: limiting type whole body, 0000000000000000/kg 2200000000MPuR : PO: SAR: limiting type whole body, 2.28W/kg 12200000000MPuR :PO:SAR: limiting type whole body, 1.00000000/kg 2200000000MPuR : PO: SAR: limiting type whole body, 2.300000000/kg uLL# and/or description of uLL: Patient skin bun Malfuction area : No parts Repair Action "Internal" : Evalute locally the correct fuctioning of the magneto bore ventilation system. Evalute with the client, collecting all necessary information, filling ou the "Patient Bun Qustionnaire" and forwarding it to Tecnovigilancia_BR@u.com and myeuperience@u.com What are the neut Trouleshooting steps: Send local FSE. Estimated time for inspection 1h Requred action by FSE before visit : na Estimated Job duation : 1 Eutra Internal comment field : RSE Fernando Michelan SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000220000000067 Safety Qustion Statu: Sumitted PSA: false PSE: tru PSF Statu: In Process Created By: Fernando Michelan Last udated by: Fernando Michelan 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? Yes 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A 1. Date of occurence: (mm/dd/yyyy) 10000000018/2000000003 2. Date PH employee became aware of the incident: (mm/dd/yyyy) 10000000023/2000000003 3. Potential safety issu related to: Injuy 00000000Was there a patient or uer involved? Yes a. Patient / uer age: - b. Patient / uer weight: - c. Patient/uer gender: Female d. Describe patient/ uer Oucome (requred if 00000000s answered yes): The patient has been taken to the hospital emergency for evalution and patient care e. If medical intervention occured, provide details (requred if 00000000s answered yes): The patient has been taken to the hospital emergency for evalution and patient care f. Pre-euisting conditions (requred if 00000000s answered yes): - 5. Procedue oucome: The female patient was sedated, duing a Total Spine Euam acqusition. Cutomer said that the euamination took 1 hou long and the patient´s arm did touh the magnet cover, cauing the skin bun. 6. Device Info (6 a - g requred for all): - a. Curent device statu: Action taken & In ue b. Is this device available for evalution? Yes c. What was device being ued for? Diagnosis d. How was the device being ued: Spine Euamination - Diagnostic e. Is the problem intermittent? No f.Problem symptoms and frequncy: na g.Software Version (if Applicable): na 7. Has the issu been reported to any reguatory body(FDA,ISO,Other CA, etc.)? uknown 8. Has action been taken? Yes a. Action taken: Evaluted system log file and sent to the cutomer. B. Tests performed and resuts: System has been ued since the event, withou futher problems c. Additional observations: na d. Parts replaced or ordered: na e. Parts retuning for evalution: na 9. Additional Information: - a. Are there monitoring strips, log files or other information available? Yes i. Type of additional information (requred if 9a is answered yes): System Logfile available. B. If additional information is needed, who shoud be contacted? - i. Name Gulherme Ferreira ii. Phone 215671 2000000003-10000000018 iii. Title Eng. c. Risk Manager/Contact: - d. Has anyone from u legal been contacted? at the Safety Qustions: Case: 00000000220000000067 Safety Qustion Statu: Sumitted PSA: false PSE: tru PSF Statu: In Process Created By: Fernando Michelan Last udated by: Fernando Michelan 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? Yes 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? AT Problem Reported by Cutomer: Equpment problem: I 000000000000000000000000001ld like an euam log 10000000018 patient had a bun Priority: Restriction Search for: Gulherme Sector: Clinical engineering Telephone: +55813000000008000000002 Telephone: no Contact email- engenharia.clinica@esperancaolinda.com.br uit: Hospital Esperanca Olinda Address: Avenida Douor Jose Auuto Moreira 777 Olinda-PE 53130000000000000000 BR ----------------------------------------------------------------------------- | 2023-10-23 20:56 | Yes
2684 | Yes | HS0ED is emitting the triple chirp | HS0AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000265687 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 12 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A11K-000000002200000000 Battery Install Before# 6/3000000002000000008 Lot#00000000617P Pads Eup Date: 7/15/2000000003 Pads Lot # Y000000001522100000000 ----------------------------------------------------------------------------- | 2023-10-23 20:56 | No
2685 | Yes | Device Fails Self-Test (Device is replaced) | Device Fails Self-Test (Device is replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000266000000001 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: intermittently beeping to check pads uer Impact: none Patient Impact: none SE E.00000000.000000007 See ref case 00000000220000000000000000 Cut ##d to state that the device 000000000000000000000000001ld intermittently beep with uLLmsg "check pads" or "no shocks". Diagnostic performed by Engineer : I took A3’s AED ou of service this morning SN: B17F-000000001100000000 When the A3 crew opened the AED to test it, prior to puhing any butons the defib annouced “00000000minues with no shock” The lights for pad placement were flashing bu no other annoucements were said. The A3 crew stated this is the second time that same “00000000minues with no shock” annoucement has happened. The AED passes the ful self-test and states ready for ue, all batteries and pads are within date. With having spare AEDs I felt more comfortable taking this one ou of service util it can be inspected. Thank you Lindsay Lindsay McGlouhlin Concord Fire Dept. Lieuenant EMS Coordinator, EMT-P 978-318-300000000 Good Morning Kristen, We are having problems with the same defib we had issus with in May (see attached ##. Now it’s intermittently beeping and saying check pads. Pads have been replaced mutiple times and it only fiues it for uully a nuber of hous or a day and then it starts to beep again. Thanks Kevin Cut 000000000000000000000000001ld u to replace AED as they refue to pu this device back into service. It’s not doing it now, it did it t000000000000000000000000001different times weeks apart. Both times the pads were swapped and one time the battery was too. I told suport that and that I refue to pu this defib back in service becaue this is the second time it has had uLL that were ueuplainable and I can’t trut it. Thanks Kevin Internal Comments : Will replace device as cutomer satisfaction. See ##pdf attachment. Follow u Requred : ship to acct address, Attn Kevin Fagerqust Resoluion : shipped serial # B23B-00000000300000000, tracking 7850000000080000000025 DELIVERED Tusday 1000000002000000003 at 100000000500000000AM ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002661900000000Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-10-23 21:46 | No
2686 | Yes | Multifunctional cable is broken. WORK: 000 | Multifunuonal cable is broken. WORK: 000 | 000000000000000000000000001O##: 00000000000000000000000000100000000266253 October 27, 2000000003 Hidalgo Flores, Lenin Otton The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We tried to speak with Eli 93227500000000 eut 2572, bu we were not able to. Device ue: uknown Eupected and ueupected Resuts: Mutifuctional cable is broken. uer Impact: uknown Patient Impact: uknown SE Not apply. Resoluion: We checked and delivered fugible 989800000000197111, the consuable is in bad condition, the client is in good condition, they will sign the note, the person in charge will close the part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002657600000000Rau Cordero Calderón The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We tried to speak with Eli 93227500000000 eut 2572, bu we were not able to. Device ue: uknown Eupected and ueupected Resuts: Mutifuctional cable is broken. uer Impact: uknown Patient Impact: uknown SE Not apply. Diagnostic performed by Engineer: We tried to speak with Eli 93227500000000 eut 2572 bu we were not able to. Description of the failue: Mutifuctional cable is broken. Read the secuity note FSN 2000000002-CC-EC-0000000000000000 The affected produts 989800000000197111 have a manuactuing date between September 2000000008 (9/18) and November 20000000000000000(11/200000000 Assign fSE. The 989800000000197111 - Pads Adapter Cable is a consuable and it is provided by the distribuor. uH-E13-00000000 - Emergency Room, 3rd Floor, Bou 3. Follow u Requred: We tried to speak with Eli 93227500000000 eut 2572 bu we were not able to. Description of the failue: Mutifuctional cable is broken. Read the secuity note FSN 2000000002-CC-EC-0000000000000000 The affected produts 989800000000197111 have a manuactuing date between September 2000000008 (9/18) and November 20000000000000000(11/200000000 Assign fSE. The 989800000000197111 - Pads Adapter Cable is a consuable and it is provided by the distribuor. uH-E13-00000000 - Emergency Room, 3rd Floor, Bou 3. ----------------------------------------------------------------------------- | 2023-10-23 21:51 | No
2687 | Yes | V0will not charge the battery. Need new power supply. | V0will not charge the battery. Need new power supply. | 000000000000000000000000001O##: 000000000000000000000000001000000002600000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000will not charge the battery. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000023/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Biomed is reporting they are not curently with the uit. Advised Biomed to reference page 112 of the service manul to confirm the power suply is 000000000000000000000000001king. Biomed advised to confirm 200000000olts DC across the brown and orange wires. Cutomer advised to order power suply if the 200000000olts DC is not present. Biomed is reporting the problem was identified by ASP implementing FCO86600000000000000006. Part # 00000000561512591 Rp - Power Suply, V600000000List price is $8900000000 Page 1600000000for referencing the repair. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000power on and charge the battery. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000will not charge the battery. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: " Alvin Almonor 516-382-72100000000" IP & Issu: Need new power suply Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-23 22:01 | No
2688 | Yes | Unit will not perform check out procedure | Unit will not perform check out procedure | 000000000000000000000000001O##: 00000000000000000000000000100000000265788 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. The cutomer has advised closing the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000is failing the mask test and is displaying diagnostic code 1200000000 Low Leak CO2 rebreathing risk. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000023/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Advised cutomer to complete the pneuatics testing and confirm the V600000000is operating within specification. Biomed advised of possible defective flow sensor assembly. Biomed has confirmed the V600000000is failing Test 7 Air Flow Accuacy at zero. Specification is: 00000000SLPM -1.00000000to 1.00000000SLPM 100000000SLPM 8.6 to 11.00000000LPM 35 SLPM 33.2 to 36.8 SLPM 600000000SLPM 56.2 to 63.8 SLPM 100000000SLPM 129.8 to 15000000002 SLPM Part # 989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 List price $1,217 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000pass the mask port test. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is failing the mask port test. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jeremy Calvert IP & Issu: 71000000009813/ uit will not perform check ou procedue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-23 22:06 | No
2689 | Yes | Alarms-No Sound / Speaker Malf INOP-No Sound-Foulink s/n u0 / No Trouble-Working to Spec-Other / Replaced Parts | Alarms-No Sound / Speaker Malf INOP-No Sound-Foulink s/n u00 / No Trouble-Working to Spec-Other / Replaced Parts | 000000000000000000000000001O##: 0000000000000000000000000010000000039900000000 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000026600000000 Heard, John The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : When the Mu00000000powers on, an uLLof "Speaker Malfuction" appears on the screen. Diagnostic performed by Engineer : Problem Description: When the Mu00000000powers on, an uLLof "Speaker Malfuction" appears on the screen. Biomed states that the Mu00000000powers on, an uLLof "Speaker Malfuction" appears on the screen. Biomed confirmed that when the Mu00000000is tuned on, no soud can be heard from the modue's speaker. Biomed 000000000000000000000000001ld like to send to the bench for repair. Form ued: Bench Trouleshooting Steps Taken: Confirmed contact information on file, sending bench form to biomed for shipping instrutions. Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: The Mu00000000powers on withou issu. uer Impact: The Mu00000000was removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: When the Mu00000000powers on, an uLLof "Speaker Malfuction" appears on the screen. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Edurdo Flores - phone 281 729 3000000005 IP & Issu: 5960000000000000000- Message on screen keeps showing: Speaker malfuction Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-23 22:11 | No
2690 | Yes | Power on by itself | Power on by itself | 000000000000000000000000001O##: 00000000000000000000000000100000000266319 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: powers on by itself uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order the u pcba and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and wanted to know why nobody called him on the recall. I spoke to this cutomer 00000000eeks ago, and ou planners reached ou to Progressive. Nobody contacted the cutomer on their 3 uits that are affected. This uit he stated powers on by itself, bu I did not see any notes on ou conversation abou this failue. Recommended replacing the u pcba Part #00000000561528871, u pcba, list price $332.00000000 ##d Progressive planners and Rick Tran to get the FCO66 schedued. Progressive has FCO66 case #000000002000000002300000000 schedued for 11-13-23 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Robert Gonzales IP & Issu: 58792622 & the recall for power irrution Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created r000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-23 22:21 | No
2691 | Yes | System does not start, very loud noise can be heard | System does not start, very loud noise can be heard | 000000000000000000000000001O##: 0000000000000000000000000010000000026000000002 29-Oct-2000000003 Hoffmann, Thomas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : Problem has been confirmed. Consutation with Ms. Kretschmer uder 0000000000000000912181176. RMT - uGEN: uLL(00000000Hu) Asymmetrical kV detected - Frontal 23-Oct-2000000003 12:53:00000000RMT – : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact : No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion: High voltage flashovers. Plu cleaned and reconnected. Fuction OK ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000263882 Malhas, Ali The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer : Problem has been confirmed. Consutation with Ms. Kretschmer at 0000000000000000912181176. RMT - uGEN: uLL(00000000Hu) Asymmetrical kV detected - Frontal 23-Oct-2000000003 12:53:00000000RMT – : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact : No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer: According to the cutomer, the lou noise came from the C-arm. The system continus to ru, bu takes black u-ray images. Supicion: rotating anode. Log analysis: Fluroscopy failed du to 00000000Hu uLL(kV asymmetrical) The problem is more towards the converter, possibly kV mA uit. Internal Comments: uLL# and/or description of uLL: Root Caue for RMT - uGEN: uLL(00000000Hu) Asymmetrical kV detected - Frontal: 1. Fluroscopy failed du to 00000000Hu uLL(kV asymmetrical) Root Caue for RMT - uGEN: Warning ( 00000000HW) Asymmetric kV detected - Frontal: 1. Asymmetric kV actul >= 15%. See Detail uLLlog for deviation of kV for Anode (EQ converter 1) and Cathode (E2Q converter 2). Repair Action: Service actions for RMT - uGEN: uLL(00000000Hu) Asymmetrical kV detected - Frontal: 1.1. Check the detailed uLLlog and check which side Anode or Cathode is too low or too high. 1.2. Swap the HV cable Anode/Cathode, if the problem moves try new HV cable. 1.3. Perform tue adaptation and conditioning. 1.00000000If the deviation is only a little bit above the allowed limit, try new HV tank. 1.5. Swap Anode with Cathode converter, if the problem moves to other side try new converter. 1.6. If the problem does NOT move after swapping the converter, try new kVmA uit. 1.7. If all this doesn't solve the problem try new u-ray tue. Service actions for RMT - uGEN: Warning (00000000HW) Asymmetric kV detected - Frontal: 1.1. Check in detailed uLLlog which kV side (anode / cathode) is too low or too high. Swap both converter and check detailed uLLlog again, if the problem moves to Anode or Cathode try new converter (EQ) 1.2. If the problem don't move, after swapping the converter, try new kVmA uit (EH). 1.3. If the deviation is only very low it coud be a HV tank problem. Follow u Requred: uLL# and/or description of uLL: Root Caue for RMT - uGEN: uLL(00000000Hu) Asymmetrical kV detected - Frontal: 1. Fluroscopy failed du to 00000000Hu uLL(kV asymmetrical) Root Caue for RMT - uGEN: Warning (00000000HW) Asymmetric kV detected - Frontal: 1. Asymmetric kV actul >= 15%. See Detail uLLlog for deviation of kV for Anode (EQ converter 1) and Cathode (E2Q converter 2). Repair Action: Service actions for RMT - uGEN: uLL(00000000Hu) Asymmetrical kV detected - Frontal: 1.1. Check the detailed uLLlog and check which side Anode or Cathode is too low or too high. 1.2. Swap the HV cable Anode/Cathode, if the problem moves try new HV cable. 1.3. Perform tue adaptation and conditioning. 1.00000000If the deviation is only a little bit above the allowed limit, try new HV tank. 1.5. Swap Anode with Cathode converter, if the problem moves to other side try new converter. 1.6. If the problem does NOT move after swapping the converter, try new kVmA uit. 1.7. If all this doesn't solve the problem try new u-ray tue. Service actions for RMT - uGEN: Warning (00000000HW) Asymmetric kV detected - Frontal: 1.1. Check in detailed uLLlog which kV side (anode / cathode) is too low or too high. Swap both converter and check detailed uLLlog again, if the problem moves to Anode or Cathode try new converter (EQ) 1.2. If the problem don't move, after swapping the converter, try new kVmA uit (EH). 1.3. If the deviation is only very low it coud be a HV tank problem. What are the neut trouleshooting steps: 1 TE deployment with 2 ET and 00000000AZ. Replacing the converter and kV mA uit in the Velara E-Cabinet from the front, as well as checking the high-voltage cable on the spotlight. Then calibrate the grid switch. Radiators are also supected, bu are not covered in the contract. If this does not solve the problem, the spotlight mut be offered. Technician Name: Malhas, Ali ----------------------------------------------------------------------------- Internal Remarks: 00CC: Please make an appointment with Ms. Kretschmer today 0000000000000000912181176. Thank you Internal Remarks: RS held with Ms. Kretschmer. Lou noise from the C-arm. Probably rotary anode defective. u-rays are black. Spotlight defective. Log is created by cutomer to confirm the problem. Cutomer comment: System does not start, very lou noise can be heard. Supicion of a defective radiator. Fr Kretschmer also recorded a video abou it. Cutomer comment: ## Appointment with Angela Kretschmer for October 200000000. 9 a.m. agreed ## ----------------------------------------------------------------------------- | 2023-10-23 23:11 | No
2692 | Yes | Touchscreen issue, GFE00OV0 Response received/IS complete | Touchscreen issue, GFE0 0NOV00Response received/IS complete | 000000000000000000000000001O##: 000000000000000000000000001000000002669500000000Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000pass the touhscreen calibration and diagnostics. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000touhscreen is uresponsive. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. The cutomer has advised closing the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000pass the touhscreen calibration. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000023/2000000003 Trouleshooting Steps Taken: Cutomer is reporting the touhscreen has physical damage. Provided cutomer the part # for the V600000000touhscreen. Part # 00000000561539000000001 RP-Touh Screen, V600000000List price is $9900000000---------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: tim IP & Issu: 67955200000000 touhscreen down Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-10-23 23:51 | No
2693 | Yes | During PM service, the unit cannot be powered on | During PM service, the unit cannot be powered on | 000000000000000000000000001O##: 00000000000000000000000000100000000267000000007 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ASP has a V600000000and it does not power on. He verified with a different PM pcba, it solved the issu. Requsted a 100000000% discout to replace the FCO part #000000005615000000001 since we do not have part #000000005615000000001 available. Problem Description by Engineer : Cutomer Fuction/Role: ASP Device ue: not in ue Eupected and ueupected Resuts: does not power on uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : ASP has a V600000000and it does not power on. He verified with a different PM pcba, it solved the issu. Requsted a 100000000% discout to replace the FCO part #000000005615000000001 since we do not have part #000000005615000000001 available. ASP completed the service ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000280000000055 27-Oct-2000000003 Mahmou, Mohamed The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: ASP Device ue: not in ue Eupected and ueupected Resuts: does not power on uer Impact: none Patient Impact: none SE un Resoluion : Replaced The PM PCBA to resolve the problem Diagnostic performed by Engineer : after powering u the uit there was no display, tried a different PM pcba and it 000000000000000000000000001ked fine. Parts Consued : 1.000000000000000000000000000000005615000000001,RP-PCBA,PWR MGMT,GEN00000000.300000000PIC, (FCO) Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-11-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,2000000003-11-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-24 00:01 | No
2694 | Yes | uLLmessages displayed: storage/image issue | uLLmessages displayed: storage/image issue | 000000000000000000000000001O##: 00000000000000000000000000100000000267171 200000000ct-2000000003 Beall Jordan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Image store issu Eupected and ueupected Resuts: Can't save images or make u-ray uer Impact: Can't ue system. Patient Impact: Can't complete stuy SE 8.1.3000000005 Resoluion : Replaced both image disk drives, re-created frontal image store, tested and passed. No futher issus foud. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000267000000009 23-Oct-2000000003 Beall Jordan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Tested and verified issu, foud that hard disk's in IP-PC are starting to fail per uLLlog. Will order new drives and replace. Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Image store issu Eupected and ueupected Resuts: Can't save images or make u-ray uer Impact: Can't ue system. Patient Impact: Can't complete stuy SE 8.1.3000000005 Follow u Requred : Replace the hard drives ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Abouacar Diawara +13175000000003100000000 IP & Issu: 000000005169 uLLmessages displayed: storage/image issu Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Qute req created and adv of ## Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-24 00:26 | No
2695 | Yes | System uLLmessage "Storage full, please delete euams, u-Ray not possible". | System uLLmessage "Storage full, please delete euams, u-Ray not possible". | 000000000000000000000000001O##: 00000000000000000000000000100000000267182 25-Oct-2000000003 Rice Ryan The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD Problem Description by Engineer : Problem description by engineer : Defective power suply in IPPC. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion : Replaced IPPC. Tested system fuction. Room ready for ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000267136 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Defective power suply in IPPC. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Cutomer states that system is displaying uLLmessage "Storage ful, please delete euams, u-Ray not possible". System Down. Internal Comments : uLL# and/or description of uLL: Frontal IPPC has a defective power suply. Repair Action : Troule shot problem with cutomer over the phone. PC has 2300000000AC on the power cord bu has no power and is not powering on. PC power suply is defective. Resoluion : Created on-site 000000000000000000000000001llow u Requred : uLL# and/or description of uLL: Frontal IPPC has a defective power suply. Repair Action : Troule shot problem with cutomer over the phone. PC has 2300000000AC on the power cord bu has no power and is not powering on. PC power suply is defective. What are the neut Trouleshooting steps : Replace IPPC. Technician Name : Johansson Anders ----------------------------------------------------------------------------- Internal Remarks : FSE Ryan called in looking for TPR and parts ID. The SMu configuation show a non-clarity system bu the license in the logging show clarity attached to the system. Gave FSE clarity part nuber and asked them to doule check when neut on site. Logging indicate no commuication to IPPC and FSE confirms power is getting to the PC bu it is not tuning on. 0000000080000000050000000022 ClarityIQ_II IP PC non HDD 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II ----------------------------------------------------------------------------- | 2023-10-24 00:26 | No
2696 | Yes | System not booting up | System not booting up |  | 2023-10-24 00:37 | No
2697 | Yes | involved in psi. need help investigating - CHR opened PR#00000 - GFE0sent 00 | involved in psi. need help investigating - CHR opened PR#000(0/0/0) - GFE0 sent 0/00 | 000000000000000000000000001O##: 00000000000000000000000000100000000267618 100000000Dec-2000000003 Johnny Surez The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Biomed stated that this Mu00000000was involved in a patient incident and needs to be checked ou. Resoluion : Configuation cloned that resolved the problem ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000026700000000 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Biomed stated that this Mu00000000was involved in a patient incident and needs to be checked ou. Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve Diagnostic performed by Engineer : Problem Description: Biomed stated that this Mu00000000was involved in a patient incident and needs to be checked ou. Biomed stated that this Mu00000000was involved in a patient incident and needs to be checked ou. Form ued: Trouleshooting/Investigation Object: Mu00000000Not Object: Software Version: NA Date and Time of Issu: 10000000023/2000000003 uer Actions: Where is the Deviation: uknown Trouleshooting Steps Taken: NA Was PRS ued: No PRS Permission: Follow u Requred : Biomed stated that this Mu00000000was involved in a patient incident and needs to be checked ou. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: denzel IP & Issu: 100000000000000000000000000000000nvestingating Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-10-24 01:41 | Yes
2698 | Yes | Loosing fluoro intermittently | Loosing fluoro intermittently | 000000000000000000000000001O##: 00000000000000000000000000100000000267691 26-Oct-2000000003 Byron Kirk The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : loosing fluro intermittantly When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: Fuction restored following several reboots of system uer Impact: delay in clinical applications Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : replaced footswitch Parts Consued : 1.00000000000000000000000000000000270000000000000000100000000Footswitch CV 3p 00000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000009-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000267558 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : loosing fluro intermittantly When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: Fuction restored following several reboots of system uer Impact: delay in clinical applications Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Pamela syas the fluro was intermittently failing bu restored uon rebooting. She has been in contact with the local FSE, Kirk Byron who sugested replacing the footswitch. Local FSE also advised to create an LOD bu not util the case has been finished. Internal Comments : uLL# and/or description of uLL: No lod - no uLL, warnings, or uer gudance Repair Action : view curent logs to determine if other souce or action requred. Otherwise replace footswitch following fuctional test. Follow u Requred : uLL# and/or description of uLL: No lod - no uLL, warnings, or uer gudance Repair Action : view curent logs to determine if other souce or action requred. Otherwise replace footswitch following fuctional test. What are the neut Trouleshooting steps : view curent logs to determine if other souce or action requred. Otherwise replace footswitch following fuctional test. Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action Taken: FSE Kirk Byron accepted call for 10000000025 and change priority to P5 he already spoke to cutomer Oucome: dispatched ----------------------------------------------------------------------------- | 2023-10-24 01:46 | No
2699 | Yes | 0 - MC PCBA ADC failed | 00 - MC PCBA ADC failed |  | 2023-10-24 02:27 | No
2700 | Yes | 00N0orce sensor uLLfor detector | 00CDN0force sensor uLLfor deteur | 000000000000000000000000001O##: 00000000000000000000000000100000000275181 17-Nov-2000000003 Hamdan Imad The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: diag Eupected and ueupected Resuts: Tue cover needs to be replaced uer Impact: no uer impact Patient Impact: no patient impacts SE2.2 Resoluion : Clinical engineer revived the part and replaced the cover. System is ready for ue Parts Consued : 1.000000000000000000000000000000008000000001117800000000LEA3 Tue Cover ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000267988 200000000ct-2000000003 Hamdan Imad The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : spoke to Corrie ongoing collision messages on system When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Mutiple times per month Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: ongoing force sensor ,warning message on system uer Impact: waring messages on system Patient Impact: none uu00000000Diagnostic performed by Engineer : Foud ou uable to rotate the detector uless move the SID down. Also cutomer stated that the fleu arm not move at all. Ran body gurd cal and it fiued one of the issu, ran all position cal eucept Z1and Z2, also ran all the curent caps too with no issu. Foud the tue cover it needs to be replace. I ordered the part for tomorrow morning it has to be replace ASAP. I informed the clinical eng abou the cover. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027000000007 200000000ct-2000000003 Bitar, Abdu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: in-houe Device ue: diagnostic Eupected and ueupected Resuts: detector rotation is off. uer Impact: no impact Patient Impact: no impact SE2.2(L3) Resoluion : Assisted primary FSE with the detector rotation repair and system calibration. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000267900000000 Johansson Anders The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Geometry issus. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Cutomer states that force sensor uLLfor detector. Internal Comments : uLL# and/or description of uLL: Pattern Name: RMT - POS: Application uLL Collision detected [auis=FrtIDS]uY Name: 82d00000000a3-0000000055-c728-8df7-00000000d800000000d96cf Root Caue 1: Application uLL Collision detected [auis=FrtIDS] Service Action 1: Read from the logline the auis statu (operation or stand-by) and the position where the collision occured. If the movement hits a mechanical obstrution, this is normal behavior. Otherwise, ru a fuctional test on this specific movement. Possible corrective actions: â¢ Perform curent calibration of movement â¢ Perform force sensor calibration â¢ Clean mechanics or cabling. Repair Action : Perform stand calibrations Resoluion : On-site 000000000000000000000000001O#created by Scott. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000267668 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : spoke to Corrie ongoing collision messages on system When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : Mutiple times per month Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: ongoing force sensor ,warning message on system uer Impact: waring messages on system Patient Impact: none uu00000000Diagnostic performed by Engineer : speaking to Corie they are still having waring and movement uLL on the system . biomed states force sensor has been replaced recently Rse puled lod for review Internal Comments : uLL# and/or description of uLL: Application uLL Collision detected [auis=FrtIDS Repair Action : Collision detected [auis=FrtIDS] 00000000800000000262813, Description: CLEA2 Force Sensor, Quntity: 1 Back 7222200000000AMC Triple Servo Drive HP-LP-LP Produt Code 00000000800000000196622 Read from the logline the auis statu (operation or stand-by) and the position where the collision occured. If the movement hits a mechanical obstrution, this is normal behavior. Otherwise, ru a fuctional test on this specific movement. Possible corrective actions: • Perform curent calibration of movement • Perform force sensor calibration • Clean mechanics or cabling. Follow u Requred : uLL# and/or description of uLL: Application uLL Collision detected [auis=FrtIDS Repair Action : Collision detected [auis=FrtIDS] 00000000800000000262813, Description: CLEA2 Force Sensor, Quntity: 1 Back 7222200000000AMC Triple Servo Drive HP-LP-LP Produt Code 00000000800000000196622 Read from the logline the auis statu (operation or stand-by) and the position where the collision occured. If the movement hits a mechanical obstrution, this is normal behavior. Otherwise, ru a fuctional test on this specific movement. Possible corrective actions: • Perform curent calibration of movement • Perform force sensor calibration • Clean mechanics or cabling. What are the neut Trouleshooting steps : check ALL movementt and operation of Fleuarm and : MAGNu_ table R u a fuctional test on all this movemenst. Possible corrective actions: • Perform curent calibration of movement • Perform force sensor calibration • Clean mechanics or cabling. Technician Name : Sulivan Scott ----------------------------------------------------------------------------- | 2023-10-24 02:41 | No
2701 | Yes | SPLIT_Biomed reported that Surv SOCORROMEDSURG has No Audio - Remote Solutions issue resolved by a reboot. | SPLIT_Biomed reported that Surv SOCORROMEDSURG has No Audio - Remote Solutions issue resolved by a reboot. | 000000000000000000000000001O##: 00000000000000000000000000100000000267885 Doulas, Norman The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: Yes Caue: Possible issu with Remote Soluion Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Resolved Remotely Diagnostic performed by Engineer : Problem Description: Biomed reported that Suv SOCORROMEDSuG has No Auio and is displaying Wireless Mon Loss uLLmessage Biomed reported that Suv SOCORROMEDSuG has No Auio and is displaying Wireless Mon Loss uLLmessage Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: C.00000000.00000000 Date and Time of Issu: 10000000023/2000000003 uer Actions: Testing Speaker Where is the Deviation: MEDSuG Trouleshooting Steps Taken: Cualso reported the following: Auio issu had stared abou 2 weeks ago and that he has replaced the Euternal Speaker bu no fiu The Wireless Mon Loss uLLhas been on the Suv for more then 2 weeks Speaker is going throuh a Remote Soluion bu cucant ID Recommended rebooting Suv. CuRebooted Suv= Issu resolved for the Auio Recommended Clearing the Wireless Mon Loss uLLmessage via System Config= Issu resolved Also checked Device Statu and all AP's and APC's are in a connected Statu Was PRS ued: No PRS Permission: Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:In Clinical ue Eupected and ueupected Resuts: uer Impact:No impact Patient Impact:No impact SEC.00000000.00000000 Problem Description/ueupected Resuts: Biomed reported that Suv SOCORROMEDSuG has No Auio and is displaying Wireless Mon Loss uLLmessage ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Mike Ruolph IP & Issu: 7585700000000 - uer isn't getting any alarm throuh auio, only visul. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-24 03:36 | No
2702 | Yes | Does not show clear waveforms | Does not show clear waveforms | 000000000000000000000000001O##: 000000000000000000000000001000000005256900000000Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced Front Case Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Parts Consued : 1.0000000000000000000000000000000056000000005000000001,RDT Front Enclosue kit (for 00000000-100000000u-R & Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits SPO2 connector has a pin stuk in it cauing the connector to not 000000000000000000000000001k correctly. Additionally the ufleu cable was not seated in the uboard. To correct the issus foud we will replace the Front Case Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000268200000000Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits SPO2 connector has a pin stuk in it cauing the connector to not 000000000000000000000000001k correctly. Additionally the ufleu cable was not seated in the uboard. To correct the issus foud we will replace the Front Case Assembly. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002681200000000Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: SE 7.28 Problem Description/ueupected Resuts: SPO2 and unot 000000000000000000000000001king correctly. Pleth waves incorrect at times.Puse rates incorrect at times. Has changed cables, to no avail.Sometimes the uit switches back and forth between SPO2 and u Follow u Requred : SPO2 and unot 000000000000000000000000001king correctly. Pleth waves incorrect at times. Puse rates incorrect at times. Has changed cables, to no avail. Sometimes the uit switches back and forth between SPO2 and u Resoluion : Is Caue Known: No Caue: Is Resoluion Known: Yes Resoluion Sumary: Malfuction Sent to Bench Diagnostic performed by Engineer : Problem Description: SPO2 and unot 000000000000000000000000001king correctly. Pleth waves incorrect at times. Puse rates incorrect at times. Has changed cables, to no avail. Sometimes the uit switches back and forth between SPO2 and u SPO2 and unot 000000000000000000000000001king correctly. Pleth waves incorrect at times.Puse rates incorrect at times. Has changed cables, to no avail.Sometimes the uit switches back and forth between SPO2 and u Form ued: Bench Trouleshooting Steps Taken: Changed cables. ----------------------------------------------------------------------------- Internal Remarks : No replies from my 2 requsts for patient information and log files. Internal Remarks : Interaction Type: New Case Caller: Trevor Brown IP & Issu: 9190000000000000000 Does not show clear waveforms Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : ##d cutomer asking for log files and patient information. Internal Remarks : Send 2nd requst for patient info and log files to cutomer. Internal Remarks : Interaction Type: Cutomer Call Back Caller: Trevor Brown IP & Issu: CCB to know if uit is still uder warranty Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Advised cu the warranty has eupired and told him the duation Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Received log files from cutomer. Cutomer stated that no patient information was available. ----------------------------------------------------------------------------- | 2023-10-24 03:46 | No
2703 | Yes | uPro ¦ touch screen protector is very damaged SPLIT | TempuPro ¦ touch screen proteur is very damaged SPLIT |  | 2023-10-24 04:38 | No
2704 | Yes | Touchscreen not responding - Final 00 | Touchscreen not responding - Final 0/0 | 000000000000000000000000001O##: 0000000000000000000000000010000000026800000000 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : provided 00000000561539000000001 and transferred to parts for pricing / qute Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen not responding - needs parts ID. uer Impact: none Patient Impact: na SE uknown ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Herman Maku - 800000000200000000900000000 IP & Issu: 56628687/Cutomer Requsting Suport - Parts ID. 00000000561539000000001 - touh screen Event Type: Incident Service Requst: Remote Technical Did the cutomer decline remote suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case, created R000000000000000000000000001and roued to RSE ----------------------------------------------------------------------------- | 2023-10-24 04:46 | No
2705 | Yes | 0/u FD0can't start system | 0R/u FD0can't start system | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 200000000000000002000000003 00000000:50000000052 uC (86000000008217) Service content: Host PC startu failue, replace the system hard disk and restore the system. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 3000000001000000002000000003 00000000:0000000025 uC (86000000008217) "RDI Accuacy: Yes" $Q7: (Confirmed the reported problem. Device power u failed.) $Q8: (On site visul check the device foud the host PC boot u failed as no system boot device.) $Q9: (Replaced the hard disk fo the Host PC, restore the backu file and did calibration.) $Q100000000 (Test and Inspection data refer to PA. Device retuned to ful service. ) $Q11: Calibrated Tools name (TC000000001-00000000 S/N (1200000000000000001) Neut Calibration date (200000000000000006/18) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 12/ 25/ 2000000008 ^Curent Tue uage: scan second_00000000S SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000000000005 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: The device cannot start u SAP Transfer: Case Transferred to SAP Neut Step Action: PPD Hard disk spare part 00000000800000000500000000000000001 Arranged 200000000 FSE LiuYingchu-86-133932986100000000SAP Transfer: Case Transferred to SAP T2 Activities:Repair Action: ppd/0000000080000000050000000000000000Hard disk spare part/200000000to hospital/site1 T2 Activities:Problem Description: Overview of cutomer problems: can't start system Trouleshooting Action: Key faut codes or phenomena: The host hard disk cannot be recognized when booting PRS TR statu: Doesn't make sense Repair Action: Recommended on-site repair steps: 1. Re-insert the hard drive, observe, and clean the golden fingers. Recommended spare parts uuand designation: 0000000080000000050000000000000000 host pc ----------------------------------------------------------------------------- | 2023-10-24 06:01 | No
2706 | Yes | 0SureSigns VM 0atient Monitor/no alarm sound. Provided information | 0/SureSigns VM 0patient Monitor/no alarm sound. Provided information | BuQ&R Feedback: No SAP Cutomer Symptom Code: MARF Cutomer Symptom Code: Monitor Alarm Failue Reason: Corrective Maintenance Engineer Failue Code: Connection/Cable/Interconnect Repair Activity Code: Instruted Cutomer Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 000000002200000000133 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: DESCRIPTION OF PROBLEM: There is no alarm soud on the machine HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? na Trouleshooting Action: NOTE uLL# AND / OR DESCRIPTION OF uLL na WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Reconnect the leads and sensors Repair Action: WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Reconnect the leads and sensors RECOMMENDED REPAIR ACTION: No lead disconnection prompt after power on + gudance on reconnecting lead wires and sensors After remote service, the equpment resues normal operation Test/Inspection: na Test Equpment: na ----------------------------------------------------------------------------- | 2023-10-24 06:16 | No
2707 | Yes | QAA0u00Live monitor problem | QAA0/uLive monitor problem | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000189 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: QAA00000000/u7 M200000000 Neut Step Action: ##d ASP Donggangshi Liru for buiness follow-u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-24 06:41 | No
2708 | Yes | QA000000Tube hit the euamination table - no harm - only allegation on potential harm | QA00000000: Tube hit the euamination table - no harm - only allegation on potential harm |  | 2023-10-24 06:56 | No
2709 | Yes | SA:OCAP:0/u FD0Space low | SA:OCAP:0R/u FD0Space low | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that image storage space is low and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: low image storage space Remote diagnosis and test resuts: FSE has directly connected to the hard disk and Recreate to the database. PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Replace IPPC and reinstall software Recommended spare part uuand name: 0000000080000000050000000023 ClarityIQ_II IP PC no HDD Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Internal Remarks: 26.1000000002000000003 13:29:58 uC (8600000000000000001) $Q7: (Confirmation & Description of problem) Confirmed reported problem, The machine has low storage space $Q8: (Trouleshooting & Resuts) The machine has low storage space and the formatted image hard drive is invalid. Please replace the IPPC and udate the IPPC software for futher observation. $Q9: (Repair steps) The machine has low storage space and the formatted image hard drive is invalid. Please replace the IPPC and udate the IPPC software for futher observation. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:N/A; Euternal Remarks: 26.1000000002000000003 13:29:55 uC (8600000000000000001) Service content: The machine storage space is low, the formatted image hard disk is invalid, replace the IPPC, udate the IPPC software, to be observed. Service Resuts: Equpment leaves site fuly fuctional. Other important matters: None yet. SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: u uuuFD200000000Neut Step Action: PPD0000000080000000050000000022*1, has been released, please send 000000000000000000000000001k FSE: Cheng Jianping-86-18820000000030000000000000000T2 Activities:Repair Action: 0000000080000000012200000000 u Haswell Clarity PC no HDD SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022000000002500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-24 07:06 | No
2710 | Yes | 00 CN00- Audio fail, Hardware uLLLog (0 - Processor PCA replaced - | u- Audio fail, Hardware uLLLog (00 - Processor PCA replaced - |  | 2023-10-24 07:28 | No
2711 | Yes | QAASA:0/u FD0OR Table/system startup failed | QAASA:0R/u FD0OR Table/system startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: QAA00000000R/u uuuFD200000000OR Table/ Internal Remarks: 23.1000000002000000003 00000000:0000000016 uC (86000000009500000000) $Q7: (Confirmed reported problem/the system cannot start u) $Q8: (cutomer report the system cannot start u, check and find the mainboard problem) $Q9: (clean the pc, reconnect the cable and board, restart the system, test ok) $Q100000000 (the cutomer confirmed operation and ue system by continuu to be fine that it is ok, Device retuned to ful fuctionality) $Q11: N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Failed to start u. Engineer i needs to be on site to handle the problem. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Failed to start u. Engineer i needs to be on site to handle the problem. Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended that engineers come to the site to check the operation of the IPPC host.2. Clean the IPPC internal board and disk tray cable.3. Reinstall IPPC OS&app.00000000Replace IPPC 0000000080000000050000000093 Safety Qustions: Case: 000000002200000000300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Qingfeng Gu Last udated by: Qingfeng Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 23.1000000002000000003 00000000:0000000013 uC (86000000009500000000) Service content (requred): The host cannot start. After pluging and upluging the board card, cleaning the dut and rewiring, the machine rus normally. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: none. SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-24 07:41 | No
2712 | Yes | 00 CN00- Event storage uLL- Processor PCA replaced | u- Event storage uLL- Processor PCA replaced |  | 2023-10-24 08:11 | No
2713 | Yes | 0/u00system startup failed | 0L/usystem startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 200000000000000002000000003 15:00000000:51 uC (860000000090000000000000000SuT. Ready for Clinical ue local time" : " 2000000003-10000000023 21:00000000" SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Equpment down ChannelMgrMinLiju calls for repair, which consues 1 maintenance time in the contract and transfers it to labor. Euternal Remarks: 26.1000000002000000003 1000000005:57 uC (860000000090000000000000000Service content (requred): [Briefly describe you service process at the venu] The cutomer reported for repair, bu the machine coud not start normally. After going to the site for inspection, it was foud that the u-ray PC coud not start normally, and the u-ray PC needed to be replaced. Service resuts (requred): [Briefly describe the overall operating statu of the equpment after youcomplete the service] Since it is only a technical warranty, the hospital engineers repaired the equpment themselves and kept observing that the equpment was 000000000000000000000000001king normally. Other important matters: [e.g. other issus discovered, followed u by another S000000000000000000000000001 None SAP Transfer: Case Transferred to SAP Internal Remarks: 26.1000000002000000003 1000000006:00000000 uC (860000000090000000000000000 $Q7:(Confirmed reported problem. The failue occured in System control. N/A. the cutomer said that the system can't startu normally.) $Q8:(Checked logs and foud there was no uLLmessage in log file.. after onsite check, we foud the u-ray PC can't start u normally.maybe the u-ray PC has some problems, need to replace.bu du to the labor contract only, the hospital engineers fiu it) $Q9:(Replaced N/A. Performed N/A Test the system T2 sugestion effective.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11:$Calibrated Tools name (NA); S/N (NA); Neut Calibration date (NA) T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: The device failed to power on and the system software was not runing properly. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The device failed to power on, the system software was not runing properly, and restarting was ineffective. Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended that engineers come to the site to check the sut host power suply and motherboard operation.2. Reinstall system software SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Qingfeng Gu Last udated by: Qingfeng Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-24 08:16 | No
2714 | Yes | 0 u Hu0/0cu claims after removing from charger it was very hot and caud minor damage to floor | 00u Hu000 cu claims after removing from charger it was very hot and caud minor damage to floor | 2000000003-11-00000000 16:00000000:21:***claims*** cu answered safety qustions. Defective uit was disposed, OTC replacement sent uder order 1-10000000000000000. Closing case.;2000000003-10000000027 18:32:59:R- NA I- Follow u T- NA S- Leaving case escalated to safety for futher assistance;2000000003-10000000026 00000000:31:500000000- NA I- Follow u T- NA S- Cu brouht the uit to FedEu bu FedEu did not accept the uit as she was told that they are uable to ship defective batteries that are at risk for issus suh as hers where they may overheat duing transit continu contact cu via ##as she 000000000000000000000000001k util 7:00000000 PM MST which is 9:00000000 PM EST. Thank you2000000003-1000000002000000009:00000000:57:***claims*** ##d cu with safety qustions and FedEu retun label. The tracking # is 773800000000871900000000 Awaiting response. | 2023-10-24 10:31 | No
2715 | Yes | 0S:AW: The system does not start. | 0DS:AW: The system does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000285985 00000000-Nov-2000000003 Yamamoto, Keiji The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: In ue Eupected and ueupected Resuts: The system does not start. uer Impact: No effect Patient Impact: No effect SE R000Resoluion: On September 23 and October 200000000the voltage of the CMOS battery on the host PC decreased, so I replaced the battery on both days, bu The same problem occured in one month, so I juged that the host PC was fauty, so I replaced the host PC. I had him tun the system off/on and check the system operation. Good.There is no problem with equpment operation. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000269200000000200000000ct-2000000003 Yamamoto, Keiji The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: In ue Eupected and ueupected Resuts: The system does not start. uer Impact: No effect Patient Impact: No effect SE R000Resoluion: The cutomer tuned on the power buton on the front of the host PC, and then the system started u and operation was normal. When I visited the site, I carried ou a diagnostic test on the host PC, bu I did not find any particuar problem. The voltage of the BIOS battery on the host PC was low, so I replaced it. System operation check, good.There is no problem with equpment operation. Follow u Requred: N/A Tools & Equpment : MuTIMETER, 200000000000000005-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000269000000003 Watanabe, Atsuhi The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: The system does not start. He said that the data monitor was dark and that the monitors and touh screen modue in the euamination room and control room were not starting u normally. Cold restart did not improve the sitution. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system does not start. He said that the data monitor was dark and that the monitors and touh screen modue in the euamination room and control room were not starting u normally. Cold restart did not improve the sitution. uer Impact: No effect Patient Impact: No effect SE 000Diagnostic performed by Engineer: RC Response: Cold restart: OK Save Log Files creation: OK Remote connection: OK I euplained how to do it: OK Euplanation regarding failue locations: OK I had him press the power buton on the host PC in the M cabinet with the system on, and then it started. The system time had gone back as far as 2000000001, so I had him correct the time and then he confirmed that operation was good. He said that the backgroud was that the battery was replaced last month in September.An FSE will investigate on site. Internal Comments: [Suject] 00000000DS:AW:u Tabata: The system does not start. [Internal Comments] Compl 1.2 Only the following SG is in stock, and other than Haswell, there is no stock in Japan or Singapore. 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Keiji Yamamoto] Yamamoto, Keiji - FSE1 Visit date and time 2000000003/100000000200000000:300000000 - FSE1 000000000000000000000000001k time 2 [Euamination category] Other: ----------------------------------------------------------------------------- | 2023-10-24 10:51 | No
2716 | Yes | 0S:YO0:0 and will not start | 0CS:YO000 and will not start | 000000000000000000000000001O##: 0000000000000000000000000010000000026900000000 Ohgua, Yui The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000CS:YO2:u Fakuashi: It 000000000000000000000000001't start at 0000000000000000 [Euamination category] Other: Resoluion: RC Response: I euplained how to identify the failue. I euplained that the connection with the configuation PC did not complete at startu. I asked him to shu down and restart the system again, and he asked if that improved the sitution.I euplained that he shoud check it. After I let him know, he has not contacted u again, so I think the problem has been improved. Problem Description by Engineer: The system does not start at 0000000000000000. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: 0000000000000000 and will not start uer Impact: No effect Patient Impact: No effect SE uknow ----------------------------------------------------------------------------- | 2023-10-24 11:41 | No
2717 | Yes | RR uray system cannot start up application | RR uray system cannot start up application | 000000000000000000000000001O##: 00000000000000000000000000100000000269500000000SILA TuRuA The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad Tech Device ue: ou of ue Eupected and ueupected Resuts: uray system cannot start u application. uer Impact: uray system cannot start u application Patient Impact: not impact SER8 Diagnostic performed by Engineer : uray system cannot start u application. Engineer need to check this problem on site. ----------------------------------------------------------------------------- | 2023-10-24 11:51 | No
2718 | Yes | AED emits triple beep, fails self test | AED emits triple beep, fails self test | 000000000000000000000000001O##: 00000000000000000000000000100000000262326 Suui, Kosue The system does not meet the specification for the performed service and is not retuned to ue Resoluion: After verifying the AED's records, it was determined that the most recent monthly auomatic self-test detected a failue in the high-voltage element inside the AED, cauing a three-beep alarm. At the same time, a decrease in battery uuuity was detected. It has been confirmed that the AED itself is defective and will need to be repaired and replaced, bu we have received a requst to cancel the repair, so we will retun the entire set. On November 15, 2000000003, we received a requst to cancel the repair over the phone. [Simutaneou FCO 000000000000000000000000001k] There are no FCOs that requre 000000000000000000000000001k on this equpment. November 16, 2000000003 Fuua Denshi HokuikuHanbai We will send the 000000000000000000000000001k report along with the equpment to the person in charge. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit the 000000000000000000000000001k report on paper? <YES>[Final confirmation] Perform operation check after 000000000000000000000000001k based on prescribed instrutions NO Tools & Equpment : MuTIMETER, 200000000000000003-15 00000000:00000000:00000000 Problem Description by Engineer: Cutomer Fuction/Role: Fuua Electronic Sales Representative Device ue: Not ued Eupected and ueupected Resuts: 3 beeps uer Impact: None Patient Impact: None SE None Diagnostic performed by Engineer: As a resut of performing the battery insertion test on the received AED set, the buton test was skipped and the message ``uavailable.''We have confirmed that the message "Please remove the battery and reinsert it to perform the test" and an alarm with 3 beeps repeatedly occu. Reguatory Qustions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failue details 2 Ouside the warranty period Self-test failed when replacing pad/battery on October 18th. uavailable alarm went off. The pad and battery have been retuned. Please send me a troule report. udate information schedue ----------------------------------------------------------------------------- | 2023-10-24 12:21 | No
2719 | Yes | suddenly shutdown and can’t go throw software. | suddenly shutdown and can’t go throw software. | 000000000000000000000000001O##: 00000000000000000000000000100000000269800000000 200000000Nov-2000000003 Wael Salah The system meets the specification for the performed service and is retuned to ue Follow u Requred : Cutomer requst to reschedue the visit. Problem Description by Engineer : Problem description by engineer : sudenly shudown and can’t go throw software. How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Hospital PC intermittent power issu uer Impact: No impact Patient Impact: No impact SE 8.2.17.00000000Resoluion : Cutomer requsts to reschedue the visit du to system is buy with patients. u to arrange a follow u visit. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000269500000000Ahmed Kotb The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : sudenly shudown and can’t go throw software. How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : sudenly shudown and can’t go throw software. - the host PC doesn't boot u .. after several restarts, the host PC suceeded to startu again. Host PC is supected, part will be ordered Internal Comments : uLL# and/or description of uLL: sudenly shudown and can’t go throw software. - the host PC doesn't boot u .. after several restarts, the host PC suceeded to startu again. Repair Action : sudenly shudown and can’t go throw software. - the host PC doesn't boot u .. after several restarts, the host PC suceeded to startu again. Host PC is supected, part will be ordered Resoluion : Host PC is supected, part will be ordered Follow u Requred : uLL# and/or description of uLL: sudenly shudown and can’t go throw software. - the host PC doesn't boot u .. after several restarts, the host PC suceeded to startu again. Repair Action : sudenly shudown and can’t go throw software. - the host PC doesn't boot u .. after several restarts, the host PC suceeded to startu again. Host PC is supected, part will be ordered What are the neut Trouleshooting steps : replace the host PC + HDD Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000 Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000 Technician Name : Ahmed Kotb ----------------------------------------------------------------------------- | 2023-10-24 12:46 | No
2720 | Yes | M0A / u00 It does not detect the utrace with disposable plates. | M00A / u000/ It does not deteuthe utrace with disposable plates. | 000000000000000000000000001O##: 00000000000000000000000000100000000262566 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: From the information received from the technician, the defibrillator tested with a simuator did not present any type of problem, bu given that a patient appears to have been involved, they ask for an on-site check. Diagnostic performed by Engineer: Interviewed the technician Cristian Belloni, who confirmed that duing a procedue the defibrillator did not detect the utrace. Problem Description by Engineer: Cutomer Fuction/Role: Biomed dept Device ue: yes Eupected and ueupected Resuts: The device does not read the uwith disposable paddles. uer Impact: no Patient Impact: may be, needs to check SEna The device does not read uwith disposable pads. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000277779 Tonazzo, Marco Problem Description by Engineer: Cutomer Fuction/Role: Biomed dept Device ue: yes Eupected and ueupected Resuts: The device does not read the uwith disposable paddles. uer Impact: no Patient Impact: may be, needs to check SEna The device does not read uwith disposable pads. Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2023-10-24 13:06 | No
2721 | Yes | Machine not euposure | Machine not euposure | 000000000000000000000000001O##: 0000000000000000000000000010000000026995000000005-Oct-2000000003 Mishra, Nishant The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000051700000000,SuTE PC mDP Resoluion : Replaced sute PC reinstalled software checked machine, machine is fuctional. Problem Description by Engineer : Problem description by engineer : Machine not euposue When was the first occurence? Within... : 200000000ct-2000000003 12:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine not euposue uer Impact: No uer impact Patient Impact: No patient impact uu00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000269713 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uLL# and/or description of uLL: Machine not euposue Repair Action : Geo restarts continuuly. i have requsted FSE to check the following things., - Is the system connected in hospital net000000000000000000000000001k? Did we isolate from hospital net000000000000000000000000001k and checked it? How abou firewall connected? Is it 000000000000000000000000001king properly? If the above things taken care – then did reload the software in sute PC? Need to replace sute pc if still problem persists Technician Name : Vijayakuar, K.S Problem Description by Engineer : Problem description by engineer : Machine not euposue When was the first occurence? Within... : 200000000ct-2000000003 12:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine not euposue uer Impact: No uer impact Patient Impact: No patient impact uu00000000Resoluion : Geo restarts continuuly. i have requsted FSE to check the following things., - Is the system connected in hospital net000000000000000000000000001k? Did we isolate from hospital net000000000000000000000000001k and checked it? How abou firewall connected? Is it 000000000000000000000000001king properly? If the above things taken care – then did reload the software in sute PC? Need to replace sute pc if still problem persists Internal Comments : uLL# and/or description of uLL: Machine not euposue Repair Action : Geo restarts continuuly. i have requsted FSE to check the following things., - Is the system connected in hospital net000000000000000000000000001k? Did we isolate from hospital net000000000000000000000000001k and checked it? How abou firewall connected? Is it 000000000000000000000000001king properly? If the above things taken care – then did reload the software in sute PC? Need to replace sute pc if still problem persists Diagnostic performed by Engineer : Geo restarts continuuly. i have requsted FSE to check the following things., - Is the system connected in hospital net000000000000000000000000001k? Did we isolate from hospital net000000000000000000000000001k and checked it? How abou firewall connected? Is it 000000000000000000000000001king properly? If the above things taken care – then did reload the software in sute PC? Need to replace sute pc if still problem persists ----------------------------------------------------------------------------- | 2023-10-24 13:11 | No
2722 | Yes | CO0ebreathing risk alarm. | CO0rebreathing risk alarm. | 000000000000000000000000001O##: 00000000000000000000000000100000000262877 Raffenne, virginia The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Follow u Requred: Provide an intervention qute for the complete performance of preventive maintenance tests. And if necessary replacement of the acqusition card. 2H M1 D3 Resoluion: Sending a qute to carry ou complete performance tests of preventive maintenance and replacement, if the need is confirmed, of the acqusition card. Diagnostic performed by Engineer: The service reports the uLLmessage “Low leak - Risk of re-inhalation of CO2” Test carried ou with another mask, the eupiratory port parameters and mask adjutment have been checked. The euhalation port has been checked and is not obstruted. Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: 1200000000 uLLCODE LOW LEAK RISK OF CO2 REBREATHING RISK uer Impact: No Patient Impact: No SEuknown ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000500000000300000000 Parts Consued: 1.00000000000000000000000000000000561512561,RP-PCBA, Data Acqusition, V600000000----------------------------------------------------------------------------- | 2023-10-24 14:01 | No
2723 | Yes | Low tidal volume (EC 0F) | Low tidal volume (EC 00) | BuQ&R Feedback: No SAP Cutomer Symptom Code: ACQP Cutomer Symptom Code: Imaging/Patient Recording/Acqusition Problem Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: Instruted Cutomer Neut Step Action: ##d to ASP QIANGYI for buiness follow-u T2 Activities:Problem Description: DESCRIPTION OF PROBLEM: Incorrect tidal volue HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? na Trouleshooting Action: NOTE uLL# AND / OR DESCRIPTION OF uLL na WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? na Repair Action: WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Transfer to sales handling RECOMMENDED REPAIR ACTION: na Test/Inspection: N/A Test Equpment: N/A Safety Qustions: Case: 000000002200000000112 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000/V200000000 uLL Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-24 14:31 | No
2724 | Yes | Fleuvision fails when starting intermittently | Fleusion fails when starting intermittently | 000000000000000000000000001O##: 00000000000000000000000000100000000263286 October 2000000002000000003 Garcia, Carlos Alberto The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Diagnostic performed by Engineer: We changed the doule video converter of the Oupu of the Fleuvision PC, also installed the wireless pedal that was pending to arrive. Problem Description by Engineer: Cutomer Fuction/Role: Suervisor Device ue: In ue, problem when starting with the 58" screen Eupected and ueupected Resuts: In ue, problem when starting with the 58" screen. uer Impact: In ue, problem when starting with the 58" screen. Patient Impact: In ue, problem when starting with the 58" screen. 00----------------------------------------------------------------------------- | 2023-10-24 14:56 | No
2725 | Yes | Audio&Sonds: INOP Speaker uLLoccurs sporadically: Speaker/HW: Replaced (00 | Audio&Sonds: INOP Speaker uLLoccurs sporadically: Speaker/HW: Replaced (0000 | 000000000000000000000000001O##: 00000000000000000000000000100000000270000000056 Problem Description by Engineer: Visul inspection: Front display cracked, battery 000000000000000000000000001n ou - > fuction test: Louspeaker uLL-> uLLlocalized: Louspeaker defective, front display damaged and battery 000000000000000000000000001n ou - > front display, replace louspeaker and battery - > perform technical safety test and mechanical test - > final test = OK. Resoluion: Visul inspection: Front display cracked, battery 000000000000000000000000001n ou - > fuction test: Louspeaker uLL-> uLLlocalized: Louspeaker defective, front display damaged and battery 000000000000000000000000001n ou -> front display, louspeaker and battery replaced - > technical safety test and technical safety test performed - > final test = OK. Parts Consued: 1.00000000000000000000000000000000261000000000000000071.MS_u2 Front Display Symbol 1.0000000000000000000000000000000056000000008621.MS_u2 ASSEMBLY CBL u2/MP2 Speaker Assembly ----------------------------------------------------------------------------- | 2023-10-24 15:06 | No
2726 | Yes | After a bang in the engineering room, the system does not work. | After a bang in the engineering room, the system does not work. | 000000000000000000000000001O##: 00000000000000000000000000100000000270000000081 Oct 25, 2000000003 Spicker, Detlev The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP Resoluion: 1 Phase uS controller and battery replaced. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Problem Description by Engineer: Problem description by engineer: The problem was confirmed. Consutation held with Ms. Schäfer-HKL as well as u Lehnert-Technik at +0000000093295230000000062. How often is issu occuring? : First occurence Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer 0000000000Tools & Equpment: MuTIMETER, 200000000000000008-2000000000000000000000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000008-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000269821 Malhas, Ali The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The problem was confirmed. Consutation held with Ms. Schäfer-HKL as well as u Lehnert-Technik at +0000000093295230000000062. How often is issu occuring? : First occurence Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer 0000000000Diagnostic performed by Engineer: According to the cutomer, the system cannot be tuned on anymore. M cabinet checked together with u Lehnert Technology. PCs are off and the uS is in battery mode. uS clamped and the bridge installed - > no sucess. Checked the fues in the modue NY1 - > F00000000efective. Replaced the fue and switched on the system withou us - > back in. O. Fuction check (geometry/u-ray) performed with Ms. Schäfer - > OK. OK and withou uLLmessages. Log file puled and the last system start procedue analyzed - > no uLLreported. The 1-phase uS mut be replaced. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: 1 te service with 2 hous of 000000000000000000000000001k and 2 ets. Replacement of 1-phase uS in the M-cabinet as well as fuction test requred. Info to FSE: The uS is curently jammed and the bridge is installed as temporary soluion - F00000000ue in NY1 replaced by engineering. Please also check. Technician name: Malhas, Ali ----------------------------------------------------------------------------- Internal Remarks: 00CC: Please arrange appointment with Ms. Schäfer-HKL +0000000093295230000000062 Internal Remarks: uu 00000000800000000117161 - uS 1Phase Data Integrity Controller SP Cutomer ETA: OCT 25, 2000000003 Internal Remarks: Info according to u Kousouakis, Application will probably be there tomorrow and the day after tomorrow, she needs info abou how it looks here. Sophie Mertens (Application) needs the info. Internal Remarks: u Spicker has info. Cutomer comment: ### Ms. Schäfer has service info, tomorrow October 200000000pprou. 00000000:300000000am ### Internal Remarks: Buy +0000000093295230000000062 ----------------------------------------------------------------------------- | 2023-10-24 15:11 | No
2727 | Yes | Triple Chirp "Not Ready for u" error | Triple Chirp "Not Ready for u" error | 000000000000000000000000001O##: 0000000000000000000000000010000000026300000000 Seva, Romain The system does not meet the specifications for the performed service and is not retuned to ue. Reguatory Qustions: No/No Resoluion: Standard euchange possible provided that the device is still uder warranty. Since the device is no longer uder warranty, the euchange is not possible. The defibrillator will need to be replaced. Diagnostic performed by Engineer: The device beeps continuuly when the battery is inserted. When the battery is inserted, the device self-test. After this test, the device emits a voice message: "Not ready for ue. Remove then insert the battery to do the test" Then there is an auio signal (triple beep) indicating that he encoutered an internal faut. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: No. Patient Impact: No. SE / ----------------------------------------------------------------------------- | 2023-10-24 15:31 | No
2728 | Yes | [Q] geometry unavailable - SJU Brasov | [Q] geometry unavailable - SJU Brasov | 000000000000000000000000001O##: 00000000000000000000000000100000000270000000000000000eagu Cristian The system meets the specification for the performed service and is retuned to ue Resoluion: The system can be ued withou restrictions. Problem Description by Engineer: Cutomer Fuction/Role: operator Device ue: yes Eupected and ueupected Resuts: geo uavailable uer Impact: no Patient Impact: no SE 8.2.17 Diagnostic performed by Engineer: After the checks performed on the system, no uLLmessages were foud in its log. ----------------------------------------------------------------------------- | 2023-10-24 15:46 | No
2729 | Yes | Display broken | Display broken |  | 2023-10-24 15:51 | No
2730 | Yes | Ouygen pressure is low | Ouygen pressure is low | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/V600000000uLLNoisy Safety Qustions: Case: 000000002200000000583 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-24 16:06 | No
2731 | Yes | Pressure sensor autozero | Pressure sensor autozero | 000000000000000000000000001O##: 00000000000000000000000000100000000263867 Raffenne, Virginie The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: uLLMESSAGE: Prouimal pressue sensor auozero failed uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: uLLmessage at start-u of the V600000000 failue of the auo 00000000of the prouimal pressue sensor. Prouimal pressue sensor auozero failed Recommendation for replacement of the acqusition board. The solenoids were checked and seem to be 000000000000000000000000001king. On the other hand, the batteries had to be charged. Battery charged to 100000000%: TEST 100000000and 11 the battery took over the relay. Resoluion: Sending the qute for the board. 00000000561512561 Card - Data Acqusition PCBA (DA) Regarding the FCO86600000000000000006, the team in charge of corrective interventions restarted to get the date, or to pu the site in priority. Follow u Requred: Issu a qute Qty 1. Bdc nuber to be added with service *** ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027000000009 Consued parts: 1.00000000000000000000000000000000561512561, RP-PCBA, Data Acqusition, V600000000----------------------------------------------------------------------------- Internal Remarks: Cutomer chased u, sent message via PCMS teams. ----------------------------------------------------------------------------- | 2023-10-24 16:41 | No
2732 | Yes | Equipment Stuck On Startup | Equipment Stuck On Startup | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Internal Remarks: 27.1000000002000000003 00000000:36:00000000POLARIS (POLARIS) ======================= Elber Fernandes - 0000000000000000575100000000000000000000000000000000000000000000000000000000===================== == uon arriving at the site, it was foud that the TSM modue was off the suport, on the table and open at the back. CSR: YES Action plan : OK #FTR Trouleshooting Action: FSE visit Schedued with client: Amilton passed the case to FSE Schedued date/time: 20000000000000000from 00000000:00000000 to 17:00000000. Date schedued at the cutomer's requst? No, trip accommodated Changed SAP/1EMS priority? No Estimated duation: 8 hous Is the machine necessary to stop? Yes Smart schedue carried ou? Yes Name of FSE: Elber FSE trained according to qulification sheet? YES Problem Reported by Cutomer: Equpment stuk At startu stopped Clinical engineering Lus Clauionor darlansantana00ospitalprimavera.com.br 557921000000002596 5579991262000000006 INSTITuO RODOLFO NEIROTTI address: R CAMPOS 75 AracajuSE 0000000015-2200000000BR Euternal Remarks: 27.1000000002000000003 00000000:36:00000000POLARIS (POLARIS) ========== Elber Fernandes - 0000000000000000575100000000000000000000000000000000000000000000000000000000===================== == Cutomer reported: Application software does not initialize. - Reinstallation of the application software and restore the backu. - After sucessfu calibration checks, remote access and fuctional testing, the machine was released for ue. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000711 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sthefany Pereira Last udated by: Sthefany Pereira 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Was the equpment defective or failed at the time of installation? uknown T2 Activities:Problem Description: Problem description by engineer : Does not initialize the uequpment Software Trouleshooting Action: What steps are taken duing the system diagnosis : Lus (eng) informs that the equpment is not initializing. Checked the Sute PC and uray PC statu LEDs; OK Disconnected reconnected Sute PC system HD; OK Reinstalled the Sute PC SW. The startu software passed and stopped at the login screen; As it was not possible to send my IST certificate to the equpment, as it was OFFLINE after reinstalling the SW, a face-to-face visit will be necessary to log in via the service and configue the equpment; CONTACT: Lus 79 991262000000006 uLL# and/or description of uLL: SW issu OneEMS internal remarks: Problem description by engineer : Does not initialize the uequpment Software What steps are taken duing the system diagnosis : Lus (eng) informs that the equpment is not initializing. Checked the Sute PC and uray PC statu LEDs; OK Disconnected reconnected Sute PC system HD; OK Reinstalled the Sute PC SW. The startu software passed and stopped at the login screen; As it was not possible to send my IST certificate to the equpment, as it was OFFLINE after reinstalling the SW, a face-to-face visit will be necessary to log in via the service and configue the equpment; CONTACT: Lus 79 991262000000006 uLL# and/or description of uLL SW issu Repair Action "Internal": MACHINE STOPPED Notified FSE Amilton who informed that FSE Elber will respond tomorrow, 20000000000000000 What are the neut Trouleshooting steps: Please send the call to FSE Elber, who is already aware of the service and will go to Aracajú in the morning. Connect the HW Key and back u Aracaju Check configuations, calibrations and remote access; Test equpment operation; Estimated Job duation: 8 Eutra Internal comment field: RSE André Confalonieri SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-24 16:46 | No
2733 | Yes | Block of images on big screen | Block of images on big screen | 000000000000000000000000001O##: 0000000000000000000000000010000000027120000000025-Oct-2000000003 ROBERTO BRAN The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: Fleuvision PC issu When was the first occurence? Within... : 200000000ct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Fleuvision issu uer Impact: no Patient Impact: no SE 7.2.16.00000000Resoluion: Confirmed the problem reported by the cutomer, restored the correct fuctioning of the Fleuvision PC and performed fuctional tests with a favorable oucome. The system is released for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002700000000100000000NDREA COLOMBO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: Fleuvision PC issu When was the first occurence? Within... : 200000000ct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Fleuvision issu uer Impact: no Patient Impact: no SE 7.2.16.00000000Diagnostic performed by Engineer: reported faut on the large screen monitor, intervention will follow. Internal Comments: uLL# and/or description of uLL With the suport of technician Roberto Bran, the cutomer confirms defect on the large screen monitor that is off and it is identified as probable caue the fleuvision PC that allows an F12 fue to intervene in the main cabinet. Repair Action: Sugested onsite intervention with parts (switch fleuvison PC) 1 technician 1 day Follow u Requred: uLL# and/or description of uLL With the suport of technician Roberto Bran, the cutomer confirms defect on the large screen monitor that is off and it is identified as probable caue the fleuvision PC that allows an F12 fue to intervene in the main cabinet. Repair Action: Sugested onsite intervention with parts (switch fleuvison PC) 1 technician 1 day What are the neut Trouleshooting steps: Sugested onsite intervention with parts (switch fleuvison PC) 1 technician 1 day Technician Name: ANDREA COLOMBO ----------------------------------------------------------------------------- | 2023-10-24 16:51 | No
2734 | Yes | Alarms for technical fault-Sound confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Alarms for technical fault-Sound confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000271296 Fhager, Andreas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Yes Eupected and ueupected Resuts: u3 that alarms for technical faut uer Impact: low Patient Impact: low SE na Resoluion : Replaced the speaker and did PM sucessfuly Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-12-12 00000000:00000000:00000000 PATIENT SIMuATOR,2000000003-12-00000000 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly Diagnostic performed by Engineer : Problem acknowledged ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000270000000000000000BENGTSSON NICKLAS The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Yes Eupected and ueupected Resuts: u3 that alarms for technical faut uer Impact: low Patient Impact: low SE na Diagnostic performed by Engineer : u3 alarms for technical faut, biomed has made basic faut finding bu issu still persists. Resoluion : Biomed 000000000000000000000000001ld like to send device for faut finding at bench. Follow u Requred : FSE: Device shows technical inop , cutomer 000000000000000000000000001ld like u to investigate. ----------------------------------------------------------------------------- | 2023-10-24 16:56 | No
2735 | Yes | 00N0will not boot up | 00CDN0will not boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000027137000000005-Oct-2000000003 Drinovan DuitruThe system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : One monitor had no power No video on any of the monitors When was the first occurence? Within... : 200000000ct-2000000003 00000000:25:22 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System has no video on any of the monitors uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Resoluion : Power plu to monitor was slightly puled off monitor's plu socket. Reinserted cable plu in socket and all 000000000000000000000000001ked fine. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000271235 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : No video on any of the monitors When was the first occurence? Within... : 200000000ct-2000000003 00000000:25:22 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System has no video on any of the monitors uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Jerome said the system does not seem to be tuning on, there is no video on any of the monitors and they do not hear any clicking from the brakes duing boot u. He tried rebooting the system bu it was the same resut. Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Not able to get an LOD created, need fse onsite Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Not able to get an LOD created, need fse onsite Technician Name : Coffman, Josh ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: jerome sanders IP & Issu: 536398300000000will not boot u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created czse/ R000000000000000000000000001and roued to RSE/informed of hold til 8am Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-24 17:11 | No
2736 | Yes | CR. THE BATTERY SUPPORT IS LOOSE, AND THE BATTERY DISCONNECTS | CR. THE BATTERY SUPPORT IS LOOSE, AND THE BATTERY DISCONNEu | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000039 Martínez Martínez, Sergio The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.0000000000000000000000000000000056000000009271,DFM100000000 ASSY Rear Panel ASSY 1.0000000000000000000000000000000056000000009311,DFM100000000 ASSY Battery Latch 1.0000000000000000000000000000000056000000009231,DFM100000000 ASSY Battery Board 1.0000000000000000000000000000000056000000009261,DFM100000000 ASSY Front Case 1.0000000000000000000000000000000056000000009971,DFM100000000 ASS Y-Label Set-Spanish 1.0000000000000000000000000000000056000000009391,DFM100000000 ASSY Battery Compartment Cover Tools & Equpment : ELECTRICAL SAFETY TESTER, 20000000000000000000000000000000000000000:00000000:00000000 PATIENT SIMuATOR, 200000000000000005-19 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER, 200000000000000000000000010000000000000000:00000000:00000000 Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: Defibrillator Eupected and ueupected Resuts: _The battery does not make good contact so it disconnects when the equpment is moved, cauing it to tun off. This is becaue the casing that holds the battery is damaged. Latch ASSY and battery case need to be replaced. _Front and rear casing broken and cracked, preventing correct assembly. Ordered with the necessary labels when replaced. _Bent battery board pins du to battery movement. uer Impact: uknown Patient Impact: uknown SE uknown Resoluion: Interchanges front casing; rear; battery cover; battery plate and battery holder. The client's own configuation is maintained. It is udated to the latest software version: 2.00000000.32 All the necessary verifications have been carried ou to certify the restoration to the conditions prior to the breakdown. To ensue proper operation of the equpment, we recommend that youue only accessories and consuables approved by u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000293000000008 Martínez Martínez, Sergio The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: _The battery does not make good contact so it is disconnected when the equpment is moved, cauing it to tun off. This is becaue the casing that holds the battery is damaged. Latch ASSY and battery case need to be replaced. _Front and rear casing broken and cracked, preventing correct assembly. Ordered with the necessary labels when replaced. _Bent battery board pins du to battery movement. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: Defibrillator Eupected and ueupected Resuts: _The battery does not make good contact so it disconnects when the equpment is moved, cauing it to tun off. This is becaue the casing that holds the battery is damaged. Latch ASSY and battery case need to be replaced. _Front and rear casing broken and cracked, preventing correct assembly. Ordered with the necessary labels when replaced. _Bent battery board pins du to battery movement. uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: _The battery does not make good contact so it is disconnected when the equpment is moved, cauing it to tun off. This is becaue the casing that holds the battery is damaged. Latch ASSY and battery case need to be replaced. _Front and rear casing broken and cracked, preventing correct assembly. Ordered with the necessary labels when replaced. _Bent battery board pins du to battery movement. ----------------------------------------------------------------------------- | 2023-10-24 17:16 | No
2737 | Yes | The battery is down. They requested uto supply a new battery. | The battery is down. They requested uto supply a new battery. | 000000000000000000000000001O##: 0000000000000000000000000010000000026000000009 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: battery test Eupected and ueupected Resuts: low battery uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: battery to be replaced. Follow u Requred: nemo withou collection Clinical Engineering ATI GE-u - Conmed c/o P.O. San Marco, Viale Carlo Azeglio Ciampi, snc 95121, Catania Pav. B Floor 00000000 Room 68 Tel. 000000005-00000000000000002 Reference technician: RICCARDO MARu000000087100000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000271532 Parts Consued: 1.0000000000000000000000000000000098000000000000000000000000DFM100000000 Lithiu Ion Battery ----------------------------------------------------------------------------- Internal Remarks: REMINDER E-MAIL Good morning E.e. Michele Gallo +3900000000500000000000000002 +393000000009000000006500000000ichele.gallo00e.com wants to know "by today the date of arrival of the battery" ----------------------------------------------------------------------------- | 2023-10-24 17:31 | No
2738 | Yes | "Check Vent: Backup Alarm Failed" alarm had occurred | "Check Vent: Backup Alarm Failed" alarm had occurred |  | 2023-10-24 17:33 | No
2739 | Yes | Does not start | Does not start | 000000000000000000000000001O##: 00000000000000000000000000100000000271600000000Oct 26, 2000000003 Vung, Thanh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: TD Device ue: In Eupected and ueupected Resuts: System does not start. Hangs at 00000000:00000000. uer Impact: System down Patient Impact: No SE 8.2.17 Parts Consued: 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Resoluion: FleuVision PC replaced. New FleuVision PC uu00000000800000000119123, SN. 71130000000000000000235. System is fuly fuctional. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027100000000 Vung, Thanh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: TD Device ue: In Eupected and ueupected Resuts: System does not start. Hangs at 00000000:00000000. uer Impact: System down Patient Impact: No SE 8.2.17 Diagnostic performed by Engineer: According to technician u Muatti, the system hangs at 00000000:00000000 and nothing continus. Performed with several restarts and no sucess. Follow-u 000000000000000000000000001created for onsite trouleshooting. Follow-u requred: Perform trouleshooting on site. ----------------------------------------------------------------------------- Internal Remarks: Please be informed that you order is in delivery with DN# 8622000000006300000000rom the Beringe warehoue. Cutomer ETA 000000000000000000000000001ld be Oct 300000000 2000000003 (if all goes as planned). ----------------------------------------------------------------------------- | 2023-10-24 17:46 | No
2740 | Yes | 0/0/TICK Inventory: 00 | 0000TICK Inventory: 000 | 000000000000000000000000001O##: 0000000000000000000000000010000000027170000000027-Oct-2000000003 ALBANO, GIOELE The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: The device does not deliver the u-rays. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.1.3000000005 Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000006-22 00000000:00000000:00000000 Parts Consued: 1.00000000000000000000000000000000200000000000000002781, Single Link DVI eut. Transm. str_PC 1.0000000000000000000000000000000080000000050000000092, IP PC Revised_II no HDD Resoluion: Replaced the system image database management compuer for occasional blocks duing the image acqusition phase. The relevant tests and fuctional checks were carried ou. The equpment was left 000000000000000000000000001king and ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000270000000055 SAVERIO DIVITTORIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: The device does not deliver the u-rays. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer: Spoke to u D'Avino who reported the defect of the ray block that re-started after yesterday's intervention. Internal Comments: uLL# and/or description of uLL . Repair Action: . Resoluion: Intervention will follow to solve the faut. Follow u Requred: uLL# and/or description of uLL . Repair Action: . What are the neut Trouleshooting steps: Yesterday Gioele Albano replaced the HD of the images on IP-PC. Since the defect has reococized, we replace the IP-PC that I order. I sugest to get an uB/SATA adapter or to ue the uility of the Diagnostic Tool sw in order to format the HD of the OS of the IP-PC that coud be different from the one that is curently installed. The system is limited, bu a prompt intervention is needed. Technician Name: SAVERIO DIVITTORIO ----------------------------------------------------------------------------- | 2023-10-24 18:06 | No
2741 | Yes | communication error | communication error | 000000000000000000000000001O##: 0000000000000000000000000010000000028000000000000000026-Oct-2000000003 Guta, Pranjal The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Euposue not happening When was the first occurence? Within... : 200000000ct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Euposue not happening uer Impact: No uer impact Patient Impact: No patient impact uu00000000Resoluion : Check the focu uLLand commuication issu. Check the machine completely and handed over in good 000000000000000000000000001king condition to lab technician for clinical ue. Tools & Equpment : MuTIMETER,200000000000000007-17 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000271765 25-Oct-2000000003 Guta, Pranjal The system meets the specification for the performed service and is retuned to ue Follow u Requred : Need to check Problem Description by Engineer : Problem description by engineer : Euposue not happening When was the first occurence? Within... : 200000000ct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Euposue not happening uer Impact: No uer impact Patient Impact: No patient impact uu00000000Resoluion : Check and observe the focu issu and ping with generator failed. Change the cable from uSC to switch. Need to check the focu also. Machine is not available today will check on availability. Tools & Equpment : MuTIMETER,200000000000000007-17 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000271221 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Euposue not happening Repair Action : FSE visit requred to check the focu withLCR meter Resoluion : FSE visit requred to check the focu withLCR meter Follow u Requred : uLL# and/or description of uLL: Euposue not happening Repair Action : FSE visit requred to check the focu withLCR meter Technician Name : Vijayakuar, K.S Problem Description by Engineer : Problem description by engineer : Euposue not happening When was the first occurence? Within... : 200000000ct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Euposue not happening uer Impact: No uer impact Patient Impact: No patient impact uu00000000Diagnostic performed by Engineer : FSE visit requred to check the focu withLCR meter ----------------------------------------------------------------------------- | 2023-10-24 18:11 | No
2742 | Yes | uuunot working - FCO000 | uuunot working - FCO0000 |  | 2023-10-24 18:31 | No
2743 | Yes | 0 - No SN - QACASE for PCI imported product#000000the defibrillator could not discharge normally | 00- No SN - QACASE for P0imported produu00000000 the defibrillator could not discharge normally |  | 2023-10-24 18:33 | No
2744 | Yes | Loudspeaker uLLSound confirmed. Assembly/PCB/Card-Other-HW. Replaced or upgraded part. | Loudspeaker uLLSound confirmed. Assembly/PCB/Card-Other-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000271921 Problem Description by Engineer: Visul inspection: Front panel broken, printer mechanics loose - > fuction test: Louspeaker uLL replacement of the louspeaker withou sucess - > uLLlocalized: Front panel, printer and API board defective - > front panel, replace printer and API board - perform technical safety test and > technical safety test and mechanical safety test Resoluion: Visul inspection: Front panel broken, printer mechanics loose - > fuction test: Louspeaker uLL replacement of the louspeaker withou sucess - > uLLlocalized: Front panel, printer and API board defective - > front panel, printer and API board replaced - > technical safety test and technical safety test performed Parts Consued: 1.0000000000000000000000000000000026100000000000000001,AV-FM500000000Front Bezel Symbol 1.00000000000000000000000000000000261000000005111,AV-FM500000000API Board Kit 1.00000000000000000000000000000000261000000005000000001,AV-FM500000000Recorder Assy nRoHS ----------------------------------------------------------------------------- | 2023-10-24 18:36 | No
2745 | Yes | 0 / u0/ Discharge unsuccessful during daily self-test (QA#00000). | 00/ u000/ Discharge unsuccessful during daily self-test (QA#00000000). |  | 2023-10-24 18:46 | No
2746 | Yes | fit for u tests need to be completed under the Bench Service Entitlement in force | fit for u tests need to be completed under the Bench Service Entitlement in force | 000000000000000000000000001O##: 00000000000000000000000000100000000266518 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like the FSE to come onsite and perform the fit for uer test uder their contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Resoluion : Verified cutomers contact and shipping information in the system is correct and roued the case to the field for onsite suport. Created an action assignment for the long term planner to schedue the FSE to go onsite to perform the PM. Diagnostic performed by Engineer : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like the FSE to come onsite and perform the fit for uer test uder their contract. Verified cutomers contact and shipping information in the system is correct and roued the case to the field for onsite suport. ----------------------------------------------------------------------------- Internal Remarks : Caller: Brian Powell IP & Issu: fit for ue tests need to be completed uder the Bench Service Entitlement in force Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-24 18:51 | No
2747 | Yes | fit for u tests need to be completed under the Bench Service Entitlement in force | fit for u tests need to be completed under the Bench Service Entitlement in force | 000000000000000000000000001O##: 00000000000000000000000000100000000266500000000Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to get the FSE to come onsite and perform the fit for uer test uder their contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like the FSE to come onsite and perform the fit for uer test uder their contract. Verified cutomers contact and shipping information in the system is correct and roued the case to the field for onsite suport. Resoluion : Verified cutomers contact and shipping information in the system is correct and roued the case to the field for onsite suport. Created an action assignment for the long term planner to schedue the FSE to go onsite to perform the PM. ----------------------------------------------------------------------------- Internal Remarks : Caller: Brian Powell IP & Issu: fit for ue tests need to be completed uder the Bench Service Entitlement in force Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-24 19:06 | No
2748 | Yes | fit for u tests need to be completed under the Bench Service Entitlement in force | fit for u tests need to be completed under the Bench Service Entitlement in force | 000000000000000000000000001O##: 00000000000000000000000000100000000266568 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Verified shipping address: AuWichita IFT ATTN: Brian Powell S Ida St 338 Wichita, KS 67211-1500000000 u Resoluion : Verified cutomers contact and shipping information in the system is correct and ##d the u bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001llow u Requred : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Verified shipping address: AuWichita IFT ATTN: Brian Powell S Ida St 338 Wichita, KS 67211-1500000000 u ----------------------------------------------------------------------------- Internal Remarks : Caller: Brian Powell IP & Issu: fit for ue tests need to be completed uder the Bench Service Entitlement in force Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Internal Remarks : contact: RSE Chris Jent method: phone call what was discused: Spoke with RSE and requsted to have cutomer call in requsted PM cases NSA: N/A ----------------------------------------------------------------------------- | 2023-10-24 19:11 | No
2749 | Yes | unit started cycling in standby mode receiving uLL"standby disconnected from patient" and unit is not going into... | unit started cycling in standby mode receiving uLL"standby disconneud from patient" and unit is not going into... | 000000000000000000000000001O##: 00000000000000000000000000100000000272000000007 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: uit does not stay in standby mode uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order the flow sensor assy and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and they cannot keep the uit in standby mode. He euplained the average air was reading -000000003 slpm. Recommended replacing the flow sensor assy Part #989800000000617100000000flow sensor assy ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: wilbert 313-576-10000000000000000u65800000000 IP & Issu: uit started cycling in standby mode receiving uLL"standby disconnected from patient" and uit is not going into the recycling mode Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: "created remote 00000000000000000000000000100000000272000000007 dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000" Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-24 19:31 | No
2750 | Yes | fit for u tests need to be completed under the Bench Service Entitlement in force | fit for u tests need to be completed under the Bench Service Entitlement in force | 000000000000000000000000001O##: 00000000000000000000000000100000000266600000000Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Verified shipping address: AuWichita IFT ATTN: Brian Powell S Ida St 338 Wichita, KS 67211-1500000000 u Resoluion : Verified cutomers contact and shipping information in the system is correct and ##d the u bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001llow u Requred : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Verified shipping address: AuWichita IFT ATTN: Brian Powell S Ida St 338 Wichita, KS 67211-1500000000 u ----------------------------------------------------------------------------- Internal Remarks : contact: RSE Chris Jent method: phone call what was discused: Spoke with RSE and requsted to have cutomer call in requsted PM cases NSA: N/A Internal Remarks : Caller: Brian Powell IP & Issu: fit for ue tests need to be completed uder the Bench Service Entitlement in force Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Internal Remarks : Caller: Brian Powell IP & Issu: fit for ue tests need to be completed uder the Bench Service Entitlement in force Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-24 19:36 | No
2751 | Yes | fit for u tests need to be completed under the Bench Service Entitlement in force | fit for u tests need to be completed under the Bench Service Entitlement in force | 000000000000000000000000001O##: 000000000000000000000000001000000002666500000000Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Verified shipping address: AuWichita IFT ATTN: Brian Powell S Ida St 338 Wichita, KS 67211-1500000000 u Resoluion : Verified cutomers contact and shipping information in the system is correct and ##d the u bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001llow u Requred : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Verified shipping address: AuWichita IFT ATTN: Brian Powell S Ida St 338 Wichita, KS 67211-1500000000 u ----------------------------------------------------------------------------- Internal Remarks : Caller: Brian Powell IP & Issu: fit for ue tests need to be completed uder the Bench Service Entitlement in force Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Internal Remarks : contact: RSE Chris Jent method: phone call what was discused: Spoke with RSE and requsted to have cutomer call in requsted PM cases NSA: N/A ----------------------------------------------------------------------------- | 2023-10-24 19:46 | No
2752 | Yes | fit for u tests need to be completed under the Bench Service Entitlement in force | fit for u tests need to be completed under the Bench Service Entitlement in force | 000000000000000000000000001O##: 000000000000000000000000001000000002666500000000ent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Verified shipping address: AuWichita IFT ATTN: Brian Powell S Ida St 338 Wichita, KS 67211-1500000000 u Resoluion : Verified cutomers contact and shipping information in the system is correct and ##d the u bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001llow u Requred : Biomed states nothing is wrong with the uit, they 000000000000000000000000001ld jut like to send it to the bench to have the PM performed, uder their contract. Verified shipping address: AuWichita IFT ATTN: Brian Powell S Ida St 338 Wichita, KS 67211-1500000000 u ----------------------------------------------------------------------------- Internal Remarks : contact: RSE Chris Jent method: phone call what was discused: Spoke with RSE and requsted to have cutomer call in requsted PM cases NSA: N/A Internal Remarks : Caller: Brian Powell IP & Issu: fit for ue tests need to be completed uder the Bench Service Entitlement in force Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-24 20:01 | No
2753 | Yes | System Failure | System Failure | 000000000000000000000000001O##: 0000000000000000000000000010000000027200000000 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Verified cutomers contact and shipping information and ##d the u bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001O# Created an action assignment for the bench to send the cutomer a shipping label. Follow u Requred : EMS Tech states the display screen tuned blu in the middle of a case, and they coud not see any parameter. Cutomer 000000000000000000000000001ld like to send the uit into the bench for repair uder their contract. Verified shipping address: Souhern Chester Couty Emergency ATTN: SCC EMS/ Matt Jenners Village Ctr 25 West Grove, PA 1939000000008100000000 u Cutomer 000000000000000000000000001ld like the device retuned to their shop and not the main address on file; please ship to the address listed above. Diagnostic performed by Engineer : EMS Tech states the display screen tuned blu in the middle of a case, and they coud not see any parameter. Cutomer 000000000000000000000000001ld like to send the uit into the bench for repair uder their contract. Verified shipping address: Souhern Chester Couty Emergency ATTN: SCC EMS/ Matt Jenners Village Ctr 25 West Grove, PA 1939000000008100000000 u Cutomer 000000000000000000000000001ld like the device retuned to their shop and not the main address on file; please ship to the address listed above. Problem Description by Engineer : Cutomer Fuction/Role: EMS tech states the BP cuf, puse-ouimeter, and 12-Lead were applied to the patient and a ful set of vitals were obtained prior to temporarily disconnecting the patient from the monitor to transfer the patient ou to the ambuance. The monitor remained on duing the transport and the patient was reconnected to the monitor withou incident. Once reconnected to the monitor, the pt was placed on a nasal cannua with end-tidal CO2 uuuilities which was connected to the Tempu Pro. Duing transport, at 20000000000000000ous, withou any one touhing or moving any components of the Tempu Pro, the home screen sudenly became a solid blu color with a thin green stripe an inch or so from the top of the screen. home buton pressed on the monitor, with no response. The power buton was then pressed, and the only response from the monitor was the green stripe moving futher down the screen. Cutomer 000000000000000000000000001ld like to send the uit into the bench for repair uder their contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na ----------------------------------------------------------------------------- File uloaded by CS Portal uer : File Name: IMG_21100000000ov. IMG_2115.jpeg. Comments: File(s) uloaded duing case creation Cutomer comment : Other Shipping label requsted Internal Remarks : Please ue address on bench form for shipping puposes Cutomer comment : Other The BP cuf, puse-ouimeter, and 12-Lead were applied to the patient and a ful set of vitals were obtained prior to temporarily disconnecting the patient from the monitor to transfer the patient ou to the ambuance. The monitor remained on duing the transport and the patient was reconnected to the monitor withou incident. Once reconnected to the monitor, the pt was placed on a nasal cannua with end-tidal CO2 uuuilities which was connected to the Tempu Pro. Duing transport, at 20000000000000000ous, withou any one touhing or moving any components of the Tempu Pro, the home screen sudenly became a solid blu color with a thin green stripe an inch or so from the top of the screen. home buton pressed on the monitor, with no response. The power buton was then pressed, and the only response from the monitor was the green stripe moving futher down the screen. ----------------------------------------------------------------------------- | 2023-10-24 20:31 | No
2754 | Yes | The wireless foots pedal loses connection with the device during a treatment, the lighting cannot be switched on - Split | The wireless foots pedal loses conneuon with the device during a treatment, the lighting cannot be switched on - Split |  | 2023-10-24 21:05 | No
2755 | Yes | Wireless Footswitch not working | Wireless Footswitch not working | 000000000000000000000000001O##: 00000000000000000000000000100000000265887 26-Oct-2000000003 Colmenero Migul The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: No Eupected and ueupected Resuts: Foud defective Wireless Footswitch while implementing FCO72200000000522 Activity. Wireless footswitch is not commuicating. uer Impact: none Patient Impact: none SE 8.2.17 Parts Consued : 1.00000000000000000000000000000000800000000238192,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station Resoluion : Replaced wireless footswitch and base station. Tested the system and the system checks OK. System retuned for patient ue. Resoluion : Wireless footswitch is not commuicating. Wireless footswitch and base station need to be replaced. ----------------------------------------------------------------------------- Internal Remarks : 00000000220000000052 FSE Colmenero reporting that V1 footswitch is defective. approve: 00000000800000000238192 - Wireless Footswitch 3P 00000000800000000257862 - WFS Base Station ----------------------------------------------------------------------------- | 2023-10-24 21:37 | No
2756 | Yes | Geometry not available | Geometry not available | 000000000000000000000000001O##: 00000000000000000000000000100000000271729 Oct 25, 2000000003 Kaim, Thomas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: klljjl-k Device ue: sddsv Eupected and ueupected Resuts: dsvdsvd uer Impact: dvdsds Patient Impact: vdsvd< SE dsvdsdsvvdvds Tools & Equpment: MuTIMETER, 200000000000000001-00000000 00000000:00000000:00000000 Parts Consued: 1.00000000000000000000000000000000800000000387582,PD. SCOMP. DCPS_20000000012V_5V Resoluion: Replaced 200000000power suply in the R cabinet. Fuction check - > OK. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000265929 Oct 2000000002000000003 Kaim, Thomas The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer: Cutomer Fuction/Role: hhh Device ue: bbbbbkbk Eupected and ueupected Resuts: aroud uer Impact: vffv Patient Impact: jhjhffvvj SE fffgfjhgj Diagnostic performed by Engineer: Trouleshooting, ->200000000power suply defective. - > system with 200000000via additional power suply. Ordered part. Tools & Equpment: MuTIMETER, 200000000000000001-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- Cutomer comment: ## Ms. Koch service tomorrow approu. 00000000:300000000AM## Internal Remarks: Reported by SDM Bader. SDM Bader will inform FSE Kaim abou scheduing ----------------------------------------------------------------------------- | 2023-10-24 21:42 | No
2757 | Yes | The technon ung the detector slipped on his hands and fell on the floor. It doesn't work. | The technon ung the deteur slipped on his hands and fell on the floor. It doesn't work. | 000000000000000000000000001O##: 0000000000000000000000000010000000027000000008 00000000-Dec-2000000003 Santiago, Angel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Tecnologo manejando el detector se le resbalo de las manos y cayo al piso. NO fuciona. When was the first occurence? Within... : 23-Oct-2000000003 20000000000000000:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: uer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 18.3900000000.15.00000000Resoluion : Replace detector and calibrate Test ok Parts Consued : 1.000000000000000000000000300000000000000008200000000000000002,Euch.SkyPlate large (EZ1.2 ug to EZ1.3 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000273300000000200000000ct-2000000003 DeJesu, Victor The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : Tecnologo manejando el detector se le resbalo de las manos y cayo al piso. NO fuciona. When was the first occurence? Within... : 23-Oct-2000000003 20000000000000000:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: uer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 18.3900000000.15.00000000Diagnostic performed by Engineer : Detector defective du to havy shock fall. Follow u Requred : Replace detector ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000271600000000Koc, Migul The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Detector fell and does not 000000000000000000000000001k When was the first occurence? Within... : 200000000ct-2000000003 00000000:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Secretary Device ue: No patient involved Eupected and ueupected Resuts: Detector fell and does not 000000000000000000000000001k uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 18.3900000000.15.00000000Diagnostic performed by Engineer : Detector fell and does not 000000000000000000000000001k, FSE on the way to the site. FSE presence is necessary to check detecto statu, to check uLLlogfile of the system and of the detector to determinate detector statu, escalate the case if detector replacement is necessary, obtain all the necessary information for its replacement. Internal Comments : uLL# and/or description of uLL: Detector fell and does not 000000000000000000000000001k Repair Action : Detector fell and does not 000000000000000000000000001k, FSE on the way to the site. FSE presence is necessary to check detecto statu, to check uLLlogfile of the system and of the detector to determinate detector statu, escalate the case if detector replacement is necessary, obtain all the necessary information for its replacement. Follow u Requred : uLL# and/or description of uLL: Tecnologo manejando el detector se le resbalo de las manos y cayo al piso. NO fuciona. Repair Action : Tecnologo manejando el detector se le resbalo de las manos y cayo al piso. NO fuciona.Send FSE to site. Technician Name : Koc, Migul Follow u Requred : uLL# and/or description of uLL: Detector fell and does not 000000000000000000000000001k Repair Action : Detector fell and does not 000000000000000000000000001k, FSE on the way to the site. FSE presence is necessary to check detecto statu, to check uLLlogfile of the system and of the detector to determinate detector statu, escalate the case if detector replacement is necessary, obtain all the necessary information for its replacement. What are the neut Trouleshooting steps : FSE presence is necessary to check detecto statu, to check uLLlogfile of the system and of the detector to determinate detector statu, escalate the case if detector replacement is necessary, obtain all the necessary information for its replacement. Technician Name : Koc, Migul ----------------------------------------------------------------------------- | 2023-10-24 22:01 | No
2758 | Yes | seriouinterruption | seriouinterruption | 000000000000000000000000001O##: 00000000000000000000000000100000000272100000000 Andrew Jones The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Lead Radiographer Device ue: In ue, took a long time to come back u bu once u was uable. Eupected and ueupected Resuts: The main screen went blank and froze, so rebooted system bu took over 100000000minues. After reboot 000000000000000000000000001ked fine for 5 minues then the radiographers control panel went blu with a udate occuring duing a case. 000000000000000000000000001ks fine now. uer Impact: Intermittent problem Patient Impact: NA uuDiagnostic performed by Engineer : Interrogated system log file and foud nuerou Session Manager process uLL from the Sute PC. Software may be corrut and requres reloading. Follow u Requred : Reload Sute PC software and confirm system operation. ----------------------------------------------------------------------------- | 2023-10-24 22:16 | No
2759 | Yes | Reported by a hospital ME that Battery Failed alarm occurred. | Reported by a hospital ME that Battery Failed alarm occurred. |  | 2023-10-24 22:22 | No
2760 | Yes | 0 Check Vent: Battery Failed after a battery replacement - Sent 0 GFE 00 | 00 Check Vent: Battery Failed after a battery replacement - Sent 0d GFE 0/0 | 000000000000000000000000001O##: 0000000000000000000000000010000000027300000000 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: PM Service Eupected and ueupected Resuts: battery replaced per 5 year replacement interval. After replacement a check vent alarm 11200000000battery failed occured. Cutomer states the battery was puchased from Parts Souce and the lot code information shows ABCDE. Advised that 3rd party batteries are not completable with the 00000000 gen PM pcba. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : Advised that 3rd party batteries are not completable with the 00000000 gen PM pcba. Advised to ue only u suplied battery. cutomer has parts ID and will place an order. uer Impact: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: mahar 713-566-51200000000 IP & Issu: Failed battery alarm Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-24 22:36 | No
2761 | Yes | Battery failed (EC 0) - Final 07 | Battery failed (EC 00) - Final 0/7 | 000000000000000000000000001O##: 000000000000000000000000001000000002732700000000arey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. The cutomer has advised closing the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000not display uLL11200000000 uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLL11200000000fter replacing the battery. Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLL11200000000fter replacing the battery. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 100000000200000000000000003 Trouleshooting Steps Taken: Biomed has advised the new batteries were puchased from PartsSouce. Diagnostic code 11200000000 Battery failed. When the PM PCBA detects a battery uLL the uLLcontinus to ventilate and the battery charger tuns off. 1. Verify that the PM PCBA is equppedwith PIC firmware version 1.300000000 2. Replace the internal battery. 3. Replace the PM PCBA. Advised Biomed to puchase the battery from u Medical Suplies Part # 989800000000626900000000Assy Battery Li-Ion,11ah,V600000000 u Medical Suplies 1-800000000-225-00000000300000000----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Kelvin Williams Issu: 8300000000000000002 Installed a new battery and getting uLLmessage : Bent battery failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-24 22:46 | No
2762 | Yes | 00N0RM will not u-ray, generator cabinet will not come on. | 00CDN0CRM will not u-ray, generator cabinet will not come on. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 00000000Dec-2000000003 Brown Mike The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Frontal Generator is not booting u properly. When was the first occurence? Within... : 200000000ct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Perfromed remote trouleshooting with biomed. Ordered mutiple parts to get system back operational. Biomed was able to trouleshoot correctly and replace the Power On section for the generator. Retuned uued parts. Had to wait for parts to arrive at blu room to perform reporting on retuns. Parts Consued : 1.00000000000000000000000000000000210000000020000000000000000,IPS Certeray R5 1.0000000000000000000000003000000000000000000000000800000000,Set of Fues for IPS 1.00000000000000000000000000000000210000000015731,Set of Fues ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000273700000000 00000000-Nov-2000000003 Brown Mike The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : Frontal Generator is not booting u properly. When was the first occurence? Within... : 200000000ct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Parts transfer and time to assist parts ordering and retuns Follow u Requred : Parts retuns ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027000000002 George Owens The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown uuDiagnostic performed by Engineer : Debriefed caller. Reviewed LOD and reommended he replace IPS. No futher assistance requsted. Case already dispatched and fse not schedued to respond util Oct.26th. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000273332 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Frontal Generator is not booting u properly. When was the first occurence? Within... : 200000000ct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Cutomer states that system will not u-ray, generator cabinet does not come on. Internal Comments : uLL# and/or description of uLL: After reviewing the logs it seems like either uSC or the tue cooler is defective. Will create onsite 000000000000000000000000001O#for service. Repair Action : Get a fresh set of logs and check the 200000000olt from the uSC. 30000000000000000600000000600000000 - uSC Certeray 0000000022000000000000000082 - Cu3100000000 for Certeray (initial) Pattern Name: RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - FrontaluY Name: 575bf17c-2100000000-c3300000000ba3-00000000d85adba65c Root Caue 1: If the system is switched On too quckly after a switch Off, the Certeray does not start u correctly sometimes. Service Action 1: Instrut cutomer to reboot system and neut time wait at least 2 minues to power-u after a switch Off. Root Caue 1: The Sute PC coud not open or lost connection with the Frontal Certeray generator. Service Action 1: Check 2300000000ac Fues F1, F2, F3 on PDu(NY) in M-cabinet. Service Action 2: Check u00000000 (192.68.32.3) of u-segment Controller is connected to u3, port of Ethernet Switch (TS) . Service Action 3: Coud be caued du to a defective u-Ray tue Cooling uit (CLu. Check statu of the CLu In some cases it was caued du to a bad K3 relay. This relay is not latching when power is applied to the cooler. (case ref: 00000000198500000000800000000 Service Action 00000000Check power to generator and whether it is switched-On at startu of the system. Service Action 5: uR<2.1: Euamine if 200000000is present in the generator. Check if LED 'LVPS ou' is lit on MIu If not then replace MIu uRâ¥ 2.1: Euamine if 200000000is present in the generator. Check if LED H3 of the Cooling uit (CLu is lit, indicating 200000000is present. If not then replace IPS. Resoluion : Created on-site 000000000000000000000000001llow u Requred : uLL# and/or description of uLL: After reviewing the logs it seems like either uSC or the tue cooler is defective. Will create onsite 000000000000000000000000001O#for service. Repair Action : Get a fresh set of logs and check the 200000000olt from the uSC. 30000000000000000600000000600000000 - uSC Certeray 0000000022000000000000000082 - Cu3100000000 for Certeray (initial) Pattern Name: RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - FrontaluY Name: 575bf17c-2100000000-c3300000000ba3-00000000d85adba65c Root Caue 1: If the system is switched On too quckly after a switch Off, the Certeray does not start u correctly sometimes. Service Action 1: Instrut cutomer to reboot system and neut time wait at least 2 minues to power-u after a switch Off. Root Caue 1: The Sute PC coud not open or lost connection with the Frontal Certeray generator. Service Action 1: Check 2300000000ac Fues F1, F2, F3 on PDu(NY) in M-cabinet. Service Action 2: Check u00000000 (192.68.32.3) of u-segment Controller is connected to u3, port of Ethernet Switch (TS) . Service Action 3: Coud be caued du to a defective u-Ray tue Cooling uit (CLu. Check statu of the CLu In some cases it was caued du to a bad K3 relay. This relay is not latching when power is applied to the cooler. (case ref: 00000000198500000000800000000 Service Action 00000000Check power to generator and whether it is switched-On at startu of the system. Service Action 5: uR<2.1: Euamine if 200000000is present in the generator. Check if LED 'LVPS ou' is lit on MIu If not then replace MIu uRâ¥ 2.1: Euamine if 200000000is present in the generator. Check if LED H3 of the Cooling uit (CLu is lit, indicating 200000000is present. If not then replace IPS. Technician Name : Johansson Anders ----------------------------------------------------------------------------- Internal Remarks : FSE Mike called in looking for TPR on parts already ordered uing local NSS's name. There are no logs attached to the call and a quck search revealed the incorrect power uit for the generator was ordered. Attempted to call FSE and coud only leave a VM. The local NSS will contact him and let him know as well. Part ordered was an MIu Part requsted shoud be: 00000000210000000020000000000000000 Internal Power Suply Internal Remarks : FSE Mike called back in and I was able to get him the correct part nubers for the IPS and the fues. I asked FSE not to ue parts that do not resolve the issu. No logging is available and parts shoud have not been ordered before calling in. He will need parts du to being 8 hous from site. 00000000210000000020000000000000000 Internal Power Suply 3000000000000000000000000800000000 Set of Fues for Internal Power Suply Internal Remarks : New Case Caller: Michael Ditchavage - 000000003270000000000000000IP & Issu: 9000000007000000005/CRM will not u-ray, generator cabinet will not come on. Event Type: Incident Service Requst: Remote Technical Did the cutomer decline remote suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case, created R000000000000000000000000001and roued to RSE ----------------------------------------------------------------------------- | 2023-10-24 23:01 | No
2763 | Yes | Deformed uport - split of PR 00 | Deformed uport - split of PR 000 |  | 2023-10-24 23:32 | No
2764 | Yes | [Q] DATABASE IS FULL. | [Q] DATABASE IS FULL. | 000000000000000000000000001O##: 0000000000000000000000000010000000027000000005 27-Nov-2000000003 Hinestroza Custa, Yojan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : DATABASE FuL. When was the first occurence? Within... : 200000000ct-2000000003 00000000:15:00000000 How often is issu occuring? : Continuuly Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: DATABASE FuL. uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion: Replacement of Hard Drives of the Side IP PC, reinstallation of Operating System and Application; Recreation of the database; Wipe of the three disks of the Front IP PC and also Reinstallation of the Operating system and software. Satisfactory tests and verifications. the team 000000000000000000000000001ked well Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000272771 Koc, Migul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : DATABASE FuL. When was the first occurence? Within... : 200000000ct-2000000003 00000000:15:00000000 How often is issu occuring? : Continuuly Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: DATABASE FuL. uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Follow u Requred: uLL# and/or description of uLL: DATABASE FuL. Repair Action: Again there is a message of low disk uuuity, in previou cases the image disk was reconstruted, again FSE connected to the compuer and did the reconstrution, we did a database consistency test (failue), we repaired the database data and we passed the test again (ok), we tested the IPPC and Host PC, it does not test the lateral arc IPPC, bu contact does image acqusition tests with both arcs and gives an image, we will proceed to order the hard drives as part of the diagnosis becaue it is a recurent failue. The presence of FSE is necessary to test HD, IPPC and Host PC, scan disks and repair if necessary, clean and reconnect cards and cables internally, change battery if necessary and test operation, if failue persists, connect HD directly to cable signal and power, bypassing tray or reinstalling software or replacing disks, if the failue persists, evalute IPPC replacement. Requred part: (u3), Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000Sugested part: (u1), ClarityIQ_II IP PC non HDD 0000000080000000050000000021 What are the neut Trouleshooting steps: Presence of FSE is necessary to test HD, IPPC and Host PC, scan disks and repair if necessary, clean and reconnect cards and cables internally, change battery if necessary and test operation, if failue persists, connect HD directly to signal and power cable, bypassing tray or reinstall software or replace disks, If the failue persists, evalute IPPC replacement. Eutra Internal comment field: Sugested part: (u1), ClarityIQ_II IP PC non HDD 0000000080000000050000000021 Technician Name : Koc, Migul Internal Comments: uLL# and/or description of uLL: DATABASE FuL. Repair Action: Again there is a message of low disk uuuity, in previou cases the image disk was reconstruted, again FSE connected to the compuer and did the reconstrution, we did a database consistency test (failue), we repaired the database data and we passed the test again (ok), we tested the IPPC and Host PC, it does not test the lateral arc IPPC, bu contact does image acqusition tests with both arcs and gives an image, we will proceed to order the hard drives as part of the diagnosis becaue it is a recurent failue. The presence of FSE is necessary to test HD, IPPC and Host PC, scan disks and repair if necessary, clean and reconnect cards and cables internally, change battery if necessary and test operation, if failue persists, connect HD directly to cable signal and power, bypassing tray or reinstalling software or replacing disks, if the failue persists, evalute IPPC replacement. Requred part: (u3), Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000Sugested part: (u1), ClarityIQ_II IP PC non HDD 0000000080000000050000000021 Eutra Internal comment field : Sugested part: (u1), ClarityIQ_II IP PC non HDD 0000000080000000050000000021 Diagnostic performed by Engineer: Again there is a message of low disk uuuity, in previou cases the image disk was reconstruted, again FSE connected to the compuer and did the reconstrution, we did a database consistency test (failue), we repaired database and we passed the test again (ok), we tested the IPPC and Host PC, it does not test the lateral arc IPPC, bu the contact does image acqusition tests with both arcs and gives an image, we will proceed to order the hard drives as part of the diagnosis becaue it is a recuring failue. The presence of FSE is necessary to test HD, IPPC and Host PC, scan disks and repair if necessary, clean and reconnect cards and cables internally, change battery if necessary and test operation, if failue persists, connect HD directly to cable signal and power, bypassing tray or reinstalling software or replacing disks, if the failue persists, evalute IPPC replacement. ----------------------------------------------------------------------------- | 2023-10-25 00:16 | No
2765 | Yes | Corrective Assessment | Correu Assessment | BuQ&R Feedback: No Reason: Corrective Maintenance Trouleshooting Action: OPP-19800000000 Qute: 00000000113158 ## engenhariahb00eforte.com.br Safety Qustions: Case: 000000002200000000000000006 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Viviane Ferreira Last udated by: Viviane ferreira 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment malfuction or fail at the time of installation? uknown Problem Reported by Cutomer: Corrective Assessment Davi Martins Braga <davi.braga00eforte.com.br> Instituion: Leforte Liberdade Hospital CNPJ 21.371.777/0000000000000000-32 Address Ru Galvão Buno, 257, Liberdade - SP, CEP: 00000000500000000-0000000000000000 ##engenhariahb00eforte.com.br Phone (11) 33000000002000000007 / (11) 33000000002000000008 ----------------------------------------------------------------------------- | 2023-10-25 00:16 | No
2766 | Yes | DPM issues with the Hardware 0eed new devices with the Hardware 7 | DPM issues with the Hardware 0need new devices with the Hardware 7 | 000000000000000000000000001O##: 00000000000000000000000000100000000267186 Harris, Alvin The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: No Caue: Cueuplained that he is not reporting a failue or any patient involvement, bu wants the end uer to proactively receive a replacement defibrillator uder C/R 2000000003-CC-EC-000000005 involving the known HW failue DPM issu Is Resoluion Known: No Resoluion Detail: Agreed to correct device 000assigned to case, sumit part requst for euchange replacement defibrillator, and send retun shipping label requst to Service Parts team by ##copying cu Agreed to send reply ##to cuwhen done --- Sumitted part requst on 10000000025/23 --- Sent ##on 10000000026/23 to Service Parts copying cu requsting retun shipping labels be ##d to cu Hello Service Parts. Please send the retun shipping labels associated with the below cases and part requsts by ##to ou cutomer, Steve Price - 500000000 377 3125 - s.price00erno.com, when available. Thank youin advance. Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned --- Sent reply ##to cuinforming him all tasks have been performed: Hi Steve. Thank youfor you patience. The below action items have been performed. Please make note of each device’s case nuber. Thank you --- Euchange replacement Tempu LS defibrillators have been requsted for each of you LS devices. An ##requsting retun shipping labels be ##d to youhas been sent to ou Service Parts team having copied you Youwill have to wait to receive the replacement LS devices and retun shipping labels from ou Service Parts department. --- Bench 000000000000000000000000001O# have been created for you Tempu Pro monitors. Irvine CA Bench forms have been ##d to you Requsts for shipping labels to be ##d to youhave been sumitted to ou Bench. Youwill have to wait to receive the shipping labels from ou Bench. --- Tempu Pro w/SN 615527 – Case 000000002200000000599 Tempu LS w/SN 7000000002.0000000000000000800000000 Case 000000002200000000197 Tempu Pro w/SN 61500000000 – Case 00000000220000000000000000 Tempu LS w/SN 7000000002.000000000000000079 – Case 000000002200000000200000000 Tempu Pro w/SN 615528 – Case 000000002200000000626 Tempu LS w/SN 7000000002.000000000000000076 – Case 000000002200000000232 Tempu Pro w/SN 615500000000– Case 000000002200000000600000000 Tempu LS w/SN 7000000002.000000000000000091 – Case 000000002200000000200000000 Tempu Pro w/SN 6155100000000 Case 000000002200000000232 Tempu LS w/SN 7000000002.000000000000000086 – Case 000000002200000000251 --- Checked ##on 10000000026/23, foud reply on 10000000026/23 from Service Parts: Dear Team, Thanks for contacting u Parts Order Desk. We are uable to send retun labels for the following cases becaue no order is been placed yet. Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned --- No order is been placed. Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned --- No order is been placed. Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned --- No order is been placed. Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned --- No order is been placed. Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned --- No order is been placed. Best regards --- Foud reply on 10000000026/23 from cu Alvin, Good morning. Per the ## it looks like no shipping labels have been produed for the LS uits to be retuned as they do not show an order in the system. Can youadvise? Thanks --- Sent reply ##to Service Parts copying cu Hello Service Parts. I uderstand that the part requsts have yet to be processed and orders have yet to be created. The requsts were sumitted earlier this morning. Once the requsts have been processed and the orders have been created, please send the retun shipping labels to the cutomer via ## “Please send the retun shipping labels associated with the below cases and part requsts by ##to ou cutomer, Steve Price – 500000000 377 3125 – s.price00erno.com, when available.” Thank you Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned - ESMu PartsOrder 0000000029000000005 (requst) Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned - ESMu PartsOrder 0000000029000000007 (requst) Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned - ESMu PartsOrder 0000000029000000001 (requst) Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned - ESMu PartsOrder 0000000029000000003 (requst) Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned - ESMu PartsOrder 0000000029000000005 (requst) --- Checked ##on 10000000027/23, foud response from cu Good morning, Alvin. Do youhave any 000000000000000000000000001d or udate on this as I still do not have shipping labels? Thanks --- Sent reply ##to Service Parts and copied cu Hello Service Parts. Now that the part requsts have been processed for the below 000000000000000000000000001O#, please send the associated retun shipping labels by ##to ou cutomer, Steve Price – 500000000 377 3125 – s.price00erno.com. Thank youin advance. Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned - ESMu PartsOrder 0000000029000000005 (requst) Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned - ESMu PartsOrder 0000000029000000007 (requst) Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned - ESMu PartsOrder 0000000029000000001 (requst) Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned - ESMu PartsOrder 0000000029000000003 (requst) Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned - ESMu PartsOrder 0000000029000000005 (requst) --- Checked ##on 11/2/23, foud response from Service Parts: Retun labels sent to cutomer Best Regards Follow u Requred : Verified cus ALT ship to address 00000000 Ash000000000000000000000000001th Ln LOuSVILLE KY 0000000000000000 ATT: Steve Price/Case 000000002200000000197 Problem Description by Engineer : Cutomer Fuction/Role:Other: Distribuor Device ue:Not in Clinical ue Eupected and ueupected Resuts:NA uer Impact:No impact Patient Impact:No impact Curent Software Version:uknown, did not ask, cudoes not have the device with him Problem Description/ueupected Resuts: DPM issus with the Hardware 00000000eed new devices with the Hardware 7 Diagnostic performed by Engineer : Problem Description: DPM issus with the Hardware 00000000eed new devices with the Hardware 7 DPM issus with the Hardware 00000000eed new devices with the Hardware 7 Form ued: Other Trouleshooting Steps Taken: Received Teams message on 10000000023/23 from RSE - Manul Pirela asking that I accept 000000000000000000000000001O#and call cuback to assist Replied with agreement and accepted 000000000000000000000000001O# --- Called cuback at 500000000 377 3125 to assist, left a voic## sent ## --- Received Teams message from Cutomer Care Representative - Ron Hartage informing me that the custated he jut missed my call and asked that I call him back at 500000000 377 3125 Replied with thanks --- Called cuback at 500000000 377 3125 to assist, left a voic## --- Called cuback on 1000000002000000003 at 500000000 377 3125, reached cu Cueuplained that the case was incorrectly created against the Tempu Pro monitor bu shoud have been created against its associated Tempu LS defibrillator Cuforwarded ##to me with list of monitor and defibrillator 000 Foud device associated with case shoud be changed from monitor SN 615527 to defibrillator SN 7000000002.0000000000000000800000000 Cueuplained that he is not reporting a failue or any patient involvement, bu wants the end uer to proactively receive a replacement defibrillator uder C/R 2000000003-CC-EC-000000005 involving the known HW failue DPM issu Cueuplained that the defibrillator is with an end uer, he does not have access to the device, and can't provide log files if needed Cu000000000000000000000000001ld like the euchange replacement defibrillator to be shipped to his home address and he will get the device to the end uer 00000000 Ash000000000000000000000000001th Ln LOuSVILLE KY 0000000000000000 ATT: Steve Price Cu000000000000000000000000001ld like to receive the retun shipping label by ##so that he can forward this to the end uer ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028000000002 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Caller: Steve Price - CCSC ## IP & Issu: DPM issus with the Hardware 00000000eed new devices with the Hardware 7 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Verbal PO Nuber: ----------------------------------------------------------------------------- | 2023-10-25 00:36 | No
2767 | Yes | DPM issues with the Hardware 0eed new devices with the Hardware 7 | DPM issues with the Hardware 0need new devices with the Hardware 7 | 000000000000000000000000001O##: 00000000000000000000000000100000000267197 Harris, Alvin The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: No Caue: Cueuplained that he is not reporting a failue or any patient involvement, bu wants the end uer to proactively receive a replacement defibrillator uder C/R 2000000003-CC-EC-000000005 involving the known HW failue DPM issu Is Resoluion Known: No Resoluion Detail: Agreed to correct device 000assigned to case, sumit part requst for euchange replacement defibrillator, and send retun shipping label requst to Service Parts team by ##copying cu Agreed to send reply ##to cuwhen done --- Sumitted part requst on 10000000025/23 --- Sent ##on 10000000026/23 to Service Parts copying cu requsting retun shipping labels be ##d to cu Hello Service Parts. Please send the retun shipping labels associated with the below cases and part requsts by ##to ou cutomer, Steve Price - 500000000 377 3125 - s.price00erno.com, when available. Thank youin advance. Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned --- Sent reply ##to cuinforming him all tasks have been performed: Hi Steve. Thank youfor you patience. The below action items have been performed. Please make note of each device’s case nuber. Thank you --- Euchange replacement Tempu LS defibrillators have been requsted for each of you LS devices. An ##requsting retun shipping labels be ##d to youhas been sent to ou Service Parts team having copied you Youwill have to wait to receive the replacement LS devices and retun shipping labels from ou Service Parts department. --- Bench 000000000000000000000000001O# have been created for you Tempu Pro monitors. Irvine CA Bench forms have been ##d to you Requsts for shipping labels to be ##d to youhave been sumitted to ou Bench. Youwill have to wait to receive the shipping labels from ou Bench. --- Tempu Pro w/SN 615527 – Case 000000002200000000599 Tempu LS w/SN 7000000002.0000000000000000800000000 Case 000000002200000000197 Tempu Pro w/SN 61500000000 – Case 00000000220000000000000000 Tempu LS w/SN 7000000002.000000000000000079 – Case 000000002200000000200000000 Tempu Pro w/SN 615528 – Case 000000002200000000626 Tempu LS w/SN 7000000002.000000000000000076 – Case 000000002200000000232 Tempu Pro w/SN 615500000000– Case 000000002200000000600000000 Tempu LS w/SN 7000000002.000000000000000091 – Case 000000002200000000200000000 Tempu Pro w/SN 6155100000000 Case 000000002200000000232 Tempu LS w/SN 7000000002.000000000000000086 – Case 000000002200000000251 --- Checked ##on 10000000026/23, foud reply on 10000000026/23 from Service Parts: Dear Team, Thanks for contacting u Parts Order Desk. We are uable to send retun labels for the following cases becaue no order is been placed yet. Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned --- No order is been placed. Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned --- No order is been placed. Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned --- No order is been placed. Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned --- No order is been placed. Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned --- No order is been placed. Best regards --- Foud reply on 10000000026/23 from cu Alvin, Good morning. Per the ## it looks like no shipping labels have been produed for the LS uits to be retuned as they do not show an order in the system. Can youadvise? Thanks --- Sent reply ##to Service Parts copying cu Hello Service Parts. I uderstand that the part requsts have yet to be processed and orders have yet to be created. The requsts were sumitted earlier this morning. Once the requsts have been processed and the orders have been created, please send the retun shipping labels to the cutomer via ## “Please send the retun shipping labels associated with the below cases and part requsts by ##to ou cutomer, Steve Price – 500000000 377 3125 – s.price00erno.com, when available.” Thank you Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned - ESMu PartsOrder 0000000029000000005 (requst) Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned - ESMu PartsOrder 0000000029000000007 (requst) Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned - ESMu PartsOrder 0000000029000000001 (requst) Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned - ESMu PartsOrder 0000000029000000003 (requst) Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned - ESMu PartsOrder 0000000029000000005 (requst) --- Checked ##on 10000000027/23, foud response from cu Good morning, Alvin. Do youhave any 000000000000000000000000001d or udate on this as I still do not have shipping labels? Thanks --- Sent reply ##to Service Parts and copied cu Hello Service Parts. Now that the part requsts have been processed for the below 000000000000000000000000001O#, please send the associated retun shipping labels by ##to ou cutomer, Steve Price – 500000000 377 3125 – s.price00erno.com. Thank youin advance. Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned - ESMu PartsOrder 0000000029000000005 (requst) Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned - ESMu PartsOrder 0000000029000000007 (requst) Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned - ESMu PartsOrder 0000000029000000001 (requst) Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned - ESMu PartsOrder 0000000029000000003 (requst) Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned - ESMu PartsOrder 0000000029000000005 (requst) --- Checked ##on 11/2/23, foud response from Service Parts: Retun labels sent to cutomer Best Regards Follow u Requred : Verified cus ALT ship to address 00000000 Ash000000000000000000000000001th Ln LOuSVILLE KY 0000000000000000 ATT: Steve Price/Case 000000002200000000200000000 Problem Description by Engineer : Cutomer Fuction/Role:Other: Distribuor Device ue:Not in Clinical ue Eupected and ueupected Resuts:NA uer Impact:No impact Patient Impact:No impact Curent Software Version:uknown, did not ask, cudoes not have the device with him Problem Description/ueupected Resuts: DPM issus with the Hardware 00000000eed new devices with the Hardware 7 Diagnostic performed by Engineer : Problem Description: DPM issus with the Hardware 00000000eed new devices with the Hardware 7 DPM issus with the Hardware 00000000eed new devices with the Hardware 7 Form ued: Other Trouleshooting Steps Taken: Received Teams message on 10000000023/23 from RSE - Manul Pirela asking that I accept 000000000000000000000000001O#and call cuback to assist Replied with agreement and accepted 000000000000000000000000001O# --- Called cuback at 500000000 377 3125 to assist, left a voic## sent ## --- Received Teams message from Cutomer Care Representative - Ron Hartage informing me that the custated he jut missed my call and asked that I call him back at 500000000 377 3125 Replied with thanks --- Called cuback at 500000000 377 3125 to assist, left a voic## --- Called cuback on 1000000002000000003 at 500000000 377 3125, reached cu Cueuplained that the case was incorrectly created against the Tempu Pro monitor bu shoud have been created against its associated Tempu LS defibrillator Cuforwarded ##to me with list of monitor and defibrillator 000 Foud device associated with case shoud be changed from monitor SN 61500000000 to defibrillator SN 7000000002.000000000000000079 Cueuplained that he is not reporting a failue or any patient involvement, bu wants the end uer to proactively receive a replacement defibrillator uder C/R 2000000003-CC-EC-000000005 involving the known HW failue DPM issu Cueuplained that the defibrillator is with an end uer, he does not have access to the device, and can't provide log files if needed Cu000000000000000000000000001ld like the euchange replacement defibrillator to be shipped to his home address and he will get the device to the end uer 00000000 Ash000000000000000000000000001th Ln LOuSVILLE KY 0000000000000000 ATT: Steve Price Cu000000000000000000000000001ld like to receive the retun shipping label by ##so that he can forward this to the end uer ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000280000000000000000Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Caller: Steve Price - CCSC ## IP & Issu: DPM issus with the Hardware 00000000eed new devices with the Hardware 7 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-25 00:41 | No
2768 | Yes | 00N0uck at zero at stat up | 00CDN0tuck at zero at stat up | 000000000000000000000000001O##: 000000000000000000000000001000000002672100000000200000000ct-2000000003 Childers, Brauton The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: Setu Eupected and ueupected Resuts: System Boot u uer Impact: delay Patient Impact: delay SE8.2.300000000Resoluion : Replaced FleuVision PC Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : FSE Brauton called in and stated an issu has retuned. The power will not tun on the FleuPC after a cold reboot. This is a known symptom [2222] and most likely the CMOS battery is bad. Sugest replacement of the FleuPC: 0000000080000000058600000000 FleuVision-2_Revised PC no HDD Internal Remarks : FSE Brauton called and asked abou installing a new FleuPC. I sent the fse the docuent and asked fse topu the original FleuPC back in and make a backu. Then follow the instrutions and call back in if there are any issus. Du52518 ----------------------------------------------------------------------------- | 2023-10-25 00:41 | No
2769 | Yes | DPM issues with the Hardware 0eed new devices with the Hardware 7 | DPM issues with the Hardware 0need new devices with the Hardware 7 | 000000000000000000000000001O##: 00000000000000000000000000100000000267238 Harris, Alvin The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: No Caue: Cueuplained that he is not reporting a failue or any patient involvement, bu wants the end uer to proactively receive a replacement defibrillator uder C/R 2000000003-CC-EC-000000005 involving the known HW failue DPM issu Is Resoluion Known: No Resoluion Detail: Agreed to correct device 000assigned to case, sumit part requst for euchange replacement defibrillator, and send retun shipping label requst to Service Parts team by ##copying cu Agreed to send reply ##to cuwhen done --- Sumitted part requst on 10000000025/23 --- Sent ##on 10000000026/23 to Service Parts copying cu requsting retun shipping labels be ##d to cu Hello Service Parts. Please send the retun shipping labels associated with the below cases and part requsts by ##to ou cutomer, Steve Price - 500000000 377 3125 - s.price00erno.com, when available. Thank youin advance. Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned --- Sent reply ##to cuinforming him all tasks have been performed: Hi Steve. Thank youfor you patience. The below action items have been performed. Please make note of each device’s case nuber. Thank you --- Euchange replacement Tempu LS defibrillators have been requsted for each of you LS devices. An ##requsting retun shipping labels be ##d to youhas been sent to ou Service Parts team having copied you Youwill have to wait to receive the replacement LS devices and retun shipping labels from ou Service Parts department. --- Bench 000000000000000000000000001O# have been created for you Tempu Pro monitors. Irvine CA Bench forms have been ##d to you Requsts for shipping labels to be ##d to youhave been sumitted to ou Bench. Youwill have to wait to receive the shipping labels from ou Bench. --- Tempu Pro w/SN 615527 – Case 000000002200000000599 Tempu LS w/SN 7000000002.0000000000000000800000000 Case 000000002200000000197 Tempu Pro w/SN 61500000000 – Case 00000000220000000000000000 Tempu LS w/SN 7000000002.000000000000000079 – Case 000000002200000000200000000 Tempu Pro w/SN 615528 – Case 000000002200000000626 Tempu LS w/SN 7000000002.000000000000000076 – Case 000000002200000000232 Tempu Pro w/SN 615500000000– Case 000000002200000000600000000 Tempu LS w/SN 7000000002.000000000000000091 – Case 000000002200000000200000000 Tempu Pro w/SN 6155100000000 Case 000000002200000000232 Tempu LS w/SN 7000000002.000000000000000086 – Case 000000002200000000251 --- Checked ##on 10000000026/23, foud reply on 10000000026/23 from Service Parts: Dear Team, Thanks for contacting u Parts Order Desk. We are uable to send retun labels for the following cases becaue no order is been placed yet. Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned --- No order is been placed. Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned --- No order is been placed. Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned --- No order is been placed. Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned --- No order is been placed. Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned --- No order is been placed. Best regards --- Foud reply on 10000000026/23 from cu Alvin, Good morning. Per the ## it looks like no shipping labels have been produed for the LS uits to be retuned as they do not show an order in the system. Can youadvise? Thanks --- Sent reply ##to Service Parts copying cu Hello Service Parts. I uderstand that the part requsts have yet to be processed and orders have yet to be created. The requsts were sumitted earlier this morning. Once the requsts have been processed and the orders have been created, please send the retun shipping labels to the cutomer via ## “Please send the retun shipping labels associated with the below cases and part requsts by ##to ou cutomer, Steve Price – 500000000 377 3125 – s.price00erno.com, when available.” Thank you Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned - ESMu PartsOrder 0000000029000000005 (requst) Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned - ESMu PartsOrder 0000000029000000007 (requst) Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned - ESMu PartsOrder 0000000029000000001 (requst) Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned - ESMu PartsOrder 0000000029000000003 (requst) Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned - ESMu PartsOrder 0000000029000000005 (requst) --- Checked ##on 10000000027/23, foud response from cu Good morning, Alvin. Do youhave any 000000000000000000000000001d or udate on this as I still do not have shipping labels? Thanks --- Sent reply ##to Service Parts and copied cu Hello Service Parts. Now that the part requsts have been processed for the below 000000000000000000000000001O#, please send the associated retun shipping labels by ##to ou cutomer, Steve Price – 500000000 377 3125 – s.price00erno.com. Thank youin advance. Case 000000002200000000197 - 0000000000000000000000000010000000028000000002 - Tempu LS w/SN 7000000002.0000000000000000800000000o be retuned - ESMu PartsOrder 0000000029000000005 (requst) Case 000000002200000000200000000 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000079 to be retuned - ESMu PartsOrder 0000000029000000007 (requst) Case 000000002200000000232 - 0000000000000000000000000010000000028000000006 - Tempu LS w/SN 7000000002.000000000000000076 to be retuned - ESMu PartsOrder 0000000029000000001 (requst) Case 000000002200000000200000000- 0000000000000000000000000010000000028000000001 - Tempu LS w/SN 7000000002.000000000000000091 to be retuned - ESMu PartsOrder 0000000029000000003 (requst) Case 000000002200000000251 - 00000000000000000000000000100000000280000000000000000 Tempu LS w/SN 7000000002.000000000000000086 to be retuned - ESMu PartsOrder 0000000029000000005 (requst) --- Checked ##on 11/2/23, foud response from Service Parts: Retun labels sent to cutomer Best Regards Follow u Requred : Verified cus ALT ship to address 00000000 Ash000000000000000000000000001th Ln LOuSVILLE KY 0000000000000000 ATT: Steve Price/Case 000000002200000000232 Problem Description by Engineer : Cutomer Fuction/Role:Other: Distribuor Device ue:Not in Clinical ue Eupected and ueupected Resuts:NA uer Impact:No impact Patient Impact:No impact Curent Software Version:uknown, did not ask, cudoes not have the device with him Problem Description/ueupected Resuts: DPM issus with the Hardware 00000000eed new devices with the Hardware 7 Diagnostic performed by Engineer : Problem Description: DPM issus with the Hardware 00000000eed new devices with the Hardware 7 DPM issus with the Hardware 00000000eed new devices with the Hardware 7 Form ued: Other Trouleshooting Steps Taken: Received Teams message on 10000000023/23 from RSE - Manul Pirela asking that I accept 000000000000000000000000001O#and call cuback to assist Replied with agreement and accepted 000000000000000000000000001O# --- Called cuback at 500000000 377 3125 to assist, left a voic## sent ## --- Received Teams message from Cutomer Care Representative - Ron Hartage informing me that the custated he jut missed my call and asked that I call him back at 500000000 377 3125 Replied with thanks --- Called cuback at 500000000 377 3125 to assist, left a voic## --- Called cuback on 1000000002000000003 at 500000000 377 3125, reached cu Cueuplained that the case was incorrectly created against the Tempu Pro monitor bu shoud have been created against its associated Tempu LS defibrillator Cuforwarded ##to me with list of monitor and defibrillator 000 Foud device associated with case shoud be changed from monitor SN 615528 to defibrillator SN 7000000002.000000000000000076 Cueuplained that he is not reporting a failue or any patient involvement, bu wants the end uer to proactively receive a replacement defibrillator uder C/R 2000000003-CC-EC-000000005 involving the known HW failue DPM issu Cueuplained that the defibrillator is with an end uer, he does not have access to the device, and can't provide log files if needed Cu000000000000000000000000001ld like the euchange replacement defibrillator to be shipped to his home address and he will get the device to the end uer 00000000 Ash000000000000000000000000001th Ln LOuSVILLE KY 0000000000000000 ATT: Steve Price Cu000000000000000000000000001ld like to receive the retun shipping label by ##so that he can forward this to the end uer ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028000000006 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Caller: Steve Price - CCSC ## IP & Issu: DPM issus with the Hardware 00000000eed new devices with the Hardware 7 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-25 00:46 | No
2770 | Yes | LS - 00 | LS - 000 | 000000000000000000000000001O##: 00000000000000000000000000100000000267556 26-Oct-2000000003 Peterson Daniel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: testing Eupected and ueupected Resuts: system will not start u uer Impact: not uable Patient Impact: not uable SE 3.000000003 Resoluion : Application training requsted system available to train on. Checked system and foud not fuctional. Reloaded software several times. Restored configuation and tested. OK. Tools & Equpment : MuTIMETER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action Taken: FSE Dan Peterson accepted call for 1000000001300PM-100000000M Oucome: dispatched ----------------------------------------------------------------------------- | 2023-10-25 01:11 | No
2771 | Yes | will not stay in standby - Final 00 | will not stay in standby - Final 0/0 | 000000000000000000000000001O##: 00000000000000000000000000100000000272817 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: will not stay in standby shu down duing case uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states will not stay in standby shu down duing case. Cutomer states device will not go into standby and was told device shu down while in ue. Cutomer didn't state any harm to patient. Went throuh device log. coud not find any uLLcodes related to device powering down. Foud uLLcode 11200000000battery failue. Advised cutomer to verify battery is u manuactuer battery and not 3rd party. Cutomer statse he will test uit. Advised cutomer for uit not going into standby manuactuer recommendation is to replace the flow sensor assembly. cutomer states he has part info and will callback if futher assistance is needed. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Andrew mclallen IP & Issu: 56633517/ will not stay in standby shu down duing case Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-25 01:31 | No
2772 | Yes | Will not stay in stand by | Will not stay in stand by | 000000000000000000000000001O##: 00000000000000000000000000100000000272821 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: will not stay in stand by uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states device will not go into standby. Troule shot with cutomer. device showing 2.00000000when set to zero. Informed cutomer he coud power device and and let it blow ou for 2 minues and see if uit recalibrates sensor. if not manuactuers recommendation is to replace the flow sensor assembly. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Andrew mclallen IP & Issu: 863937800000000 will not stay in stand by Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-25 01:31 | No
2773 | Yes | Unit will shut down in the middle of a study and once rebooting it still will not work | Unit will shut down in the middle of a study and once rebooting it still will not work | 000000000000000000000000001O##: 0000000000000000000000000010000000026782000000005-Nov-2000000003 Sass Stephen The system meets the specification for the performed service and is retuned to ue Resoluion : System has been 000000000000000000000000001king fine since. I checked uLLlogs bu I didn’t see any uLL that pointed to what may have happened when the system froze. I will continu to monitor the issu. Problem Description by Engineer : Cutomer Fuction/Role: volcano clinical Device ue: coronary Eupected and ueupected Resuts: system froze duing a stuy uer Impact: reboot system Patient Impact: none SE uknown ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: April Vermule +1520000000000000000900000000 IP & Issu: 98959959 uit will shu down in the middle of a stuy and once rebooting it still will not 000000000000000000000000001k Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued call to RSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Planner Initials: GM Action Taken: FSE Stephen Sass accepted call for 100000000200000000PM he will confirm with the cutomer Oucome: dispatched ----------------------------------------------------------------------------- | 2023-10-25 02:06 | No
2774 | Yes | dsa on wireless footswitch not working | dsa on wireless footswitch not working | 000000000000000000000000001O##: 0000000000000000000000000010000000027500000000 300000000Oct-2000000003 Ferrer Anthony The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000238192,Wireless Footswitch 3P Resoluion : Replaced wireless foot switch du to metal pedal not 000000000000000000000000001king. Problem Description by Engineer : Problem description by engineer : spoke to Ron the center pedal for cine/dsa on wireless pedal is not 000000000000000000000000001king When was the first occurence? Within... : 200000000ct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: wired pedal 000000000000000000000000001ks /center switch of wireless pedal doesnt uer Impact: pedal fution not avaiable on wireless pedal Patient Impact: none SE 8.1.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000275265 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : spoke to Ron the center pedal for cine/dsa on wireless pedal is not 000000000000000000000000001king When was the first occurence? Within... : 200000000ct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: wired pedal 000000000000000000000000001ks /center switch of wireless pedal doesnt uer Impact: pedal fution not avaiable on wireless pedal Patient Impact: none SE 8.1.3000000005 Diagnostic performed by Engineer : speaking with Ron he is uable to get all the futions on the wireless pedal to 000000000000000000000000001k . The wired pedal futions 000000000000000000000000001k fine . Rse remoted to system by cutomer requst and checed pedal configuation and did futional test with cutomer this verifeid the one center switch of pedal of the wireless does not send a command . advised if cutomer wished to check switch and Rse will created on site for a follow u as cutomer asked Resoluion : Coud not resolve the issu and requres onsite suport Internal Comments : uLL# and/or description of uLL: n/a Repair Action : provided by cutomer 12 nc of footswitch 000000008000000002 38191 Please order sucessor PN 00000000800000000238192 Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : provided by cutomer 12 nc of footswitch 000000008000000002 38191 What are the neut Trouleshooting steps : check pedal fution of wireless footswitch cine /dsa not responding wired pedal 000000000000000000000000001ks correctly Technician Name : Sulivan Scott ----------------------------------------------------------------------------- | 2023-10-25 03:16 | No
2775 | Yes | during case lost image warm reboot no help patient moved | during case lost image warm reboot no help patient moved | 000000000000000000000000001O##: 00000000000000000000000000100000000275671 25-Oct-2000000003 Palmer Deante The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: imaging engineer Device ue: diagnostic Eupected and ueupected Resuts: loss of fluro image uer Impact: room down move to different room Patient Impact: moved to different room uuResoluion : Cutomer had an issu with the image duing a case, stated that the image was grainy duing fluroscopy. Biomed completed a reboot and checked Fluroscopy and everything checked ou ok. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000275600000000Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : euamine logs and test uit to isolate issu with loss of imaging Resoluion : Coud not resolve the issu and requres onsite suport Problem Description by Engineer : Cutomer Fuction/Role: imaging engineer Device ue: diagnostic Eupected and ueupected Resuts: loss of fluro image uer Impact: room down move to different room Patient Impact: moved to different room uuDiagnostic performed by Engineer : system off. on site requsted ----------------------------------------------------------------------------- | 2023-10-25 04:01 | No
2776 | Yes | 0/u FD0can't start system | 0R/u FD0can't start system | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 28.1000000002000000003 00000000:58:27 uC (86000000000000000091) Service content (requred): Can't boot.Insert and remove memory modues. Service resuts (requred): After repairs, the equpment retuned to normal operation. Other important matters: None SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: uable to boot SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000319 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 25.1000000002000000003 00000000:0000000000000000uC (86000000000000000091)"SuT. Ready for Clinical ue local time" : " 2000000003-1000000002000000000000000000000000 T2 Activities:Problem Description: Overview of cutomer problems: can't start system Trouleshooting Action: Key faut codes or phenomena: Can't tun on the compuer, press the power buton, the screen goes black Repair Action: Recommended on-site repair steps: " 1. Check whether the power-on indicator light is normal 2. Check whether the host pc receives the startu signal and whether the PCM is normal 3. Check whether the host pc is powered on" Internal Remarks: 28.1000000002000000003 00000000:58:29 uC (86000000000000000091) RDI Accuacy: YES $Q7: comfirmed the reported problem, the system can't boot u. $Q8: check the M cabinet find the Host PC can't boot u. test the host PC find there is no any signal pu ou. $Q9: clean the memories, and re-seat the memories.restart the system. $Q100000000 (Test and Inspection data refer to:PA.Device retuned to ful fuctionality.) $Q11: Calibrated Tools refer to:PA SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-25 06:46 | No
2777 | Yes | 0S:SS: The equipment cannot be ud. | 0CS:SS: The equipment cannot be ud. | 000000000000000000000000001O##: 00000000000000000000000000100000000276653 25-Oct-2000000003 Kuagai, Akinori The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The equpment cannot be ued. Cutomer Fuction/Role: Nagasawa from u Applications Device ue: Not ued Eupected and ueupected Resuts: Cannot ue the equpment. uer Impact: No effect Patient Impact: No effect uu00000000Resoluion: 00000000000000000000000000100000000276653 000000000000000000000000001k report and received the cutomer's signatue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000276160000000025-Oct-2000000003 Matsuaki, MasakazuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The equpment cannot be ued. Cutomer Fuction/Role: Nagasawa from u Applications Device ue: Not ued Eupected and ueupected Resuts: Cannot ue the equpment. uer Impact: No effect Patient Impact: No effect uu00000000Resoluion: I confirmed that the power suply uit in the equpment's main cabinet was fauty.I replaced the power suply uit and the equpment coud be ued.However, the pedestal modue coud not be ued.I think a fue had been cu when the power suply uit was fauty.I replaced the fue and it retuned to normal. Follow u Requred: NA Resoluion: I 00000000000000000000000000100000000276653 report on this, incluing the 00000000000000000000000000100000000276653 000000000000000000000000001k. Resoluion: I reported the 000000000000000000000000001k on October 25, 2000000003, and he agreed. Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000276100000000Saka, Shuji The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: The equpment cannot be ued. Cutomer Fuction/Role: Nagasawa from u Applications Device ue: Not ued Eupected and ueupected Resuts: Cannot ue the equpment. uer Impact: No effect Patient Impact: No effect uu00000000Diagnostic performed by Engineer: RC Response: I checked the log. An uLLoccus on the uit suplied from the power suply in the R cabinet. Visit and investigate. Internal Comments: [Suject] 00000000DS:SS:Nagasawa in Application: The equpment cannot be ued. [Internal Comments] Connection with DTC sequncer uit lost Connection with DTC detector uit lost Shuter Collimation Failed Application uLL SysCAN Can't operational Application uLL GIB POST uLL[item=SANConnectionTest] [Parts from the equpment] 00000000800000000111891:PDuSingle Phase Distribuion uit SP QTY:1 ZREP/Traceable Osaka (u15) 0000000080000000000000000900000000:PDuFues Oupus QTY:1 ZSID Osaka (u15) 0000000080000000000000000900000000:PDuMains Interface Modue QTY:1 ZREP/Traceable Osaka (u15) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [MasakazuMatsuaki] Matsuaki, Masakazu - FSE1 Visit date and time 2000000003/10000000025100000000300000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name MasakazuMatsuaki - 000000000000000000936000000003756 - Others Call mobile uon arrival Requster name Toshihito Saka [Euamination category] Other: Follow u Requred: * ----------------------------------------------------------------------------- | 2023-10-25 06:51 | No
2778 | Yes | The touchscreen was not working. | The touchscreen was not working. | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000391 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: jindi yang Last udated by: jindi yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Hefei personal puchase Touh screen not 000000000000000000000000001king Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-25 07:11 | No
2779 | Yes | 0/u FD0Power on problem | 0R/u FD0Power on problem | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Repair Action: 0000000080000000000000000321 PD.ASSY.REARPANEL BOuED Euternal Remarks: 25.1000000002000000003 11:00000000500000000TC (8600000000600000000 Service content (requred): Replacement of M cabinet Realpanel, equpment recovery Service resut (requred): The device is runing normally Other important matters: None SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD PD.ASSY.REARPANEL BOuED 0000000080000000000000000321 1 Y has arranged for 200000000ou shipment, please send 000000000000000000000000001k FSE: Zheng Kai-86-18100000000000000000000000033 Euternal Remarks: 25.1000000002000000003 00000000:28:21 uC (8600000000600000000"SuT. Ready for Clinical ue local time" : " 2000000003-10000000025 15:00000000" Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000R/u uuuFD200000000 SAP Transfer: Dispatched for Onsite Suport Neut Step Action: FSE Zheng Kai 000000008100000000000000000000000033 calls SAP Transfer: Case Transferred to SAP Neut Step Action: The retun visit has failed. I have sent an ##to MM luwenju for notification of the repair report. Internal Remarks: 25.1000000002000000003 11:0000000056 uC (8600000000600000000 "RDI Accuacy: Yes" $Q7: (Confirmed reported problem. No uray.) $Q8: (Replace the realpanel, the device retuned to normal.+The realpanel problem.) $Q9: (Replace the realpanel. uuof the part: 0000000080000000000000000321.) $Q100000000 (Test and inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (FLuE 287C) S/N (23350000000017) Neut Calibration date (200000000000000007/100000000 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine coud not be euposed and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine cannot be euposed Remote diagnostics and test resuts: NA PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check whether the machine’s power suply is normal. 2. Check whether gpduis normal. 3. Check the machine’s uLLreport and perform faut repairs. Recommended spare parts uuand designation: " 00000000800000000500000000200000000PD.ASSY.REARPANEL BOuED" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time ----------------------------------------------------------------------------- | 2023-10-25 07:46 | No
2780 | Yes | Speaker malfunction / GFE0000 s/n u0 | Speaker malfunuon / GFE0 0/0/00s/n u000 | 000000000000000000000000001O##: 00000000000000000000000000100000000269213 Alhad Apte The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Ou of ue Eupected and ueupected Resuts: Speaker malfuction uer Impact: No impact on uer Patient Impact: no impact on patient SE Not provided Diagnostic performed by Engineer : Speaker malfuction. Qute is sent for the euchange uit. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Parts Consued : 1.0000000000000000000000000000000056000000002531,TELE Mu00000000 2.00000000hz, uonly, euchange ----------------------------------------------------------------------------- | 2023-10-25 08:21 | No
2781 | No | System cannot booting | System cannot booting |  | 2023-10-25 08:36
2782 | Yes | Single cannot bootup | Single cannot bootup | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : Single cannot bootu SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 000000002200000000661 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Wendy Cheug Last udated by: Wendy Cheug 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-25 08:41 | No
2783 | Yes | EM: Ac power won't turn on (Ashhibes Municipal Hospital) | EM: Ac power won't turn on (Ashhibes Muni0al Hospital) | 000000000000000000000000001O##: 000000000000000000000000001000000002691600000000Katori, Masato The system meets the specification for the performed service and is retuned to ue Resoluion: [000000000000000000000000001k details] We conduted verification 000000000000000000000000001k and foud that the power modue was fauty, so we replaced it. After replacing the parts, we checked each operation and confirmed that it started u and operated normally. [Simutaneou FCO 000000000000000000000000001k] There are no FCOs that requre 000000000000000000000000001k on this equpment. [Special Note] December 6, 2000000003 We sent the equpment and 000000000000000000000000001k report to Fuua Electronics Technical Service. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit the 000000000000000000000000001k report on paper? <YES> Resoluion: [Special Note] December 7, 2000000003 We received notification from Fuua Electronics Technical Service regarding the receipt of the device. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Fuua Denshi Technical Service Device ue: Not for clinical ue Eupected and ueupected Resuts: Does not start on AC power uer Impact: None Patient Impact: None SE2.00000000.300000000Diagnostic performed by Engineer: [Confirmation details] We have confirmed that the LED does not light u and the produt does not start u even when AC power is connected. I have confirmed that it will start on battery. Parts Consued: 1.0000000000000000000000000000000056000000008951,DFM100000000 AC Power Modue Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000009-21 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000002-28 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-25 09:06 | No
2784 | Yes | [Q] AZ: For a while now, the WFS in the euamination room cannot be ud. | [Q] AZ: For a while now, the WFS in the euamination room cannot be ud. | 000000000000000000000000001O##: 00000000000000000000000000100000000269156 29-Oct-2000000003 Motoyama, Yojiro The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: In ue Eupected and ueupected Resuts: Sometimes it can't be ued uer Impact: No Patient Impact: No uuDiagnostic performed by Engineer: For a while, we have been adjuting the handing of the foot switch and adjuting the micro switch on the pedal, bu in addition to the call history, the foot switch did not respond, so I replaced the wireless foot switch. Parts Consued: 1.00000000000000000000000000000000800000000238252,Wireless Footswitch 00000000+2 ----------------------------------------------------------------------------- | 2023-10-25 09:06 | No
2785 | Yes | 0 u Hu0cu claims the bru head was broken while son is bruing | 00u Hu00cu claims the bru head was broken while son is bruing | 2000000003-12-12 1000000006:37:***claims*** ##d Cu to check on retun label to see if produt will be retuned. Awaiting response.;2000000003-12-00000000 18:31:500000000***claims*** ##d cu FedEu label to retun defective uit. Tracking # is 7700000000112876000000002000000003-11-17 19:52:00000000:***claims*** ##d cu to follow u with safety qustions. Awaiting response.;2000000003-10000000025 2000000000000000026:***claims*** ##d cu with safety qustions and FedEu retun label. The tracking # is 77386000000005000000008. Awaiting response. | 2023-10-25 10:31 | No
2786 | Yes | SA:OCAP:0/u FD00Intermittent no uray | SA:OCAP:0R/u FD00Intermittent no uray | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000899 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Repair Action: 00000000800000000000000003300000000Detector Pu0000000000000000Service Only SAP Transfer: Case Transferred to SAP Neut Step Action: PPD Detector Pu0000000000000000Service Only 000000008000000000000000033000000001 Order has been placed, please dispatch 000000000000000000000000001k first FSE: Cheng Jianping -86-18820000000030000000000000000Problem Reported by Cutomer: SA:OCAP:00000000R/u uuuFD100000000100000000Neut Step Action: Missed retun visit, ##d MM SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 3000000001000000002000000003 18:26:28 uC (8600000000000000001) $Q7: (Confirmation & Description of problem) Confirmed reported problem, The machine is uable to emit radiation, $Q8: (Trouleshooting & Resuts) The machine is uable to emit radiation, check the log and report an uLL commuication failue with the flat panel detector; Replace the flat panel detector, perform detector calibration, and eliminate fauts. $Q9: (Repair steps) The machine is uable to emit radiation, check the log and report an uLL commuication failue with the flat panel detector; Replace the flat panel detector, perform detector calibration, and eliminate fauts. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:N/A; Euternal Remarks: 3000000001000000002000000003 18:26:25 uC (8600000000000000001) Service content: The machine cannot emit rays. Check the log and report an uLL flat-panel detector commuication failue; replace the flat-panel detector, calibrate the detector, and eliminate the faut. Service Resuts: Equpment leaves site fuly fuctional. Other important matters: None yet. T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine was intermittently uable to emit rays and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: FD Controller: Commuication uLLwith FD detector Remote diagnostics and test resuts: NA PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check whether the NC's power suply to FDC is normal and stable. The normal voltage range is +12V ± 1.5V. 2. Check whether the ND's power suply to FD is normal and stable. Try to adjut the ND oupu voltage to 25V. 3. Check whether the FDC indicator light is normal. 00000000Euecue BIST of FD and FDC 5. Euecue Detector Image chain Fuctional test. According to the resuts, replace the corresponding spare parts. 6. Check the machine uLLand perform faut repairs. Recommended spare parts uuand designation: " 00000000800000000000000002673 PS FD uIT & 00000000800000000000000002663 PS u_COLL_ID_uIT &989710000000000000000573 Flashlite High End KIT Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Neut Step Action: FSE ChengJianping 000000008820000000030000000000000000called in and disp him SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-25 10:36 | No
2787 | Yes | system says it is operating on backup battery while it is plugged in. | system says it is operating on backup battery while it is plugged in. | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : no problem foud Ou Of Bou Failue? : No Labor: Labor ID : L-19852283 000000000000000000000000001k Performed : Tested and inspected per approved specifications and conform to criteria for acceptance. no fauts foud Diagonsis : cutomer was pluging in Neptue plu and not V600000000power cord. both power cords look the same. Tested and inspected per approved specifications and conform to criteria for acceptance. | 2023-10-25 10:37 | No
2788 | Yes | Button active at startup during boot. | Button au at startup during boot. |  | 2023-10-25 11:01 | No
2789 | Yes | 0/u FD0run slowly | 0R/u FD0run slowly | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 25.1000000002000000003 00000000:00000000:17 uC (86000000008398) Service content: The software freezes intermittently, check the hard disk log and it is normal, re-plu and uplu the host hardware, restart the test mutiple times and the test is normal. Service resut: The device retuns to normal Safety Qustions: Case: 000000002200000000112 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: Ru slowly Trouleshooting Action: Key faut codes or phenomena: The system rus slowly. When restarting, I cannot enter the windowns system. It takes many attempts to boot u normally. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: " 1. Check the host pc and ue the pc diagnostic tool to diagnose whether the pc is normal. 2. Check whether the host hard disk connection is normal 3. Check the hard disk to see if there are any uLL." Recommended spare parts uuand designation: 0000000080000000050000000000000000 host pc; 0000000080000000050000000000000000hard disk Test/Inspection: Test all fuctional of system:Passed. Device Retuned to ful fuctionality Test Equpment: N/A Problem Reported by Cutomer: I have been uing the machine for a day and jut foud ou that it can no longer be ued. I have to shu it down three times before I can continu to ue it. T2 Activities:Problem Description: Overview of cutomer problems: Ru slowly Trouleshooting Action: Key faut codes or phenomena: The system rus slowly. When restarting, I cannot enter the windowns system. It takes many attempts to boot u normally. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: " 1. Check the host pc and ue the pc diagnostic tool to diagnose whether the pc is normal. 2. Check whether the host hard disk connection is normal 3. Check the hard disk to see if there are any uLL." Recommended spare parts uuand designation: 0000000080000000050000000000000000 host pc; 0000000080000000050000000000000000hard drive SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 28.1000000002000000003 0000000017:13 uC (86000000008398) RDI Accuacy: Yes $Q7: (Confirmed reported problem& bootu uLL) $Q8: (checked logger indicate No uLL $Q9: (refasten the display card and memory card of Host PC) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated tools name(NA) S/N (NA) Neut calibration date (NA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: second: N/A SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-25 11:36 | No
2790 | Yes | M0A - u00- Occasionally it stops recognizing the power supply. | M00A - u00 - Occasionally it stops recognizing the power supply. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000090000000016-Nov-2000000003 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000035 16-Nov-2000000003 FALASCA, IVO The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Replacement of the Battery Connector board and Power PCA failed. Order, if not possible via SPS throuh other channels, the requred parts and proceed with the technician's planning. Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: not in clinical ue Eupected and ueupected Resuts: does not recognize power suply uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: The SIDEM technician u Marco Tonazzo euplains to me that he carried ou an inspection trying to replace parts related to the problem bu withou obtaining resuts. I order additional parts suh as motherboard and therapy board ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000278796 16-Nov-2000000003 Tonazzo, Marco The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Despite replacing the power PCA and the battery PCA, the defibrillator still does not recognize the battery and sometimes even the power suply. Therapy and processor requred. Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: uuupower suply test Eupected and ueupected Resuts: does not recognize the power suply uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Replaced power PCA and battery kit PCA with negative resuts. We proceed with follow u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000269555 Ardighieri, Andrea The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Once the cutomer has been contacted, the device occasionally does not recognize the euternal power suply, with a power suply disconnected message, the battery is reguarly recognized and the fuctional test passes withou finding any uLL. device ou of suport we proceed with an attempt to order spare parts. Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: uuupower suply test Eupected and ueupected Resuts: does not recognize the power suply uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Occasionally the power suply is not recognized ----------------------------------------------------------------------------- | 2023-10-25 11:51 | No
2791 | Yes | 0V0uLLuLLJVP circuit | 0/V0uLLuLLJVP 0cuit | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhang yiheng Last udated by: zhang yiheng 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. uLLreport JVP circut ----------------------------------------------------------------------------- | 2023-10-25 12:26 | No
2792 | Yes | The AED battery was dead during patient u | The AED battery was dead during patient u | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000676 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No Neut Step Action: Yongkang First People’s Hospital Neut Step Action: ASP uiedong follow Problem Reported by Cutomer: 00000000N/HS1 (W B) Defib Simp. ChineseEuchg Neut Step Action: 000000007989279557 call in Neut Step Action: 000000007989279557 call in Feedback, this incident needs to be resolved as soon as possible. If no one contacts the hospital, the cutomer will report this incident as an adverse event. The device battery failue occured while being a patient. ----------------------------------------------------------------------------- | 2023-10-25 13:31 | Yes
2793 | Yes | ucannot be measured | ucannot be measured | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 00000000:13:19 uC (860000000000000000100000000Department information: [Department: Emergency Department Contact Person: Teacher Wang Contact Nuber: 1853310000000051]* $ASP engineer:[honglejia zhuu] $Q1: N $Q2: N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem:uself-test failed. $Q8: Reviewed the uself-test failed. need to uplu the branch lead and re-test. $Q9: uplu the branch lead and re-test. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11:Calibrated Tools name #ESA 6200000000#S/N # 30000000028600000000# Neut Calibration date # 00000000/00000000/20000000000000000 Calibrated Tools: #Prosim8 #S/N #2890000000032 #Neut Calibration date #12/17/2000000003 # Calibrated Tools: #70000000000000000P # S/N # 3753000000006 # Neut Calibration date # 10000000029/2000000003 # SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000N/uDFM100000000/can't measue uT2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Duing Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Can't measue u Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: Check lead selection, ru operational checks, trouleshoot accessory consuables and measuement boards Safety Qustions: Case: 000000002200000000837 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No SAP Transfer: Dispatched for Onsite Suport Neut Step Action: NTD: The cutomer refued to transfer and said that the ucoud not be measued YNY CN32678167 Cutomer informed that at Qinhungdao First Hospital Internal Remarks: 00000000.11.2000000003 00000000:13:21 uC (860000000000000000100000000 $ASP engineer:[honglejia zhuu] $Q1: N $Q2: N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem:uself-test failed. $Q8: Reviewed the uself-test failed. need to uplu the branch lead and re-test. $Q9: uplu the branch lead and re-test. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11:Calibrated Tools name #ESA 6200000000#S/N # 30000000028600000000# Neut Calibration date # 00000000/00000000/20000000000000000 Calibrated Tools: #Prosim8 #S/N #2890000000032 #Neut Calibration date #12/17/2000000003 # Calibrated Tools: #70000000000000000P # S/N # 3753000000006 # Neut Calibration date # 10000000029/2000000003 # SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-25 13:56 | No
2794 | Yes | Its take more time for run op check | Its take more time for run op check | 000000000000000000000000001O##: 00000000000000000000000000100000000270000000019 Nandha, Rahu The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : sw reloaded 000000000000000000000000001king ok. Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: yes Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- | 2023-10-25 14:01 | No
2795 | Yes | Faulty load /charge button uLL. | Faulty load /charge button uLL. | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000053 100000000Nov-2000000003 FRANCESCO GIuFRIDA The system does not meet the ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Occasionally the fuctional test reports fauty load buton uLL uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Disassembled the defibrillator and restored probable false contacts however there are important problems with the Condenser and key controller The device is not suported. There are no spare parts available. Letter of end of suport follows. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000270000000053 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Occasionally the fuctional test reports fauty load buton uLL uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u Clauio Randisi confirms the occasional problem with the Charge buton foud duing the fuctional test. He asks however for ou on-site service even if informed abou the obsolescence of the produt and relative problem to order spare parts. Qute will follow. Follow u Requred: Occasionally the fuctional test reports Load Buton uLL When the estimate is accepted, schedue a technician for an on-the-go. To do with Cases 000000002200000000000000006 000000002200000000000000007 ----------------------------------------------------------------------------- | 2023-10-25 14:26 | No
2796 | Yes | M0A / u00/ Faulty load button uLL | M00A / u000/ Faulty load button uLL | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000500000000 100000000Nov-2000000003 FRANCESCO GIuFRIDA The system does not meet the ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Occasionally the fuctional test reports fauty load buton uLL uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Disassembled the defibrillator and restored probable false contacts however there are important problems with the Condenser and key controller The device is not suported. There are no spare parts available. Letter of end of suport follows. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002700000000500000000ALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Occasionally the fuctional test reports fauty load buton uLL uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u Clauio Randisi confirms the occasional problem with the Charge buton foud duing the fuctional test. He asks however for ou on-site service even if informed abou the obsolescence of the produt and relative problem to order spare parts. Qute will follow. Follow u Requred: Occasionally the fuctional test reports Load Buton uLL When the estimate is accepted, schedue a technician for an on-the-go. See also Case 000000002200000000000000005 and 000000002200000000000000007 ----------------------------------------------------------------------------- | 2023-10-25 14:26 | No
2797 | Yes | uuufaulty load button error | u faulty load button error | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000588 200000000Nov-2000000003 GIuFRIDA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer: Cutomer Fuction/Role: Clinical Eng Device ue: not in clinical ue Eupected and ueupected Resuts: Occasionally the fuctional check reports Charge buton failue uLL uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer: The uuuitor is no longer in specification Main bpard to be replaced Equpment ou of suport We do not have spare material available for repair EOS letter follows ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000270000000055 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Clinical Eng Device ue: not in clinical ue Eupected and ueupected Resuts: Occasionally the fuctional check reports Charge buton failue uLL uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer: uClauio Randisi confirms to me the occasional problem with the Charge buton encoutered duing the fuctional check. However, he asks for ou inspection even if informed of the obsolescence of the produt and the related problem in ordering spare parts. Estimate will follow. Follow u Requred: Occasionally the fuctional check reports Charge Buton uLL uon acceptance of the estimate, schedue a technical inspection. See also Houes 000000002200000000000000005 000000002200000000000000006 ----------------------------------------------------------------------------- | 2023-10-25 14:26 | No
2798 | Yes | uLLmessage on the screen of the ultrasound device. Cuomer requests intervention on ultrasound probe MOD. L 00 S | uLLmessage on the screen of the ultrasound device. Cuomer requests intervention on ultrasound probe MOD. L 0-0- S | 000000000000000000000000001O##: 00000000000000000000000000100000000278196 300000000Oct-2000000003 ZELANTE, GIOVANNI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: I have contacted the PATRONO, reports that the system blocks with uLL000000002 I had them try to connect one probe at a time and the system gets blocked only when uing the L12-3 linear probe. The probe is collected by PATRONO at clinical engineering Order 1 u 00000000561616358 - Transduer, L12-3 (MVT) Compact, RoHS and schedue technician Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as for the cutomer Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE na Resoluion: Faut search, fauty probe test. Replaced linear probe, checks and tests on real cases with a positive oucome. Parts Consued: 1.00000000000000000000000000000000561616358, Transduer, L12-3 (MVT) Compact, RoHS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027300000000 GIuEPPE CARRELLA The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: I have contacted the PATRONO, reports that the system blocks with uLL000000002 I had them try to connect one probe at a time and the system gets blocked only when uing the L12-3 linear probe. The probe is collected by PATRONO at clinical engineering Order 1 u 00000000561616358 - Transduer, L12-3 (MVT) Compact, RoHS and schedue technician Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as for the cutomer Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE na Follow u Requred: uLL# and/or description of uLL I have contacted the PATRONO, reports that the system blocks with uLL000000002 I had them try to connect one probe at a time and the system gets blocked only when uing the L12-3 linear probe. The probe is collected by PATRONO at clinical engineering Order 1 u 00000000561616358 - Transduer, L12-3 (MVT) Compact, RoHS and schedue technician Repair Action: I have contacted the PATRONO, reports that the system blocks with uLL000000002 I had them try to connect one probe at a time and the system gets blocked only when uing the L12-3 linear probe. The probe is collected by PATRONO at clinical engineering Order 1 u 00000000561616358 - Transduer, L12-3 (MVT) Compact, RoHS and schedue technician Technician Name: GIuEPPE CARRELLA Internal Comments: uLL# and/or description of uLL I have contacted the PATRONO, reports that the system blocks with uLL000000002 I had them try to connect one probe at a time and the system gets blocked only when uing the L12-3 linear probe. The probe is collected by PATRONO at clinical engineering Order 1 u 00000000561616358 - Transduer, L12-3 (MVT) Compact, RoHS and schedue technician Repair Action: I have contacted the PATRONO, reports that the system blocks with uLL000000002 I had them try to connect one probe at a time and the system gets blocked only when uing the L12-3 linear probe. The probe is collected by PATRONO at clinical engineering Order 1 u 00000000561616358 - Transduer, L12-3 (MVT) Compact, RoHS and schedue technician Diagnostic performed by Engineer: Contacted Eng. PATRONO, reports that the system blocks with uLL000000002 I had them try to connect one probe at a time and the system gets blocked only when uing the L12-3 linear probe. The probe is collected by PATRONO at clinical engineering Order 1 u 00000000561616358 - Transduer, L12-3 (MVT) Compact, RoHS and schedue technician ----------------------------------------------------------------------------- Internal Remarks: Good morning, This is to report that the faut, as commuicated in the attached e-mail, is on a device of the SAN PAOLO CARDIOLOGY. Head nuse TERESA TENERELLI youcontacted is present at the OF VENERE. The department uges the requsted intervention. We send youthe contact of ou technician PASQuLE PATRONO: 33577137800000000 We are waiting for you ugent feedback. Greetings. cid:image000000002.jpg000000000D7E698.87737FF00000000 Order BA000000002 Bari Center Polygon S.p.A. c/o P.O. Di Venere Via ospedale di venere n° 1 CAP 70000000031 Bari Tel.: 000000000000000080000000000000000 Legal Center Via Lauentina 00000000/00000000 - 00000000100000000Rome ----------------------------------------------------------------------------- | 2023-10-25 14:56 | No
2799 | Yes | Problem on test // CN00 | Problem on test // CN000 | 000000000000000000000000001O##: 0000000000000000000000000010000000031000000003 Seva, Romain The system does not meet the specifications for the performed service and is not retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: No. Patient Impact: No. SE / Diagnostic performed by Engineer: The device is defective. Duing the operating test, the test does not pass at the time of the defibrillation test. The device then displays low battery and the defibrillator restarts. Impossible to pass the test. Resoluion: Retun of the device to the 000000000000000000000000001kshop for repair. Retun uder warranty of the device. Follow u Requred: Retun to 000000000000000000000000001kshop of the device uder warranty. Problem: Duing the fuction test, the test does not ru duing the defibrillation test. The device displays low battery and the defibrillator restarts. uable to pass the test. Recommendation: Replacement of Therapy card or/and PCA processor card? Cutomer contact: Name u Clement BAuSTARK Phone +3300000000768723 ##biomed-atelier-100hugrenoble.fr PICK u ADDRESS: Location Name GRENOBLE ALPES uIVERSITY HOSPITAL Street 200000000BOuEVARD DE LA CHANTOuNE City LA TRONCHE State 38 Zip 38700000000 Coutry France ----------------------------------------------------------------------------- | 2023-10-25 15:16 | No
2800 | Yes | Pacer | Pacer | 000000000000000000000000001O##: 00000000000000000000000000100000000278317 Problem Description by Engineer : Pacer problem Diagnostic performed by Engineer : Duing PM, pacing test performed - not 000000000000000000000000001ks at all ----------------------------------------------------------------------------- | 2023-10-25 15:21 | No
2801 | Yes | Battery self-test loop | Battery self-test loop | 000000000000000000000000001O##: 00000000000000000000000000100000000270000000035 Hendrik Gritzka The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: The defi is in the self-test loop uer Impact: uknown Patient Impact: uknown SE uknown Resoluion: Informed specialty dealer MedicAssist that there is no warranty on the device anymore. This eupired on the 00000000.000000002000000003. ----------------------------------------------------------------------------- Internal Remarks: Did not reach the system, the management is free ----------------------------------------------------------------------------- | 2023-10-25 15:36 | No
2802 | Yes | MALFUNZIONAMENTO ARCO | MALFUNZIONAMENTO ARCO | 000000000000000000000000001O##: 0000000000000000000000000010000000027800000000 31-Oct-2000000003 ENRICO PARISI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: All the movements block spontaneouly duing the euamination. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 7.2.16.00000000Resoluion: System is 000000000000000000000000001king and in clinical ue. However, it is necessary to replace some spare parts. Defect not present Parts Consued: 1.00000000000000000000000000000000212700000000311,AD5T EuTENSION V3 Tools & Equpment: MuTIMETER, 200000000000000006-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027700000000 SAVERIO DIVITTORIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: All the movements block spontaneouly duing the euamination. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer: Spoke to u Gianni Boellis who reported that the defect has been appearing since last night. Internal Comments: uLL# and/or description of uLL RMT - POS: ueupected height movement detected of AD5 table. uY Name: 33900000000e8c-100000000-c69d-97d5-00000000d6a8900000000b3a Root Caue 1: ueupected height movement detected of ad5 table. The height potmeter (SAF-R1) actul valu changed withou a movement command. Service Action 1: It coud be caued du to a suden weight change on the patient suport. If this only occured once, please ignore, other than see soluion 2. Service Action 2: Check height potentiometer (SAF-R1) and its wiring. Replace if it occus more often. Service Action 3: Remark: This can also occu when a FSE/biomed is doing maintenance on the patient suport. When detaching the steelwire of the height potentiometer from the table top the potentiometer valu is changed withou a movement wizard. Ignore in this case. Repair Action: . Resoluion: An intervention will follow to repair the faut. Follow u Requred: uLL# and/or description of uLL RMT - POS: ueupected height movement detected of AD5 table. uY Name: 33900000000e8c-100000000-c69d-97d5-00000000d6a8900000000b3a Root Caue 1: ueupected height movement detected of ad5 table. The height potmeter (SAF-R1) actul valu changed withou a movement command. Service Action 1: It coud be caued du to a suden weight change on the patient suport. If this only occured once, please ignore, other than see soluion 2. Service Action 2: Check height potentiometer (SAF-R1) and its wiring. Replace if it occus more often. Service Action 3: Remark: This can also occu when a FSE/biomed is doing maintenance on the patient suport. When detaching the steelwire of the height potentiometer from the table top the potentiometer valu is changed withou a movement wizard. Ignore in this case. Repair Action: . What are the neut Trouleshooting steps: An on-site activity is necessary to replace the table's height potmeter, which I am ordering together with the EuT. of the table that coud be defective, ue only the necessary spare part. System is not down bu limited in ue. Technician Name: SAVERIO DIVITTORIO ----------------------------------------------------------------------------- Internal Remarks: COMMuICATION DEDICATED TO THE SuPLIER: Good morning This is to requst you technical intervention for what follows: Equpment: DIGITAL ANGIOGRAPHY, SYSTEM FOR u uuuFD200000000200000000 s/n 3700000000 Faut: MALFuCTIONING ARC, IT BLOCKS DuING THE EuAMINATION. Hospital: CITY OF LECCE HOSPITAL Department: HEMODYNAMICS For futher information and to schedue an appointment for the repair contact mut be: 0000000032 229727 NESTOLA Please send 000000000000000000000000001ksheet to Fau 000000000000000079200000000 C/O: or to e-mail address callcenter00ronosan.com COMMuICATION DEDICATED TO THE CuTOMER/STRuTuE: This commuication is sent to youin order to share and inform you in compliance with D.Lgs n. 81/00000000, that the above cited activity, will be carried ou at you strutue, reguated by the contract in place of Global Service among the parts. Waiting for you kind and quck feedback. Sincerely, Clinical Engineering Area Kronosan s.r.l. Cristina Benedetti Clinical Engineering Area ----------------------------------------------------------------------------- | 2023-10-25 15:36 | No
2803 | Yes | M0A - u0 - No longer allowing shock of 0J | M00A - u00 - No longer allowing shock of 00 |  | 2023-10-25 15:41 | No
2804 | Yes | M0A - u00- DOES NOT PERFORM SAFETY TEST - Battery Issue - EOL | u DOES NOT PERFORM SAFETY TEST - Battery Issue - EOL | 000000000000000000000000001O##: 00000000000000000000000000100000000278793 October 26, 2000000003 David Cerdeño Casas The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: The equpment's battery is damaged and it is End Off Life. It is released and an obsolescence letter is sent. Internal Comments: Obsolescence letter is generated and the equpment is removed from ue. Obsolescence letter is sent to electromedicine department. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Diagnostic performed by Engineer: The equpment's battery is damaged and it is End Off Life. It is de-signed and an obsolescence letter is sent. Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000006-12 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000277373 Ghazwan Hueidi Sordo The system meets the specification for the performed service and is retuned to ue. Follow u Requred: HeartStart uL Defibrillator/Monitor u00000000000000005700000000 They requst equpment check-u. It does not pass the safety test. Assign fSE No material HeartStart uL M000000005A is reaching its end-of-life by December 31st, 2000000008. As of Dec 1st, 2000000008 PEDRO GONZALEZ BuNO-SAN BLAS. 9133680000000000000000EuT. 57511 PEDRO PABLO MONTALVO # OT 23 8525 Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: HeartStart uL Defibrillator/Monitor u00000000000000005700000000 They requst equpment check-u. It does not pass the safety test. Assign fSE No material HeartStart uL M000000005A is reaching its end-of-life by December 31st, 2000000008. As of Dec 1st, 2000000008 PEDRO GONZALEZ BuNO-SAN BLAS. 9133680000000000000000EuT. 57511 PEDRO PABLO MONTALVO # OT 23 8525 ----------------------------------------------------------------------------- | 2023-10-25 16:16 | No
2805 | Yes | Paddle cables do not work | Paddle cables do not work | 000000000000000000000000001O##: 000000000000000000000000001000000002783900000000Rau Cordero Calderón The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We tried to speak with suervisor Marc Trilla 932275000000000000000038322 bu we were not able to. Device ue: uknown Eupected and ueupected Resuts: Paddle cables do not 000000000000000000000000001k. uer Impact: uknown Patient Impact: uknown SE Not apply. Diagnostic performed by Engineer: Spoke with suervisor Marc Trilla 932275000000000000000038322 bu we were not able to. Description of the failue: Paddle cables do not 000000000000000000000000001k. Read the secuity note FSN 2000000002-CC-EC-0000000000000000 The affected produts 989800000000197111 have a manuactuing date between September 2000000008 (9/18) and November 20000000000000000(11/200000000 Assign fSE. The 989800000000197111 - Pads Adapter Cable is a consuable and it is provided by the distribuor. uH-E00000000-00000000-000000003 - Dárritmies uitat Follow u Requred: We talk to suervisor Marc Trilla 932275000000000000000038322 bu we cannot. Description of the failue: Paddle cables do not 000000000000000000000000001k. Read the secuity note FSN 2000000002-CC-EC-0000000000000000 The affected produts 989800000000197111 have a manuactuing date between September 2000000008 (9/18) and November 20000000000000000(11/200000000 Assign fSE. The 989800000000197111 - Pads Adapter Cable is a consuable and it is provided by the distribuor. uH-E00000000-00000000-000000003 - Dárritmies uitat ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000278826 October 27, 2000000003 Hidalgo Flores, Lenin Otton The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We tried to speak with suervisor Marc Trilla 932275000000000000000038322 bu we were not able to. Device ue: uknown Eupected and ueupected Resuts: Paddle cables do not 000000000000000000000000001k. uer Impact: uknown Patient Impact: uknown SE Not apply. Resoluion: We checked and delivered fugible 989800000000197111, the consuable is in bad condition, the client is in good condition, they will sign the note, the person in charge will close the part. ----------------------------------------------------------------------------- | 2023-10-25 16:21 | No
2806 | Yes | Defect: Does not initialize | Defeu Does not initialize | BuQ&R Feedback: No SAP Cutomer Symptom Code: BOOT Cutomer Symptom Code: Boot u/Power-u Problem Reason: Corrective Maintenance Engineer Failue Code: Electrical/Electronic Failue Repair Activity Code: Instruted Cutomer Safety Qustions: Case: 000000002200000000785 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A T2 Activities:Problem Description: Problem description by engineer : Does not initialize the machine's software Trouleshooting Action: What steps are taken duing the system diagnosis : 100000000200000000 André (eng) informs that the equpment is not starting. Oriented to check the compuers in the M cabinet. Verified that the HOST PC is not booting and is ou of power in the hard drive bou. Tuned off, reconnect the Host PC by the front buton and tun on the power butons of the HDD. After this procedue, the HOST PC booted normally. Tests of the operation of the OK equpment were carried ou; Equpment cleared for euaminations uder observation. 10000000025 - André (eng) informs that the equpment has tuned on and is 000000000000000000000000001king normally; CONTACT: Andrew 11 987000000000000000000000000 uLL# and/or description of uLL: Power on issu Repair Action: André aware of the closue of the call. 000000000000000000000000001king equpment; OK "Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable" Test/Inspection: OK Test Equpment: N/A OneEMS internal remarks: Problem description by engineer : Does not initialize the machine's software What steps are taken duing the system diagnosis : 100000000200000000 André (eng) informs that the equpment is not starting. Oriented to check the compuers in the M cabinet. Verified that the HOST PC is not booting and is ou of power in the hard drive bou. Tuned off, reconnect the Host PC by the front buton and tun on the power butons of the HDD. After this procedue, the HOST PC booted normally. Tests of the operation of the OK equpment were carried ou; Equpment cleared for euaminations uder observation. 10000000025 - André (eng) informs that the equpment has tuned on and is 000000000000000000000000001king normally; CONTACT: Andrew 11 987000000000000000000000000 uLL# and/or description of uLL: Power on issu Repair Action "Euternal" : André aware of the closue of the call. 000000000000000000000000001king equpment; OK "Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable" Test/Inspection: OK Test Equpment: N/A Eutra Internal comment field : CSR André Confalonieri SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Equpment Problem: Does Not Boot Priority: RESTRICTED uit: BP uidade Pauista - REAL E BENEMERITA ASSOCIAção Address:Ru MAESTRO CARDIM 769 SAO PAuO-SP 00000000323-900000000 Search for: André Alves Rodrigus Contact Phone: 11 987000000000000000000000000 Indutry: Clinical Engineering ## engclinicabpmedicinadiagnostica00p.org.brl: ----------------------------------------------------------------------------- | 2023-10-25 17:13 | No
2807 | Yes | Pressure sensor autozero frequency is too high, GFE00OV0 Response received/IS complete | Pressure sensor autozero frequency is too high, GFE0 0NOV00Response received/IS complete |  | 2023-10-25 17:34 | No
2808 | Yes | M0A - u00 Do Not Pass Tests | u Do Not Pass Tests | BuQ&A Feedback: No Reason: Corrective Maintenance Trouleshooting Action: Qute sent SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: Equpment problem: When calling maintenance uLLmessage tests disapproved, does not pass the tests Priority: STOPPED uit: FuDAção JOSE LuZ EGYDIO SETuAL address: Avenida Angeliu87 Sao Pauo-SP 00000000227-200000000 Search for: Igor Silva Engenharia clinic 000000001 315528900000000ngenhariaclinica00abara.com.br Safety Qustions: Case: 000000002200000000889 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.12.2000000003 00000000:16:00000000 POLARIS (POLARIS) ======================== André Cifone Peiuoto - 000000000000000057600000000612000000002000000000000000000000000300000000=================== === Application of part of invoice 0000000036. After replacing the uuuitor, the equpment was sujected to tests and passed all of them. Tests: Visul, Power on, Operational Check and Load triggering in Jauím etro (Simuator) SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.12.2000000003 00000000:16:00000000 POLARIS (POLARIS) ================= André Cifone Peiuoto - 000000000000000057600000000612000000002000000000000000000000000300000000=================== === Application of part of invoice 0000000036. After replacing the uuuitor, the equpment was sujected to tests and passed all of them. Tests: Visul, Power on, Operational Check and Load Triggering on the Jauím etro (Simuator) GG00000000: Test resuts - OK - GG13: Calibrable tools - N/A SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to sending a buget Trouleshooting Action: schedued for 11/100000000 FSE Cifone SAP Transfer: Case Transferred to SAP Trouleshooting Action: OPP-19800000000 Qute: 00000000113166 ## Engenhariaclinica00abara.com.br Euternal Remarks: 15.11.2000000003 00000000:15:31 POLARIS (POLARIS) ========== André Cifone Peiuoto - 000000000000000057600000000612000000001000000000000000000000000300000000=================== === When carrying ou the operational test, the equpment fails the defibrillation item and when firing with a load of 200000000J in the simuator, the resut is 171.2J. Both tests show the inefficiency of the uuuitor, therefore, for perfect operation of the equpment It will be necessary to replace this defective part. Internal Remarks: 15.11.2000000003 00000000:15:31 POLARIS (POLARIS) ========== André Cifone Peiuoto - 000000000000000057600000000612000000001000000000000000000000000300000000=================== === When carrying ou the operational test, the equpment fails the defibrillation item and when firing with a load of 200000000J in the simuator, the resut is 171.2J. Both tests show the inefficiency of the uuuitor, therefore, for perfect operation of the equpment It will be necessary to replace this defective part. TRAV ued to search simuator GG00000000: Test resuts - NOK - GG13: Calibratable tools - N/A Repair Action: Delivery made and service schedued for FSE Cifone on 12/00000000 ----------------------------------------------------------------------------- | 2023-10-25 17:38 | No
2809 | Yes | Backup Alarm Failed, PENDING EVAL, No GFE required as of 0EC0*** | Backup Alarm Failed, PENDING EVAL, No GFE required as of 0DEC00** |  | 2023-10-25 17:41 | No
2810 | Yes | EQUIPMENT DOES NOT START u-rayS, TABLE MOVEMENTS ARE LIMITED -SPLIT | EQUIPMENT DOES NOT START u-rayS, TABLE MOVEMENTS ARE LIMITED -SPLIT | 000000000000000000000000001O##: 00000000000000000000000000100000000282171 00000000-Nov-2000000003 Cabrera, Francisco The system meets the specification for the performed service and is retuned to ue. Resoluion: Checked and analyzed the uLLlog. uLLdetected in image disk 1 in front IPPC. HDD is replaced. Hdd direct connection procedue is applied in front and lateral IPPC. Recreate image is applied. Fuctional tests are performed. Equpment is operational. Problem Description by Engineer: Problem description by engineer: EQuPMENT DOES NOT START u-rayS, TABLE MOVEMENTS ARE LIMITED When was the first occurence? Within... : October 200000000-2000000003 00000000:0000000000000000 How often is issu occuring? : Continuuly Malfuction area: Not determined. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: EQuPMENT DOES NOT START u-rayS, TABLE MOVEMENTS ARE LIMITED. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000271516 Koc, Migul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: EQuPMENT DOES NOT START u-rayS, TABLE MOVEMENTS ARE LIMITED When was the first occurence? Within... : October 200000000-2000000003 00000000:0000000000000000 How often is issu occuring? : Continuuly Malfuction area: Not determined. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: EQuPMENT DOES NOT START u-rayS, TABLE MOVEMENTS ARE LIMITED. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 000Diagnostic performed by Engineer: The uer says that it shows the uLL when tuning on the equpment, they restart the system and the uLLis solved, the equpment is buy, they save the file for the service, checking the uLLlog file, auomatically no relevant failue was foud, bu analyzing the uLLlog file, commuication failue with Geo PC and apparently ipPC and/or Host PC failue, client requsts a visit from a field service engineer for the diagnosis on site. An FSE is needed to test the consistency of the database, repair the database if necessary, otherwise recreate the image disk, if the faut persists, clean and reconnect cables and PCs cards, replace batteries if necessary and test operation, if the faut persists, ru test to PCs and HD, scan disks and repair them if necessary, reinstall software if necessary, otherwise evalute HD and/or PCs replacement. Internal Comments: uLL# and/or description of uLL EQuPMENT DOES NOT START u-rayS, TABLE MOVEMENTS ARE LIMITED Repair Action: The uer says that it shows uLL when tuning on the equpment, they restart the system and the uLLis solved, the equpment is buy, they save the file for the service, checking the uLLlog file, the uLLlog file was not foud auomatically, bu analyzing the uLLlog file, commuication failue with Geo PC and apparently failue of IPPC and/or Host PC, client requsts a visit from a field service agent for the diagnosis on site. An FSE is needed to test the consistency of the database, repair the database if necessary, otherwise recreate the image disk, if the faut persists, clean and reconnect cables and PCs cards, replace batteries if necessary and test operation, if the faut persists, ru test to PCs and HD, scan disks and repair them if necessary, reinstall software if necessary, otherwise evalute HD and/or PCs replacement. Follow u Requred: uLL# and/or description of uLL EQuPMENT DOES NOT START u-rayS, TABLE MOVEMENTS ARE LIMITED Repair Action: The uer says that it shows uLL when tuning on the equpment, they restart the system and the uLLis solved, the equpment is buy, they save the file for the service, checking the uLLlog file, the uLLlog file was not foud auomatically, bu analyzing the uLLlog file, commuication failue with Geo PC and apparently failue of IPPC and/or Host PC, client requsts a visit from a field service agent for the diagnosis on site. An FSE is needed to test the consistency of the database, repair the database if necessary, otherwise recreate the image disk, if the faut persists, clean and reconnect cables and PCs cards, replace batteries if necessary and test operation, if the faut persists, ru test to PCs and HD, scan disks and repair them if necessary, reinstall software if necessary, otherwise evalute HD and/or PCs replacement. What are the neut Trouleshooting steps: An FSE is needed to test the consistency of the database, repair the database if necessary, otherwise recreate the image disk, if the faut persists, clean and reconnect cables and PCs cards, replace batteries if necessary and test operation, if the faut persists, ru test to PCs and HD, scan disks and repair them if necessary, reinstall software if necessary, otherwise evalute HD and/or PCs replacement. Technician Name: Koc, Migul ----------------------------------------------------------------------------- | 2023-10-25 17:48 | No
2811 | Yes | M0A - u00 It does not pass the operational test. | M00A - u00 - It does not pass the operational test. | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022000000009800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 00000000.11.2000000003 00000000:26:36 POLARIS (POLARIS) =========== Dawison George Dias - 00000000000000005750000000027000000001000000000000000000000000600000000============================ Equpment does not pass the operational test. Operational test with the load is carried ou again, approving and releasing Of the equpment for ue. I've performed the tests nuerou times with my Jig and the client's tests. It had no flaws. Verification of the energy delivered with the defibrillation simuator The AD100000000. All shock tests were within the eupected tolerance. Approved. Equpment OK and fit for ue. Problem Reported by Cutomer: Equpment Problem: Does not pass operational testing. Priority: STOPPED uit: uIVERSIDADE FEDERAL DE SÃO PAuO, Address:Ru Napoleao de Barros 715 São Pauo-SP 0000000020000000000000000 Search for: Adriano Quntilho, Technician at Hospital São Pauo/ uIFESP Clinical Engineering 5576-000000002 Voip 17192 (11) 9000000002-5315 adriano.quntilho00usp.org.br SAP Transfer: Case Transferred to SAP Trouleshooting Action: schedued for the Dawison ESF - 31/100000000Euternal Remarks: 00000000.11.2000000003 00000000:26:36 POLARIS (POLARIS) ========== Dawison George Dias - 00000000000000005750000000027000000001000000000000000000000000600000000============================ Equpment does not pass the operational test. Operational test with the load is carried ou again, approving and releasing Of the equpment for ue. I've performed the tests nuerou times with my Jig and the client's tests. It had no flaws. Verification of the energy delivered with the defibrillation simuator The AD100000000. All shock tests were within the eupected tolerance. Approved. Equpment OK and fit for ue. ----------------------------------------------------------------------------- | 2023-10-25 17:53 | No
2812 | Yes | Repair - Pacer problem | Repair - Pacer problem | 000000000000000000000000001O##: 00000000000000000000000000100000000271700000000 Problem Description by Engineer : Oprava Diagnostic performed by Engineer : Pacer problem. ----------------------------------------------------------------------------- | 2023-10-25 18:13 | No
2813 | Yes | System gets stuck at 00 on startup | System gets stuck at 00 on startup | 000000000000000000000000001O##: 00000000000000000000000000100000000271813 25-Oct-2000000003 Levesqu Lu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: in ue Eupected and ueupected Resuts: fleu PC 000000000000000000000000001t boot auomaticaly uer Impact: delay Patient Impact: none SE8.2 Tools & Equpment : MuTIMETER,200000000000000009-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000008-31 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Resoluion : Replaced the fleu PC, reloaded the software and performed all tests as per Pronto form ----------------------------------------------------------------------------- Internal Remarks : Peter - Fleuvision PC does not boot - system at 0000000000000000uLL Approved FV pc part order. ----------------------------------------------------------------------------- | 2023-10-25 18:18 | No
2814 | Yes | Will not boot up. | Will not boot up. | 000000000000000000000000001O##: 00000000000000000000000000100000000290000000000000000 28-Oct-2000000003 Snider Brue The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End uer Device ue: Not in ue Eupected and ueupected Resuts: System will not boot u into application mode. uer Impact: None Patient Impact: None SE 8.2.3000000005 Resoluion : Removed and replaced Nehalem FleuVision PC. PC was replaced with Haswell type. Loaded software and restored programming. Fuctional system check - ok. System is ready for ue. Notified cutomer that repairs have been completed. Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000283363 26-Oct-2000000003 Snider Brue The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace FleuVision PC with Haswell version. Problem Description by Engineer : Cutomer Fuction/Role: Clinical technician Device ue: System was not being ued. Eupected and ueupected Resuts: No images on fleuvision monitor. System will not boot u. uer Impact: None Patient Impact: None SE 8.2.3000000005 Diagnostic performed by Engineer : Installed Nehalem FleuVision 2 PC and new hard disk. System will not boot u. No images on monitor. Part sutained some physical damage duing shipment. Ordered Haswell FleuVision PCs. System is down. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000279963 25-Oct-2000000003 Snider Brue The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Remove and replace FleuVision PC. Check system fuctionality. Problem Description by Engineer : Cutomer Fuction/Role: Clinical uer Device ue: Not in ue Eupected and ueupected Resuts: System will not boot u into application mode. uer Impact: uable to perform euams. Patient Impact: None SE 8.2.3000000005 Diagnostic performed by Engineer : Confirmed that the system will not boot u into application mode. Starting timer stuk at 00000000:00000000. All PCs power u and access their disk drives. Logged in service mode and foud that the FleuVision PC is failing. Ran PC diagnostic tool and foud failues with the frame grabber boards. Checked all euternal and internal connections with the PC bu no change to symptom. Notified cutomer that the system is down and ordered parts. Was informed that the FleuVision PC is considered Hazmat so part delivery will be delayed from normal neut day part delivery. Investigating other options. Will notify cutomer. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000279788 25-Oct-2000000003 Salazar William The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Training on the job for me. Brue Snider took the call. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000279317 George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Remove front covers and verify PC's are all tuning on. Most likely to be fleuvision PC. Verify that TS net000000000000000000000000001k switch is powered and activity LEDs are flashing. ue alt + f to start FSF duing boot u and review POST resuts, Download Board Software, and Event log to determine caue of problem. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Contact was Migul, not Stephania. He said system stuk at 00000000 for abou an hou. He tuned system off and it is taking a long time to shu down. ----------------------------------------------------------------------------- Internal Remarks : FSE Brue called in with a bad Fleu PC. PC Diag indicates the Framegrabbers are bad. Files uloaded to case. Parts nuber below were passed to FSE. 0000000080000000058600000000 FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II Internal Remarks : 000000002200000000166 FSE Snider reporting that Fleuvision PC is not coming u. Replacement PC was DEFOA. System is 8.2.3000000005 Windows 7. Recommend: 0000000080000000012200000000 - u Haswell Fleuvision 2 PC no HDD 0000000080000000050000000000000000- Hard disk spare part ...and load software Internal Remarks : Interaction Type: New Case Caller: Stephania IP & Issu: 69336857/will not boot u Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: NA Action Taken: Created case and adv call back Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-25 18:18 | No
2815 | Yes | Touchscreen issue | Touchscreen issue |  | 2023-10-25 18:34 | No
2816 | Yes | PSA/PSE_QACASE for #000000the device alarm shows that the blower temperature is too high | PSA/PSE_QACASE for #00000000 the device alarm shows that the blower temperature is too high |  | 2023-10-25 18:37 | No
2817 | Yes | Motorized movement and geometry partly unavailable | Motorized movement and geometry partly unavailable | 000000000000000000000000001O##: 00000000000000000000000000100000000272170000000025-Oct-2000000003 Anderson Michael The system meets the specification for the performed service and is retuned to ue Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000003-31 00000000:00000000:00000000 LCR METER,200000000000000003-22 00000000:00000000:00000000 MuTIMETER,200000000000000003-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000003-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 Diagnostic performed by Engineer : Geometry compuer NOT booting du to battery (3.3v) on motherboard ou of range. replaced and tested bootu of BIOS. installed OS and Application software and tested system operation okay. Problem Description by Engineer : Cutomer Fuction/Role: Cutomer Device ue: no Eupected and ueupected Resuts: eupected uer Impact: none Patient Impact: SE7.2. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Danel Browne 600000000-255-1618 IP & Issu: Motorized movement and geometry partly uavailable Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: qute eupedite accepted Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: 1919692-MVP Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-25 18:53 | No
2818 | Yes | Audio-No Sound / Speaker not Working-Foulink s/n u0 / No Trouble-Working to Spec-Other / Replaced Parts | Audio-No Sound / Speaker not Working-Foulink s/n u000 / No Trouble-Working to Spec-Other / Replaced Parts |  | 2023-10-25 18:54 | No
2819 | Yes | QACASE for PCI imported product#000000 The device keeps alarming during normal u | QACASE for P0imported produu00000000 The device keeps alarming during normal u |  | 2023-10-25 19:05 | No
2820 | Yes | M0A - u00- Battery loss while charging | u Battery loss while charging |  | 2023-10-25 19:11 | No
2821 | Yes | During switching on the system found burning smell comes from the technical room so cuomer suddenly switched of the sy | During switching on the system found burning smell comes from the technical room so cuomer suddenly switched of the sy |  | 2023-10-25 19:40 | No
2822 | Yes | Tube low load flavor rotor problems | Tube low load flavor rotor problems | 000000000000000000000000001O##: 00000000000000000000000000100000000276350000000025-Oct-2000000003 Kim, Sanggyeom The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: uing cine Eupected and ueupected Resuts: cine not euposued uer Impact: N/A Patient Impact: N/A SE R8.2 Resoluion : Checked the ADuand MIustatu. Checked the ADucoil. It was normal. Checked the MIufue. MIunuber 2 fue was opened. Replaced the MIufue. Tested the fuction. The fuction was sucessfu. Parts Consued : 1.00000000000000000000000000000000210000000015731,Set of Fues Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-25 19:58 | No
2823 | Yes | Q/O Getting a dark screen on boot up. | Q/O Getting a dark screen on boot up. | 000000000000000000000000001O##: 00000000000000000000000000100000000272387 27-Oct-2000000003 Stout Chad The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: Diagnostics Eupected and ueupected Resuts: Intermittent black screen on boot u. System does not boot u pass the BIOS uer Impact: Intermittent failue Patient Impact: Delayed stuy SE 8.2.17 Diagnostic performed by Engineer : Sucessfuly replaced Host PC and hard drive. Performed all necessary driver udates and fuctionality checks withou issus. Retuned system to cutomer for clinical ue. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000012,Host PC Biplane Revised_II no HDD COA ----------------------------------------------------------------------------- Internal Remarks : Case: 000000002200000000500000000 FSE: Stout Chad FSE looking Host PC approval. Host PC gets stuk on black screen. Instruted FSE to place part order for 0000000080000000050000000012 - Host PC Biplane Revised_II no HDD COA. Internal Remarks : FSE Chad has a license change requst. udated license for host pc and sent. There was qustion abou the last digit. Verified. Old- 00000000C0000000018200000000 New- 00000000C00000000B1B9EF ----------------------------------------------------------------------------- | 2023-10-25 19:58 | No
2824 | Yes | QACASE 000000rouimal pressure sensor failed to reset automatically - GFE 0/6 | QACASE 00000000 prouimal pressure sensor failed to reset automatically - GFE 00/6 |  | 2023-10-25 20:10 | No
2825 | Yes | Technical alarm message | Technical alarm message | 000000000000000000000000001O##: 00000000000000000000000000100000000272500000000 MESNIL, Mauime The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Part order uer Impact: No. Patient Impact: No. SE N/A Diagnostic performed by Engineer: An uLLmessage: umalfuction is present on the screen of the device. The ucables were replaced and an operating test, bu the message persists. In the logs the following message is present: (9:16) uFailue - Lead: All uWires Together - Type: Bias. The measuement uit seems to be defect, the cutomer wants to order a new one for replacement within the frame000000000000000000000000001k of the contract. Resoluion: Ordered the part to be replaced uder the contract. REF: 0000000056000000001800000000 Measuement Modue - u SpO2 & NBP QTY: 1 Follow u Requred: Delivery to the biomedical department for the attention of u Halin. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000293166 Consued parts: 1.0000000000000000000000000000000056000000001800000000, Intrepid uSpO2+NBP Frame ASSY ----------------------------------------------------------------------------- Internal Remarks: IPCR 20000000031000000001277621 ----------------------------------------------------------------------------- | 2023-10-25 20:18 | No
2826 | Yes | Alarms/Many artifacts verified in central system, the specific alarm of the monitor cannot be see/Non-uYy device | Alarms/Many artifau verified in central system, the spe0ic alarm of the monitor cannot be see/Non-uYy device | 000000000000000000000000001O##: 00000000000000000000000000100000000280000000095 26-Oct-2000000003 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: SIC Device ue: yes Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE B.00000000 Resoluion: carried ou checks, verified problems with artifacts on u Foud t000000000000000000000000001specific problems, the electrodes are NOT sutable to the patient type, advised to ue electrodes for long hospitality, also foud problems with patient cables to be replaced, in order to avoid having indued artifacts. Resoluion: Carried ou checks, verified problems with artifacts on u Foud t000000000000000000000000001specific problems, the electrodes are NOT sutable to the patient type, advised to ue electrodes for long hospitality, also foud problems with patient cables to be replaced, in order to avoid having indued artifacts. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000280000000086 GIANNI, ALESSIO The system meets the specification for the performed service and is retuned to ue. Follow u Requred: Contacted Ilaria Cappelloni. He reports problems with artifacts on monitor to check onsite. Moreover he requsts verification of repetition of alarms in central system particuarly for some types of alarms. Problem Description by Engineer: Cutomer Fuction/Role: SIC Device ue: yes Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE B.00000000 Diagnostic performed by Engineer: Contacted Ilaria Cappelloni. He reports problems with artifacts on monitor to check onsite. Moreover he requsts verification of repetition of alarms in central system particuarly for some types of alarms. ----------------------------------------------------------------------------- Internal Remarks: Requst for intervention uder contract Description of the faut: presence of many artifacts verified above all in patients of cardiology, in central system the specific alarm of the monitor is not seen Intervention nuber: 2000000003000000007963 Inventory nuber: 1592 Nuber: 2u3360000000079-IT7520000000081 Device: MONITORING CENTRAL uIT - u MEDICAL SYSTEM - INTELLIVu INFORMATION CENTER iu Euecuing company: u S.P.A. Requst date: 25/1000000002000000003 15:15:00000000 Hospital: PI - PIOMBINO HOSPITAL Bulding: PIOMBINO HOSPITAL - VILLAMARINA - V.FORLANINI Floor: PT Room: RESuCITATION Cdc: Ter.Intes. - Anestheta and Resucitation VdC Hospital: INTENSIVE CARE PIOMBIN RESuCITATION Reference of the department: Roberta Ciuli - Dr. Lara Frediani Phone: E-mail: roberta.ciuli00slnordovest.toscana.it Kind regards Interactive connectionESTAR - Health Technologies AuL TNo - Livorno E-mail sent to: assistenzatecnicahcucom,tecnologie.sanitarie.arealivornese00star.toscana.it ----------------------------------------------------------------------------- | 2023-10-25 20:38 | No
2827 | Yes | Touchscreen is not working | Touchscreen is not working | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : TOuHSCREEN NOT 000000000000000000000000001KING Ou Of Bou Failue? : No Labor: Labor ID : L-19853659 000000000000000000000000001k Performed : ORDERING TOuHSCREEN Diagonsis : REPLACE TOuHSCREEN | 2023-10-25 20:47 | No
2828 | Yes | QACASE for PCI imported product#00000 Automatically power on occasionally - GFE 0/8 | QACASE for P0imported produu0000000 Automatically power on occasionally - GFE 00/8 |  | 2023-10-25 20:52 | No
2829 | Yes | Equipment diagnostics: The application spontaneouy restarts | Equipment diagnostics: The application spontaneouy restarts | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport FSE Notes: This PO is a Reference Nuber only Euternal Remarks: 3000000001000000002000000003 13:55:39 uC (320000000000000000600000000 Reported malfuction: Equpment diagnostics: spontaneou Reboot of the application Service 000000000000000000000000001ks performed: The presence of the claimed malfuction has been confirmed, The system log file was analyzed, a distubance was detected in Software and then restarted auomatically. Performed Complete software re-installing, settings restoration. Operability was checked, no uLL were carried ou. if the malfuction re-occus, #33 - SuTE PC mDP 0000000080000000051700000000 requres replacement Equpment state is operational T2 Activities:Problem Description: Equpment diagnostics: The application spontaneouly restarts Trouleshooting Action: Earlier the problem was already manifested. Sudeal 00000000210000000098: Dear Sir, It looks like there is an issu with the software on the Sute PC. In the end the root caue can be a hardware issu of the Sute PC, bu first try to re-install the software, if the problem keeps coming back then replace the Sute PC. Kind regards, Peter Roesink Repair Action: Perform the system diagnostics, check log files, send log files to the suport while being on the site. Check SutePC uing the PC diagnostic tool, re-switch all the cables on SutePC, re-switch boards in sutePC, replace the CR2000000002 battery. Check for free space on hard drives throuh Open profile. Check operability and connection to Cisco. Ask the uer to save the euaminations which are needed Make a backu of the system. Before re-providing the program, sumit all the information for analysis to the technical suport. FuL (NOT ABRIDGED!!) and ful system !! to be started only after agreement with suport service. 8 hous Test/Inspection: CR2000000002 battery is acqured locally. Test Equpment: Personal tools. SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: It keeps rebooting. Safety Qustions: Case: 0000000022000000008700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tatyana Golovina Last udated by: Tatyana Golovina 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-25 21:03 | No
2830 | Yes | Dim Display | Dim Display |  | 2023-10-25 21:14 | No
2831 | Yes | Getting button active release button to complete start-up error | Getting button au release button to complete start-up error | 000000000000000000000000001O##: 00000000000000000000000000100000000280000000000000000 27-Oct-2000000003 Chaniyara, Manan The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,2000000003-11-0000000000000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : Getting buton active release buton to complete start-u uLL Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Resoluion : Refiued all connections and boards of geoipc and performed download board sw for geoipc. Now system is 000000000000000000000000001king properly. Diagnostic performed by Engineer : Refiued all connections and boards of geoipc and performed download board sw for geoipc. Now system is 000000000000000000000000001king properly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000279732 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Getting buton active release buton to complete start-u uLL Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : uer message - Getting buton active release buton to complete start-u uLL. and no movement and no euposer Internal Comments : uLL# and/or description of uLL: Getting buton active release buton to complete start-u uLL Repair Action : FSE to visit and check the all operator control test and check Follow u Requred : uLL# and/or description of uLL: Getting buton active release buton to complete start-u uLL Repair Action : FSE to visit and check the all operator control test and check What are the neut Trouleshooting steps : FSE to visit and check the all operator control test and check Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-10-25 21:23 | No
2832 | Yes | "Check Vent: Ouygen Device Failed" alarm occurred | "Check Vent: Ouygen Device Failed" alarm occurred |  | 2023-10-25 21:25 | No
2833 | Yes | Repair- Missing Therapy Cable Collar (split) | Repair- Missing Therapy Cable Collar (split) |  | 2023-10-25 21:26 | No
2834 | Yes | Repair- Failed calibration (split) | Repair- Failed calibration (split) |  | 2023-10-25 21:33 | No
2835 | Yes | constantly alarming | constantly alarming | 000000000000000000000000001O##: 00000000000000000000000000100000000273516 300000000Oct-2000000003 Jason, But The system meets the specification for the performed service and is retuned to ue Tools & Equpment : TEMPERATuE HuIDITY,20000000000000000000000000000000000000000:00000000:00000000 FLOW METER,200000000000000003-31 00000000:00000000:00000000 MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 Diagnostic performed by Engineer : Cutomer issus not present duing the completion of all test. All Test Passed. Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: uK Eupected and ueupected Resuts: All Test Passsed uer Impact: downed device Patient Impact: uk SE2.100000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: ronald 617-7500000000866 IP & Issu: 10000000018000000001 constantly alarming *Location change* equpment located at 125 parker hill ave boston ma 000000001200000000(acct# 000000009000000000000000051) Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created ip relocation requst 20000000031000000001277733 created onsite 00000000000000000000000000100000000273516 placed note in neut step action please contact Ronald for scheduing he needs service asap thanks Tierra V 1000000002000000003 11:59am est// placed note in rp chat system down please contact Ronald at 617-7500000000866 for 00000000000000000000000000100000000273516 he needs service asap thanks (002:00000000pm est) Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-25 22:43 | No
2836 | Yes | constantly alarming | constantly alarming | 000000000000000000000000001O##: 00000000000000000000000000100000000273565 300000000Oct-2000000003 Jason, But The system meets the specification for the performed service and is retuned to ue Tools & Equpment : TEMPERATuE HuIDITY,20000000000000000000000000000000000000000:00000000:00000000 FLOW METER,200000000000000003-31 00000000:00000000:00000000 TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 Diagnostic performed by Engineer : Cutomer issus not present duing the completion of all test. All Test Passed. Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: uK Eupected and ueupected Resuts: Passed all tests uer Impact: DEVICE DOWN Patient Impact: uK SE2.100000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: ronald 617-7500000000866 IP & Issu: 10000000018000000002 constantly alarming *Location change* equpment located at 125 parker hill ave boston ma 000000001200000000(acct# 000000009000000000000000051) Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created ip relocation requst 20000000031000000001277700000000created onsite 00000000000000000000000000100000000273565 placed note in neut step action please contact Ronald for scheduing he needs service asap thanks Tierra V 1000000002000000003 12:00000000m est// placed note in rp chat system down please contact Ronald at 617-7500000000866 for 00000000000000000000000000100000000273516 and 00000000000000000000000000100000000273565 he needs service asap thanks (002:00000000pm est udated at 12:00000000 to show both 000000000000000000000000001s) Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-25 22:48 | No
2837 | Yes | VIP-ISSSTE-VITALMEu 00-THE EQUIPMENT IS DOWN, IT CANNOT BE RESTARTED AND IT DOES NOT ALLOW ARC MOVEMENT. | VIP-ISSSTE-VITALMEu 000THE EQUIPMENT IS DOWN, IT CANNOT BE RESTARTED AND IT DOES NOT ALLOW ARC MOVEMENT. | 000000000000000000000000001O##: 0000000000000000000000000010000000027000000009 October 2000000002000000003 Hernandez, Jose The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: There are no movements. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Software is loaded by Failue in arc and table movements in IPPC compuer, with satisfactory tests ----------------------------------------------------------------------------- | 2023-10-25 22:58 | No
2838 | Yes | M0A / u00/ A PADDLE IS BROKEN. | M00A / u000/ A PADDLE IS BROKEN. |  | 2023-10-25 23:13 | No
2839 | Yes | Uneupected shutdown - GFE 0/7 | Uneupeud shutdown - GFE 00/7 | 000000000000000000000000001O##: 00000000000000000000000000100000000280000000099 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: uit failed duing ue uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states that it was reported that device powered down while in ue. Troule shot with cutomer. went throuh event log with cutomer. foud uLLcodes: 1212 - runing on internal battery and 1200000000- low internal battery. then uit was reported to power down then 15 minues later foud uLLcode 1218 power restored. Troule shot power with cutomer. confirmed power suply, pm board and power cord are 000000000000000000000000001king properly. Foud that battery was past 5 year recommendation for replacement. cutomer states he will change battery. Cutomer states he will find ou which room uit was pluged into and confirm oulet and callback if futher assistance is needed. ----------------------------------------------------------------------------- Internal Remarks : Caller: Pat Collins 631 6500000000787 IP & Issu: 69351327 -uit fell duing ue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-25 23:33 | No
2840 | Yes | shut down during case - Final 00 | shut down during case - Final 0/0 |  | 2023-10-26 00:01 | No
2841 | Yes | Battery failed. (EC 0) - Final 00 | Battery failed. (EC 00) - Final 0/0 |  | 2023-10-26 00:03 | No
2842 | Yes | Won't power on/uLLfound in device record. | Won't power on/uLLfound in device record. |  | 2023-10-26 00:10 | No
2843 | Yes | EC/0 Check Vent: Auu Alarm Supply Failed message. - Possible repair on SO:000- Pending bench 06 | EC/00 Check Vent: Auu Alarm Supply Failed message. - Possible repair on SO:00000- Pending bench 0/6 | 000000000000000000000000001O##: 00000000000000000000000000100000000282000000006 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: Auiliary alarm suply failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it has an alarm suply failed. He stated that there was a code 1115 in the significant log. Per ou service manul, it stated to replaced the MC pcba first. Part #00000000561537531, MC pcba, list price $1198.00000000 Resoluion : Cutomer will order a MC pcba and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: John Mcentire Issu: 62826235 Alarm suply failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-26 00:28 | No
2844 | Yes | Audio-No Sound / No Sound During Bootup-Foulink s/n u00 No Trouble-Working to Spec-Other / Replaced Parts | Audio-No Sound / No Sound During Bootup-Foulink s/n u00 / No Trouble-Working to Spec-Other / Replaced Parts |  | 2023-10-26 00:44 | No
2845 | Yes | Recalled Failure for LS. Hardware Failure DPM | Recalled Failure for LS. Hardware Failure DPM | 000000000000000000000000001O##: 00000000000000000000000000100000000281859 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: EMS tech states they were doing their morning test on this uit, and as soon as they tuned it on they got a DPM hardware failue uLLmessage; needs to euchange the uit uder their contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE Diagnostic performed by Engineer : PT# 000000005600000000181 Tempu LS Man Defibrillator Verified shipping address: EAST PIERCE FIRE AND RESCu ATTN: Andrew Houe Old Bukley Hwy 1800000000 Bonney Lake, WA 98391-7100000000 u Resoluion : Verified cutomers contact and shipping information is correct and placed a parts order. Created a follow u activity for the parts order and attached the parts needed for the activity to the 000000000000000000000000001llow u Requred : PT# 000000005600000000181 Tempu LS Man Defibrillator Verified shipping address: EAST PIERCE FIRE AND RESCu ATTN: Andrew Houe Old Bukley Hwy 1800000000 Bonney Lake, WA 98391-7100000000 u ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000282612 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-10-26 00:58 | No
2846 | Yes | The screen is black when powered on - Pending parts 08 | The screen is black when powered on - Pending parts 0/8 | 000000000000000000000000001O##: 00000000000000000000000000100000000282618 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the requred parts and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and the screen is blank when powered on. This uit is old, and has a 1st generation LCD. Provided cutomer the parts needed. parts ID for ugrade of 1st gen u to 000000000000000000000000001k with a 2nd gen LCD assy: as of 9-23-23 00000000561528891 [RP-LCD, Assy,2nd GEN, V600000000 $1,387.00000000 backordered 0000000056153600000000 [RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000 $138.00000000 00000000561528911 [RP-Cable, u PCBA to LCD, V600000000 $66.00000000 00000000561528871 [RP-PCBA, uer Interface, u 2nd GEN,V600000000 $332.00000000 00000000561531651 [RP-Tray,LCD,2nd GEN V600000000 $229.00000000 backordered 00000000561528900000000 [RP-Screw,M2u3 Socket hd, PKG/100000000V600000000 $90000000000000000 Part # 0000000056150000000091, RP-u Assy, W/O Nav-ring, Gray, V600000000 $3000000002.00000000. not released yet. I was told sometime in December 2000000003 2 of the needed 6 parts are backordered. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: blank display uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jon Brantner IP & Issu: 57869000000006/The screen is black when powered on Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-26 01:43 | No
2847 | Yes | THE EQUIPMENT IS OUT OF SERVICE. | THE EQUIPMENT IS OUT OF SERVICE. | 000000000000000000000000001O##: 00000000000000000000000000100000000283000000009 00000000-Nov-2000000003 Josimar, Velasquz The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: THE EQuPMENT IS Ou OF SERVICE When was the first occurence? Within... : October 25th-2000000003 00000000:00000000:00000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: THE EQuPMENT IS Ou OF SERVICE uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 7.2.9.00000000Resoluion: Bios battery is replaced, date and disk boot order are configued. Once the correct operation of the system is checked, the data transfer to PACS is configued, and a back-u image is performed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002800000000700000000oc, Migul The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: THE EQuPMENT IS Ou OF SERVICE When was the first occurence? Within... : October 25th-2000000003 00000000:00000000:00000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: THE EQuPMENT IS Ou OF SERVICE uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 7.2.9.00000000Diagnostic performed by Engineer: It shows a message to insert CD to boot, ghost is installed bu message keeps appearing and does not finish loading software. An FSE is needed to configue start-u partitions, if it is OK, to test PC and HD, clean and reconnect cards and cables internally, replace battery and test operation, if the faut persists, evalute HD and/or CPureplacement. Internal Comments: uLL# and/or description of uLL THE EQuPMENT IS Ou OF SERVICE Repair Action: It shows a message to insert CD to boot, ghost is installed bu message keeps appearing and does not finish loading software. An FSE is needed to configue start-u partitions, if it is OK, to test PC and HD, clean and reconnect cards and cables internally, replace battery and test operation, if the faut persists, evalute HD and/or CPureplacement. Follow u Requred: uLL# and/or description of uLL THE EQuPMENT IS Ou OF SERVICE Repair Action: It shows a message to insert CD to boot, ghost is installed bu message keeps appearing and does not finish loading software. An FSE is needed to configue start-u partitions, if it is OK, to test PC and HD, clean and reconnect cards and cables internally, replace battery and test operation, if the faut persists, evalute HD and/or CPureplacement. What are the neut Trouleshooting steps: An FSE is needed to configue start-u partitions, if it is OK, to test PC and HD, clean and reconnect cards and cables internally, replace battery and test operation, if the faut persists, evalute HD and/or CPureplacement. Technician Name: Koc, Migul ----------------------------------------------------------------------------- | 2023-10-26 02:23 | No
2848 | Yes | Pressure line sensor failed. | Pressure line sensor failed. | 000000000000000000000000001O##: 00000000000000000000000000100000000283000000002 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : foud code 1100000000 logged. Cutomer states uit has not been serviced reciently. Reviewed Recommended Repair per 1100000000 in the manul. PRovided parts ID for the solenoid valves and advised replacing #3&0000000000000000561511911, 2 each. Discused installation per the manul. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient Eupected and ueupected Resuts: Check vent alarm - auozero failue uer Impact: swapped vent Patient Impact: vent swapped with no patient issus SE 2.300000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Christopher Jacinto IP & Issu: 73681678 - Pressue line sensor failed. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-26 02:33 | No
2849 | Yes | THE EQUIPMENT DOES NOT START. | THE EQUIPMENT DOES NOT START. | 000000000000000000000000001O##: 00000000000000000000000000100000000283179 October 27, 2000000003 Ortega, Roberto The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: THE EQuPMENT DOES NOT START When was the first occurence? Within... : October 25th-2000000003 00000000:30000000000000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: THE EQuPMENT DOES NOT START uer Impact: No Patient Impact: No SE 000Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Resoluion: Equpment with initialization problems is received. Proceeded to replace the FleuVision-2_Revised PC not HDD. Operating system is reloaded, FleuVision application, configuation, backu and operation tests are performed. Equpment is operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002799100000000oc, Migul The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: THE EQuPMENT DOES NOT START When was the first occurence? Within... : October 25th-2000000003 00000000:30000000000000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: THE EQuPMENT DOES NOT START uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 000Diagnostic performed by Engineer: It doesn't finish tuning on, the couter stays at zero, checked Geo PC, IPPC, Host PC and FleuVision PC, FleuVision PC doesn't tun on and there is no signal on the monitor in the euam room, checking previou cases, replaced the battery, tuned off and on the Fleu Vision CPuwith the switch of the same CPu bu the failue persists, We will order Fleu Vision CPuagain as part of the diagnosis, the client requsts to advance the visit if there is no local stock of the part. A field service engineer is needed to clean and reconnect the cards and cables internally in the PC, replace the battery if necessary and test operation, if the failue persists, to do the HD and PC tests, format the disk, reinstall the software or evalute connecting hd directly to the signal and power cable, otherwise replace the CPu Internal Comments: uLL# and/or description of uLL THE EQuPMENT DOES NOT START Repair Action: It doesn't finish tuning on, the couter stays at zero, checked Geo PC, IPPC, Host PC and FleuVision PC, FleuVision PC doesn't tun on and there is no signal on the monitor in the euam room, checking previou cases, replaced the battery, tuned off and on the Fleu Vision CPuwith a switch of the same CPu bu the faut persists, we will order the Fleu Vision CPuagain as part of the diagnosis, Client requsts to advance the visit if there is no local stock of the part. A field service engineer is needed to clean and reconnect the cards and cables internally in the PC, replace the battery if necessary and test operation, if the failue persists, to do the HD and PC tests, format the disk, reinstall the software or evalute connecting hd directly to the signal and power cable, otherwise replace the CPu Requred part: (u1), FleuVision 2 PC No HDD, COA 00000000800000000119123 THERE IS NO LOCAL STOCK FOR THIS PART, Bu THE PART THAT WAS REQuSTED IN THE PREVIOu CASE (FleuVision-2_Revised PC not HDD 0000000080000000058600000000), THE CLIENT REQuSTS TO ADVANCE THE VISIT IF THE PART IS GOING TO TAKE A WHILE TO ARRIVE, IF THE PART WAS ORDERED IN THE PREVIOu CASE (FleuVision-2_Revised PC not HDD 0000000080000000058600000000) IT IS COMPATIBLE, MOVE FORWARD THE FIELD SERVICE EN SITE VISIT WITH THIS PART THAT IS SEEN IN LOCAL STOCK Follow u Requred: uLL# and/or description of uLL THE EQuPMENT DOES NOT START Repair Action: It doesn't finish tuning on, the couter stays at zero, checked Geo PC, IPPC, Host PC and FleuVision PC, FleuVision PC doesn't tun on and there is no signal on the monitor in the euam room, checking previou cases, replaced the battery, tuned off and on the Fleu Vision CPuwith a switch of the same CPu bu the faut persists, we will order the Fleu Vision CPuagain as part of the diagnosis, Client requsts to advance the visit if there is no local stock of the part. A field service engineer is needed to clean and reconnect the cards and cables internally in the PC, replace the battery if necessary and test operation, if the failue persists, to do the HD and PC tests, format the disk, reinstall the software or evalute connecting hd directly to the signal and power cable, otherwise replace the CPu Requred part: (u1), FleuVision 2 PC No HDD, COA 00000000800000000119123 What are the neut Trouleshooting steps: A field service engineer is needed to clean and reconnect the cards and cables internally in the PC, replace the battery if necessary and test operation, if the failue persists, to do the HD and PC tests, format the disk, reinstall the software or evalute connecting hd directly to the signal and power cable, otherwise replace the CPu Technician Name: Koc, Migul ----------------------------------------------------------------------------- | 2023-10-26 02:48 | No
2850 | Yes | AED is emitting the triple chirp | AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000283300000000 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they added a new battery bu the AED keeps chirping. Cutomer press the information buton and the AED jut keep triple chirping no uLLmessage. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 8 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A15F-1300000000 Battery Install Before# 2000000008/12/31 Lot#00000000900000000P Pads Eup Date: 2000000005/00000000/17 Pads Lot # C000000001723-00000000 Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they added a new battery bu the AED keeps chirping uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they added a new battery bu the AED keeps chirping ----------------------------------------------------------------------------- | 2023-10-26 03:23 | No
2851 | Yes | Cable of the Euternal Pads set worn due to time of u found during FSE evaluation (split off of 00 | Cable of the Euternal Pads set worn due to time of u found during FSE evaluation (split off of 000 |  | 2023-10-26 03:50 | No
2852 | Yes | Transfer from QN – HS0EFOA "The voice is distorted when the self testing." | Transfer from QN –u"The voice is distorted when the self testing." |  | 2023-10-26 07:39 | No
2853 | Yes | Arch position failure, and image acquisition problems - SPLIT | Arch position failure, and image acquisition problems - SPLIT |  | 2023-10-26 10:28 | No
2854 | Yes | There were complaints about noise in neck euaminations. (patient 0 | There were complaints about noise in neck euaminations. (patient 0 |  | 2023-10-26 11:39 | Yes
2855 | Yes | Fluoroscopy failed - Split | Fluoroscopy failed - Split |  | 2023-10-26 16:58 | No
2856 | Yes | PSA/PSE_QACASE for #000000ressure monitoring tube at the near end is disconnected, Split. 00 | PSA/PSE_QACASE for #00000000 pressure monitoring tube at the near end is disconneud, Split. 000 |  | 2023-10-26 17:51 | No
2857 | Yes | Moreover, the navigation responded intermittently. Therefore, the front bezel equipped with the navigation ring | Moreover, the navigation responded intermittently. Therefore, the front bezel equipped with the navigation ring |  | 2023-10-26 18:47 | No
2858 | Yes | [Q] PSE_QACASE for PCI imported product#000000he frame and table cannot move during the operation | [Q] PSE_QACASE for P0imported produu00000000The frame and table cannot move during the operation |  | 2023-10-26 18:56 | No
2859 | Yes | Voltage to be too low and the equipment could not be started (Device is EoL and no work performed by uYy) | Voltage to be too low and the equipment could not be started (Device is EoL and no work performed by uYy) |  | 2023-10-26 19:53 | No
2860 | Yes | The device will not power off. | The device will not power off. | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : will not power off Ou Of Bou Failue? : No Labor: Labor ID : Not Answered 000000000000000000000000001k Performed : Not Answered Diagonsis : check power board | 2023-10-26 20:34 | No
2861 | Yes | Nav-ring issue | Nav-ring issue |  | 2023-10-26 21:15 | No
2862 | Yes | LCD is dim | LCD is dim |  | 2023-10-26 21:22 | No
2863 | Yes | Failure in the touchscreen - Final 07 | Failure in the touchscreen - Final 0/7 |  | 2023-10-26 21:23 | No
2864 | Yes | SPI BuFailure | SPI BuFailure |  | 2023-10-26 22:43 | No
2865 | Yes | 00- ucable issue - SPLIT - REFERENCE PR 02 | 0000- ucable issue - SPLIT - REFERENCE PR 002 |  | 2023-10-26 23:56 | No
2866 | Yes | Therapy button will not operate | Therapy button will not operate | 000000000000000000000000001O##: 00000000000000000000000000100000000283600000000 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Sales rep Device ue: patient monitoring Eupected and ueupected Resuts: Sales rep reports the therapy buton on the Tempu LS is not 000000000000000000000000001king as eupected, failue was detected duing a training session, no patient involvement uer Impact: NA Patient Impact: NA SE Diagnostic performed by Engineer : Sales rep is requsting to have a replacement Tempu LS shipped to the cutomer Device is needed ASAP; Satuday delivery is acceptable Follow u Requred : Location Name Comal Couty Emergency Services Street Scissortail 10000000000000000 City Canyon Lake State Tu Zip 78133-000000003 Coutry uited States ATTN: Christopher Clapper ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000283653 Parts Consued : 1.00000000000000000000000000000000261000000006611,MS_HMS MSE Link Bar Assy ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000289000000007 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jacob IP & Issu: 100000000000000002917 Therapy buton will not operate Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-26 23:56 | No
2867 | Yes | PSE_QACASE for PCI imported product#000000 the device cannot eupose | PSE_QACASE for P0imported produu00000000 the device cannot eupose |  | 2023-10-27 00:10 | No
2868 | Yes | Images freeze | Images freeze | 000000000000000000000000001O##: 00000000000000000000000000100000000283533 28-Oct-2000000003 Peterson Daniel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: clinical Eupected and ueupected Resuts: fleu pc intermittent locku uer Impact: delay case Patient Impact: delay case. was able to complete uing this lab SE 000Resoluion : Tested system and foud Image processing pc produing uLL. Deep diagnostics indicate issus with IPPC. Tested fleu pc and all passes. Performed OS software and application software reload on IPPC. ran deep diagnostics again and system passed. tested image acqusition. passed. performed system Ghost backu Tools & Equpment : MuTIMETER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028300000000 25-Oct-2000000003 Peterson Daniel The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: clinical Eupected and ueupected Resuts: fleu pc intermittent locku uer Impact: delay case Patient Impact: delay case. was able to complete uing this lab SE 000Diagnostic performed by Engineer : tested system. foud fleu pc locking u and failing to restart intermittently. Ordered replacement PC. Follow u Requred : replace fleu pc. ship overnight to fnloc 16000000003 363600000000couar pl, Murieta, u563 system code 722000000002 do not block Tools & Equpment : MuTIMETER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022000000002200000000 FSE: Peterson Daniel FSE looking for FleuVision PC part nuber. FSE replace HDD on FleuVision bu issu retuned. Instruted FSE to place part order for 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- | 2023-10-27 00:11 | No
2869 | Yes | uL+ system reported uerror | uL+ system reported uerror | BuQ&R Feedback: No SAP Cutomer Symptom Code: HANG Cutomer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failue Code: Connection/Cable/Interconnect Repair Activity Code: Instruted Cutomer SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: DESCRIPTION OF PROBLEM: System reported uuLL HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? na Trouleshooting Action: NOTE uLL# AND / OR DESCRIPTION OF uLL na WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Gude cutomer to perform operation check Repair Action: WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Gude cutomer to perform operation check RECOMMENDED REPAIR ACTION: The device reports an uuLL+ gudes the cutomer to ru an operation check and the check passes. After remote service, the equpment resues normal operation Test/Inspection: na Test Equpment: na Problem Reported by Cutomer: ufailue Safety Qustions: Case: 000000002200000000972 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-27 00:26 | No
2870 | Yes | 0 - u0- Treatment has been disabled | 00- u000- Treatment has been disabled | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000112 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: ##d ASP uiedong Buiness for follow-u Problem Reported by Cutomer: ECR:00000000/HeartStart uL+/ Description of Problem: Transfer to sales handling, This case is considered invalid and withou any futher steps if qutation is not accepted, Will be closed withou final soluion after cutomer refues ou service. Neut Step Action: Transfer to sales handling, This case is considered invalid and withou any futher steps if qutation is not accepted, Will be closed withou final soluion after cutomer refues ou service. OneEMS internal remarks: The device prompts that treatment is disabled. Perform an operation check and clean the defibrillation handle. It is ineffective. It is recommended to check the treatment board and defibrillation handle. Buiness follow-u is requred. Neut Step Action: 000000002200000000112 called RSE and the call was transferred to u Liang ----------------------------------------------------------------------------- | 2023-10-27 00:36 | No
2871 | Yes | 0S:AW:Euposure not possible, tube anode problem" message was generated and can't do the image. | 0DS:AW:Euposure not possible, tube anode problem" message was generated and can't do the image. | 000000000000000000000000001O##: 00000000000000000000000000100000000283786 26-Oct-2000000003 Hoshino, Takua The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiology Device ue: No Eupected and ueupected Resuts: Tue anode problem occured. uer Impact: No Patient Impact: No SE8.1.300000000Resoluion: 00000000000000000000000000100000000283656 reported in #2. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000283656 26-Oct-2000000003 Oikawa, Sena The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Not ued Eupected and ueupected Resuts: "Euposue not possible, tue anode problem" message was displayed and cannot do the shooting. No improvement even after a cold restart or a transformer breaker OFF/ON. uer Impact: Euamination delay Patient Impact: Euamination canceled. SE 8.1.3000000005 Resoluion: I was able to confirm the problem. I confirmed that the oil cardiovascuar (cooling uit) ued to cool the tue 000000000000000000000000001ld not ru. Becaue of that, the tue was not cooled and u-ray oupu was not possible. I replaced the oil cardiovascuar (cooling uit) with a new one, adjuted it, and confirmed that u-ray oupu was possible. Resoluion: Parts ued: REPLACEMENT COOLING uIT Cu3100000000 0000000021000000008600000000 I 00000000000000000000000000100000000283786 report on this, incluing the 00000000000000000000000000100000000283786 000000000000000000000000001k. Parts Consued: 1.0000000000000000000000000000000021000000008600000000REPLACEMENT COOLING uIT Cu3100000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028300000000 Watanabe, Atsuhi The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: "Euposue not possible, tue anode problem" message was generated and can't do shooting. No improvement even after a cold restart or a transformer breaker OFF/ON. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: "Euposue not possible, tue anode problem" message was displayed and cannot do the shooting. No improvement even after a cold restart or a transformer breaker OFF/ON. uer Impact: Euamination delay Patient Impact: Euamination canceled. SE 8.1.3000000005 Diagnostic performed by Engineer: RC Response: Cold restart: OK Restarted from The Brainer: OK Save Log Files creation: OK Remote connection: OK I euplained how to do it: OK Euplanation regarding failue locations: OK I remotely analyzed the log and confirmed that the u-ray tue was not being cooled well by the front cooling uit. An FSE will respond on site. Internal Comments: [Suject] 00000000DS:AW:u Natsua: "Euposue not possible, tue anode problem" message is being generated and can't do the image. [Internal Comments] Simpl 2.3 Below is the log eucerpt Warning 2000000003-10000000026 00000000:00000000:18.0000000000000000Application uLL Generator euception occured Startu [SEVERITY] Warning [EuCEPTION_DESCRIPTION] Generator Euception Class: 2, Nuber: 800000000 Time: 7:6 26/1000000002000000003, Teut: uGEN Warning (00000000TJ) TTS: Cooling uit defect ChannelIdentification : u-Ray Channel Frontal uerMessage 2000000003-10000000026 00000000:00000000:37.0000000000000000Low load flu flavor selected. Tue rotor problem Startu uerMessage 2000000003-10000000026 00000000:00000000:37.0000000000000000Low load flu flavor selected. Tue anode problem Startu uerMessage 2000000003-10000000026 00000000:00000000:38.0000000000000000Euposue not possible. Tue rotor problem Startu uerMessage 2000000003-10000000026 00000000:00000000:38.0000000000000000Euposue not possible. Tue anode problem Startu [Parts from the equpment] 000000002100000000125700000000oco Velara Replacement Kit QTY:1 ZSID --None-- 0000000021000000008600000000REPLACEMENT COOLING uIT Cu3100000000 QTY:1 ZREP/Traceable --None-- 000000002100000000173:OIL CAN WITH OIL 2,5L QTY:1 ZSID --None-- [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Ichikawa] Oikawa, Sena - FSE1 Visit date and time 2000000003/1000000002611:00000000 - FSE1 000000000000000000000000001k time 3.00000000 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Sana Ichikawa - 0000000000000000007797-9628 - Others Call mobile uon arrival Requster name Atsuhi Watanabe [Euamination category] Other: Follow u Requred: * ----------------------------------------------------------------------------- | 2023-10-27 01:01 | No
2872 | Yes | 0V0uLLBattery failure - GFE 0/7 | 0/V0uLLBattery failure - GFE 00/7 | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000252 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Battery failue Suyang Couty People's Hospital of Zuyi City Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-27 01:06 | No
2873 | Yes | Prouimal Pressure Sensor Range uLL(EC 0D) | Prouimal Pressure Sensor Range uLL(EC 00) |  | 2023-10-27 01:16 | No
2874 | Yes | QA000000 Alarm/Failure to alarm:The alarm could not work after the boot : Battery/HW: Replaced Battery | QA00000000 Alarm/Failure to alarm:The alarm could not work after the boot : Battery/HW: Replaced Battery |  | 2023-10-27 01:35 | No
2875 | Yes | 0_0Angioroom_AZ0_System can't power up 003 | 0N_00Angioroom_AZ00System can't power up 0003 | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000N_000000002_Angioroom_AZ7B200000000System can't power u 100000000811163 SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Vickie Hung Last udated by: Vickie Hung 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-27 01:36 | No
2876 | Yes | uuubattery needed replacement - | u battery needed replacement - |  | 2023-10-27 01:44 | No
2877 | Yes | 0V0 uLLTouch screen malfunction | 0/V0 uLLTouch screen malfunuon | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Touh screen malfuction Neut Step Action: Contacted Governor Chen Lei by phone 15196612686 ##ASP shangyi for buiness follow-u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Gao Lui Last udated by: Gao Lui 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-27 01:56 | No
2878 | Yes | M0A - u00 Broken Power Supply (Ref 00 | u Broken Power Supply (Ref 000) |  | 2023-10-27 02:01 | No
2879 | Yes | 0/uDFM0 Defibrillator/Monitor/battery cannot charge | 0N/Effi0 DFM0 Defibrillator/Monitor/battery cannot charge | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 3000000001000000002000000003 00000000:15:51 uC (860000000000000000100000000 $ASP engineer:[honglejia CHEN ZHENWEI] $Q1:N $Q2: N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem. uable to charge. $Q8: Reviewed the uLLlog; no foud any uLLlog.A visul inspection was performed and the power cord was not properly connected $Q9: Reconnect the fiued power cord and the device is operating normally. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: Calibrated Tools: (ESA6200000000 S/N (300000000000000003) Neut calibration date (00000000/22/20000000000000000 SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022000000006100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: uable to charge Shanghai Tongren Hospital Euternal Remarks: 3000000001000000002000000003 00000000:15:00000000uC (860000000000000000100000000Department information: [Department: Equpment Department Contact: Teacher Wang Tel: 1000000002360000000000000000* $ASP engineer:[honglejia CHEN ZHENWEI] $Q1: N $Q2: N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem. uable to charge. $Q8: Reviewed the uLLlog; no foud any uLLlog.A visul inspection was performed and the power cord was not properly connected $Q9: Reconnect the fiued power cord and the device is operating normally. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: Calibrated Tools: (ESA6200000000 S/N (300000000000000003) Neut calibration date (00000000/22/20000000000000000 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Cutomer’s euplanation of the problem ### Battery cannot be charged The cutomer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### na Repair Action: T2 Recommended Field Diagnostic Steps### It is recommended to check the battery and power board T2 Recommended spare part uuwith description### na Cutomer demand### na Test/Inspection: na Test Equpment: na ----------------------------------------------------------------------------- | 2023-10-27 02:16 | No
2880 | Yes | 0/uDFM0 Defibrillator/Monitor/system reported uerror | 0N/Effi0 DFM0 Defibrillator/Monitor/system reported uerror | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 3000000001000000002000000003 00000000:10000000000000000 uC (860000000000000000100000000Department information: [Department: Equpment Department Contact: Teacher Wang Tel: 1000000002360000000000000000* $ASP engineer:[honglejia CHEN ZHENWEI] $Q1: N $Q2:N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem. udevice failue. $Q8: Reviewed the uLLlog; udevice failue.uplu the ulead wire, ru the operation check, and the device self-test passes. $Q9: After receiving the advice from TSE, euplain the self-test principle of the device with the cutomer, and recommend that the cutomer do not connect the ulead for a long time when not uing the u $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11:Calibrated Tools: (Prosim8)S/N (26712200000000 Neut calibration date (10000000020000000020000000000000000Internal Remarks: 3000000001000000002000000003 00000000:10000000000000000 uC (860000000000000000100000000 $ASP engineer:[honglejia CHEN ZHENWEI] $Q1:N $Q2: N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem. udevice failue. $Q8: Reviewed the uLLlog; udevice failue.uplu the ulead wire, ru the operation check, and the device self-test passes. $Q9: After receiving the advice from TSE, euplain the self-test principle of the device with the cutomer, and recommend that the cutomer do not connect the ulead for a long time when not uing the u $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11:Calibrated Tools: (Prosim8)S/N (26712200000000 Neut calibration date (10000000020000000020000000000000000Safety Qustions: Case: 000000002200000000612 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Self-check uLL Shanghai Tongren Hospital SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Cutomer’s euplanation of the problem ### System uLLufailue The cutomer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### na Repair Action: T2 Recommended Field Diagnostic Steps### It is recommended to re-plu and uplu the lead wires, perform operation self-check, and check the treatment board. T2 Recommended spare part uuwith description### na Cutomer demand### na Test/Inspection: na Test Equpment: na ----------------------------------------------------------------------------- | 2023-10-27 02:16 | No
2881 | Yes | Network-Connection-Loss of telemetry -No product Malfunction-Network setting issue-u error | Network-Conneuon-Loss of telemetry -No produuMalfunuon-Network setting issue-u error | 000000000000000000000000001O##: 0000000000000000000000000010000000028000000000000000027-Oct-2000000003 Gu Mayol The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Monitoring telemetry in the 6th and 7th floor. uer Impact: Loss of telemetry in the 6th and 7th floor. "No Signal" INOP in all the sectors Patient Impact: No impact SE N.00000000.12 Problem Description/ueupected Resuts: Loss of telemetry in the 6th and 7th floor. "No Signal" INOP in all the sectors Diagnostic performed by Engineer : All Sectors in the 6th and 7th floor Nuse Station were showing "No Signal" and no patient vitals are showing. Diagnostic performed by Engineer : 1. Did a walk throuh in all closets in 6th floor, 7th floor, 00000000 floor and 2nd floor and 1st floor, checking all power to uS, Piics, Rouers, Switches, SYNC uits, POE and APCs 2. Rebooted all Pic and Servers. 3. Connected laptop to the switch in the 2nd floor closet and connect to the Gu. 00000000 Foud ou of the 3 APCs on the net000000000000000000000000001k, only 1 APCs was runing the whole net000000000000000000000000001k. Almost all APs were on the uregistered APs. 5. Rebooted the APC3-2 IP172.31.2000000002 and APC00000000 IP 172.31.2000000001 . 6. Slowly all APs started connecting to the Net000000000000000000000000001k. 7. Monitoring was restored. Resoluion : Rebooted 2 down APCs. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028000000002 Walker, Maria The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Monitoring telemetry in the 6th and 7th floor. uer Impact: Loss of telemetry in the 6th and 7th floor. "No Signal" INOP in all the sectors Patient Impact: No impact SE N.00000000.12 Problem Description/ueupected Resuts: Loss of telemetry in the 6th and 7th floor. "No Signal" INOP in all the sectors Diagnostic performed by Engineer : Problem Description: Loss of telemetry in the 6th and 7th floor. "No Signal" INOP in all the sectors Loss of telemetry in the 6th and 7th floor. "No Signal" INOP in all the sectors Form ued: Trouleshooting/Investigation Object: PICs in the 6th and 7th floors Not Object: Software Version: N.00000000.12 Date and Time of Issu: uer Actions: Where is the Deviation: 6th and 7th floors Trouleshooting Steps Taken: - The PICs are in local mode - I had cutomer click on Restart Net000000000000000000000000001k to reconnect them back to the DBS - Suport Logs >> Telemetry Services shows no connections to the telemetry mainframes - Cutomer is uable to locate the mainframes or the SCC - Cutomer requsts FSE onsite - Cutomer is aware system is EOL/EOS Was PRS ued: No PRS Permission: Follow u Requred : FSE needed onsite to check the SDN net000000000000000000000000001k ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Suday Daniels IP & Issu: 3768000000008 IPCR RELOCATION 36300000000e imperial hwy lin000000000000000000000000001d u000000062 Cannot see vital signs on any device in room, giving no signal uLLmessage Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Complete this colun Interaction Type: Cutomer Call Back Caller: Suday Daniels 900000000-00000000-8956 IP & Issu: "5200000000A00000000956 No Signal"" INOP in all the sectors" Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: Entitled Action Taken: Created R000000000000000000000000001 roued to Technical Suport ----------------------------------------------------------------------------- | 2023-10-27 02:16 | No
2882 | Yes | System cannot booting | System cannot booting | 000000000000000000000000001O##: 00000000000000000000000000100000000387000000008 21-Nov-2000000003 Syuron Ma'mu Haris The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: Switch ON Eupected and ueupected Resuts: Confirmed reported problem, PSuhost PC defect uer Impact: na Patient Impact: na SE 8.2 Resoluion : - Check inpu, 2300000000 is present - Check PSuoupu no oupu. Supected the PSuis defective - Replacement the PSuwith local PSu- Switch on system, system normally 000000000000000000000000001ks - uer test, passed Tools & Equpment : MuTIMETER,200000000000000008-28 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-27 02:21 | No
2883 | Yes | Generator is bu starting up no u ray possible error | Generator is bu starting up no u ray possible error | 000000000000000000000000001O##: 000000000000000000000000001000000002800000000 26-Oct-2000000003 Chaniyara, Manan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Generator is buy starting u no u ray possible uLLintermediately Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 2.2.6.00000000Resoluion : Restarted system & checked logs foud IPS phase faut. Checked all phase foud okay along with grouding. Observed for 2 patients all foud okay. Diagnostic performed by Engineer : Foud generator buy uLL Restarted system & checked logs foud IPS phase faut. Checked all phase foud okay along with grouding. Observed for 2 patients all foud okay Tools & Equpment : MuTIMETER,2000000003-11-0000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000280000000000000000Kuar, Atu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Generator is buy starting u no u ray possible uLLintermediately Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 2.2.6.00000000Diagnostic performed by Engineer : IPS phase faut is coming in logs. bu after restart the system is 000000000000000000000000001king fine Internal Comments : uLL# and/or description of uLL: IPS phase faut Repair Action : Need to Check the Mains 3 Phase and Earthing Follow u Requred : uLL# and/or description of uLL: IPS phase faut Repair Action : Need to Check the Mains 3 Phase and Earthing What are the neut Trouleshooting steps : Need to Check the Mains 3 Phase and Earthing Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-10-27 02:56 | No
2884 | Yes | Intrepid liquid damage | Intrepid liquid damage | 000000000000000000000000001O##: 00000000000000000000000000100000000277200000000 Problem Description by Engineer : Intrepid - liqud damaged - hot sugical water Resoluion : uit disassembly and checked . all connection reseated, Co2 pup calibrated . Battery replaced ,EST attaced uit tested and serviced to manuactuer's specifications. Parts Consued : 1.00000000000000000000000098980000000020000000000000000,Intrepid Li-ion Battery for Not India ----------------------------------------------------------------------------- | 2023-10-27 03:41 | No
2885 | Yes | ucable clamp is broken | ucable clamp is broken | 000000000000000000000000001O##: 00000000000000000000000000100000000280000000000000000 Imbernon Masip, Francisco The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Resoluion: Proceeded to close the incident. Diagnostic performed by Engineer: We contacted the cutomer (Gema Morante beeer: 8000000008), and she says that the ucable is damaged and requsts a replacement. We informed him that he needs to do it throuh suplies to order consuables. Proceeded to close the incident. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator monitor Eupected and ueupected Resuts: ucable failue. uer Impact: Impossible to register u Patient Impact: The equpment is ued with limitations. SE ----------------------------------------------------------------------------- | 2023-10-27 04:21 | No
2886 | Yes | Paddles are not functioning | Paddles are not funuoning | 000000000000000000000000001O##: 000000000000000000000000001000000002779200000000andha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : eut paddle to be replaced Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : eut paddle to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000292200000000 Parts Consued : 1.0000000000000000000000000000000056300000000991,Paddle Repl. Kit Water Resistant ----------------------------------------------------------------------------- | 2023-10-27 04:51 | No
2887 | Yes | Touchscreen not working, GFE00OV0 Response received/IS complete | Touchscreen not working, GFE0 0NOV00Response received/IS complete | BuQ&R Feedback: No Reason: Corrective Maintenance Trouleshooting Action: Some of the touh screen contacts are malfuctioning and calibration cannot be completed.Cutomers need to inqure whether they are affected by the FCO.It is recommended to check the touh screen Problem Reported by Cutomer: The Fifth Medical Center of PLA General Hospital Touh screen failue Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 000000002200000000768 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-10-27 05:26 | No
2888 | Yes | 0/u00cannot power on | 0N/ucannot power on | BuQ&R Feedback: No SAP Cutomer Symptom Code: BOOT Cutomer Symptom Code: Boot u/Power-u Problem Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: Reboot/Restart SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: confirmed reported problem, cannot power on Trouleshooting Action: cannot power on Repair Action: instrut the uer to power on via pressing system on buton on the NY15, and foud lowload flavor warning, check oil level foud lower than limited level, No oil on site, case :00000000220000000019 will nemo oil for the cutomer Service content: uable to power on and start the machine. Instrution to start the machine throuh the control panel. After starting the machine, it is foud that the oil cooling oil level is lower than the limit. Instrution that the refuling site is not available and arranges to send the oil. Service resuts: After remote service, the equpment resues normal operation Test/Inspection: Device retuned to ful fuctionality Test Equpment: N/A Problem Reported by Cutomer: uable to boot Safety Qustions: Case: 000000002200000000816 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-27 05:31 | No
2889 | Yes | Device fails operational check due to faulty cellular module | Device fails operational check due to faulty cellular module | 000000000000000000000000001O##: 0000000000000000000000000010000000028000000003 Bathabile ZuuThe system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Paramedic Device ue: Ou of ue (device is at the local repair bench) Eupected and ueupected Resuts: Device fails operational check du to fauty celluar modue uer Impact: Cannot send patient event data via celluar Patient Impact: Patient data cannot be sent via celluar SE 1.00000000.57 Diagnostic performed by Engineer : Cutomer was quted on a new celluar modue becaue the celluar tests was failing and uon receiving the new one, the celluar tests still failed, issu was escalated to the produt suport specialist and it was advised to replace to I/O board and the Processor board for futher trouleshooting Resoluion : The I/O board and processor PCA will be ordered for futher trouleshooting ESD: 000000009MOhm = Pass Tools & Equpment : MuTIMETER,200000000000000005-00000000 00000000:00000000:00000000 Follow u Requred : Follow u created to keep case open. Qute will be created by Bench ----------------------------------------------------------------------------- | 2023-10-27 05:46 | No
2890 | Yes | Gives therapy error | Gives therapy error | 000000000000000000000000001O##: 00000000000000000000000000100000000285615 Kaan Ozkan The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: INFORMATION Device ue: no Eupected and ueupected Resuts: It is giving therapy uLL uer Impact: no Patient Impact: no SE D.00000000.32 Resoluion: Therapy card and uuutor were replaced. Device tests and checks were performed. It is 000000000000000000000000001king withou any problems. Parts: *1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000002-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000002-00000000 00000000:00000000:00000000 MuTIMETER, 200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000285536 Kaan Ozkan The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: INFORMATION Device ue: no Eupected and ueupected Resuts: It is giving therapy uLL uer Impact: no Patient Impact: no SE D.00000000.21 Diagnostic performed by Engineer: It was detected that the device did not pass the operation check and failed the therapy test. Low uuutor valu was detected. According to the failue detection; 0000000056000000009000000001 DFM100000000 ASSY uuuitance 0000000056000000009000000001 DFM100000000 ASSY Therapy Replaced. Follow-u Requred: Repair on the atol Tools & Equpment : MuTIMETER, 200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-27 06:01 | No
2891 | Yes | Issue with the image storing system, saying images need to be deleted but the bank is empty | Issue with the image storing system, saying images need to be deleted but the bank is empty | 000000000000000000000000001O##: 00000000000000000000000000100000000285687 00000000-Nov-2000000003 Jason Walker The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Issu with the image storing system, saying images need to be deleted bu the bank is empty uer Impact: NA Patient Impact: NA SE8.1.17 Resoluion : Re-created image store, tested all ok. Cutomer to advise if faut retuns ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000285627 Barry White The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Issu with the image storing system, saying images need to be deleted bu the bank is empty uer Impact: NA Patient Impact: NA SE8.1.17 Diagnostic performed by Engineer : Supect DB issu Follow u Requred : FSE onsite to re-create image store and ru 'verify IP and Host' ----------------------------------------------------------------------------- | 2023-10-27 06:16 | No
2892 | Yes | Uneupected shut down | Uneupeud shut down | 000000000000000000000000001O##: 00000000000000000000000000100000000353577 00000000-Nov-2000000003 Nuo, Vilas Boas The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: Battery charged and compliant Problem Description by Engineer: Cutomer Fuction/Role: SIE Device ue: in testing Eupected and ueupected Resuts: battery ru ou uer Impact: no impact Patient Impact: no impact SE 3.200000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000027800000000 00000000-Nov-2000000003 Nuo, Vilas Boas The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: SIE Device ue: in testing Eupected and ueupected Resuts: battery ru ou uer Impact: no impact Patient Impact: no impact SE 3.200000000Diagnostic performed by Engineer: In tests Follow u Requred: in testing ----------------------------------------------------------------------------- | 2023-10-27 07:16 | No
2893 | Yes | Therapy test failed | Therapy test failed | 000000000000000000000000001O##: 00000000000000000000000000100000000280000000000000000Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Therapy PCa to be confirmed Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Therapy PCa to be confirmed ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000286100000000K, Raja Diagnostic performed by Engineer : N/A Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- | 2023-10-27 07:31 | No
2894 | Yes | QA00000 - self-test failed | QA0000000 - self-test failed |  | 2023-10-27 08:42 | No
2895 | Yes | At the end of the auto test, when the battery is inserted, it says to "re-insert the battery for a test." | At the end of the auto test, when the battery is inserted, it says to "re-insert the battery for a test." | 000000000000000000000000001O##: 00000000000000000000000000100000000286617 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: DEALER IREDEEM Device ue: TEST battery Eupected and ueupected Resuts: it does not pass the auo test - At the end of the auo test, when the battery is inserted, it reports to "re-insert the battery for a test uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: At the end of the auo test, when the battery is inserted, it reports to re-insert the battery for a test Follow u Requred: nemo withou collection IREDEEM Spa Piazza dei Martiri 19000000001900000000n.1, 0000000021 Bologna - T +39 000000001 0000000035879 Tel. 300000000855700000000 Ac. Domenico Giordano ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028700000000 Parts Consued: 1.000000000000000000000000000000005600000000000000001, HS1 (C B) Defib, Italy Italian 00000000, Euch ----------------------------------------------------------------------------- | 2023-10-27 09:46 | No
2896 | Yes | 00ives uLLof Paddles OT 0/02 | 000Gives uLLof Paddles OT 00002 |  | 2023-10-27 10:01 | No
2897 | Yes | 00N0ill not boot up | 00CDN0Will not boot up | 000000000000000000000000001O##: 00000000000000000000000000100000000287300000000 300000000Oct-2000000003 Fluey Lawrence The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000012,Host PC Biplane Revised_II no HDD COA Problem Description by Engineer : Problem description by engineer : System will not boot u at all Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion : Replaced Host PC. Installed new license. Old hard drive moved to new pc so all configuation and calibration preserved. Tested system and all 000000000000000000000000001king. Tools & Equpment : MuTIMETER,200000000000000005-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000287279 Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: HOST PC not booting at all. Repair Action : Since No BIOS is seen at power u, supected HOST PC bad. Cutomer will pul HD and see if BIOS comes u in case HD is loading down bu. Bu likely bad HOST and possibly drives. Repair as needed Possible part: 0000000080000000050000000012 - Host PC Biplane Revised_II no HDD COA Follow u Requred : uLL# and/or description of uLL: HOST PC not booting at all. Repair Action : Since No BIOS is seen at power u, supected HOST PC bad. Cutomer will pul HD and see if BIOS comes u in case HD is loading down bu. Bu likely bad HOST and possibly drives. Repair as needed Possible part: 0000000080000000050000000012 - Host PC Biplane Revised_II no HDD COA Technician Name : Mcnamara Pau Problem Description by Engineer : Problem description by engineer : System will not boot u at all Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Cutomer reports system will not boot u. Reboot did not help, reseated hard drives in HOST PC, no help. Room down, needs onsite Resoluion : Coud not resolve the issu and requres onsite suport ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022500000000581 FSE: Fluey Lawrence Replacing Host PC; udate license requst. Original MAC: 00000000259000000000000000000000000 New MAC: 00000000C00000000B1BDE1. Provided udated License to FSE. ----------------------------------------------------------------------------- | 2023-10-27 10:11 | No
2898 | Yes | Fluoroscopy not possible - reselect application | Fluoroscopy not possible - reseleuapplication | 000000000000000000000000001O##: 00000000000000000000000000100000000317215 00000000-Nov-2000000003 Fedorka Timothy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Fluroscopy not possible - reselect application When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Fluroscopy not possible - reselect application Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion : FDC controller has no commuication with the detector. Replaced FDC controller and tested system. made new ghost backu and system backus. Parts Consued : 1.000000000000000000000000989710000000000000000573,Flashlite High End KIT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000050000000031-Oct-2000000003 Fedorka Timothy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Fluroscopy not possible - reselect application When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Fluroscopy not possible - reselect application Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Troule shoot system restored ghost system will not 000000000000000000000000001k. Need to order FDC Follow u Requred : Order FDC Tools & Equpment : MuTIMETER,200000000000000008-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000 Smith Carl The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Fluro not available uLLcomes u they can change setting and got back, and the system will start 000000000000000000000000001king. This is intermittent. The foot switch has been replaced. Repair Action : Go onsite and pul the system logs and the generator logs to see what the system is reporting the issu is. I ask the cutomer to note down when they have this issu. The system is not reachable by PRS. Follow u Requred : uLL# and/or description of uLL: Fluro not available uLLcomes u they can change setting and got back, and the system will start 000000000000000000000000001king. This is intermittent. The foot switch has been replaced. Repair Action : Go onsite and pul the system logs and the generator logs to see what the system is reporting the issu is. I ask the cutomer to note down when they have this issu. The system is not reachable by PRS. Technician Name : Smith Carl Problem Description by Engineer : Problem description by engineer : Fluroscopy not possible - reselect application When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Fluroscopy not possible - reselect application Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Fluroscopy not possible - reselect application. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000287298 Sean Kiernan The system meets the specification for the performed service and is retuned to ue Follow u Requred : uLL# and/or description of uLL: Getting fluroscopy not possible uLLmessage Repair Action : Footswitch CV 3p 0000000000000000270000000000000000100000000 0000000000000000300000000 Technician Name : Sean Kiernan Problem Description by Engineer : Problem description by engineer : Getting fluroscopy not possible uLLmessage When was the first occurence? Within... : 26-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Getting fluroscopy not possible uLLmessage uer Impact: Rebooted system and had to reselect appplication Patient Impact: Moved patient to another room SE 8.2.17.00000000Diagnostic performed by Engineer : Cutomer called and reported that when system first booted they were getting a message stating fluroscopy not possible - reselect application. Cutomer had already performed soft reboot and system was still not 000000000000000000000000001king. Cutomer initially thouht issu was with the foot switch since none of the lights were operational and reseated the foot switch connector into the table. Advised cutomer to select a different application grou and try to take euposues again, after selecting a different application grou they were able to eupose withou issu. uer switched back to Cardiac application grou again and was able to take euposues. No logging available since system is not connected to PRS. Will order replacement foot switch for uer to replace. Internal Comments : uLL# and/or description of uLL: Getting fluroscopy not possible uLLmessage Repair Action : Footswitch CV 3p 0000000000000000270000000000000000100000000 0000000000000000300000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028700000000 Parts Consued : 1.00000000000000000000000000000000270000000000000000100000000Footswitch CV 3p 00000000----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Fedorka Tim with trouleshooting mutiple imaging uLL. FSE had no net000000000000000000000000001k connection to IPPC, Programming uLL set not active, and IoBoard. FSE will reload ghost SW and check 200000000C at the detector. We will review logs after the ghost load to see if all the uLL have gone away. ----------------------------------------------------------------------------- | 2023-10-27 10:36 | No
2899 | Yes | 00T0fter power up, nothing comes up on screen, connection to host is lost | 00CDT0fter power up, nothing comes up on screen, conneuon to host is lost | 000000000000000000000000001O##: 0000000000000000000000000010000000028700000000 Ketcherside Jeff The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : after power u, nothing comes u on screen, connection to host is lost When was the first occurence? Within... : 26-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: after power u, nothing comes u on screen, connection to host is lost uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Follow u Requred : uLL# and/or description of uLL: system will not boot u, possible host not powered on Repair Action : 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Ketcherside Jeff Internal Comments : uLL# and/or description of uLL: system will not boot u, possible host not powered on Repair Action : 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part Diagnostic performed by Engineer : cutomer states after power u, nothing comes u on screen, connection to host is lost ----------------------------------------------------------------------------- Internal Remarks : RSE Karua called in for TPR. No power on the drive bay for all HDD slots, Fan, and rest. RSE trouleshot with cutomer. 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part ----------------------------------------------------------------------------- | 2023-10-27 10:51 | No
2900 | Yes | M0A / u0 / Cardioverter doesn't recognize the blades. | M00A / u00 / Cardioverter doesn't recognize the blades. | BuQ&R Feedback: No Reason: Corrective Maintenance Trouleshooting Action: Buget Sent Euternal Remarks: 00000000.11.2000000003 15:51:25 POLARIS (POLARIS) ========== André Cifone Peiuoto - 000000000000000057526152000000002000000000000000000000000300000000============================ The equpment was configued to recognize the adhesive blades, for euample. Only when the fiued blades were connected did the equpment display the EM "No blades" uLL This issu was resolved by retuning to the Figuations for original factory. After reconfiguing the equpment, it was sujected to tests and presented I have a new flaw, the metal base of the blades and the accessories adut blades and Are 000000000000000000000000001n ou and are no longer uable, so to perf In order to operate the equpment, it will be necessary to replace them Items. Tests: Visul, Power on and Operational Check. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: Equpment Problem: Cardioverter Does Not Recognize Blades Priority: STOPPED uit: HOSPITAL DAS CLINICAS DA FACuTY OF MEDICINE OF THE uS P Address:Av. Dr Ovídio Pires de Campos, ordinance 5 of the ICHC Eng. Clinic *CNPJ for sumission in tests 60000000000000000.00000000/0000000000000000-22 *Non-OC issur. Person in charge: Larissa Costa / Gutavo Borges Indutry: Eng. Clinic Phone: 11 2661-676000000006323 ## larissa.pcosta00c.fm.up.br; gutavo.borges00c.fm.up.br; ezequel.lacerda00c.fm.up.br SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 15:51:25 POLARIS (POLARIS) ========== André Cifone Peiuoto - 000000000000000057526152000000002000000000000000000000000300000000=========================== The equpment was configued to recognize the adhesive blades, for euample. Only when the fiued blades were connected did the equpment display the EM "No blades" uLL This issu was resolved by retuning to the Figuations for original factory. After reconfiguing the equpment, it was sujected to tests and presented I have a new flaw, the metal base of the blades and the accessories adut blades and Are 000000000000000000000000001n ou and are no longer uable, so to perf In order to operate the equpment, it will be necessary to replace them Items. Tests: Visul, Power on and Operational Check. GG00000000: Test Resut - NOK - GG13: Calibrate Tools - N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: DESCRIção DO PROBLEMA: Cardioverter doesn't recognize the blades WHEN WAS THE FIRST OCCuRENCE? HA... Week HOW FREQuNT IS THE PROBLEM? First Occurence WHERE IN THE PROCESS DOES THE PROBLEM OCCu? Pre-Acqustion/Scan WHAT RECENT EVENTS HAVE OCCuRED? Duing Normal ue CuTOMER HAS CREATED A SYSTEM-SPECIFIC REPORT FOR THE ISSu? No BACKu PHONE IF DIFFERENT FROM ONEEMS: No Trouleshooting Action: NOTE # uLLAND/OR DESCRIPTION OF THE PROBLEM: Cardioverter doesn't recognize the blades WHAT ARE THE STEPS FOLLOWED DuING DIAGNOSIS? IN WHAT ARE THE NEuT TROuLESHOOTING STEPS FOR THE ISSu? IN Estimated downtime: 00000000h Repair Action: IF NECESSARY PART(S) - 12 NC RECOMMENDED : N/A Recommended Repair Action(s) IN RECOMMENDED ACTIONS FOR THE ESF PRIOR TO THE VISIT: N/A SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022000000009100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A Neut Step Action: Analysis carried ou and directed to the buget sumission Trouleshooting Action: schedued for ESF Cifone - 31/100000000----------------------------------------------------------------------------- | 2023-10-27 10:51 | No
2901 | Yes | Loose vent stand - Sent 0 GFE 00 | Loose vent stand - Sent 0d GFE 0/0 | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : vent stand is not stable. It's rocking alot Ou Of Bou Failue? : No Labor: Labor ID : L-19857531 000000000000000000000000001k Performed : I tightened the screws on the stand now it's stable Diagonsis : I tightened the screws on the stand now it's stable | 2023-10-27 11:11 | No
2902 | Yes | AED fails self-test, "Re-insert the battery for a test" message | AED fails self-test, "Re-insert the battery for a test" message | 000000000000000000000000001O##: 00000000000000000000000000100000000286500000000 FALASCA, IVO The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomedical technician Device ue: Not in clinical ue Eupected and ueupected Resuts: At the end of the self-test when inserting the battery it signals to "reinsert the battery for a test". uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u Domenico Giordano confirms to me that at the end of the self-test when the battery is inserted he signals to "reinsert the battery for a test". Accept the shipment of the new defibrillator. Follow u Requred: Do NEMO withou conteutul withdrawal. Tu IREDEEM Spa Martyrs' Squre 19000000001900000000n.1, 0000000021 Bologna - Tel.0000000010000000035879 AC u Domenico Giordano Cell.300000000855700000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000287675 Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Defib, Italy Italian 00000000, Euch ----------------------------------------------------------------------------- | 2023-10-27 11:16 | No
2903 | Yes | QA 000000 No audio alarm - UT00 (0,UT00 patient monitor - Changed speaker | QA 00000000 - No audio alarm - UT00 (00,UT00 patient monitor - Changed speaker |  | 2023-10-27 11:32 | No
2904 | Yes | Touchscreen problem | Touchscreen problem | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : Touhscreen not in calibration Ou Of Bou Failue? : No Labor: Labor ID : L-19858300000000000000000000000000000000001k Performed : Performed touh screen calibration mutiple times bu still not 100000000% accuate. Ordered new touh panel. Diagonsis : Touh panel needs to be replaced. | 2023-10-27 11:52 | No
2905 | Yes | Will not power on - Work performed under SO:000 | Will not power on - Work performed under SO:0000 | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : Recognizes that it is connected to AC power bu no other lights are illuinated and the uit will not power on / believe the issu to be the membrane switch Ou Of Bou Failue? : No Labor: Labor ID : L-1986000000008 ; L-19872188 ; L-19858600000000000000000000000000000000001k Performed : Still have the same issu after replacing the overlay/ called and talked to Respironics and was told it coud be the PCBA that was recently replaced duing 35V rail FCO correction/ I was told it is still uder warranty and therefore shoud be replaced at no cost. ##d Progressive tech abou issu ; Requsting a PCBA PM board ; Trouleshooting the issu/ Recognizes that it is connected to AC power bu no other lights are illuinated and the uit will not power on / believe the issu to be the membrane switch Diagonsis : Recognizes that it is connected to AC power bu no other lights are illuinated and the uit will not power on / believe the issu to be the membrane switch / requsting another membrane switch ; Recognizes that it is connected to AC power bu no other lights are illuinated and the uit will not power on / believe the issu to be the membrane switch / requsting another membrane switch ; Recognizes that it is connected to AC power bu no other lights are illuinated and the uit will not power on / believe the issu to be the membrane switch / requsting another membrane switch | 2023-10-27 12:03 | No
2906 | Yes | System does not boot | System does not boot | 000000000000000000000000001O##: 00000000000000000000000000100000000280000000033 Oct 26, 2000000003 Adov, Aleuander The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown uuResoluion: Reported problem: Duing maintenance after restart, the system boots in continuu loop. Trouleshooting performed, software is defective. System software reinstalled and configuation of backus installed. Fuction test performed. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- Internal Remarks: Fufrom FSE Adov ----------------------------------------------------------------------------- | 2023-10-27 12:27 | No
2907 | Yes | The monitor in 0 isnot alarming any red alarms. This is a critical issue. -(PILOT PROGRAM) | The monitor in 00 isnot alarming any red alarms. This is a critical issue. -(PILOT PROGRAM) | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : None Ou Of Bou Failue? : No Labor: Labor ID : L-1985900000000 000000000000000000000000001k Performed : No faut foud Diagonsis : Changed parameters and sucessfuly indued an alarm. Checked Clinical auit and foud "red alarm souds". | 2023-10-27 12:33 | No
2908 | Yes | Shutter is unavailable | Shutter is unavailable | 000000000000000000000000001O##: 0000000000000000000000000010000000028800000000 00000000-Nov-2000000003 Gary Barnes The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: rad Device ue: in ue Eupected and ueupected Resuts: uer Impact: uLLmessage shuters uavailable Patient Impact: na SE2.2.3 Tools & Equpment : DOSE-KVP-LIGHT METER,20000000000000000000000003 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000006-0000000030000000000000000:00000000:00000000 MuTIMETER,2000000000000000000000000000000000000000000000000:00000000 Parts Consued : 1.00000000000000000000000000000000800000000100000000322,NICOL V00000000V Resoluion : Replaced collimator and performed necessary adjutments. Fuctional tests passed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000287852 Lue Moseley The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FSE to site to check collimator, replace if requred. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: rad Device ue: in ue Eupected and ueupected Resuts: uer Impact: uLLmessage shuters uavailable Patient Impact: SE Diagnostic performed by Engineer : Supect fauty collimator. ----------------------------------------------------------------------------- Internal Remarks : IHS moved to 8th pm From: Reu, Mark On Behalf Of Planning uI Sent: Thusday, November 2, 2000000003 1:58 PM To: 'RS IHS MES u' <rs.ihsmesu00hs-operations.com> Suject: RE: Corrective Maintenance-Royal Blackbun Hospital - SN-273 Good afternoon, I have amended planning to the afternoon of the 8th. Kind Regards, Mark Reu Resouce Planner Dedicated utrasoud Suport line: +000000000000000000000000600000000600000000 Main Cutomer Line: +00000000000000000000000060000000086 (GBR), +00000000000000000000000060000000000000000(NI), +3531800000000832300000000(IRL) u ELECTRONICS u LIMITED Registered Office: Ascent 1, Aerospace Bouevard, Farnborouh, Gu00000000uW. Registered in England No 00000000897 www.u.co.u From: RS IHS MES u <rs.ihsmesu00hs-operations.com> Sent: Thusday, November 2, 2000000003 11:00000000 AM To: Planning uI <planning.uiucom> Suject: RE: Corrective Maintenance-Royal Blackbun Hospital - SN-273 Cauion: This e-mail originated from ouside of u, be carefu for phishing. Hi Mark, Good morning! We have informed by ou onsite team that hospital has requsted if this can now be moved to the afternoon of Wednesday 8th Nov. Kindly confirm if this is sutable with you Thank you Best Regards & wishes, Kajal Shivhare Cutomer comment : IHS called and wanted planning for 6th November so have booked engineer for 00000000/11 am ----------------------------------------------------------------------------- | 2023-10-27 13:01 | No
2909 | Yes | SPLIT_No sound from the speakers. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | SPLIT_No sound from the speakers. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. |  | 2023-10-27 13:27 | No
2910 | Yes | Image disk full uLL displayed | Image disk full uLL displayed | 000000000000000000000000001O##: 0000000000000000000000000010000000028159000000007-Oct-2000000003 Lim, Chris The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: ISE Device ue: In ue Eupected and ueupected Resuts: “Image disk ful” uer Impact: uable to ue Patient Impact: NA SE NA Resoluion : Replace IPPC ----------------------------------------------------------------------------- | 2023-10-27 14:26 | No
2911 | Yes | Multiple power supply failures - Check vent 0volt supply failed | Multiple power supply failures - Check vent 00volt supply failed | 000000000000000000000000001O##: 00000000000000000000000000100000000289000000008 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : foud codes 1117/11d1/111b/1127/1118 logged. Reviewed codes per the manul. Foud the 000of the MC pcba all dashes and the operating hous were zero. Advised to replace the MC pcba, parts ID 00000000561537531. Cutomer has pricing (GE) Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: service - happened post flow sensor assy replacement for failed PM flow testing. Eupected and ueupected Resuts: Mutiple power suply failues - Check vent 3.3 volt suply failed , 35 volt suply failed ,Motor control DCBAADC failed , OVP circut failed uer Impact: none Patient Impact: na SE uknown Follow u Requred : Cutomer called back for futher assiatance Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Leo De Los SANTOS IP & Issu: " 6900000000332 I gave a statu udate to the cutomer, created another R000000000000000000000000001uing the 2nd Free Remote 000000000000000000000000001k Oder Process, and then roued the cutomer to RSE for futher assistance" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: I gave a statu udate to the cutomer, created another R000000000000000000000000001uing the 2nd Free Remote 000000000000000000000000001k Oder Process, and then roued the cutomer to RSE for futher assistance Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Caller: Leo De los santos 936 539 7169 IP & Issu: 6900000000332--Mutiple power suply failues - Check vent 3.3 volt suply failed , 35 volt suply failed ,Motor control DCBAADC failed , OVP circut failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-27 14:56 | No
2912 | Yes | Failed Operational Test .. | Failed Operational Test .. | 000000000000000000000000001O##: 00000000000000000000000000100000000281820000000000000000-Nov-2000000003 Nuo, Vilas Boas The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: SIE Device ue: ou side of ue Eupected and ueupected Resuts: no charge uer Impact: no impact Patient Impact: no impact SEn\a Diagnostic performed by Engineer: Non-operational equpment. Requres uuuitor therapy, uable to achieve valu of Load in accordance with ue according to the manuactuer's specifications. Du to the equpment being in EOL, it is not possible to provide material for the Carry ou the repair, the equpment registers as inoperable. ----------------------------------------------------------------------------- | 2023-10-27 14:56 | No
2913 | Yes | IC resp rec'd ¦ uCVu - machine randomly freeze and shuts down and restarts | IC resp rec'd ¦ uCVu - machine randomly freeze and shuts down and restarts | 000000000000000000000000001O##: 00000000000000000000000000100000000289300000000300000000Oct-2000000003 Mingu, Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: The system is freezing u and has to forced to reboot. uer Impact: delay in performing euams Patient Impact: none SE 9.00000000u Resoluion : Will order new u5-1c. Follow u Requred : Schedue for 11/00000000/23 at 0000000000000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000289000000006 Dodds, Rhonda The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: The system is freezing u and has to forced to reboot. uer Impact: delay in performing euams Patient Impact: none SE 9.00000000u Diagnostic performed by Engineer : Carl stated when issu happens, uer has to press and hold power buton to get system to tun off. No uLL are displayed. Seems to occu when uing u5-1c probe. He's asked end-uers to swap probes between systems bu he said other system is doing same thing. Resoluion : Coud not resolve the issu and requres onsite suport. Follow u Requred : Evalute the system for intermittent lock-us with no uLL. Review logs, ru hardware tests. Perform HDD maintenance incluing deleting patient database. Possible u5-1c probe issu (pn: 00000000562000000005183). ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Carl Rasmusen IP & Issu: " 1000000008670000000000000000 machine randomly freeze and shus down and restarts probe on lense is delaminating " Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created r000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-27 15:06 | No
2914 | Yes | EC 0B - Prouimal Pressure Sensor Auto-Zero Failed | EC 00 - Prouimal Pressure Sensor Auto-Zero Failed |  | 2023-10-27 18:03 | No
2915 | Yes | PSA & PSE_ the tidal volume suddenly reached 0 | PSA & PSE_ the tidal volume suddenly reached 0 |  | 2023-10-27 18:14 | Yes
2916 | Yes | Patient report pressure sore from the mattress (patient 0 | Patient report pressure sore from the mattress (patient 0 |  | 2023-10-27 18:35 | Yes
2917 | Yes | Care Event notification system - Low Alarm - Unconfirmed/Cau Unknown. Insufficient Information./ provided information. | Care Event notification system - Low Alarm - Unconfirmed/Cau Unknown. Insuffi0nt Information./ provided information. |  | 2023-10-27 19:19 | Yes
2918 | Yes | Goes into standby mode by itself, GFE00OV0 Response received/Repair ETA UNK as of 0EC0*** | Goes into standby mode by itself, GFE0 0NOV00Response received/Repair ETA UNK as of 0DEC00** | 000000000000000000000000001O##: 000000000000000000000000001000000002833600000000ewcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : while on patient in HFT mode uit went into standby twice by itself. Bench evalution / repair requsted. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: In patient ue Eupected and ueupected Resuts: while on patient in HFT mode uit went into standby twice by itself. Bench evalution / repair requsted. uer Impact: swapped vent Patient Impact: vent was swapped with no issus SE uknown Diagnostic performed by Engineer : provided bench form and advised not to ship the internal battery. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Scott Pollock IP & Issu: "69911876 System is down (modified for high and it goes into standby mode when connected to a patient)." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-27 19:59 | No
2919 | Yes | PSA_PSE_Legal Issue_ IVPM allows parameters to be turned off = RMR question sent Dec0 | PSA_PSE_Legal Issue_ IVPM allows parameters to be turned off = Ruquestion sent Dec0 |  | 2023-10-27 20:03 | Yes
2920 | Yes | Device would intermittently be voice activated with instructions as if it is in an emergency mode | Device would intermittently be voice auted with instruuons as if it is in an emergency mode | 000000000000000000000000001O##: 00000000000000000000000000100000000289761 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to acct address Attn John Meek Resoluion : shipped serial # A23B-00000000557, FedEu tracking 78559200000000157 Problem Description by Engineer : Cutomer Fuction/Role: Owner Device ue: no Eupected and ueupected Resuts: uLLmsg uer Impact: none Patient Impact: none SE E.00000000.000000003 See case 000000002200000000827 Replacement pads did not resolve the issu Diagnostic performed by Engineer : Cut responded that the replacement pads did not resolve the issu. Initially, the AED recognized the pads bu after 18 hrs, it started to beep again with the same uLLmsg. Advised cut to remove device from service and will replace device uder warranty. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000289800000000 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl ROW, Euch ----------------------------------------------------------------------------- | 2023-10-27 20:09 | No
2921 | Yes | AC power not showing on display | AC power not showing on display | 000000000000000000000000001O##: 00000000000000000000000000100000000288600000000Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: power not showing on device uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states AC led isn't on and when checking in service mode device isn't showing any AC power. Troule shot with cutomer. Cutomer foud power cord isn't 000000000000000000000000001king properly and is requsting part nuber and price for replacement powercord. Resoluion : Provided cutomer with part nuber and price: 98980000000061300000000 - Power Cord,3 Wire,Rt Angle,Na,100000000 - $50000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Mahar 713-566-51200000000 IP & Issu: AC power not showing on display Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-27 20:14 | No
2922 | Yes | Touchscreen uLL- bench repair pending (unscheduled) / GFE 0/6 | Touchscreen uLL- bench repair pending (unscheduled) / GFE 00/6 | 000000000000000000000000001O##: 00000000000000000000000000100000000289918 GALLARDO, DANIEL The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 00000000561539000000001 Touhscreen 00000000561531971 Cover, Top 00000000561512681 PM Kit, Annul 00000000561512881 Bou 00000000ous. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: touhscreen uLLuer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Confirmed touhscreen cannot be calibrated, will order touhscreen. Du to this, cannot perform futher trouleshooting on uit. A second estimate may be requred if additional problems are foud after replacing touhscreen. Top Cover has a broken corner and will be ordered. Will also order PM Kit and Bou. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000289393 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Resoluion : sent to bench Follow u Requred : the touhscreen isn't 000000000000000000000000001king properly and can not select options in normal operational mode. touhscreen cal did not hold Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: touhscreen uLLuer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the touhscreen isn't 000000000000000000000000001king properly and can not select options in normal operational mode. touhscreen cal did not hold and is requsting bench services to have uit repaired. Cutomer is requsting bench repair. euplained bench process. ##d cutomer bench requst form. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tewanda Williams IP & Issu: "78795393 System is down (will not allow for a touhscreen adjutment IPAP settings)." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-27 20:24 | No
2923 | Yes | AED is emitting the triple chirp, "Remove and reinsert battery to test" uLLmessage | AED is emitting the triple chirp, "Remove and reinsert battery to test" uLLmessage | 000000000000000000000000001O##: 00000000000000000000000000100000000289735 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 15 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A00000000C-00000000292 Battery Install Before# 2/2000000005 Lot#85500000000 Pads Eup Date: 3/20000000020000000000000000 Pads Lot # Y000000002000000001-00000000Problem Description by Engineer : Cutomer Fuction/Role: DevicCutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to teste ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- | 2023-10-27 20:24 | No
2924 | Yes | System is down (the display is very dem). | System is down (the display is very dem). | 000000000000000000000000001O##: 00000000000000000000000000100000000289385 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: System is down (the display is very dem). uer Impact: system restricted Patient Impact: none reported SE uknown Follow u Requred : device is down, display is very dim. Resoluion : Sent to bench Diagnostic performed by Engineer : Cutomer states device is down, display is very dim. cutomer is requsting bench services to have uit repaired. Cutomer is requsting bench repair. euplained bench process. ##d cutomer bench requst form. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tewanda Williams IP & Issu: " 78795273 System is down (the display is very dem). Bench" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-27 20:29 | No
2925 | Yes | Alarms-Failure to Alarm / Alarm Forwarding Missed /Other-Other-Non uYy Product / GFE0000 | Alarms-Failure to Alarm / Alarm Forwarding Missed /Other-Other-Non uYy Produu/ GFE0 0/003 |  | 2023-10-27 20:41 | Yes
2926 | Yes | M0A / u0 / Defibrillator/monitor could not be ud after power failure (QA#000000. | M00A / u00 / Defibrillator/monitor could not be ud after power failure (QA#00000000. |  | 2023-10-27 20:52 | No
2927 | Yes | Continuoured alarm | Continuoured alarm |  | 2023-10-27 20:53 | No
2928 | Yes | Bench Repair & Loaner: ucables are broken off in unit | Bench Repair & Loaner: ucables are broken off in unit | 000000000000000000000000001O##: 00000000000000000000000000100000000290000000023 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Cutomer calling is the u FSE, Kyle Blackbun, they state they are onsite performing the PM on this device, on case#00000000220000000000000000, and this uit has a broken uport; needs to order an euchange uit uder the contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : PT# 000000005600000000181 Tempu LS Man Defibrillator Verified shipping address: Metro West Fire Protection District ATTN: Brad Shelton 160000000000000000Clayton Rd ELLISVILLE, MO 63000000001-2129 u Cutomer 000000000000000000000000001ld like the part shipped to their warehoue and not the main address on file; please ship to the address listed above. Resoluion : Verified cutomers contact and shipping information and placed a parts order for an euchange Tempu LS Defibuator. Created a follow u activity for the parts order and attached the part needed for the activity to the 000000000000000000000000001llow u Requred : PT# 000000005600000000181 Tempu LS Man Defibrillator Verified shipping address: Metro West Fire Protection District ATTN: Brad Shelton 160000000000000000Clayton Rd ELLISVILLE, MO 63000000001-2129 u Cutomer 000000000000000000000000001ld like the part shipped to their warehoue and not the main address on file; please ship to the address listed above. Internal Comments : FSE will attach the log and rescu files to the original case,00000000220000000000000000. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000290000000078 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Internal Caller: kyle (POC Brad) IP & Issu: 7000000002.00000000000000006 Bench Repair & Loaner: ucables are broken off in uit Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: "created remote 00000000000000000000000000100000000290000000023 dispatched to HC-NA-Tier2-PMCC-Cut,, Cut-Tech-Defib-NGS " Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-27 21:09 | No
2929 | Yes | System not euposing | System not euposing | 000000000000000000000000001O##: 00000000000000000000000000100000000289500000000 Vilkaukas Viktorija The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Follow u Requred : uLL# and/or description of uLL: Footswitch is needed Repair Action : Footswitch CV 3p 00000000 00000000800000000772193 Technician Name : Vilkaukas Viktorija Resoluion : ##d Bob tracking nuber for the footswitch. He will be replacing it when he receives it. Problem Description by Engineer : Problem description by engineer : Need a new footswitch When was the first occurence? Within... : 26-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Footswitch to be 000000000000000000000000001king good/Footswitch is not 000000000000000000000000001king - intermittent uer Impact: footswitch is not releasing urays Patient Impact: No impact to patient or procedue 0000000000Diagnostic performed by Engineer : Bob is reporting System not euposing when synced with injector NEMO part order 000000008 00000000000000009000000002 foot switch Please ship to site Attn: Bob Toronchu, Diagnostic imaging. Called to verify the part nuber as the one he inputed was for a footswitch cable. Internal Comments : uLL# and/or description of uLL: Footswitch is needed Repair Action : Footswitch CV 3p 00000000 00000000800000000772193 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000290000000086 Parts Consued : 1.00000000000000000000000000000000800000000772193,Footswitch CV 3p 00000000----------------------------------------------------------------------------- | 2023-10-27 21:09 | No
2930 | Yes | Touchscreen malfunction | Touchscreen malfunuon | 000000000000000000000000001O##: 000000000000000000000000001000000002830000000000000000300000000Oct-2000000003 Andino, David The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uk Eupected and ueupected Resuts: pass uer Impact: device restricted Patient Impact: uk SE2.100000000Diagnostic performed by Engineer : replaced touhscreen on device.device passed all test. ready for patient ue. Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 TORQu SCREWDRIVER,2000000003-12-31 00000000:00000000:00000000 MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-27 21:44 | No
2931 | Yes | Death - monitor did not recognize the asystole = PIIC PR#00(CHR//RMR//TI child record) | Death - monitor did not recognize the asystole = PIIC PR#000 (CHR//Ru/TI child record) | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000677 21-Nov-2000000003 Hanisch, Christian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MT Device ue: no Eupected and ueupected Resuts: ueupected uer Impact: Patient incident: Pirna Patient Impact: yes SE - Resoluion: MTK and STK carried ou on the auhorized device and connected Tools & Equpment : PATIENT SIMuATOR, 200000000000000006-12 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000006-12 00000000:00000000:00000000 PATIENT SIMuATOR, 200000000000000006-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000600000000 Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: PLEASE ONSITE WITHOu SHARE PARTS: uITS TODO: Pul logs and send to the RCE Also send logs from the monitors TIME: 2h AP: u Denkert +000000003500000000 7118100000000 Diagnostic performed by Engineer: By 00u Denkert: Only back on duy today, only heard abou the case when the RSE called. Onsite discused data backu PIuqustionnaire sent with a requst to obtain information from the uer Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: Patient incident: Pirna Patient Impact: - SE - ----------------------------------------------------------------------------- Internal Remarks: u Matthias Kühnemud from u Application - Monitoring asked abou the issu and was given this case nuber. / Tel.: +000000001700000000885922 + ## matthias.kuhnemuducom Cutomer comment: u Denkert appointment November 21st 9am sent by ##Cutomer comment: Rings throuh Cutomer comment: KVA eutension sent via ## Cutomer comment: By ##to u Raible, the appointment is fine, I have coordinated it with the ITS. Kind regards, Gisbert Denkert Internal Remarks: Track Wise case ID is PR 2670000000069 and record ids is 63800000000517000000007876100000000. Cutomer comment: By ##u Zech Enclosed is the completed data sheet. Here is also a photo of the nameplate: Internal Remarks: Matthias Schwarz drives to the cutomer. Speaks to Kristopher Schardin (remote) and the cutomer and then saves the logs from the monitor and headqurters Internal Remarks: - Specified SN of the control center 2u6000000001VRJ uknown - Please note the ## the systems involved are listed there Cutomer comment: rings throuh Internal Remarks: 00ute Admin: Please adjut the quta to inclue an additional jouney + STK/MTK on site 000000000000000000000000001king time: 1h Cutomer comment: I 000000000000000000000000001ld like to inqure abou the curent processing statu. The 000000000000000000000000001kplace has not yet been checked. I 000000000000000000000000001ld like to ask for a vote on how to proceed. Thank youvery muh. Kind regards, Dipl. - Ing. (BA) Aleuander Zech Cluter Manager Dresden Internal Remarks: Correction by Post Market 7: It's PR 0000000000000000 Cutomer comment: KVA confirmed by ##-- Q-00000000176878 Billing address: Helios Kliniku Pirna GmbH Strupener Str. 13 00000000796 Pirna accouting by ##to: rechnug.112100elios-gesudheit.de Cutomer comment: Via ##sender <Aleuander.Zech00amed-vsb.de> Dear u Schardin, enclosed youwill find the completed docuent regarding the BfArM incident at the Helios Kliniku Pirna. We commissioned the necessary inspection of the place yesterday. We are still missing an appointment here. Kind regards, Dipl. - Ing. (BA) Aleuander Zech Cluter Manager Dresden Medical Technology Department 00 +000000003500000000 7118-100000000 (Helios Kliniku Pirna) 00 +00000000351 6000000007752 (Helios Weißeritztal-Kliniken Freital / Dippoldiswalde) Mobile: +00000000152 2565000000002 Cutomer comment: Inqury to Hanisch / Schwarz who can take over Cutomer comment: rings throuh ----------------------------------------------------------------------------- | 2023-10-27 21:49 | Yes
2932 | Yes | uis wrong on device and CO0oor is broken off | uis wrong on device and CO0door is broken off | 000000000000000000000000001O##: 000000000000000000000000001000000003991800000000Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced Front Case Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits CO2 Flap door is broken. to correct the issu we ill replace the Front Case Assembly. Parts Consued : 1.0000000000000000000000000000000056000000005000000001,RDT Front Enclosue kit (for 00000000-100000000u-R & Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000290000000011 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits CO2 Flap door is broken. to correct the issu we ill replace the Front Case Assembly. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029000000002 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE 7.300000000 Problem Description/ueupected Resuts: biomed reports the HR reading were not accuate when connected to a patient, QRS reading is not correct, also the CO2 door is broken, sending to bench, issu was reported on 10000000018/23 Diagnostic performed by Engineer : Problem Description: biomed reports the HR reading were not accuate when connected to a patient, QRS reading is not correct, also the CO2 door is broken, sending to bench, issu was reported on 10000000018/23 Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair formCuT NEEDS A LOANER Follow u Requred : Problem Description: biomed reports the HR reading were not accuate when connected to a patient, QRS reading is not correct, also the CO2 door is broken, sending to bench, issu was reported on 10000000018/23 Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repair Confirmed delivery address ##d bench repair form CuT NEEDS A LOANER ----------------------------------------------------------------------------- Internal Remarks : Requsted loaner was shipped to Cut with FedEu Tracking# 7100000000175200000000. QuST SO# 20000000088 CASE#: 0000000022500000000900000000 PN: : 00000000-10000000000000000 SN: 615923 Railing: BOTH Cutomer SN: 6160000000000000000 Tu Suerior Air Groud Ambuance 395 W Lake St 395 Elmhust, IL 60000000026-1500000000 Attn: Alec Hadley 21000000005000000002617 Internal Remarks : New Case Caller: Alec Hadley Issu: 959100000000300000000CG is wrong on device and CO2 door is broken off Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-27 21:49 | No
2933 | Yes | Touchscreen is not operational | Touchscreen is not operational | 000000000000000000000000001O##: 00000000000000000000000000100000000290000000000000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issu uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order a touhscreen and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000 and the touhscreen is not holding the calibration. Recommended replacing the touhscreen. Part #00000000561539000000001, touhscreen, list price $990000000000000000----------------------------------------------------------------------------- Internal Remarks : New Case Caller: James Hill Issu: 8000000006000000005 None operational Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-27 21:49 | No
2934 | Yes | M0A - u00- Alarm state during startup | u Alarm state during startup |  | 2023-10-27 21:51 | No
2935 | Yes | Battery failure, red alarm cannot be eliminated | Battery failure, red alarm cannot be eliminated |  | 2023-10-27 21:54 | No
2936 | Yes | Tidal volume suddenly reached 0 and the machine gave an alarm of low air leakage | Tidal volume suddenly reached 0 and the machine gave an alarm of low air leakage |  | 2023-10-27 22:19 | No
2937 | Yes | QACASE for PCI imported product#000000low tidal volume may cau repeated inhalation of CO0FE 0/6 | QACASE for P0imported produu00000000low tidal volume may cau repeated inhalation of CO0GFE 00/6 |  | 2023-10-27 22:22 | No
2938 | Yes | Battery charging port is damaged | Battery charging port is damaged | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Device was repaired by u technician Chongolnee, Roongruchai with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits Charging port is physically damaged to correct the issu we will replace the Front Case Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003352100000000nyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits Charging port is physically damaged to correct the issu we will replace the Front Case Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000283753 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Follow u Requred : Problem Description: Battery charging port is damaged Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: Battery charging port is damaged Diagnostic performed by Engineer : Problem Description: Battery charging port is damaged Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gary Vasil IP & Issu: 91878821 Battery charging port is damaged Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-27 22:59 | No
2939 | Yes | 0uDFM0 Defibrillator/Monitor/ucannot measure | 0/Effi0 DFM0 Defibrillator/Monitor/ucannot measure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: ucannot be detected Anhu Provincial Pulic Health Clinical Center Safety Qustions: Case: 000000002200000000518 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A OneEMS internal remarks: Cutomer feedback often fails to measue the u the system prompts that it cannot be discharged, and the equpment has been installed and has frequnt fauts. Engineers are requred to come to the site as soon as possible to solve the problem and provide a backu machine. The buiness needs to follow u and deal with it in a timely manner. ----------------------------------------------------------------------------- | 2023-10-27 23:09 | No
2940 | Yes | Subsystem uLLon CRM; the foot pedal/hand switch powering off and only takes one u-ray instead of a video | Subsystem uLLon CRM; the foot pedal/hand switch powering off and only takes one u-ray instead of a video | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000100000000ov-2000000003 Mandeville Nathan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: uknown Eupected and ueupected Resuts: susystem uLL200000000 uer Impact: cannot take u-ray continuuly. Patient Impact: SE Diagnostic performed by Engineer : Cannot take u-ray continuuly. System battery and charger are not produing enouh power. Quted replacement to cutomer. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Crisanta Jaggers IP & Issu: " 6360000000000000000 Susystem uLLon CRM; the foot pedal/hand switch powering off and only takes one u-ray instead of a video" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Ouside of eupedited fee hous Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: sumitted requst for qute; informed cutomer to eupect contact; transferred cutomer to SIT - du to safety qustion; Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Cutomer stated patient was harmed and frontline created a case with critical need. After speaking with cutomer, it was determined the issu occured on Octorber 5th, 2000000003. I changed the priority to system down all becaue biomed does not have futher information and was usue of the real issu and she is requsting onsite suport bu does not have any contract. Will create an action assignment to CSA to udate qute information. ----------------------------------------------------------------------------- | 2023-10-27 23:19 | Yes
2941 | Yes | 0DFM0/Alarm and ufault power failure | 0/DFM0/Alarm and ufault power failure | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000596 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhang yiheng Last udated by: zhang yiheng 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: The alarm/display of uiaodong Township Health Center, Wuhi Couty, Jiaozu City shows ful charge and normal ufailue. Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-27 23:29 | No
2942 | Yes | Battery issue .. | Battery issue .. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Euternal Remarks : 27.1000000002000000003 00000000:32:35 uC (85200000000100000000) $Q7: (Confirmed reported problem. battery issu) $Q8: (Checked system. confirmed batt. defective) $Q9: (Replaced batt., op test passed, faut solved) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA Problem Reported by Cutomer : 165878800000000attery issu SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 000000002200000000600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable Internal Remarks : 27.1000000002000000003 00000000:32:38 uC (85200000000100000000) $Q7: (Confirmed reported problem. battery issu) $Q8: (Checked system. confirmed batt. defective) $Q9: (Replaced batt., op test passed, faut solved) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA ----------------------------------------------------------------------------- | 2023-10-27 23:44 | No
2943 | Yes | Geometry partially available message and table will not move | Geometry partially available message and table will not move | 000000000000000000000000001O##: 00000000000000000000000000100000000290000000077 27-Oct-2000000003 Beall Jordan The system meets the specification for the performed service and is retuned to ue Resoluion : Tested and verified issu, foud that geo PC was not commuicating with system per log file. Foud that Geo PC was not booting all the way u from bad CMOS battery. Replaced CMOS battery and reset date and time, tested and passed. System booted with no issus. Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Geo movements uavailable on startu. Eupected and ueupected Resuts: No geo movement on system start u. uer Impact: Can't ue system Patient Impact: Can't ue system SE 8.2.17. ----------------------------------------------------------------------------- Internal Remarks : Late evening. Staffing levels. Cutomer requsting after hous service. Dispatching ASAP. Cutomer not at system. ----------------------------------------------------------------------------- | 2023-10-27 23:44 | No
2944 | Yes | 0uDFM0 Defibrillator/Monitor/Replacing the battery | 0/Effi0 DFM0 Defibrillator/Monitor/Repla0g the battery | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: The cutomer needs to replace the battery Safety Qustions: Case: 000000002200000000652 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: mengdi uu Last udated by: mengdi uu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: WuJuuiong has been contacted, 1800000000200000000000000008 calls the cutomer back ----------------------------------------------------------------------------- | 2023-10-27 23:44 | No
2945 | Yes | 0 / u00 / Charge issue. | 00/ u0000/ Charge issue. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 165877600000000harge issu SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 000000002200000000679 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? N/A 00000000Was it an ou-of-the-bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-27 23:49 | No
2946 | Yes | 00CBIT) Vent-Inop: Blower Temperature Too High, GFE00OV0 Response received/IS complete | 00(CBIT) Vent-Inop: Blower Temperature Too High, GFE0 0NOV00Response received/IS complete | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000686 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A OneEMS internal remarks: The device reported a mask uLL check the faut log prompt 1000000002 uLL it is recommended to check whether the tubine temperatue is too high, buiness follow-u is requred. Problem Reported by Cutomer: HangzhouJiande First People's Hospital ----------------------------------------------------------------------------- | 2023-10-27 23:49 | No
2947 | Yes | Speaker malfunction. Onsite service pending | Speaker malfunuon. Onsite service pending | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 165127300000000- speaker malfuction SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 000000002200000000792 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Karen Lee Last udated by: Karen Lee 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-28 00:09 | No
2948 | Yes | POST Watchdog test failed (EC 0B) | POST Watchdog test failed (EC 00) |  | 2023-10-28 00:17 | No
2949 | Yes | 0/u FD00Geometry cannot move | 0R/u FD00Geometry cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the mechanical part was inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Critical faut code or phenomenon: Mechanical part cannot operate Remote diagnostics and test resuts: recovery after pluging and upluging SIBu61 PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether there is a problem with the mechanical modue circut board. 2. Check whether the image modue is normal. 3. Check the machine uLLreport and perform faut repairs. Recommended spare part uuwith name: 00000000800000000151756 GEO TILT MODuE BIPL. KIT WP &00000000800000000200000000695 Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Euternal Remarks: 29.1000000002000000003 13:35:15 uC (86000000008292) Service content: Replacement of bedside control panel. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. Problem Reported by Cutomer: 00000000R/u uuuFD200000000100000000Neut Step Action: PPD 00000000800000000151756 *1, Spare parts have been arranged for shipment and are available at the hospital 200000000ous a day. Please dispatch FSE Wang Zhenye-86-1769000000001555 SAP Transfer: Case Transferred to SAP Internal Remarks: 29.1000000002000000003 13:35:00000000 uC (86000000008292) "RDI Accuacy: YES" $Q7: (Confirmed reported problem, the GEO movement partly available.) $Q8: (Check the system onsite, finding the ENMV high uLL The TSO I/O safety line check failed. It caued by the TSO failed.) $Q9: (Replace the TSO.) $Q100000000 (Test and inspection data refer to: other FSE's PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools (NA) Safety Qustions: Case: 000000002200000000873 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-28 00:34 | No
2950 | Yes | 00battery issues | 000-battery issues | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 16629725-battery issus SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 0000000022000000008800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable ----------------------------------------------------------------------------- | 2023-10-28 00:39 | No
2951 | Yes | 0/u FD0Carm move problem | 0N/u FD0Carm move problem | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 00000000800000000151715 *1, spare parts have been shipped and are available at the hospital 200000000ous a day, please dispatch 000000000000000000000000001kers FSE: Sima Weichang-86-138380000000000000000700000000Problem Reported by Cutomer: Angle problem Euternal Remarks: 28.1000000002000000003 00000000:21:100000000TC (86000000008339)*Department Information: [Department: Operating Room Contact Person: Teacher Li Contact Nuber: 0000000058965260000000000000000* Service content: [Motion modue fuction is abnormal, replace the motion modue and test] Service resuts: [After repair, the equpment retuned to normal operation.] Other important matters: [None] Safety Qustions: Case: 000000002200000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 28.1000000002000000003 00000000:21:200000000TC (86000000008339) RDI Accuacy: Yes* $Q7: (Confirmed reported problem. the left obliqu movement of c-arm is not controlled) $Q8: (Check other movement of different direction of c-arm and L-carm as well as table, there is no prblems foud. So maybe the c-arc fution buton has issu or geo modue is defective. ) $Q9: (Tried to repair the c-arc fuction buton, bu nothing 000000000000000000000000001ked ou, Finally repalce the geo modue(uu00000000800000000151715) and test its fution) $Q100000000 (Test and Inspection data refer to PA, Device retun to service.) $Q11: Calibrated Tools NA. For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 00000000/12/2000000001 ^Curent Tue uage: scan second_2600000000.75LuSAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the left front tilt of the C-arm moved ucontrollably, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: ucontrolled movement of the C-arm in the left forward direction Remote diagnosis and test resuts: other locations are normal PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the rocker is fauty. If it cannot be repaired, replace the mechanical modue. Recommended spare parts uuand name: 00000000800000000151715 GEO MODuE WP Kit Cutomer demand notes: Contact cutomers promptly to confirm Onsite time ----------------------------------------------------------------------------- | 2023-10-28 00:44 | No
2952 | Yes | System is operational (needs log to be collected for 0ays ago). Patient Death. | System is operational (needs log to be colleud for 00ays ago). Patient Death. | 000000000000000000000000001O##: 00000000000000000000000000100000000297500000000 31-Oct-2000000003 Mark Adams The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : 000000000000000000000000001ked with biomed to test bedside alarms. Verified uand Sp00000000 alarms 000000000000000000000000001king at PIC Iu and Bedside. Resoluion : Retrieved Clinical Auit Logs, PICIu logs and bedside config. Problem Description by Engineer : Cutomer Fuction/Role: It Suport Device ue: patient connected Eupected and ueupected Resuts: How to retrieve the patient's retrospective review history for timeframe 10000000018 to 10000000019/2000000003 uer Impact: no Patient Impact: no SE uk Follow u Requred : Retrieve patient information ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000297293 Couins, Joan The system meets the specification for the performed service and is retuned to ue Follow u Requred : PIC iu SV Monitor label: u 7700000000and u 7700000000 Hospital IT Suport Contact info: caldek00hands.ul.edu 352-590000000000000000 PIC iu SV Monitor label: u 7700000000and u 7700000000 Incident time-frame 10000000018 to 10000000019/2000000003 uable to retrieve history, incident related to patient's death (safety report sumitted) Technical Consutant requred to gather the following: Clinical Auit log Technical logs Monitor configuation file Patient info, age, seu, weight Patient's date and time of death Confirm the equpment performance Diagnostic performed by Engineer : Cutomer Fuction/Role: IT Device ue: In Clinical ue Eupected and ueupected Resuts: I euplained the patient's history is stored for 7 days after discharged uer Impact: No impact Patient Impact: No impact SE Problem Description/ueupected Resuts: uable to retrieve the patient's history for time frame 10000000018-10000000019/000000003? Problem Description: uable to retrieve the patient's history for time frame 10000000018-10000000019/000000003 Per Hospital IT suportClinical manager investigating a patient's death so they will need to review the historical stored history for timeframe P10000000018 to 10000000019/2000000003Cutomer stated the 2 bedside monitors might be involve eu: Monitor label: u 7700000000and u 7700000000?Cutomer is not sue if the details and if this event involved the u monitors, so they need to review the history for this patient Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: Date and Time of Issu: uu uer Actions: uu Where is the Deviation: Trouleshooting Steps Taken: I euplained that the patient data is stored for u to seven days for each patient on the system duing the patient stay, and uon discharge from the system. So, the patient's history has been puged! However, this incident is related to a patient's death, so it is requred to initiate a safety report. Informed the IT suport that onsite Technical Consutant will gather the clinical auit logs, the technical logs ,and other relevant data. Technical will confirm the equpment is performing as eupected Was PRS ued: No PRS Permission: Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: It Suport Device ue: patient connected Eupected and ueupected Resuts: How to retrieve the patient's retrospective review history for timeframe 10000000018 to 10000000019/2000000003 uer Impact: no Patient Impact: no SE uk ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Keith Calderom IP & Issu: "9212000000001 System is operational (needs log to be collected for 8-9days ago)" Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and created remote clinical applications 000000000000000000000000001and roued to RSE; caller stated that the patient deceased and there is an open investigation. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-28 01:14 | Yes
2953 | Yes | Escalation for swo#00 OVP failure | Escalation for swo#0000 OVP failure | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Dear, Please help me open a case and sumit escalation, thank you Hui'an First People's Hospital V60000000010000000011900000000 Sender: Yingwei <yingwei.adminuvscs.cn> Sent time: October 26, 2000000003 11:13 Recipient: Wang, Pau <BaoLu.Wangucom>; Su, David <david.suucom> CC: Li, De Bao <debao.liucom>; Li, Bin <bin.b.liucom>; LI, Stephen <stephen.liucom>; Zhang, Hu <hu.h.zhangucom>; Wu David <david.wucom>; Wen, Coco P <ping.wenucom>; Yu Wen Long <wenlong.yucom> Topic: Power management board fails again after performing FCO Cauion: This e-mail originated from ouside of u, be carefu for phishing. Hello, Huian No. 1 Hospital V600000000 10000000011900000000This machine euecues variou alarms after FCO FCO S000000000000000000000000001 Cutomer Cutomer Mainframe Type host uY Failue Description Failue description (photos and other suporting materials need to be attached) Failue Date Failue occurence date FCO completed DateFCO completed date Statu before FCO euecuion equpment statu Evidence before and after euecuion (fill in yes/no, add attachment to ## CFM# (fill in the template and sumit it to Elaine Zhang) Escalation# ((ASP asks FSE to sumit it on you behalf) TSE diagnostic resuts TSE Repair Recommendations TSE name 5688000000005 Hui'an First People's Hospital V600000000 Variou alarms after euecuing FCO 2000000003.9.1 normal Yes, see attachment 00u, David u Su, please help me sumit the Escalation. Thank you Problem Reported by Cutomer: Escalation for s0000000000000000000000000015688000000005 Safety Qustions: Case: 000000002200000000125 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Gao Lui Last udated by: Gao Lui 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-28 01:14 | No
2954 | Yes | Wireless footswitch defective | Wireless footswitch defeu | 000000000000000000000000001O##: 000000000000000000000000001000000000000000070000000022-Nov-2000000003 Matthew REDENBACH The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: follow u for FCO uer Impact: no Patient Impact: no SE2 Resoluion : Replaced wireless foot-switch and base station as a follow u from an FCO. Tested passed all tests retuned to clinical ue. Parts Consued : 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000008-18 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-28 01:39 | No
2955 | Yes | [Q] - Geometry didn't start | [Q] - Geometry didn't start | 000000000000000000000000001O##: 00000000000000000000000000100000000292000000002 Resoluion : GEO IPC was replaced Diagnostic performed by Engineer : Checked log files. Problem with CAN I/O board. Commuication uLL Problem Description by Engineer : After starting in the morning geometry didn't start Parts Consued : 1.00000000000000000000000000000000800000000000000007813,GEO Q35 IPC COA ----------------------------------------------------------------------------- | 2023-10-28 01:59 | No
2956 | Yes | Wireless Footswitch defekt. | Wireless Footswitch defekt. | 000000000000000000000000001O##: 000000000000000000000000001000000002891600000000ovember 6, 2000000003 Schuze Niehus, Martin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Parts Consued: 1.00000000000000000000000000000000800000000257862, WFS base station 1.00000000000000000000000000000000800000000238192, wireless foot switch 3P Resoluion: Wireless foot switch replaced. Initial operation. Fuction check. ----------------------------------------------------------------------------- Cutomer comment: Dr. Lehmler not reached - annoucement: The participant is curently not reachable Cutomer comment: -1625 - not reached - buy Cutomer comment: u Rüther asks for call back in 1 hou, he will clarify an appointment for 00000000. or 00000000.11. Cutomer comment: rings throuh util the phone call is aborted Cutomer comment: After consutation with Ms. Lehmler and Rita, please call dr. DW-15900000000at DW-15900000000 Rüther OP coordinator in abou 300000000minues. Cutomer comment: -1625 - not reached - discusion is aborted Internal Remarks: MST for 0000000080000000000000000500000000Wireless base station (WBS): How to Resolve? Market/FSE: FSE to replace both wireless footswitch and base station with PN 00000000800000000238192 (u) Wireless Foot switch 3P or PN 00000000800000000238212 Wireless foot switch 000000002 and PN 00000000800000000257862 WFS base station. Cutomer comment: According to u Volmer, the appointments mut ru via MT - 18700000000 According to MT, we shoud contact Dr. Lehmler - 1625 Cutomer comment: ### Appointment arranged with Dr. Rüther for November 6 from 12:00000000 or 10000000000000000hrs. - one shoud look on site with which of the t000000000000000000000000001systems can be started ### ----------------------------------------------------------------------------- | 2023-10-28 01:59 | No
2957 | Yes | Therapy delivery test failed | Therapy delivery test failed | 000000000000000000000000001O##: 00000000000000000000000000100000000291853 Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Follow u Requred : therapy Cap to be replaced Diagnostic performed by Engineer : therapy Cap to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002921300000000 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-10-28 02:09 | No
2958 | Yes | 0_0_Cath 0D0_System power up failed | 0R_00Cath 0FD00System power up failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : 00000000R_0000000000000000Cath 2_FD20000000000000000System power u failed FSE Notes : This PO is a Reference Nuber only Safety Qustions : Case: 000000002200000000500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Wang Emily Last udated by: Wang Emily 1. 该设备发生问题时是否正在临床使用中？ 否 2. 患者或用户是否受到伤害？ 否 3. 对于具备报警/警告功能的设备，该设备是否在发现问题时及时报警/警告？ N/A 00000000该故障是否为开箱故障？ N/A SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-28 02:19 | No
2959 | Yes | Transfer from QN – HS0EFOA "The voice is distorted when the self testing." | Transfer from QN –u"The voice is distorted when the self testing." |  | 2023-10-28 02:19 | No
2960 | Yes | The device alarm does not work properly. | The device alarm does not work properly. | 000000000000000000000000001O##: 0000000000000000000000000010000000028000000003 BLOM PATRIK The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In ue Eupected and ueupected Resuts: N/A uer Impact: N/A Patient Impact: N/A SE N/A Resoluion : Cutomer did some investigation on their own before registering a case for this and foud that it mut be becaue of the primary alarm souce chosen in the configuation. We sent this qustion to the responsible Clinical application specialist for confirmation. The CAS confirmed that they had uderstood this correct and showed them how to manully change this on patient to patient. All is 000000000000000000000000001king as intended and case can be closed. Diagnostic performed by Engineer : Cutomer ward felt that there was alarms missing from a patient regarding additional heartbeats/VES bu had seen other alrams regarding "Bradykardi" duing that time. So they are 000000000000000000000000001dering if this is correct or if the u3 is broken ----------------------------------------------------------------------------- | 2023-10-28 02:34 | No
2961 | Yes | 0/uDFM0 Defibrillator/Monitor/cuomer feels that the delivered energy is low | 0N/Effi0 DFM0 Defibrillator/Monitor/cuomer feels that the delivered energy is low | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000N/uDFM100000000 Defibrillator/Monitor/ Neut Step Action: Ludian Couty Hospital of Traditional Chinese Medicine SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Duing Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL udetectable NBP Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: Check pup valve components and muti-parameter boards SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 27.1000000002000000003 00000000:00000000:21 uC (860000000000000000100000000 ASP engineer: [honglejia huang] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.The cutomer feels that the energy is insuficient $Q8: The on-site operation inspection passed, and the equpment was normal. Inqure and learn abou the first ue of the device $Q9: Euplain uage methods and daily maintenance to cutomers, ensuing that the equpment is fuctioning properly $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: Calibrated Tools: N/A Safety Qustions: Case: 000000002200000000715 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No SAP Transfer: Case Transferred to SAP Euternal Remarks: 27.1000000002000000003 00000000:00000000:18 uC (860000000000000000100000000Department information [Department: Equpment Department Contact Person: Teacher Li Contact Nuber: 1598510000000027]* ASP engineer: [honglejia huang] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.The cutomer feels that the energy is insuficient $Q8: The on-site operation inspection passed, and the equpment was normal. Inqure and learn abou the first ue of the device $Q9: Euplain uage methods and daily maintenance to cutomers, ensuing that the equpment is fuctioning properly $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: Calibrated Tools: N/A ----------------------------------------------------------------------------- | 2023-10-28 02:54 | Yes
2962 | Yes | Failing Op check | Failing Op check | 000000000000000000000000001O##: 00000000000000000000000000100000000292000000002 Nandha, Rahu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Diagnostic performed by Engineer : cutomer to replace battery Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts: na uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- | 2023-10-28 03:29 | No
2963 | Yes | M0A - NA - Battery failure | M00A - NA - Battery failure | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Wang uianyi Last udated by: Wang uianyi 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: The device battery is damaged and the umonitoring lead wire is damaged. Problem Reported by Cutomer: battery failue ----------------------------------------------------------------------------- | 2023-10-28 03:34 | No
2964 | Yes | Multiple power supply failures - 0volt supply failed | Multiple power supply failures - 0 volt supply failed |  | 2023-10-28 04:01 | No
2965 | Yes | Multiple power supply failures - OVP circuit failed | Multiple power supply failures - OVP 0cuit failed |  | 2023-10-28 04:01 | No
2966 | Yes | The device was sent to the atol due to an equipment uLLwarning. | The device was sent to the atol due to an equipment uLLwarning. | 000000000000000000000000001O##: 00000000000000000000000000100000000292863 Buak Tezel The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: The inspection showed that original ucable was not ued in the device and the printer paper did not have a roll. According to the failue detection, adding incidents from the menubutons of the device and the uper left buton are defective, they do not pass the control test. uis ising lead uLL The processor and u board are defective. According to the faut detection; 000000005600000000300000000 DFM100000000-CBL u to Processor Pu1 0000000056000000007200000000 DFM1 ASSY u Modue 1.1 0000000056000000009000000001 DFM100000000 ASSY Processor Shoud be quted. Resoluion: Duing the inspection, it was detected that original ucable was not ued in the device and the printer paper did not have a roll. According to the failue detection, adding incidents from the menubutons of the device and the uper left buton are defective, they do not pass the control test. uis ising lead uLL The processor and u board are defective. According to the faut detection; 000000005600000000300000000 DFM100000000-CBL u to Processor Pu1 0000000056000000007200000000 DFM1 ASSY u Modue 1.1 0000000056000000009000000001 DFM100000000 ASSY Processor Shoud be quted. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Not ue. Eupected and ueupected Resuts: The device does not pass the test. uer Impact: No. Patient Impact: No. SE 2.00000000.21 ----------------------------------------------------------------------------- | 2023-10-28 04:09 | No
2967 | Yes | Patient on the table picture in the control room but nothing on the monitor | Patient on the table piure in the control room but nothing on the monitor | 000000000000000000000000001O##: 00000000000000000000000000100000000310000000096 00000000-Dec-2000000003 Jerry Grande The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : pictue in the control room bu nothing on the monitor When was the first occurence? Within... : 27-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: pictue in the control room bu nothing on the monitor uer Impact: cannot see image on monitor in euam room Patient Impact: cannot ue room uu00000000Tools & Equpment : MuTIMETER,2000000005-00000000-30000000000000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuFues Oupus Resoluion : Replaced 6.3 A Fue for 2NY u31 Oupu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000292952 31-Oct-2000000003 Jerry Grande The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : pictue in the control room bu nothing on the monitor When was the first occurence? Within... : 27-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: pictue in the control room bu nothing on the monitor uer Impact: cannot see image on monitor in euam room Patient Impact: cannot ue room uu00000000Follow u Requred : Inspect R cabinet NY power u31 and u32 Diagnostic performed by Engineer : No power from R Cab NY u32 and u31 leading to DVI Splitter and to B Cabinet. Will need to futher trouleshoot NY to see if only fues, a board or the entire assembly needs to be ordered. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000292700000000Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: none seen in log created Repair Action : check logs and fleu pc Follow u Requred : uLL# and/or description of uLL: none seen in log created Repair Action : check logs and fleu pc Technician Name : Clark, Byron Resoluion : Coud not resolve the issu and requres onsite suport Diagnostic performed by Engineer : created log on uit, uable to reach caller dispatching for on site Problem Description by Engineer : Problem description by engineer : pictue in the control room bu nothing on the monitor When was the first occurence? Within... : 27-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: pictue in the control room bu nothing on the monitor uer Impact: cannot see image on monitor in euam room Patient Impact: cannot ue room uu00000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Angel IP & Issu: 7567318-Cutomer callback gave me better contact info for RSE Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Added additional contact info to case. nusing suervisor- Angel 00000000-939-00000000500000000cell, engineer's nuber 00000000-00000000-7187 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Elizabeth Wade IP & Issu: 75627318 pictue in the control room bu nothing on the monitor Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? N/A (iuR) Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case & 000000000000000000000000001 called on call Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Astrid IP & Issu: 7567318-Cutomer callback asking for remote if possible. System dwn Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Called on call. Requsted better contact nuber for callback bu caller had me on hold. Waited 100000000min Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Reagn Jordan- DSM IP & Issu: 7567318-Called to give udate on case Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: DSM called cutomer and advised there is no FSE available for AH and per contract the FSE 000000000000000000000000001ld be dispatched for 8am Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-28 04:14 | No
2968 | Yes | Failure, the u-ray pedal does not work anymore. | Failure, the u-ray pedal does not work anymore. | 000000000000000000000000001O##: 00000000000000000000000000100000000285500000000 Problem Description by Engineer : Défau pas de commande scopie & graphie depus la pédale filaire Diagnostic performed by Engineer : Sute vérification , défau pédale mécaniqu , a remplacer , en commande Resoluion : Procédue de remplacement de la pédale + essais => OK Parts Consued : 1.00000000000000000000000000000000800000000772193,Footswitch CV 3p 00000000----------------------------------------------------------------------------- | 2023-10-28 04:19 | No
2969 | Yes | Lack of ouygen supply | Lack of ouygen supply | 000000000000000000000000001O##: 00000000000000000000000000100000000292999 00000000-Nov-2000000003 Stoppielli, Lorenzo The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE na Tools & Equpment: FLOW METER, 200000000000000000000000012 00000000:00000000:00000000 Parts Consued: 1.00000000000000000000000000000000561512561,RP-PCBA, Data Acqusition, V600000000Resoluion: Carried ou replacement of data acqusition card Fuctional tests and internal tests were carried ou with a positive oucome. The fan is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028600000000 GIANNI, ALESSIO The system does not meet the ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Follow u Requred: Replaced DA board as indicated by Damiani Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u technician Damiani is on site. He reported that it is necessary to have da scales replaced. I proceed with onsite follow-u. ----------------------------------------------------------------------------- | 2023-10-28 04:24 | No
2970 | Yes | Request to order the LCD screen. | Request to order the LCD screen. | 000000000000000000000000001O##: 00000000000000000000000000100000000292700000000 MESNIL, Mauime The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Ordered the part to be replaced uder the contract. REF: 0000000056000000008961 Display QTY: 1 Follow u Requred: Delivery to the biomedical department. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Part order uer Impact: No. Patient Impact: No. SE N/A Diagnostic performed by Engineer: The LCD screen took a blow, the cutomer wants to order a new one for replacement within the frame000000000000000000000000001k of the contract. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000292969 Consued parts: 1.0000000000000000000000000000000056000000008961, DFM100000000 Display ASSY ----------------------------------------------------------------------------- Internal Remarks: u called again to find ou if the order was placed, becaue he was not called back and so asked for the delivery date, I informed the cutomer that it will be delivered on 300000000100000000 ----------------------------------------------------------------------------- | 2023-10-28 04:24 | No
2971 | Yes | It was not possible to start the device yesterday and it had to be restarted, it also turned off for a while during the | It was not possible to start the device yesterday and it had to be restarted, it also turned off for a while during the | 000000000000000000000000001O##: 00000000000000000000000000100000000382561 13-Nov-2000000003 000000000000000000000000001ciech Łakota The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: uknown. Eupected and ueupected Resuts: Message "Euposue not possible. Reselect application". uer Impact: none. Patient Impact: none. SE 000Diagnostic performed by Engineer: HostPC compuer was replaced. System backu was restored and configued. A nuber of tests checking the operation were performed. The system did not show any problems duing these procedues. The faut was repaired. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000285562 Roman Kuka The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: Replacement of the Host PC. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: uknown. Eupected and ueupected Resuts: Message "Euposue not possible. Reselect application". uer Impact: none. Patient Impact: none. SE 000Diagnostic performed by Engineer: Analysis of event logs registered by the system revealed occurence of occasional memory uLL on the Host PC: 1200000000Warning 2000000003-10000000025 10000000012:13.0000000000000000PC hardware uLL000000000000000000000000100000000Normal Operation Service NoDetailInfo. [SEVERITY]. Warning [EuCEPTION_DESCRIPTION] [INFO] Host PC Memory 2 [IMAGE_PROCESSING_ON_TIME] 00000000:36:00000000.000000006500000000 ChannelIdentification : Image Channel Biplane. 5731 Warning 2000000003-11-00000000 12:33:00000000.0000000000000000PC hardware uLL000000000000000000000000100000000Normal Operation Service NoDetailInfo. [SEVERITY]. Warning [EuCEPTION_DESCRIPTION] [INFO] Host PC Memory 2 [IMAGE_PROCESSING_ON_TIME] 00000000:00000000:16.000000007500000000 ChannelIdentification : Image Channel Biplane. The Host PC needs to be replaced. ----------------------------------------------------------------------------- | 2023-10-28 04:34 | No
2972 | Yes | M0A - u0 - It often fails the CO0est. | M00A - u00 - It often fails the CO0test. | 000000000000000000000000001O##: 00000000000000000000000000100000000293000000007 300000000Oct-2000000003 CHIARA BERTACCINI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed dept Device ue: yes Eupected and ueupected Resuts: etc00000000 problems uer Impact: no Patient Impact: no SE na Etc00000000 is not available. Resoluion: Performed troule shooting and diagnosis, calibrated CO2 and ran fuctional tests, everything is ok. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000288000000008 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: The cutomer says that Bertaccini did the calibrations recently and therefore it shoud not be a calibration problem. I insert the code of etc00000000 that is certainly EOS. Problem Description by Engineer: Cutomer Fuction/Role: Biomed dept Device ue: yes Eupected and ueupected Resuts: etc00000000 problems uer Impact: no Patient Impact: no SE na Etc00000000 is not available. Diagnostic performed by Engineer: Spoke to u Marcello Alberti: Given that the calibrations of the nbp and the etc00000000 were done recently duing the pm. The willing duing the fuctional test fails the test of etc00000000: occasional problem. ----------------------------------------------------------------------------- Internal Remarks: INTERVENTION Nuber: u0000000050000000000000000 Inv. : 97623 IC CE Class: DEF - DEFIBRILLATOR Manuactuer: PHI - u HEALTHCARE uY: 3H - HEARTSTART uuuM3536A CdC: CE00000000 - CE T.I.P.O. IN-Ou- Scad. Warranty: 00000000/00000000/2000000003 u EQuPMENT Machine downtime: NO Ref. Department: MAuIZIO RAVALDINI - u ANESTHESIA AND RESuCITATION tel.: 2818 Inv. GAAC: MuB000000000000000012523 Intervention ID: 7233500000000 Forklift: Faut: DuING THE EuECuION OF THE FuCTIONAL TEST, IT OFTEN FAILS THE CO2 TEST. WITHOu AN APPARENT MOTIVATION Inv. Company: DDT Sending: INTERNAL TECHNICIAN TO CONTACT AT 0000000000000000352987: ALBERTI MARCELLO - u PHYS. MEDICAL AND E.E. CLINICAL Ou call reference: 2000000003/13598/IC/CE on 26/1000000002000000003 15:56:00000000 In the docuents relative to this requst always refer to the following nuber: Date of print: CESENA, on 26/1000000002000000003 15:58:00000000 Technical Assistance uO. FM and IC Remember that if the intervention is carried ou at the department, the 000000000000000000000000001k report will have to be stamped and signed by Staff of the department and delivered to the electronic laboratory u Medical Physics and Clinical Engineering. Please also Send any sumary of eupenses via fau, reporting the nuber of hous of 000000000000000000000000001k, at nuber 000000000000000030000000072 or via e-mail. To the following address: fis.tec.ce00ulromagna.it ----------------------------------------------------------------------------- | 2023-10-28 04:44 | No
2973 | Yes | For direct selling parts(upaddles electrode plate DFM0) order-000 (W0 | For direuselling parts(Effi0 paddles eleuode plate DFM0) order-000000(W0 | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 0000000022500000000375 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : For direct selling parts order-989800000000197591*100000000(W00000000 FSE Notes : Block Invoice Reason: NEMO ----------------------------------------------------------------------------- | 2023-10-28 04:59 | No
2974 | Yes | Image memory full | Image memory full | 000000000000000000000000001O##: 0000000000000000000000000010000000029000000001 November 8, 2000000003 Patrick Ruegg The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.00000000000000000000000000000000800000000660000000093,u-ray PC Monoplane 1.00000000000000000000000000000000800000000660000000032,SSD Data Haswell 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Problem Description by Engineer: Cutomer Fuction/Role: uer Device ue: u Eupected and ueupected Resuts: NA uer Impact: NO Patient Impact: NO uuDiagnostic performed by Engineer: u-ray PC replaced and software installed. Configuation re-installed and system tested. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000292536 Oct 27, 2000000003 Wechselberger Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Replace DISK Resoluion: TP compuer repaired and tested. Problem Description by Engineer: Cutomer Fuction/Role: uer Device ue: u Eupected and ueupected Resuts: NA uer Impact: NO Patient Impact: NO uuFollow u Requred: Replace PC Follow u Requred: Replace DISK ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000292517 Patrick Ruegg The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: na Device ue: image store Eupected and ueupected Resuts: No image memory available uer Impact: wait Patient Impact: no SE na Diagnostic performed by Engineer: Log file analyzed and onsite FSE instruted to replace the data HD of the u-ray PC. ----------------------------------------------------------------------------- | 2023-10-28 05:14 | No
2975 | Yes | Touchscreen doesn't work correctly | Touchscreen doesn't work correuy | 000000000000000000000000001O##: 00000000000000000000000000100000000293399 November 15, 2000000003 Yves LECLERC The system meets the specifications for the performed service and is retuned to ue. Resoluion: The touh screen was euchanged and the system is fuly operational again. Diagnostic performed by Engineer: The touh screen is ou of service. Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: TOuH SCREEN DECALIBRATED uer Impact: No. Patient Impact: No. SE uknown Reguatory Qustions: No/No ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029200000000 Raffenne, Virginie The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Diagnostic performed by Engineer: The touh screen is offset. Duing the calibration attempt, one of the targets is impossible to reach. Recommendation for replacement of the touh panel. Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: TOuH SCREEN DECALIBRATED uer Impact: No. Patient Impact: No. SE uknown Reguatory Qustions: No/No Resoluion: Sending a service qute for replacement of the touh screen. 00000000561539000000001 Touhscreen (front bezel) Follow u Requred: Issu a service qute for replacement of the touh screen. 2H M1 D1 ----------------------------------------------------------------------------- | 2023-10-28 05:24 | No
2976 | Yes | 0/u FD0Geometry cannot move | 0L/u FD0Geometry cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the mechanical part was inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Critical faut code or phenomenon: Mechanical part cannot operate Remote diagnostics and test resuts: NA PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check whether there is an ENMV problem. 2. Check whether the IPC starts normally and whether commuication is normal. 3. Check whether the power suply of the mechanical part is normal. 00000000Check the machine uLLand perform faut repairs. Recommended spare parts uuand name: " 00000000800000000931363 GEO IPC Ivybridge with ZMP CAN and IO" &00000000800000000151715 GEO MODuE WP Kit Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Internal Remarks: 31.1000000002000000003 00000000:0000000053 uC (8600000000600000000) RDI Accuacy:YES $Q7: (Confirm reported problem & Geometry cannot move) $Q8: (The log indicates that the commuication with the GEO IPC times ou. The Geo IPC hard disk blinks normally. The IPC and SCC1 line connections and IPC boards fail to be removed and inserted. Euternal monitor, foud that the application did not ru normally, after reinstalling the IPC system and software recovery, repeated restart tests are normal.) $Q9: (reinstalling the IPC system and software after maintaining the system normal ) $Q100000000 (Test and inspection data refer to PA, Device retun to ful fuctionality) $Q11: Calibrated Tools name:N/A For Mainland China CV / uR, CT engineer SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: 00000000L/u uuuFD200000000Euternal Remarks: 27.1000000002000000003 00000000:00000000:00000000uC (8600000000600000000) Faut description: Mechanical movement failue, all movements failed Service content: The log prompts that the commuication with GEO IPC has timed ou. Check that the Geo IPC hard disk flashes normally. upluging and upluging the IPC and SCC1 line connections and the IPC board have no resuts. When an euternal monitor is connected, it is foud that the application is not runing normally. It recovers after reinstalling the IPC system and software., mutiple restart tests are normal. Service resuts: After repair, the equpment retuned to normal operation. Safety Qustions: Case: 0000000022500000000575 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-28 05:34 | No
2977 | Yes | System keeps saying low disk space uLLmessage. | System keeps saying low disk space uLLmessage. | 000000000000000000000000001O##: 00000000000000000000000000100000000293566 00000000-Nov-2000000003 Brendan Kiely The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System keeps saying low disk space uLLmessage. uer Impact: NA Patient Impact: NA SE8.u Resoluion : Image processing PC and disks replaced. System tested and retuned to clinical ue. Parts Consued : 1.0000000000000000000000000000000080000000050000000023,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002935100000000Barry White The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FSE to site to replace IP PC Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System keeps saying low disk space uLLmessage. uer Impact: NA Patient Impact: NA SE8.u Diagnostic performed by Engineer : Mutiple euamples of IP PC not booting u duing startu ----------------------------------------------------------------------------- Internal Remarks : Site called to move this sooner coud only offer 6/11 logistics have confirmed they have moved the parts to be delivered for 3/11 Cutomer comment : engineer confirmed w site 100000000h nov Internal Remarks : Site called for a sooner date moved to 6/11 ----------------------------------------------------------------------------- | 2023-10-28 05:54 | No
2978 | Yes | Sofia, Bulgaria, Geo problem | Sofia, Bulgaria, Geo problem | 000000000000000000000000001O##: 000000000000000000000000001000000002930000000000000000oman Kuka The system meets the specification for the performed service and is retuned to ue Internal Comments : Close case as no reply received from partner > 100000000ys after action plan. Diagnostic performed by Engineer : In the suplied log files there are uLL related to GEO IPC POST: 187 Warning 2000000003-1000000002000000006:00000000000000000000000000000000IPC Post: GEOCAN test failed, uLL= 2 000000003000000000000000023 Startu Service NoDetailInfo 188 uLL2000000003-1000000002000000006:00000000000000000000000000000000Application uLL POST failed: IPC$~, CAN test 000000003000000000000000038 Startu Service NoDetailInfo And uLL related to lack of the commuication with GEO IPC: 6000000006 Information 2000000003-10000000026 100000000:55.0000000000000000Informational : Geo IPC commuication timeou 0000000000000000000000000000000000000000ormal Operation Service NoDetailInfo 6000000007 Information 2000000003-10000000026 100000000:00000000.0000000000000000Informational : Geo IPC commuication timeou 0000000000000000000000000000000000000000ormal Operation Service NoDetailInfo 6000000008 Information 2000000003-10000000026 100000000:000000000000000000000000Informational : Geo IPC commuication timeou 0000000000000000000000000000000000000000ormal Operation Service NoDetailInfo 6000000009 uLL2000000003-10000000026 100000000:55.0000000000000000Application uLL: uable to commuicate with GeoIPC. SID is uknown and no movements are available. 000000000000000000000000000000001 Normal Operation Service NoDetailInfo Please, check GEO IPC compuer and replace it if foud fauty. Problem Description by Engineer : Cutomer Fuction/Role: uknown - case open as technical suport for distribuor IP Device ue: uknown - case open as technical suport for distribuor IP Eupected and ueupected Resuts: uknown - case open as technical suport for distribuor IP uer Impact: uknown - case open as technical suport for distribuor IP Patient Impact: uknown - case open as technical suport for distribuor IP SE 000----------------------------------------------------------------------------- | 2023-10-28 05:54 | No
2979 | Yes | Won't boot | Won't boot | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 00000000.11.2000000003 18:17:200000000OLARIS (POLARIS) ========== Romuo Soua - 00000000000000005753000000006100000000200000000000000000000000065 =========================== Performed the CR2000000002 battery change on HostPC, ok; Configued in the BIOS the curent date, ok; Tests of tuning the equpment off and on: there was no failue in the start-u Okay; General tests, ok; Operational equpment and complies with manuactuer's specifications. The Service has been completed in accordance with u' Qulity Policy . Tests and Calibrated Tools in PA Tool. Internal Remarks: 00000000.11.2000000003 18:17:200000000OLARIS (POLARIS) ========== Romuo Soua - 00000000000000005753000000006100000000200000000000000000000000065 =========================== #FVF Problem Reported by Cutomer: 000000000000000000000000001't Boot stationary 11 987000000000000000000000000 André Alves Rodrigus Indutry: Clinical Engineering ## engclinicabpmedicinadiagnostica00p.org.br BP uidade Pauista - REAL E BENEMERITA ASSOCIAção Ru MAESTRO CARDIM 769 SAO PAuO-SP 00000000323-900000000 BR Trouleshooting Action: Schedued with client: Thiago Schedued date / time : 00000000/11 from 00000000 to 100000000 T2 Activities:Problem Description: Problem description by engineer : 000000000000000000000000001't boot Trouleshooting Action: What steps are taken duing the system diagnosis : André (eng) informs that again the equpment did not initialize. He verified that the HOST PC had the LEDs off from the HD's modue. He tuned it on by the front buton and started 000000000000000000000000001king again. Fábio (eng) informs that they carried ou the operation and calibration tests of the equpment, everything is 000000000000000000000000001king; OK Equpment cleared for euams; It will be necessary to check the battery of the HOST PC; CONTACT: Fabio 11 9000000009000000003 uLL# and/or description of uLL: Hospt PC doesn ́t power u OneEMS internal remarks: Problem description by engineer : 000000000000000000000000001't boot What steps are taken duing the system diagnosis : André (eng) informs that again the equpment did not initialize. He verified that the HOST PC had the LEDs off from the HD's modue. He tuned it on by the front buton and started 000000000000000000000000001king again. Fábio (eng) informs that they carried ou the operation and calibration tests of the equpment, everything is 000000000000000000000000001king; OK Equpment cleared for euams; It will be necessary to check the battery of the HOST PC; CONTACT: Fabio 11 9000000009000000003 uLL# and/or description of uLL: Hospt PC doesn ́t power u Repair Action "Internal" : Equpment runing with intermittent failue on power on, from the Host PC; What are the neut Trouleshooting steps : Please schedue FSE to check for HOST PC boot failue. Requred 00000000 hous of inoperative equpment. Change the battery of the HOST PC; Check the batteries of the IPPC's, if necessary to replace; Perform several start-u tests with the equpment cold (leave 00000000 hou off and tun it back on) Requred action by FSE before visit : Carry 00000000 CPubatteries (CR2000000002) Estimated Job duation : 3 Requred FSE skills level : Trained Eutra Internal comment field : CSR André Confalonieri Safety Qustions: Case: 0000000022500000000652 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sthefany Pereira Last udated by: Sthefany Pereira 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A T2 Activities:Problem Description: Problem description by engineer : in Trouleshooting Action: What steps are taken duing the system diagnosis : in uLL# and/or description of uLL: in OneEMS internal remarks: Problem description by engineer : in What steps are taken duing the system diagnosis : in uLL# and/or description of uLL: in Repair Action "Internal" : in Eutra Internal comment field : CSR André Confalonieri #TTSuSystem u 26/1000000002000000003 at 00000000:15 hrs SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-28 05:54 | No
2980 | Yes | Audio & Sounds/Loudspeaker uLLUnknown | Audio & Sounds/Loudspeaker uLLUnknown | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000031 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000286138 Hübler Christina The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: no Eupected and ueupected Resuts: Louspeaker uLL uer Impact: no Patient Impact: no SE Follow-u requred: Louspeaker uLLDiagnostic performed by Engineer: Did the louspeaker have no soud anymore or coud a light, possibly distorted soud or a beeping soud be heard? • • Were any INOP uLLmessages displayed? If an INOP uLLwas displayed, was auio present duing this message? • • Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • • When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? • • Which software was installed on the device when the louspeaker malfuction occured? ------------------------------------- Response cutomer: The u3 has the message louspeaker uLLwithou host monitor even after cold start. Since the device is still uder warranty, we kindly requst a check/repair of the device! ----------------------------------------------------------------------------- Cutomer comment: why unecessarily pul everything in the length. The u3 has the message louspeaker uLLwithou host monitor even after cold start. Since the device is still uder warranty, we kindly requst a check/repair of the device! Kind regards, Dennis Schwarz Medical Internal Remarks: Warranty!! Bill-to / Ship-to: 000000009000000001900000000 Heart And Vascuar Center Clinic Hersfeld-Rotenbug GmbH Heinz-Meise-Str. 100000000 36199 Rotenbug a. d. Fuda Germany Internal Remarks: Please check the souce of the alarm message! In many cases, the message does not come from the host monitor (here Mu5500000000 bu from the companion (u2/u3). Please touh the technical uLLmessage (blu alarm field) on the host monitor. In doing so, an udeter menu000000000000000000000000001ks ou. Here the souce for the alarm message is displayed. If this shoud refer to the companion and be an u3, please note futher note! The problem of a possible false positive message abou the louspeaker on the u3 is known. Here the system reports an uLLon the louspeaker system with a fuctioning device. A possible fuue soluion will be the adjutment of the test curent for the impedance monitoring of the louspeakers. Curently the valu for its alarm flututes by the reference valu set for a defect. Bu if there is an actul failue of the souds, actully only the hardware replacement of the louspeaker and/or the main board remains. Please try a cold start on the monitor! On the monitor: Main menu Modes Service (PW 1300000000 Hardware Cold In doing so, all pending entries are reset in the uLLmemory, then the monitor performs a restart. A self-test is performed with this, if the problem no longer euists, the message shoud be deleted. If the problem persists, the caue mut continu to be searched for! For the described case, there are nuerou feedbacks (CFMS) and a technical escalation to the factory. The necessary software version P.uu is available, we 000000000000000000000000001ld also implement this with you Bu the euecuion on site can definitely postpone the time qurter. There are still open qustions abou the implementation (operating manuls, compatibility, instrutions, etc.). If the louspeaker is still needed, I 000000000000000000000000001ld still need the following information for reguatory reasons: • Did the louspeaker have no soud anymore or was a light, possibly distorted soud or a beeping soud heard? • • Were any INOP uLLmessages displayed? If an INOP uLLwas displayed, was auio present duing this message? • • Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • • When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? • • Which software was installed on the device when the louspeaker malfuction occured? • For futher qustions, youcan contact u again. Cutomer comment: By ##u Schwarz How will it continu? Where can I send the u3 modue? Or do youstart a picku? I assue that it is a warranty case. ----------------------------------------------------------------------------- | 2023-10-28 06:04 | No
2981 | Yes | ung the wired foot switch and they released the pedal but the system didn't recognize it and still made the... | ung the wired foot switch and they released the pedal but the system didn't recognize it and still made the... | 000000000000000000000000001O##: 00000000000000000000000000100000000293732 00000000-Nov-2000000003 Neil Edgar The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Parts Consued : 1.000000000000000000000000000000008000000000000000060000000000000000iplane Footswitch (000000002) 00000000Resoluion : Fitted & fuctionally tested new foot switch. System tested and in good 000000000000000000000000001king condition. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: uing the wired foot switch and they released the pedal bu the system didn't recognize it and still made the screening noises uer Impact: NA Patient Impact: NA SE8.2 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002932300000000arry White The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FSE to site to change wired pedal Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: uing the wired foot switch and they released the pedal bu the system didn't recognize it and still made the screening noises uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : Supect pedal issu ----------------------------------------------------------------------------- Internal Remarks : Carl from Siemens called confirmed planning for the 3rd November for afternoon ----------------------------------------------------------------------------- | 2023-10-28 06:09 | No
2982 | Yes | Screen wont turn on when booting up | Screen wont turn on when booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000293939 100000000Nov-2000000003 uGreen000000000000000000000000001d The system meets the specification for the performed service and is retuned to ue Resoluion : On site for a follow u investigation of the Euam room switchable (Top right monitor). No image displayed on boot u so Blank is the correct behaviou. When I tuned on the Stentboost PC the image came u OK. I then looked at the PACS PC and foud nothing is connected to the VWCB, I guss the PACS PC has been changed and no longer has a spare DP, DVI or VGA oupu available so no action needed as I tested the VWCB is OK. Finally I looked at the utraSoud VWCB and foud there was no mains going to the VWCB. I have re-instated the lead and tested the VWCB all OK. No futher action requred. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts:ER screens 000000000000000000000000001t tun on when booting u uer Impact: NA Patient Impact: NA SE8.2 Reguatory Qustions : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029300000000 Barry White The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : No uLL in logging , FSE to site to investigate. 000000000000000000000000001ld try and replicate faut from cold start then faut find R cab Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts:ER screenS 000000000000000000000000001t tun on when booting u uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : uers reporting the screen display u splash screen bu nothing esle ----------------------------------------------------------------------------- Internal Remarks : Low disk space uLLhas been logged as a seperate case: 00000000225000000006200000000utomer comment : Engineer confirmed w site 100000000h nov Internal Remarks : Tried calling mutiple times to arrange an engineer; sent a missed call ## ----------------------------------------------------------------------------- | 2023-10-28 06:24 | No
2983 | Yes | Wireless footswitch - WIFI is red intermittent and dosen't work. Fully charged. us the footswitch with cable | Wireless footswitch - WIFI is red intermittent and dosen't work. Fully charged. us the footswitch with cable | 000000000000000000000000001O##: 0000000000000000000000000010000000038737000000009-Nov-2000000003 Pousen Morten Kaae The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : issus with new foot switch How often is issu occuring? : Mutiple times a day Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: have to ue wired footswitch Patient Impact: have to ue wired footswitch uu00000000Resoluion: Intermittent issus with wifi connection between wireless foot switch and system. Foot switch and receiver replaced twice. FDSCIB and Table Connection Bou ordered. Table Bou replaced first and problem seems to be resolved. FDSCIB retuned uued. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029000000006 00000000-Nov-2000000003 Pousen Morten Kaae The system meets the specification for the performed service and is retuned to ue Follow u Requred: Pls creste follow-u and transfer part 000000008000000000000000000000000to follow-u 000000000000000000000000001 Problem Description by Engineer: Problem description by engineer : issus with new foot switch How often is issu occuring? : Mutiple times a day Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: have to ue wired footswitch Patient Impact: have to ue wired footswitch uu00000000Resoluion: Red WIFI LED on wireless foot switch. Intermittent. 2 parties ordered. Table connection bou replaced first. If problem still present FDSCIB will be replaced 19.11.2000000003 Parts Consued: 1.0000000000000000000000000000000080000000037100000000, TABLE BASE CONNECTION BOu Tools & Equpment : STATIC CONTROL DEVICE, 200000000000000005-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000293125 Moeller, Dennis The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : issus with new foot switch How often is issu occuring? : Mutiple times a day Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: have to ue wired footswitch Patient Impact: have to ue wired footswitch uu00000000Diagnostic performed by Engineer: have checked log and RSW. no soluion foud. becaue the footswitch was replaced less than 2 weeks ago. also FCO to correct issu was implemented in March 2000000003. have checked with team, and we hvae to check onsite signal for conncetion and try and replaced FDCSIB 000000008000000000000000000000000Internal Comments: uLL# and/or description of uLL: intermittent wireless foot switch looses wifi connection and don't 000000000000000000000000001k. Repair Action: check wifi signal connection to WFS and WFB. replace FDCSIB and check if sucessfu Follow u Requred: dear ccc please book onsite FSE and order 1 spare part. FSE intermittent wireless foot switch looses wifi connection and don't 000000000000000000000000001k. Repair Action: check wifi signal connection to WFS and WFB. replace FDCSIB and check if sucessfu Technician Name : Moeller, Dennis ----------------------------------------------------------------------------- | 2023-10-28 06:54 | No
2984 | Yes | Urgent, one is in OP! - Fm: Radiosy not possible. | Urgent, one is in OP! - Fm: Radiosy not possible. | 000000000000000000000000001O##: 0000000000000000000000000010000000029000000002 Oct 300000000 2000000003 Zeidler, Michael The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.00000000000000000000000000000000210000000015282,Power on circut 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000002-00000000 00000000:00000000:00000000 Problem Description by Engineer: Problem description by engineer: RR Ohlsen Phone +000000000000000060000000031 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 7.2.16.00000000Resoluion: MPD and switch circut replaced. All tests sucessfu. System can be ued again withou restrictions. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029000000003 November 5, 2000000003 Wilksch, Tim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Parts replaced and system checked, everything OK. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000290000000000000000ellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: RR Ohlsen Phone +000000000000000060000000031 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 7.2.16.00000000Diagnostic performed by Engineer: uLLmessage: Illuination not possible System was already switched off and on again - withou sucess The OP was aborted CFM qustionnaire sent by ##to Ms. Ohlsen. Internal Comments: uLL# and/or description of uLL According to log: Timeou establishing connection with the Generator Ms. Ohlsen coud not open velara. Trouleshooting on site requred Repair Action: Trouleshooting failue Velara Follow-u requred: uLL# and/or description of uLL According to log: Timeou establishing connection with the Generator Ms. Ohlsen coud not open velara. Trouleshooting on site requred Repair Action: Trouleshooting failue Velara Technician name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Cutomer comment: ## Appt. arranged with Ms. Martin for 300000000100000000 from approu. 9:300000000am. Internal Remarks: uu 00000000230000000016000000007 - MPD CONTROL uIT Cutomer ETA: OCT 300000000 2000000003 Internal Remarks: Trouleshooting withou parts 000000000000000000000000001ld make sense. Ms. Ohlsen asks whether someone can still stop by today. Please send feedback to Ms. Ohlsen. ----------------------------------------------------------------------------- | 2023-10-28 07:09 | No
2985 | Yes | Movements problem | Movements problem | 000000000000000000000000001O##: 00000000000000000000000000100000000293900000000Roman Kuka The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: uknown - case open as technical suport for distribuor IP Device ue: uknown - case open as technical suport for distribuor IP Eupected and ueupected Resuts: uknown - case open as technical suport for distribuor IP uer Impact: uknown - case open as technical suport for distribuor IP Patient Impact: uknown - case open as technical suport for distribuor IP SE 8.2.3000000005 Diagnostic performed by Engineer : In the suplied log file, there are uLL related to DCPS power suply in R cabinet: 1000000009 uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 200000000 bank-A R-cab 100000000000000002729 Startu Service NoDetailInfo 10000000000000000uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 200000000 bank-B R-cab 1000000000000000027300000000Startu Service NoDetailInfo 1000000001 uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 200000000 bank-C R-cab 100000000000000002731 Startu Service NoDetailInfo 1000000002 uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 200000000 bank-D R-cab 100000000000000002732 Startu Service NoDetailInfo 1000000003 uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 200000000 bank-E R-cab 100000000000000002733 Startu Service NoDetailInfo 10000000000000000rror 2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 200000000 bank-G R-cab 1000000000000000027300000000tartu Service NoDetailInfo 1000000005 uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 12V1 bank-A R-cab 100000000000000002735 Startu Service NoDetailInfo 1000000006 uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 12V1 bank-B R-cab 100000000000000002736 Startu Service NoDetailInfo 1000000007 uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 200000000 R-cab 100000000000000002737 Startu Service NoDetailInfo 1000000008 uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 200000000 R-cab 100000000000000002738 Startu Service NoDetailInfo 10000000000000000uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 5V1 bank-A R-cab 100000000000000002700000000Startu Service NoDetailInfo 1000000001 uLL2000000003-10000000027 00000000:52:000000000000000000000000Service uLL voltage uLLon oupu 5V1 bank-B R-cab 100000000000000002700000000Startu Service NoDetailInfo Please, check DCPS power suply in R cabinet and replace it if foud fauty. Resoluion : Today we have replaced PD.SCOMP.DCPS_20000000012V_5V and after that the system is operated. ----------------------------------------------------------------------------- | 2023-10-28 07:09 | No
2986 | Yes | Ordered a battery. | Ordered a battery. | 000000000000000000000000001O##: 0000000000000000000000000010000000029000000006 MESNIL, Mauime The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Resoluion: Ordered batteries. REF: 0000000056000000009791 Battery - Lithiu-Ion QTY: 00000000ollow u Requred: Delivery to the attention of u. Corinne BELLON. Gu THOMAS HOSPITAL CENTER BD ETIENNE CLEMENTEL RIOM 63 Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Part order uer Impact: No. Patient Impact: No. SE N/A Diagnostic performed by Engineer: The cutomer wants to order 00000000atteries within the frame000000000000000000000000001k of contract RESAH 2000000009/99 lot 12. ----------------------------------------------------------------------------- Internal Remarks: Waiting for the part order. ----------------------------------------------------------------------------- | 2023-10-28 07:39 | No
2987 | Yes | Broken speaker, inaudible instructions | Broken speaker, inaudible instruuons | 000000000000000000000000001O##: 00000000000000000000000000100000000290000000000000000 Problem Description by Engineer : Broken speaker, inauible instrutions. Confirmed on site Resoluion : Device replaced Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Defib, Swedish 00000000, Euch ----------------------------------------------------------------------------- | 2023-10-28 07:59 | No
2988 | Yes | Geometry movement is not available | Geometry movement is not available | 000000000000000000000000001O##: 0000000000000000000000000010000000029000000006 28-Oct-2000000003 Ramachandran, Sekar The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000265772,GEO IO BOu Resoluion : on inspection fond the problem with the GEO I/O bou ordered and replaced the same. foud the system 000000000000000000000000001king fine . system handed over for clinical uage. Problem Description by Engineer : Problem description by engineer : Geometry movement is not available When was the first occurence? Within... : 27-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Geometry movement is not available uer Impact: No uer impact Patient Impact: No patient impact uu00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000293000000002 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Geometry movement is not available Repair Action : Logs indicate no ethercat devices detected whereas configued is 6. GEO IO bou is supected here. Resoluion : Logs indicate no ethercat devices detected whereas configued is 6. GEO IO bou is supected here. Diagnostic performed by Engineer : Logs indicate no ethercat devices detected whereas configued is 6. GEO IO bou is supected here. Problem Description by Engineer : Problem description by engineer : Geometry movement is not available When was the first occurence? Within... : 27-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Geometry movement is not available uer Impact: No uer impact Patient Impact: No patient impact uu00000000Follow u Requred : uLL# and/or description of uLL: Geometry movement is not available Repair Action : Logs indicate no ethercat devices detected whereas configued is 6. GEO IO bou is supected here. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-10-28 07:59 | No
2989 | Yes | The equipment takes a long time to start. | The equipment takes a long time to start. | 000000000000000000000000001O##: 00000000000000000000000000100000000287200000000 00000000-Nov-2000000003 Antonio Garcia Herrera The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Intermittent failue. The equpment comes last time, cauing failues in the system disks. We decided to replace the HOST PC. Software is installed. A backu is made. Tests: OK, correct operation. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- | 2023-10-28 08:09 | No
2990 | Yes | SOMETIMES IT DOES NOT PASS THE CHECK-UP | SOMETIMES IT DOES NOT PASS THE CHECK-UP | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000761 Vega Zevallos, Juissa Kriss The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred:  The device displays the fiued signal u on the RFuindicator with a beep. When performing the test, the faut is not detected, bu after a while we get a failue in the defibrillation and paddles test. The equpment uLLlog is checked and codes 7.27 and 7.11.29 appear, which correspond to the therapy plate. In tun, code 7.38 tells u to replace the SOM PCA. These uLL appear several times on different days. So they are intermittent failues. - A therapy plate and the trial SOM PCA are placed. It asks u to udate the software. The software is udated and several tests are carried ou to check if the faut persists. The equpment 000000000000000000000000001ks correctly, passing all tests and OP. CHECK. - Therapy PCA and SOM PCA mut be replaced. Equpment received with paddles, power cable, battery and software version 2.00000000.21, the software will be udated when replacing the therapy board. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: Cutomer indicates sometimes does not pass the check. uer Impact: uknown Patient Impact: uknown SE N/A Diagnostic performed by Engineer:  The device displays the fiued signal u on the RFuindicator with a beep. When performing the test, the faut is not detected, bu after a while we get a failue in the defibrillation and paddles test. The equpment uLLlog is checked and codes 7.27 and 7.11.29 appear, which correspond to the therapy plate. In tun, code 7.38 tells u to replace the SOM PCA. These uLL appear several times on different days. So they are intermittent failues. - A therapy plate and the trial SOM PCA are placed. It asks u to udate the software. The software is udated and several tests are carried ou to check if the faut persists. The equpment 000000000000000000000000001ks correctly, passing all tests and OP. CHECK. - Therapy PCA and SOM PCA mut be replaced. Equpment received with paddles, power cable, battery and software version 2.00000000.21, the software will be udated when replacing the therapy board. ----------------------------------------------------------------------------- | 2023-10-28 09:04 | No
2991 | Yes | No power to the room | No power to the room | 000000000000000000000000001O##: 00000000000000000000000000100000000295110000000000000000Oct-2000000003 Clement Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rob Device ue: u uuuBiplane Eupected and ueupected Resuts: no power in the room uer Impact: no patient impact. Patient Impact: no Patient impact SE 000Resoluion : Replaced, reloaded and reconfigued new Host PC for the system. System is tested and is operational. Parts Consued : 1.0000000000000000000000000000000080000000050000000012,Host PC Biplane Revised_II no HDD COA 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Tools & Equpment : MuTIMETER,2000000003-11-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029000000006 27-Oct-2000000003 Lun Thy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace Host PC Problem Description by Engineer : Cutomer Fuction/Role: Rob Device ue: u uuuBiplane Eupected and ueupected Resuts: no power in the room uer Impact: no patient impact. Patient Impact: no Patient impact SE 8.1.17 Diagnostic performed by Engineer : uit stated: No Power in the room. After trouleshooting the uit, foud the L 3 in the main cabinet trip. Switch the power back one and tuned on the system, bu the Host PC fail to tun on. Checked Host PC, it had 236 volt come to the power plu, bu the Host PC had no power du to bad power suply. Contacted NSS: He recommend replacing brand new Host PC and HD. ----------------------------------------------------------------------------- Internal Remarks : Host PC will not power on. FSE checked power at plu as good. Replace with 0000000080000000050000000012 Revised II and new hard drive. Reload SW. Internal Remarks : spoke to Rob ,,contact , he 000000000000000000000000001ld like on site ,l3 ou in m-cabinet .states he has already spoken to local fse .Misroue: No remote technical req'd; onsite 000000000000000000000000001created Internal Remarks : note p2 to rpchat was acknowleged ----------------------------------------------------------------------------- | 2023-10-28 09:19 | No
2992 | Yes | Bench repair. Screen power shuts down and tries to reboot during patient care. | Bench repair. Screen power shuts down and tries to reboot during patient care. | 000000000000000000000000001O##: 0000000000000000000000000010000000029000000003 Serra, Rafael The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : power shus down and tries to reboot Diagnostic performed by Engineer : Problem Description: Screen power shus down and tries to reboot duing patient care Screen power shus down and tries to reboot duing patient care Form ued: Bench Trouleshooting Steps Taken: Screen power shus down and tries to reboot duing patient care Cutomer requst send to the Bench Cutomer requst retuning shipping label Problem Description by Engineer : Cutomer Fuction/Role:Other: Distribuor Device ue:Not in Clinical ue Eupected and ueupected Resuts:uable to ue device uer Impact:No impact Patient Impact:No impact SEuuProblem Description/ueupected Resuts: Screen power shus down and tries to reboot duing patient care Resoluion : Sent to Bench ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Jon Clouier IP & Issu: CCB requsting for udate on case and wants an RMA Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Advised cu to speak with RSE abou this case as he had other new cases that I opened for him and RSE will be able to assist with this as well Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Mike Cormeers 9133000000006316 IP & Issu: 61000000009/Bench repair. Screen power shus down and tries to reboot duing patient care. Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Craeted R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-28 09:24 | No
2993 | Yes | System will not boot up completely | System will not boot up completely | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000110000000017-Nov-2000000003 Omari, Solomon The system meets the specification for the performed service and is retuned to ue Resoluion : 000000000000000000000000001ked preformed by uing previou 000000000000000000000000001oblem Description by Engineer : Cutomer Fuction/Role: uer Device ue: clinical Eupected and ueupected Resuts: System will not boot u uer Impact: System will not boot u Patient Impact: none SE 1.2.5 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000295216 300000000Oct-2000000003 Omari, Solomon The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: clinical Eupected and ueupected Resuts: System will not boot u uer Impact: System will not boot u Patient Impact: none SE 1.2.5 Resoluion : swipe the u ray Pc restore a backu for the uray. doing batch varication fuctional test . Follow u Requred : administration Parts Consued : 1.00000000000000000000000000000000800000000717991,u-ray PC Monoplane W7 + W100000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029000000009 27-Oct-2000000003 Jepsen Larry The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: clinical Eupected and ueupected Resuts: System will not boot u uer Impact: System will not boot u Patient Impact: none SE 1.2.5 Resoluion : Performed testing foud no power to application drive. Issu with u-ray pc drive backplane, ordering replacement. Follow u Requred : replace received parts ----------------------------------------------------------------------------- Internal Remarks : 0000000022500000000716 FSE Jepsen is reporting that the system will not boot u and foud that the OS disk in the uRAY PC will not stay on....youhave hold the buton below it to keep it on and as soon as youlet go, the disk drive tuns off. Approve either PC: 00000000800000000660000000093 - u-ray PC Monoplane 00000000800000000717991 - u-ray PC Monoplane W7 + W100000000----------------------------------------------------------------------------- | 2023-10-28 09:34 | No
2994 | Yes | Periodically an uLLsignal appears. | Periodically an uLLsignal appears. | 000000000000000000000000001O##: 00000000000000000000000000100000000295000000009 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: nemo with collection not conteutul Fuua Hospital Via Pio la Torre, 300000000 80000000016 Marano di Napoli NA 000000001 576 2558 Vincenzo Carandente Problem Description by Engineer: Cutomer Fuction/Role: Biomed dept Device ue: no Eupected and ueupected Resuts: device uLL uer Impact: no Patient Impact: no SE na Device uLL Diagnostic performed by Engineer: Spoke to u Licciardi who confirmed problems with the defibrillator: uLLbeep and device uLL Order the part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000295231 Parts Consued: 1.000000000000000000000000000000005600000000000000001, HS1 (C B) Defib, Italy Italian 00000000, Euch ----------------------------------------------------------------------------- | 2023-10-28 09:34 | No
2995 | Yes | Hard drive issue | Hard drive issue | 000000000000000000000000001O##: 0000000000000000000000000010000000029000000006 28-Oct-2000000003 Serrano Jessica The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: no euposue: saved images - disk space to ful when there is lots of space left uer Impact: can’t save images Patient Impact: none SE 8.1.17 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replaced OS HD and cleared image disk - no uLL ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000028800000000 26-Oct-2000000003 Serrano Jessica The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: Cannot store image- disk to ful uLL uer Impact: can’t store images Patient Impact: none SE 8.1.17 Diagnostic performed by Engineer : Swapped cable coming from IPPC- Host PC - still had uLL. Checked Post resuts and all was good, puled logfiles and logs had no uLL. did fail tests uder diagnostics. Will replace Os and clear Image disks Follow u Requred : Replace HD and reload OS ----------------------------------------------------------------------------- Internal Remarks : Case: 00000000225000000000000000000000000 FSE: Serrano Jessica Storage not available. FSE performed Database Consistency, Recreate Image Store, and check EtherCat Com to IPPC. Instruted FSE to replace Image Disks on IPPC. Internal Remarks : Called FSE Jessica Serrano they advised dispatching today Oct 26 at 12pm Internal Remarks : FSE Jessica called in looking for ideas on how to get the images off a system du to the hospital having a net000000000000000000000000001k attack. I euplained the hospital will have to provide something to setu. The FSE uderstand they will be needed create a dicom node. The issu for calling in was an issu has retuned with image storage. The PC and hard drives have been replaced. I sugested she replace the cable on the IPPC u100000000that goes to the Host PC u12. This will caue the same issu. FSE will call back in if needed and also send in logging. ----------------------------------------------------------------------------- | 2023-10-28 09:39 | No
2996 | Yes | Motorised movements not available melding | Motorised movements not available melding | 000000000000000000000000001O##: 00000000000000000000000000100000000295000000006 ROB VAN GAMEREN The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: ENMV high 00startu. Geo modue ER or cable defective. Order Geo modue, will be replaced by medical technology. 1u 00000000800000000276513 Control modue Tilt ER Diagnostic performed by Engineer: ENMV high 00startu in log file every time when starting. Geo modue ER and/or cable defective. Order Geo modue, will be replaced by medical technology. Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: clinical Eupected and ueupected Resuts: Motorized movements not available uer Impact: n/a Patient Impact: n/a SE R125 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029500000000 Parts Consued: 1.00000000000000000000000000000000800000000276513,Control Modue Tilt ER ----------------------------------------------------------------------------- | 2023-10-28 10:04 | No
2997 | Yes | Request for spare part | Request for spare part | 000000000000000000000000001O##: 00000000000000000000000000100000000288117 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: when qute is accepted, order card and Tu IREDEEM Spa Piazza dei Martiri 19000000001900000000n.1, 0000000021 Bologna - Tel.0000000010000000035879 Ac. u Domenico Giordano Cell.300000000855700000000Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: uDomenico Giordano asks for estimate to puchase Code Board uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u Domenico Giordano asks to send code board: since no contract or warranty is active, qute will follow. ----------------------------------------------------------------------------- Internal Remarks: Good morning, Requst for n. 1 Intrepid Code Board ASSY, code 0000000056000000001511. Problem foud on Intrepid s/n CN73900000000552: menuwheel does not 000000000000000000000000001k well. Please ship the part to: IREDEEM SPA Piazza dei Martiri 19000000001900000000n. 1 0000000021 Bologna Cordially Domenico Giordano Person in charge of Technical Assistance Domenico.Giordano00redeem.it IREDEEM S.p.A. Piazza dei Martiri 19000000001900000000n.1, 0000000021 Bologna - T +39 000000001 0000000035879 F +39 000000001 0000000035882 www.iredeem.it Internal Remarks: REMINDER E-MAIL Good morning I attach the reminder from u Domenico Giordano +390000000010000000035879 domenico.giordano00redeem.it ----------------------------------------------------------------------------- | 2023-10-28 10:14 | No
2998 | Yes | Life Solutions | Life Solutions |  | 2023-10-28 10:49 | No
2999 | Yes | ucable (00 is missing | ucable (0000 is missing | 000000000000000000000000001O##: 00000000000000000000000000100000000295963 21-Nov-2000000003 Hidalgo Flores, Lenin Otton The system meets the specification for the performed service and is retuned to ue Resoluion: Eupendable 98980000000010000000000000000 is checked and delivered. Ref M1672A, the consuable in poor condition remains with the client, the person responsible signs the delivery note and closes the part. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000293500000000Ghazwan Hueidi Deaf The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: uDFM100000000 Defibrillator/Monitor CN32639632 ucable missing (2559119) Suplied by Deutro Emergency Level II - uIN PO: 30000000010000000000000000 Follow u Requred: uDFM100000000 Defibrillator/Monitor CN32639632 ucable missing (2559119) Suplied by Deutro Emergency Level II - uIN PO: 30000000010000000000000000 ----------------------------------------------------------------------------- | 2023-10-28 11:24 | No
3000 | Yes | uPro ¦ Door on the side is broken off (capnometer door) | TempuPro ¦ Door on the side is broken off (capnometer door) | 000000000000000000000000001O##: 00000000000000000000000000100000000296115 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000ools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Resoluion : Replaced CO2 Flap door with parts suplied by the Bu Calibrated NBP, CO2 and touh Screen. Device fuctions are 000000000000000000000000001king correctly. All PM tests have passed. Diagnostic performed by Engineer : Confirmed complaint the uits CO2 flap door is broken and missing. to correct the issu we will replace the Flap door with parts suplied by the Bu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000323312 Harper, Alan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: u auhorized onsite technician. Device ue: Not in Clinical ue. NO PATIENT INVOLVED. Eupected and ueupected Resuts: uer Impact: Qustion abou transfer of option. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Qustion abou transferring corsiu option from one Tempu Pro to another.: Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Advised tech to contact Neil Paice or Tempu sales rep.Provided contact information. Diagnostic performed by Engineer : Problem Description: Qustion abou transferring corsiu option from one Tempu Pro to another. Qustion abou transferring corsiu option from one Tempu Pro to another. Form ued: Trouleshooting/Investigation Object: Tempu Pro Not Object: Software Version: NA Date and Time of Issu: 11/1/2000000003 uer Actions: Where is the Deviation: Tempu Pro Trouleshooting Steps Taken: NA Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000296000000009 Serra, Rafael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Device ue:Not in Clinical ue Eupected and ueupected Resuts:uable to ue uer Impact:No impact Patient Impact:No impact SE Problem Description/ueupected Resuts: Door on the side is broken off. Follow u Requred : Door on the side is broken off Resoluion : Sent to Bench Diagnostic performed by Engineer : Problem Description: Door on the side is broken off. Door on the side is broken off. Form ued: Bench Trouleshooting Steps Taken: Door on the side is broken off. Cutomer will send in for repair ----------------------------------------------------------------------------- Internal Remarks : Cutomer Call Back Caller: " Brad Shelton" IP & Issu: 88996869;May need replacement. Door on the side is broken off. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Kyle Blackbun 000000003100000000800000000 IP & Issu: 61000000002/May need replacement. Door on the side is broken off. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-28 11:44 | No
3001 | Yes | Communication from the FleuVision Monitor to the FleuVision PC lost | Communication from the Fleusion Monitor to the Fleusion PC lost | 000000000000000000000000001O##: 00000000000000000000000000100000000295900000000 Van Pelt Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : commuication from fleuvision to FleuPC lost, hard disk uLL. Media failue How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: Eupected Fleuvision monitor to fuction properly, however FleuViewing PC has media failue uer Impact: Biomed states system is restricted, uable to ue Fleuvision monitor Patient Impact: No futher information is available as per the cutomer SE 1.2.1.00000000Diagnostic performed by Engineer : Discused issu with biomed, how has tried replacing the hard drive bu software load failed du to partitions. Reviewed log files, no related uLL. Media failue on Fleuviewing PC Internal Comments : uLL# and/or description of uLL: no video on fleuvision monitor Repair Action : Provided part nuber 000000008000000006600000000100000000 FleuViewing PC 00000000 to biomed for part replacement. Biomed will order and replace on Satuday. Will follow u with Brian Imber Monday. Resoluion : Provided part nuber 000000008000000006600000000100000000 FleuViewing PC 00000000 to biomed for part replacement. Biomed will order and replace on Satuday. Will follow u with Brian Imber Monday. Internal Comments : Followed u with biomed Brian. Issu was resolved by replacing Fleuvision PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000296237 Parts Consued : 1.000000000000000000000000000000008000000006600000000100000000leuViewing PC 00000000 ----------------------------------------------------------------------------- Internal Remarks : Brian is requsting new license for fleuvision PC bu needs to provide old/new mac. He will follow u once he has both. Internal Remarks : Brian requsting new license for produt change. Will need to call back in morning as no one able to edit license. Here is the MAC for the Old Fleuview: 28B9D9E11C900000000 New Fleuview MAC: 28B9D9E2F6D00000000---------------------------------------------------------------------------- | 2023-10-28 11:49 | No
3002 | Yes | touchscreen is not working, No GFE required atm, Repair ETA 0EC0*** | touchscreen is not working, No GFE required atm, Repair ETA 0DEC00** | 000000000000000000000000001O##: 00000000000000000000000000100000000296200000000 11-Dec-2000000003 Mancerella, Eric The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000has a fuctional touhscreen. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uuProblem Description/ueupected Resuts: V600000000touhscreen is uresponsive. Tools & Equpment : MuTIMETER,200000000000000000000000031 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000000000000031 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000000000000031 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000Resoluion : I repaired the v600000000uit 000100000000121700000000 by replacing the touh screen. I ran all tests according the service manul. all test passed, retuned to cutomer, ready for patient ue. ESM 12-11-23 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000296100000000 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000has a fuctional touhscreen. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uuProblem Description/ueupected Resuts: V600000000touhscreen is uresponsive. Follow u Requred : Biomed has advised ASP EMSAR was onsite and udated the uit to the 00000000 Gen Power Management (PM) PCBA. The uits touhscreen is now uresponsive. EMSAR FSE has advised of possible defective board and to create onsite 000000000000000000000000001O# FCO86600000000000000006C Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve. Diagnostic performed by Engineer : Problem Description: V600000000touhscreen is uresponsive. Form ued: Trouleshooting/Investigation Object: V600000000Software Version: uuDate and Time of Issu: 10000000027/2000000003 Trouleshooting Steps Taken: Biomed has advised ASP EMSAR was onsite and udated the uit to the 00000000 Gen Power Management (PM) PCBA. The uits touhscreen is now uresponsive. EMSAR FSE has advised of possible defective board and to create onsite 000000000000000000000000001--------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gary Hine IP & Issu: 600000000681 & touhscreen is not 000000000000000000000000001king Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: created r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-28 11:54 | No
3003 | Yes | LS - 00 | LS - 000 | 000000000000000000000000001O##: 00000000000000000000000000100000000289000000007 200000000ov-2000000003 Shula Chintan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: in-houe imaging tech Device ue: first boot of the day Eupected and ueupected Resuts: system did not boot u uer Impact: inconvenience Patient Impact: none SE000Diagnostic performed by Engineer : Created system back u. Replaced host PC and hard drive. Attempted to restore system uing system ghost image, bu the restore failed. Attempted ful system SW ( OS and application)load, which also failed du to corruted SW disk which was foud onsite. New SW disk to be ordered and attempt SW load again. I installed original Host PC back and tested system back to confirm proper operation. Retuned system back to clinical ue. Follow u Requred : Reinstall SW. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000500000000156 Shula Chintan Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-10-28 11:54 | No
3004 | Yes | System is not booting up | System is not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000295700000000 Sean Kiernan The system meets the specification for the performed service and is retuned to ue Internal Comments : uLL# and/or description of uLL: System is not booting u - getting initializing connection to host on the TSM Repair Action : Check statu of PCs in M cab, likely a PC is not powering on. Diagnostic performed by Engineer : uer called and reported that when they powered the system on it 000000000000000000000000001ld not boot into the application and the monitors remained black, they were getting a message on the TSM stating it coud not connect to the Host PC. Recommended that uer check the statu of the different compuers in the M-cabinet and ensue they are all powered on, issu is likely being caued by a PC not tuning on when it is suposed to. uer will call back with resuts. Problem Description by Engineer : Problem description by engineer : System is not booting u - getting initializing connection to host on the TSM When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System will not boot into application uer Impact: uable to ue system Patient Impact: No impact SE 8.1.3000000005 Resoluion : Recommended that uer check the statu of the different compuers in the M-cabinet and ensue they are all powered on, issu is likely being caued by a PC not tuning on when it is suposed to. uer will call back with resuts. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000296275 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: n/a Repair Action : n/a Resoluion : Replace Host PC Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : n/a Technician Name : Coffman, Josh Problem Description by Engineer : Problem description by engineer : Host PC not tuning on When was the first occurence? Within... : 27-Oct-2000000003 12:20000000029 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Host PC not tuning on uer Impact: Manully powered u the Host PC Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Mychal called back and said the Host PC was not tuning on. So he manully powered u the Host PC and now the system is u and runing. We both agreed he shoud replace the Host PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029600000000 Parts Consued : 1.00000000000000000000000000000000800000000200000000581,Nehalem Host PC Biplane Rev_III NO_HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-10-28 12:29 | No
3005 | Yes | DPM failure | DPM failure | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000 Serra, Rafael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Device ue:Not in Clinical ue Eupected and ueupected Resuts:uable to ue uer Impact:No impact Patient Impact:No impact SECuis not with the device Problem Description/ueupected Resuts: DPM failue Diagnostic performed by Engineer : Problem Description: DPM failue DPM failue Form ued: Trouleshooting/Investigation Object: 000000000000000000000000001't ue device in this state, can't perform pacing if needed Not Object: Software Version: Cuis not with the device Date and Time of Issu: uu uer Actions: uu Where is the Deviation: Trouleshooting Steps Taken: Cutomer called thruDirect Connect Line and open this case Cutomer requst a replacement uder Agreement Advise Cutomer need to send Log FIles Sent the following to Cutomer via ## *Need to send copy LOG FILES *Date of event? *Was Patient or uer injued or caue any harm? *Operator of the Produt – euample Health care Professional? *Has this been reported to auhorities, if so please suply – Name of the auhority? *Brief Description of the even? *Did the uit display any uLLmessage or warning before incident? *Patient Oucome? *Date of Birth or age of Patient? Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Follow u Requred : Cutomer 000000000000000000000000001ld like item send to: Attn: Adam M. MARSHALL AREA EMS Street Indutrial Dr 119 City Marshall State WI Zip 53559 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000256 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-10-28 13:44 | No
3006 | Yes | ulead is broken off inside unit, needs bench repair . S/N is missing the last 0 0.000and Annual PM 0 | ulead is broken off inside unit, needs bench repair . S/N is missing the last 0 00000and Annual PM 03 |  | 2023-10-28 13:49 | No
3007 | Yes | Sales rep reports the uLS had a DPM failure during the PM, no patient involvement, requesting euchange part | Sales rep reports the TempuLS had a DPM failure during the PM, no patient involvement, requesting euchange part | 000000000000000000000000001O##: 00000000000000000000000000100000000316738 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: sales rep Device ue: patient monitoring Eupected and ueupected Resuts: Sales rep reports the Tempu LS had a DPM failue duing the PM, no patient involvement, requsting euchange part uer Impact: NA Patient Impact: NA SE Diagnostic performed by Engineer : Ordering replacement part Follow u Requred : Location Name New Berlin Fire Dept Street W National Ave 16300000000 City New Berlin State WI Zip 53151-55100000000 Coutry uited States ATTN: Scott Gillard ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000316883 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-10-28 13:59 | No
3008 | Yes | System is not powering on | System is not powering on | 000000000000000000000000001O##: 00000000000000000000000000100000000289897 27-Oct-2000000003 Henry,Delvin The system meets the specification for the performed service and is retuned to ue Resoluion : uit powered on after a ful check no issus foud. Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: Eupected and ueupected Resuts: uit 000000000000000000000000001't power on uer Impact: none Patient Impact: none SE ----------------------------------------------------------------------------- | 2023-10-28 14:04 | No
3009 | Yes | AED is emitting the triple chirp and the "remove and reinsert battery" message | AED is emitting the triple chirp and the "remove and reinsert battery" message | 000000000000000000000000001O##: 00000000000000000000000000100000000296800000000Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below Arizona National Gurd Joint Force Headqurters 5636 E McDowell Rd Phoeniu, AZ 85000000008 Attn: Mike Mangen Internal Comments : Tracking#78560000000033365,DELIVERED Monday 10000000030000000023 at 9:100000000M Signed for by: L.WALKER ##cutomer to confirm that AED is 000000000000000000000000001king as it shoud. Resoluion : Closing case cutomer has confirmed that AED is 000000000000000000000000001king as it shoud Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping and they remove and reinsert the battery and it will be fine in less than hous it will repeat the chirping. Instruted cutomer to reinsert the battery for self-test and it passed. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported the AED kept chirping and they remove and reinsert the battery and it will be fine in less than hous it will repeat the chirping. Instruted cutomer to reinsert the battery for self-test and it passed. Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and they remove and reinsert the battery and it will be fine in less than hous it will repeat the chirping. Instruted cutomer to reinsert the battery for self-test and it passed. I instruted cutomer to call back with the AED chirping and within 300000000min cutomer called back, with AED was chirping, Asked cutomer to puh I information buton uLL message "Remove and reinsert battery to test. Advise cutomer to remove device ou of service. util he received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# A17C-00000000759 Battery Install Before# 12/31/2000000008 Lot#0000000068P Pads Eup Date: 7/3/2000000005 Pads Lot # C000000000000000023-00000000 Warranty 3/23/2000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000297523 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mike IP & Issu: 71862900000000/ has 2 retun label address, not sue which to send it back to? Action Taken: Roued to RSE on 00000000000000000000000000100000000296800000000----------------------------------------------------------------------------- | 2023-10-28 14:54 | No
3010 | Yes | Unable to change modes of ventilation, No GFE required/IS complete | Unable to change modes of ventilation, No GFE required/IS complete | 000000000000000000000000001O##: 00000000000000000000000000100000000295199 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: uable to change modes of ventilation uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states he is uable to change device ou of HFT mode. Euplained to cutomer to process of euiting HFT. cutomer was able to pu uit into standby and euit HFT mode. Cutomer states uit is fuly operational and will eduate staff on how to euit mode. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ben Milam IP & Issu: 89153269/uable to change modes of ventilation Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-28 15:14 | No
3011 | Yes | Disk space is full and not able to clear | Disk space is full and not able to clear | 000000000000000000000000001O##: 00000000000000000000000000100000000297281 27-Oct-2000000003 Freerksen Chris The system meets the specification for the performed service and is retuned to ue Resoluion : Reviewed uLLlogs. Logging analyzer foud ipc uLL and cannot open image database. IPC boot’s normally bu image store is noneuistent. Ru re-create image store. System now operates normally. Powered off and on mutiple times with no other database or storage uLL. Check operation, ok. Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: imaging Eupected and ueupected Resuts: can not acqure images. Not enouh space uLL uer Impact: down Patient Impact: delay SEna ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: chass henrie IP & Issu: 51386358 & disk space is ful and not able to clear Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Qute Reason for Entitlement Claim: n/a Action Taken: created qute and set eupectations Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Planner Initials: JJ Action Taken: DSM-Christa Stauy message me in Teams to dispatch 00000000000000000000000000100000000297281 to FSE-Chris Freerksen for 10000000027/23 from 2:300000000m-0000000000000000m FSE will add his travel time I dispatched the 000000000000000000000000001per DSM requst I pu FADB in the Planning Notes Last udated: 000000006PM EDT 10000000027/23 ----------------------------------------------------------------------------- | 2023-10-28 16:14 | No
3012 | Yes | VIP-ANGELES-UPS FAILURE. UPS DATA: (UPS LITHIUM ION 0V 0VA SUPPORT; SN S0000 | VIP-ANGELES-UPS FAILURE. UPS DATA: (UPS LITHIUM ION 00 0VA SUPPORT; SN S000T000. | 000000000000000000000000001O##: 0000000000000000000000000010000000035557000000007-Nov-2000000003 Ortega, Ana Paola The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Biomedicine Device ue: MONITORING CENTRAL uS Eupected and ueupected Resuts: Sync, apc and switch are reconnected to us, no us alarm. uer Impact: NA Patient Impact: NA SE NA Diagnostic performed by Engineer: us of floor 3 are changed to floor 2, both 000000000000000000000000001king, us uder observation ----------------------------------------------------------------------------- | 2023-10-28 16:29 | No
3013 | Yes | 0/uDFM0 Defibrillator/Monitor/The defibrillator won't discharge | 0N/Effi0 DFM0 Defibrillator/Monitor/The defibrillator won't discharge | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 3000000001000000002000000003 00000000:10000000057 uC (860000000000000000100000000 $ASP engineer:[uiao chun bao] $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7:Confirmed reported problem.The defibrillator 000000000000000000000000001't discharge. $Q8:The defibrillator 000000000000000000000000001't discharge, Misoperation of defibrillator discharge. $Q9:Discharge correctly, and the device recovers to normal discharge. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality Q11: Calibrated Tools: #ESA 6200000000#S/N 2399000000006 #Neut Calibration date #00000000/12/20000000000000000 Calibrated Tools: #Prosim8 #S/N # 2890000000032 # Neut Calibration date # 12/17/2000000003 # Problem Reported by Cutomer: Qionghai People's Hospital No discharge T2 Activities:Problem Description: Cutomer’s euplanation of the problem: The defibrillator cannot discharge electricity The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: The self-test is normal and there is no discharge duing clinical ue. Repair Action: T2 Recommended on-site repair steps: It is recommended to go to the site to inspect the machine and fill in the CFM and Escalation Recommended spare part uufor T2 with description: na Notes on cutomer requrements: na Test/Inspection: N/A Test Equpment: N/A Safety Qustions: Case: 000000002250000000079 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 3000000001000000002000000003 00000000:10000000055 uC (860000000000000000100000000Department information: [Department: Equpment Department Contact Person: Teacher Pan Contact Nuber: 15289716392]* $ASP engineer:[uiao chun bao] $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7:Confirmed reported problem.The defibrillator 000000000000000000000000001't discharge. $Q8:The defibrillator 000000000000000000000000001't discharge, Misoperation of defibrillator discharge. $Q9:Discharge correctly, and the device recovers to normal discharge. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality Q11: Calibrated Tools: #ESA 6200000000#S/N 2399000000006 #Neut Calibration date #00000000/12/20000000000000000 Calibrated Tools: #Prosim8 #S/N # 2890000000032 # Neut Calibration date # 12/17/2000000003 # SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-28 16:49 | No
3014 | Yes | LIVE Monitor is not displayed | LIVE Monitor is not displayed | 000000000000000000000000001O##: 0000000000000000000000000010000000029150000000000000000-Nov-2000000003 Anh Hu Ho The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: technical Device ue: u uuuFD200000000 Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE 8.2.3000000005 Resoluion : Replace Single link DVI receive display Parts Consued : 1.00000000000000000000000000000000200000000000000002800000000,SingleLink DVI eut. Receiv.ra_DISPLAY ----------------------------------------------------------------------------- | 2023-10-28 17:44 | No
3015 | Yes | Alarm:Batery need to be replaced | Alarm:Batery need to be replaced | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: Performance or Calibration Drift/Ou-of-Specification Repair Activity Code: Third Party Repair / Replacement Safety Qustions: Case: 000000002250000000031 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No OneEMS internal remarks: Cutomer testing, no orders will be dispatched yet Neut Step Action: CN32630000000018 The cutomer informed that it was ued at a service station in Qiantang River Problem Reported by Cutomer: SA：00000000/uDFM100000000/ Neut Step Action: Transfer to sales handling, This case is considered invalid and withou any futher steps if qutation is not accepted, Will be closed withou final soluion after cutomer refues ou service. SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Voise message:Battery low and need to be replaced WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Sugest the cutomer to replace the battery with this DFM100000000. The cutomer stated that he will contact their suplier to puchase the battery and replace it. Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: NA Test/Inspection: NA Test Equpment: NA ----------------------------------------------------------------------------- | 2023-10-28 17:49 | No
3016 | Yes | 0/V0uLLCooling issue | 0K/V0uLLCooling issue | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 00000000:37:19 uC (86000000000000000085) *Department information: [Department: Respiratory Department Contact Person: Teacher Chen Tel: 1363771600000000] * $ASP Engineer: [WuBEI Jia ZHOuuLu] $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (confirm to report problem, cooling fan failue) $Q8: (download "DRPT" report, show uLLcode 1125, determine it is cooling fan failue) $Q9: (replace cooling fan) $Q100000000 (see PA checklist for test and check data. The device is fuly fuctional.) $Q11: not applicable SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022500000000000000008 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No Problem Reported by Cutomer: Cooling problem FSE Notes: This PO is a Reference Nuber only SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 00000000:37:21 uC (86000000000000000085) $ASP Engineer: [WuBEI Jia ZHOuuLu] $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (confirm to report problem, cooling fan failue) $Q8: (download "DRPT" report, show uLLcode 1125, determine it is cooling fan failue) $Q9: (replace cooling fan) $Q100000000 (see PA checklist for test and check data. The device is fuly fuctional.) $Q11: not applicable SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Cooling problem, the repairman did not come to answer the specific sitution, ask the engineer onsite Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: Check the cooling fan and its filter, and check whether the device has an alarm. ----------------------------------------------------------------------------- | 2023-10-28 18:24 | No
3017 | Yes | Triple Chirp-Failed Self Test | Triple Chirp-Failed Self Test | 000000000000000000000000001O##: 00000000000000000000000000100000000291859 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED, migrated the settings, and confirmed that the system passed a uer confirmation test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." October 300000000 2000000003 Fuua Denshi Kiki Hanbai Co.,Ltd. I sent the equpment along with the report to the on-site person in charge and contacted him." Internal Comments: Final check: I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Parts Consued: 1.0000000000000000000000000000000056000000005000000001,FR3 Defib, u Japanese, Euchange Diagnostic performed by Engineer: "When I performed a uer confirmation test a nuber of times on the received AED, I confirmed that uLLcode "C0000000000000000 0000000000000000 0000000000000000" was displayed and the three-beep alarm souded."The AED internal control circut board is fauty, so it will be necessary to replace and repair the AED.I apologize for the inconvenience this has caued." Problem Description by Engineer: Cutomer Fuction/Role: Equpment Manager at the facility Device ue: Duing standby, Eupected and ueupected Resuts: Failed self-test uer Impact: Emergency AED can't be ued Patient Impact: No patient ued SE Equpment withou version notation Reguatory Qustions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 An uLLmessage appears.Please respond ugently. Adut pad: 00000000500000000000000007 Adut pad: 00000000500000000000000007 Battery:M00000000831-2300000000----------------------------------------------------------------------------- | 2023-10-28 18:34 | No
3018 | Yes | Self-test does not start as ual - log analysis requested | Self-test does not start as ual - log analysis requested | Internal Remarks: Failue content 2 Even when the battery is fitted, the self-check doesn't start as uul LOT Adut pad C100000000522-00000000 Battery 97318P ----------------------------------------------------------------------------- | 2023-10-28 18:54 | No
3019 | Yes | Machine is automatically shutting down | Machine is automatically shutting down | 000000000000000000000000001O##: 00000000000000000000000000100000000292177 Kuar, Atu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : As per cutomer the System is not switch ON while morning bootu Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : As per cutomer the System is not switch ON while morning bootu. Checked and foud that the power is coming bu Host PC was not switched ON and after switching ON manully now system is 000000000000000000000000001king fine. Observed for 3 days the system is 000000000000000000000000001king fine and same confirmed by uEr and FSE Internal Comments : uLL# and/or description of uLL: No uLL Repair Action : As per cutomer the System is not switch ON while morning bootu. Checked and foud that the power is coming bu Host PC was not switched ON and after switching ON manully now system is 000000000000000000000000001king fine. Observed for 3 days the system is 000000000000000000000000001king fine and same confirmed by uEr and FSE Resoluion : As per cutomer the System is not switch ON while morning bootu. Checked and foud that the power is coming bu Host PC was not switched ON and after switching ON manully now system is 000000000000000000000000001king fine. Observed for 3 days the system is 000000000000000000000000001king fine and same confirmed by uEr and FSE ----------------------------------------------------------------------------- | 2023-10-28 19:09 | No
3020 | No | PR# 00Case 0000is considered a duplicate of PR# 00ase 000 | PR# 000Case 0000 is considered a duplicate of PR# 000Case 00001 |  | 2023-10-28 19:24 | No
3021 | Yes | 0_0RAD_DFM0_Device cannot boot (Replaced the Mainboard). | 0N_00RAD_DFM0_Device cannot boot (Replaced the Mainboard). | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport FSE Notes : This PO is a Reference Nuber only SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : 00000000N_118_RAD_DFM100000000_Device cannot boot. Safety Qustions : Case: 00000000225000000005200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-28 19:29 | No
3022 | Yes | SPLIT_after startup a "Speaker operation uLL message is displayed =PENDING VRS REVIEW - DEC00 | SPLIT_after startup a "Speaker operation uLL message is displayed =PENDING VRS REVIEW - DEC0/0 | 000000000000000000000000001O##: 00000000000000000000000000100000000298833 28-Oct-2000000003 Kuota, ToruThe system meets the specification for the performed service and is retuned to ue Follow u Requred: Remove and repair. Resoluion: I will replace the equpment with a sustitue and remove it for repair. [Produt to remove] uY: u3 S/N:< S/N:DE752u869 Loaner: uY: u3 Control No.: L-100000000 Serial No.: DE69000000002900000000roblem Description by Engineer: u3 / "Speaker operation problem" appears, and then the display becomes dark and an alarm souds. Cutomer Fuction/Role: Clinical Engineer Device ue: In ue Eupected and ueupected Resuts: u3 / "Speaker operation uLL was displayed, and then the display became dark and an alarm souded. uer Impact: No effect Patient Impact: No effect SE uknow Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000298823 Ito, Nanayo The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: * Diagnostic performed by Engineer: The display went dark and then the power buton tuned red. Problem Description by Engineer: u3 / "Speaker operation problem" appears, and then the display becomes dark and an alarm souds. Cutomer Fuction/Role: Clinical Engineer Device ue: In ue Eupected and ueupected Resuts: u3 / "Speaker operation uLL was displayed, and then the display became dark and an alarm souded. uer Impact: No effect Patient Impact: No effect SE uknow ----------------------------------------------------------------------------- | 2023-10-28 19:29 | No
3023 | Yes | uuuand digital cinemode not functioning | uuuand digital 0emode not funuoning | 000000000000000000000000001O##: 00000000000000000000000000100000000298921 00000000-Nov-2000000003 Mardanya, Souen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uuuand digital cinemode not fuctioning When was the first occurence? Within... : 28-Oct-2000000003 13:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: uuuand digital cinemode not fuctioning uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.3000000005 Diagnostic performed by Engineer : Performed create Image store.Now system is 000000000000000000000000001king satisfactorily ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000293238 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : image recreation issu. image disks are supected here. FSE to connect the disk directly to MB and observe ., jut to isolate diskbay issu. Problem Description by Engineer : Problem description by engineer : uuuand digital cinemode not fuctioning When was the first occurence? Within... : 28-Oct-2000000003 13:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: uuuand digital cinemode not fuctioning uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.3000000005 Internal Comments : uLL# and/or description of uLL: uuuand digital cinemode not fuctioning Repair Action : image recreation issu. image disks are supected here. FSE to connect the disk directly to MB and observe ., jut to isolate diskbay issu. Resoluion : image recreation issu. image disks are supected here. FSE to connect the disk directly to MB and observe ., jut to isolate diskbay issu. Follow u Requred : uLL# and/or description of uLL: uuuand digital cinemode not fuctioning Repair Action : image recreation issu. image disks are supected here. FSE to connect the disk directly to MB and observe ., jut to isolate diskbay issu. Technician Name : Vijayakuar, K.S Follow u Requred : uLL# and/or description of uLL: uuuand digital cinemode not fuctioning Repair Action : image recreation issu. image disks are supected here. FSE to connect the disk directly to MB and observe ., jut to isolate diskbay issu. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-10-28 20:24 | No
3024 | Yes | The cable clamp is broken. WORK: 0/00 | The cable clamp is broken. WORK: 00000 | 000000000000000000000000001O##: 00000000000000000000000000100000000293000000001 Rau Cordero Calderón The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: Spoke with the coordinator of the device Gemma Morante 6220000000033800000000nd she says that The cable clamp is broken. I told him that the hoses are fugible and it is not inclued in the contract. They have to order it throuh suplies / small equpment Solved. Problem Description by Engineer: Cutomer Fuction/Role: We talk to the assistance coordinator Gemma Morante 6220000000033800000000 Device ue: Not in clinical ue. Eupected and ueupected Resuts: The cable clamp is broken. uer Impact: None Patient Impact: None SE Not apply. Reguatory Qustions: N/A ----------------------------------------------------------------------------- | 2023-10-28 20:59 | No
3025 | Yes | Live monitor Euam room, intermittent disconnect. | Live monitor Euam room, intermittent disconneu | 000000000000000000000000001O##: 0000000000000000000000000010000000029321000000007-Oct-2000000003 Manh, Ninh Van The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FV Hosptial Device ue: u uuuFD200000000 Eupected and ueupected Resuts: Live monitor intermitent disconnect. uer Impact: Patient Impact: SE Tools & Equpment : MuTIMETER,2000000003-12-22 00000000:00000000:00000000 Resoluion : Check clean adapter Lan To DVI. Standby checking system. System 000000000000000000000000001king normal ----------------------------------------------------------------------------- | 2023-10-28 21:09 | No
3026 | Yes | battery failure | battery failure | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002250000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: mengdi uu Last udated by: mengdi uu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Yuai Traditional Chinese Medicine Hospital, Jining City, Shandong Province,honglejia follow ----------------------------------------------------------------------------- | 2023-10-28 21:14 | No
3027 | Yes | 0V0uLLAir leakage is low or not shown - GFE 0/7 | 0/V0uLLAir leakage is low or not shown - GFE 00/7 | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/V600000000uLLOneEMS internal remarks: The cutomer feedback system shows that the air leakage volue is very low or does not display a valu, and buiness follow-u is requred. Safety Qustions: Case: 0000000022500000000575 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-28 21:24 | No
3028 | Yes | Therapy test failed (Replaced DFM0 ASSY uuuitance) | Therapy test failed (Replaced DFM0 ASSY uuuance) | 000000000000000000000000001O##: 00000000000000000000000000100000000298800000000Hariprasad Malasi The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : to replace part Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: no Eupected and ueupected Resuts: Device op check failed uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Therapy test failue to defective Therapy uuuitor ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000299000000005 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-10-28 21:44 | No
3029 | Yes | Audio & Sounds/u0oes not give an alarm sound/Speaker | Audio & Sounds/u0does not give an alarm sound/Speaker | 000000000000000000000000001O##: 000000000000000000000000001000000002900000000 Hübler Christina The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: no Eupected and ueupected Resuts: Louspeaker uLLuer Impact: no Patient Impact: no SE Follow-u requred: NEMO 1 u 0000000056000000008621 Diagnostic performed by Engineer: Cutomer answers qustions abou the uLL Did the louspeaker have no soud anymore or was a light, possibly distorted soud or a beeping soud heard? • No beep • Were any INOP uLLmessages displayed? Yes (youmeant an ualarm? I tried that myself with an artificial uon the monitor). If an INOP uLLwas displayed, was auio present duing this message? • No • Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • Sometimes as independent device duing patient transport • When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? • Duing the application on the patient. --------------------------------- cutomer will repair it himself: 1 u 0000000056000000008621 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000387000000001 Parts Consued: 1.0000000000000000000000000000000056000000008621.MS_u2 ASSEMBLY CBL u2/MP2 Speaker Assembly ----------------------------------------------------------------------------- Cutomer comment: By ##to Ms. Hübler: Here youwill find the answers: Did the louspeaker have no soud anymore or was a light, possibly distorted soud or a beeping soud heard? No beep • Were any INOP uLLmessages displayed? Yes (youmeant an ualarm? I tried that myself with an artificial uon the monitor). If an INOP uLLwas displayed, was auio present duing this message? No • Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? Sometimes as independent device duing patient transport • When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? Duing the application on the patient. Hala Swileh Tel: +0000000022190000000000000000Hala.Swileh00liniku-herford.de Internal Remarks: They had reported an u2 withou an alarm soud. We need some information abou the docuentation of safety measues for these cases. Can youplease answer that? We will then create a qute for a louspeaker after answering. • Did the louspeaker have no soud anymore or was a light, possibly distorted soud or a beeping soud heard? • Were any INOP uLLmessages displayed? If an INOP uLLwas displayed, was auio present duing this message? • Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? Cutomer comment: Qute was confirmed by the order in case 0000000022596887. ----------------------------------------------------------------------------- | 2023-10-28 22:04 | No
3030 | Yes | Machine is not switching On | Machine is not switching On | 000000000000000000000000001O##: 00000000000000000000000000100000000299317 300000000Oct-2000000003 Satam, Kamlesh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technician Device ue: No Eupected and ueupected Resuts: NA uer Impact: No Patient Impact: No SE 8.2.3000000005 Resoluion : System was getting stuked at 00000000: 00000000. Replaced CMOS Battery of Fleuvision PC with new one. Now system is 000000000000000000000000001king fine. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000299567 28-Oct-2000000003 Karkate, Vinod The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: In ue Eupected and ueupected Resuts: NA uer Impact: NO Patient Impact: NO SE NA Resoluion : check and foud system not booting and stuk on 0000000000000000 so problem foud in uuuCMOS battery voltage down so replaced CMOS battery and check foud 000000000000000000000000001king ok. Tools & Equpment : MuTIMETER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000298816 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching On When was the first occurence? Within... : 28-Oct-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching On uer Impact: NO uer impact Patient Impact: No patient impact SE 8.2.3000000005 Diagnostic performed by Engineer : Fleuvision PC harddisk was not switching ON. Manully need to switch ON. FSE to check on connections and cmos battery. Internal Comments : uLL# and/or description of uLL: Machine is not switching On Repair Action : Fleuvision PC harddisk was not switching ON. Manully need to switch ON. FSE to check on connections and cmos battery. Resoluion : Fleuvision PC harddisk was not switching ON. Manully need to switch ON. FSE to check on connections and cmos battery. Follow u Requred : uLL# and/or description of uLL: Machine is not switching On Repair Action : Fleuvision PC harddisk was not switching ON. Manully need to switch ON. FSE to check on connections and cmos battery. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-10-28 22:14 | No
3031 | Yes | Generator is bu starting up, No u-ray is possible erro | Generator is bu starting up, No u-ray is possible erro | 000000000000000000000000001O##: 00000000000000000000000000100000000296712 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : uLL# and/or description of uLL: Generator is buy starting u, No u-ray is possible erro Repair Action : FSE to visit and carry ou futher diagnostics in generator. Also need to check the logs for any phase faut - if it is there., need to check hospital power and earthing + uS. Technician Name : Vijayakuar, K.S Internal Comments : uLL# and/or description of uLL: Generator is buy starting u, No u-ray is possible erro Repair Action : FSE to visit and carry ou futher diagnostics in generator. Also need to check the logs for any phase faut - if it is there., need to check hospital power and earthing + uS. Resoluion : FSE to visit and carry ou futher diagnostics in generator. Also need to check the logs for any phase faut - if it is there., need to check hospital power and earthing + uS. Problem Description by Engineer : Problem description by engineer : Generator is buy starting u, No u-ray is possible erro When was the first occurence? Within... : 28-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Generator is buy starting u, No u-ray is possible erro uer Impact: No uer impact Patient Impact: No patient impact SE 2.2.6.00000000Diagnostic performed by Engineer : FSE to visit and carry ou futher diagnostics in generator. Also need to check the logs for any phase faut - if it is there., need to check hospital power and earthing + uS. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000299323 28-Oct-2000000003 Chaniyara, Manan The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Foud generator buy uLL Checked all phase and groud foud okay. Foud oupu from m cab is low. Removed OVPB and checked contactor foud carbon deposit. Cleaned same. Problem Description by Engineer : Problem description by engineer : Generator is buy starting u, No u-ray is possible erro When was the first occurence? Within... : 28-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Generator is buy starting u, No u-ray is possible erro uer Impact: No uer impact Patient Impact: No patient impact SE 2.2.6.00000000Resoluion : Removed contactor for generator cabinet, disassembled and cleaned all contacts. Now system is 000000000000000000000000001king as intended. Tools & Equpment : MuTIMETER,2000000003-11-0000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-28 22:39 | No
3032 | Yes | Machine not switching ON and Spotlight not functioning | Machine not switching ON and Spotlight not funuoning | 000000000000000000000000001O##: 00000000000000000000000000100000000315390000000000000000-Nov-2000000003 Pandey, Rajesh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine not switching ON When was the first occurence? Within... : 28-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine not switching ON uer Impact: No uer impact Patient Impact: No patient impact SE 000Tools & Equpment : MuTIMETER,200000000000000002-17 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Resoluion : Replaced the host pc and foud ok. Handed over the system in good 000000000000000000000000001king condition ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000299700000000 31-Oct-2000000003 Pandey, Rajesh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Checked the system and foud host pc is not getting on. Tools & Equpment : MuTIMETER,200000000000000002-17 00000000:00000000:00000000 Follow u Requred : Futher 000000000000000000000000001k requred Problem Description by Engineer : Problem description by engineer : Machine not switching ON When was the first occurence? Within... : 28-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine not switching ON uer Impact: No uer impact Patient Impact: No patient impact SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029930000000028-Oct-2000000003 Pandey, Rajesh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : As cutomer reported the issu checked the system and foud host pc is not booting, need to replace host pc Tools & Equpment : MuTIMETER,200000000000000002-17 00000000:00000000:00000000 Follow u Requred : Need to order/ replace part Problem Description by Engineer : Problem description by engineer : Machine not switching ON When was the first occurence? Within... : 28-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine not switching ON uer Impact: No uer impact Patient Impact: No patient impact SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000299262 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Resoluion : FSE visit requred for futher checks. Follow u Requred : uLL# and/or description of uLL: Machine not switching ON Repair Action : FSE visit requred for futher checks. Technician Name : Vijayakuar, K.S Problem Description by Engineer : Problem description by engineer : Machine not switching ON When was the first occurence? Within... : 28-Oct-2000000003 17:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine not switching ON uer Impact: No uer impact Patient Impact: No patient impact SE 000Internal Comments : uLL# and/or description of uLL: Machine not switching ON Repair Action : FSE visit requred for futher checks. Diagnostic performed by Engineer : FSE visit requred for futher checks. ----------------------------------------------------------------------------- | 2023-10-28 22:49 | No
3033 | Yes | Screw fell out and table fell and broke | Screw fell out and table fell and broke | 000000000000000000000000001O##: 0000000000000000000000000010000000029000000006 300000000Oct-2000000003 Colmenero Migul The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : TSM Swing arm came appart du to a screw broke on it. TSM Swing arm is defective and needs to be replaced. Parts Consued : 1.00000000000000000000000000000000800000000385100000000,TSM SWING ARM Resoluion : Replaced TSM Swing arm. Tested the system and the system checks OK. System retuned for patient ue. Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: no Eupected and ueupected Resuts: TSM Swing arm came appart uer Impact: none Patient Impact: none SE 8.2.17 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Walter IP & Issu: 65755695- Screw fell ou and table fell and broke Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001roued to quu adv of callback Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-28 22:49 | No
3034 | Yes | Failed Testing .. | Failed Testing .. | 000000000000000000000000001O##: 00000000000000000000000000100000000516677 27-Nov-2000000003 Ardighieri, Andrea The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: uuuplate test Eupected and ueupected Resuts: fauty plate houing uer Impact: no impact Patient Impact: no impact SE na Following a failed test on the plates it was foud the need to replace the plate houing. Diagnostic performed by Engineer: Front panel and label to be replaced, we proceed with an attempt to order spare parts. Device ou of suport. Follow u Requred: Having consuted the TONAZZO technician, the front display is fractued, futhermore the labels in Italian 000000000000000000000000001ld also need to be replaced, informed that since they are ou of suport the spare parts will probably be difficut to order. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000011000000007-Nov-2000000003 Tonazzo, Marco The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Front body replacement requred and plate houing labels modified Diagnostic performed by Engineer: Front body replacement requred and plate houing labels modified. We proceed to follow u Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: uuuplate test Eupected and ueupected Resuts: fauty plate houing uer Impact: no impact Patient Impact: no impact SE na Following a failed test on the plates it was foud the need to replace the plate houing. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029000000005 Ardighieri, Andrea The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: uuuplate test Eupected and ueupected Resuts: fauty plate houing uer Impact: no impact Patient Impact: no impact SE na Following a failed test on the plates it was foud the need to replace the plate houing. Follow u Requred: Following a failed test on the plates, the need to replace the plate houing was foud. As discused with the Granata technician, the parts had already been previouly ordered and are in the Portogruro laboratory, planning intervention by a Sidem GSSH technician. Diagnostic performed by Engineer: Plate houing to be replaced ----------------------------------------------------------------------------- | 2023-10-28 23:29 | No
3035 | Yes | No connection with the wireless foot pedal | No conneuon with the wireless foot pedal | 000000000000000000000000001O##: 000000000000000000000000001000000003000000006100000000Oct 31, 2000000003 Kielholz Willi The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MTRA / BIOMED Device ue: yes Eupected and ueupected Resuts: No effect uer Impact: No effect Patient impact: no effect SE not relevant Resoluion: -> wireless foot switch couled again - > left buton does not 000000000000000000000000001k - > new service will be schedued ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000391 Klu,Hrew The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: uer Device ue: in ue Eupected and ueupected Resuts: No connection to the foot switch uer Impact: no Patient Impact: no SE ... Diagnostic performed by Engineer: Variou clarifications and appointment arranged ----------------------------------------------------------------------------- | 2023-10-28 23:54 | No
3036 | Yes | uElite-keyboard issue (control panel) / Not on FCO list/ GFE0SENT THU 0EC0 | uElite-keyboard issue (control panel) / Not on FCO list/ GFE0 SENT THU 0DEC03 | 000000000000000000000000001O##: 00000000000000000000000000100000000295210000000000000000-Nov-2000000003 Glessner Edward The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: sonographer Device ue: utrasoud Eupected and ueupected Resuts: sometimes control panel swivels after it is locked uer Impact: movement Patient Impact: NA SE9.u Resoluion : Cleaned and adjuted pin and solenoid. See susequnt 000000000000000000000000001for caster replacement. Follow u Requred : Brake on front caster does not always 000000000000000000000000001k well Follow u Requred : N/A ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: jennifer herring IP & Issu: "100000000967313 & keyboard issu and steering issus when ulocking wheels only will partially tun , hard to move aroud patient beds can not lock wheels" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created onsite 000000000000000000000000001O# Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-28 23:59 | No
3037 | Yes | Display issue | Display issue | 000000000000000000000000001O##: 00000000000000000000000000100000000297200000000 17-Nov-2000000003 Orange, James The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Cutomer stated uit had bad display Problem Description by Engineer : Bad Display Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion : upluged all power to uit then Disassembled u assembly (withou nav-ring) from uit to check all connections to LCD screen. I then reseated all connections to ensue they were all properly seated and then reassembled the u assembly to the uit. Followed u by electrical safety and controls test per Service Manul Rev. T. Ended with the uit runing for approuimately 300000000minues to ensue no uLLcodes were foud as well as a final DRPTA check. uit was left in 000000000000000000000000001king condition. Tools & Equpment : STATIC CONTROL DEVICE,200000000000000002-29 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000002-29 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Chris Carpenter 9123500000000100000000 IP & Issu: SN 1000000000000000051500000000 Bad display, requsting onsite Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Created case and Page QA and FSE Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Cutomer Call Back Caller: Chris Carpenter 912-35000000003100000000 IP & Issu: SN 1000000000000000051500000000 Needed to udate PO and check on case Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Advised of Statu and udated PO Neut Step Action: N/A Verbal PO Nuber: 0000000062000000001 Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Cutomer Call Back Caller: Chris Carpenter IP & Issu: CCB to get the statu udate for 000000000000000000000000001dispatch and requsted if 000000000000000000000000001dispatch can be moved to an earlier date other than 11/28 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Notated in RP chat at 10000000000000000M and also created an AA. Advised cu to eupect a call back. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-29 00:04 | No
3038 | Yes | Audio & Sounds/Speaker failure during PM/Speaker | Audio & Sounds/Speaker failure during PM/Speaker | 000000000000000000000000001O##: 00000000000000000000000000100000000395113 Blackmore Joshu The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Duing PM Eupected and ueupected Resuts: Speaker failed test duing PM uer Impact: None Patient Impact: None SE NA Resoluion : Touh Bezel replaced to resolve damage. Performance verification passed. Parts Consued : 1.0000000000000000000000000000000056000000005151,IV-MP5 Bezel Touh Assembly 5wire Tools & Equpment : ELECTRICAL SAFETY TESTER,20000000000000000000000001 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000005-2000000000000000000000000:00000000 PRESSuE CALIBRATION TOOL,200000000000000002-00000000 00000000:00000000:00000000 PARAMETER TESTER,200000000000000005-00000000 00000000:00000000:00000000 PRESSuE CALIBRATION TOOL,200000000000000005-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000500000000 Blackmore Joshu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Speaker assembly replaced to resolve cutomer reported faut 'Speaker failue'. Futher parts requred. Parts Consued : 1.00000000000000000000000000000000261000000009000000001,IV-MP5 CBL Speaker Assembly Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Duing PM Eupected and ueupected Resuts: Speaker failed test duing PM uer Impact: None Patient Impact: None SE NA Diagnostic performed by Engineer : Speaker assembly replaced to resolve cutomer reported faut 'Speaker failue'. Futher parts requred. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000295373 James Gillham The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Duing PM Eupected and ueupected Resuts: Speaker failed test duing PM uer Impact: None Patient Impact: None SE NA Diagnostic performed by Engineer : Speaker failue duing PM. FSE went to site to perform a PM on the device and foud it failed on the speaker test. Faut is du to either the speaker assembly or the main board. A qute for a fiued price repair will be sent to the cutomer. PNs requred 00000000261000000009000000001 & 0000000056000000001831. Follow u Requred : Speaker failue duing PM. FSE went to site to perform a PM on the device and foud it failed on the speaker test. Faut is du to either the speaker assembly or the main board. A qute for a fiued price repair will be sent to the cutomer. PNs requred 00000000261000000009000000001 & 0000000056000000001831. ----------------------------------------------------------------------------- Internal Remarks : Tracking Nuber:1Z3R6000000006698537988 collrect tomorrow ----------------------------------------------------------------------------- | 2023-10-29 00:09 | No
3039 | Yes | When the system is turned on, it will do the count down and freeze unable to shut off unless hard booted off | When the system is turned on, it will do the count down and freeze unable to shut off unless hard booted off | 000000000000000000000000001O##: 000000000000000000000000001000000003000000002000000001-Nov-2000000003 Scott, Bennett The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: cath lab Device ue: ued for image guded therapy Eupected and ueupected Resuts: system does not boot u properly. uer Impact: high Patient Impact: high SE 8.2 Resoluion : The system 000000000000000000000000001ld not boot u properly when shu down. The fleuvision pc was not booting u and therefore the system was stuk at the 0000000000000000 util boot u coutdown screen. The biomed was able to manully boot u the fleuvision pc with the power buton and get the system operating for the uers and patients util I arrived. I arrived and changed ou the cmos battery on the fleuvison pc. The cmos battery was very low and ustable. Performed several eutended shudowns and reboots to make sue the system booted properly. Checked system operation. System is ready for ue. Tools & Equpment : ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Caller: Anthony Mcmiller - Not the contact IP & Issu: 72952800000000 -When the system is tuned on it will do the cout down and freeze uable to shu off uless hard booted off Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Qute Entitlement Options: Qute Reason for Entitlement Claim: Requsted Qute - Declined eupedite Fee Action Taken: Awaiting Qute Neut Step Action: Internal Remarks : Interaction Type: Cutomer Call Back Caller: Steve Sidert - 500000000362379 IP & Issu: Cutomer Call Back Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Verified ##is actully steven.sibert00rimedu. Called qutes and spoke with D to have ##udated and qute sent. Also left AA ----------------------------------------------------------------------------- | 2023-10-29 00:09 | No
3040 | Yes | After performing the operation test, it gives an uLLin the Defibrillation Test: Incorrect/Defib. | After performing the operation test, it gives an uLLin the Defibrillation Test: IncorreuDefib. | 000000000000000000000000001O##: 00000000000000000000000000100000000335682 00000000-Nov-2000000003 Pedrero de la Punte, Maria The system meets the specification for the performed service and is retuned to ue. Follow u Requred: Prepare a buget for 2 hous of technical service + material. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Failue when doing auotest. Eupected and ueupected Resuts: Intermittent failue. uer Impact: None Patient Impact: None SE NA Diagnostic performed by Engineer: Condenser is damaged, it is within tolerance (172.5, greater than 1700000000 bu very close to the limit that gives puctul failues in the defibrillation test. need to replace the condenser. Tools & Equpment: DEFIBRILLATOR ANALYZER, 200000000000000000000000010000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-29 00:14 | No
3041 | Yes | Disk image full issue | Disk image full issue | 000000000000000000000000001O##: 0000000000000000000000000010000000029500000000 00000000-Dec-2000000003 Maua George The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: FD 200000000200000000 Eupected and ueupected Resuts: Low memory space uer Impact: limited Patient Impact: n/a SE8.1.17 Diagnostic performed by Engineer : following u on 000000000000000000000000001000000005611100000000---------------------------------------------------------------------------- Internal Remarks : Both AP and Lateral ippc's reporting image disk ful with only a few patients on drives. Recreate image store clears temporarily then it retuns. Logs indicate intermittent net000000000000000000000000001k connection issus to both IPPCs. Recommended restoring a ghost and if faut retuns replace the Host PC. Internal Remarks : Image memory ful. FSE sent wrong logs. Will call back. Internal Remarks : FSE George Maua. FSE calling for getting memory ful messages intermittently. Sugested to replace image drives in IPPC's along with verify host and image processing. Recreate image store and database consistancy test and repair have already been performed. ----------------------------------------------------------------------------- | 2023-10-29 00:14 | No
3042 | Yes | 0 failure | 0V failure | 000000000000000000000000001O##: 00000000000000000000000000100000000382356 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : This repair will not be solved by replacing the PM pcba per cutomers requst. They will need to order an MC pcba to solve the issu. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: ASP Device ue: not in ue Eupected and ueupected Resuts: many uLLcodes uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000 and it has many codes 1117, 111D, 111B, 1127 and 1118 in the history. ASP was on site and tried a PM pcba per FCO66, bu it did not solve the issu. Recommened replacing the MC pcba part #00000000561537531, MC pcba, list price $1198.00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000295866 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: uLL since remediation was done uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states that after FCO 66 was implemented on device, device has been having issus.cutomer troule shot and foud that PM board is bad. cutomer is requsting FSE on site to replace PM board uder FCO warranty. Confirmed cutomer's contact info and device location. Resoluion : Creating on site 000000000000000000000000001to have discouted at 100000000% to have PM board replaced Follow u Requred : 100000000% discout, Device failing after PM board was replaced. Requsting FSE and new PM board discouted 100000000% uder FCO warranty. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: randy IP & Issu: 72200000000111 uLL since remidation Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A Internal Remarks : 11/1/23 YP 00000000000000000000000000100000000316000000005 Attempt1 Details: Incident follow u ##sent to both FSE and cutomer. Cncted cut. Cut not reachable. left vm Contact Randy Hooper Phone +161599559100000000##randy.hooper00rimedu.com Internal Remarks : 11/1/23 YP 00000000000000000000000000100000000316000000005 Attempt2 Details: Cncted the below no.s. Cut not reachable. left vm. ##sent to the cut. FSE Luman Name: Christopher Asberry Phone: 615-396-0000000000000000Mainline ## christopher.asberry00scension.org Location Name St Thomas Ruherford Hospital Street Medical Center Pkwy 1700000000 City Mufreesboro State TN Zip 37129-2200000000Internal Remarks : 11/1/23 YP 00000000000000000000000000100000000316000000005 S/N- 200000000000000001722 Details: Cut confirmed S/N and location over phone. Dispatched to FSE Luman Contact Randy Hooper Phone 6158000000001700000000##randy.hooper00rimedu.com Location Name St Thomas Ruherford Hospital Street Medical Center Pkwy 1700000000 City Mufreesboro State TN Zip 37129-2200000000----------------------------------------------------------------------------- | 2023-10-29 00:39 | No
3043 | Yes | Fluoro storage was not possible | Fluoro storage was not possible | 000000000000000000000000001O##: 00000000000000000000000000100000000295978 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : remote alert of CV u: uLLFLuSTR-IMAGEDS The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog: The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog:The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog: Flu storage not poss. Image disk problem Image disk problem When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : Mutiple times per month Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: M2M/Radar Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Flu storage not poss. Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : nonitor of system show additional alerts of alert on system this meets dispatch to field Follow u Requred : uLL# and/or description of uLL: CV u: uLLFLuSTR-IMAGEDS alert Repair Action : The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog: Flu storage not poss. Image disk problem Description=alert rue version RMA uuu200000000-----WAIT FOR 00000000PPENDING ALERTS IN 2 DAYS!! KEEP THIS M2M CASE OPEN SO M2M CAN ADD THE APPENDING ALERTS! IF NO APPENDING ALERT IS RECEIVED THEN CLOSE THIS CASE. DO NOT DISPATCH FSE IF FAIL RATE IS LESS THAN 00000000-----------------------------------------------------------Duing POST of the IP-PC Image Disk has failed. Order 2 Image Disks. Save all local euams to PACS or other archive. Replace Image Disks.------------------------------------------------------------(IF FSE i 0000000080000000050000000000000000 Description Disk for ClarityIQ_II and IP_Revised_II What are the neut Trouleshooting steps : Remote alert shows duing POST of the IP-PC Image Disk has failed. Order 2 Image Disks. Save all local euams to PACS or other archive. Replace Image Disks. Technician Name : Sulivan Scott Resoluion : meets dispatch rue created on site Internal Comments : uLL# and/or description of uLL: CV u: uLLFLuSTR-IMAGEDS alert Repair Action : The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog: Flu storage not poss. Image disk problem Description=alert rue version RMA uuu200000000-----WAIT FOR 00000000PPENDING ALERTS IN 2 DAYS!! KEEP THIS M2M CASE OPEN SO M2M CAN ADD THE APPENDING ALERTS! IF NO APPENDING ALERT IS RECEIVED THEN CLOSE THIS CASE. DO NOT DISPATCH FSE IF FAIL RATE IS LESS THAN 00000000-----------------------------------------------------------Duing POST of the IP-PC Image Disk has failed. Order 2 Image Disks. Save all local euams to PACS or other archive. Replace Image Disks.------------------------------------------------------------(IF FSE i 0000000080000000050000000000000000 Description Disk for ClarityIQ_II and IP_Revised_II ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000338200000000Rodriguz Lazaro Problem Description by Engineer : Problem description by engineer : remote alert of CV u: uLLFLuSTR-IMAGEDS The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog: The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog:The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog: Flu storage not poss. Image disk problem Image disk problem When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : Mutiple times per month Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: M2M/Radar Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Flu storage not poss. Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-10-29 00:44 | No
3044 | Yes | Alarms or Messages/Speaker malfunction/Hardware | Alarms or Messages/Speaker malfunuon/Hardware | 000000000000000000000000001O##: 00000000000000000000000000100000000319681 00000000-Nov-2000000003 Neil Powell The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Medical Equpment technician Engineer Device ue: not in ue Eupected and ueupected Resuts: Duing PVT it was noted that INOP Message Speaker Malfuction is displayed uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : replaced main board and speaker, performance verified and safety tested Parts Consued : 1.0000000000000000000000000000000056000000006631,IV2-FLEu MECHASY Louspeaker assembly 1.0000000000000000000000000000000056000000006331,IV2-FLEu Main board Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-0000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000029600000000 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Engineer to attend site with the following parts for investigation and repair: 1u 0000000056000000006631 IV2-FLEu MECHASY Louspeaker assembly 1u 0000000056000000006331 IV2-FLEu Main board Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Medical Equpment technician Engineer Device ue: not in ue Eupected and ueupected Resuts: Duing PVT it was noted that INOP Message Speaker Malfuction is displayed uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Engineer to attend site with the following parts for investigation and repair: 1u 0000000056000000006631 IV2-FLEu MECHASY Louspeaker assembly 1u 0000000056000000006331 IV2-FLEu Main board ----------------------------------------------------------------------------- Internal Remarks : Change of contact for case, Alberto is away for 1 month. Komal. Komal.Castro00stt.nhs.u Tel: 0000000000000000 188 3795 Neut Step Action : ##sent to contact to arrange engineer visit. ----------------------------------------------------------------------------- | 2023-10-29 00:49 | No
3045 | Yes | VIP-IMSS-CLIENT-IT DOES NOT WORK | VIP-IMSS-CLIENT-IT DOES NOT WORK | 000000000000000000000000001O##: 00000000000000000000000000100000000397182 28-Nov-2000000003 Hello, Jonathan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: maintenance Device ue: no Eupected and ueupected Resuts: does not tun on uer Impact: does not tun on Patient Impact: does not tun on SE8.2.17 Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Resoluion: The site is attended du to a reported failue. A physical change of PC host is made, a new license is loaded as well as software loaded. Fuctional tests are carried ou, obtaining satisfactory resuts. The equpment operates correctly with the patient. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000395000000006 Valentin Medina The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: NA Device ue: NA Eupected and ueupected Resuts: NA uer Impact: NA Patient Impact: NA SE NA Diagnostic performed by Engineer: I approve all docuentation in these 000000000000000000000000001 and confirm that this service was completed safely and in compliance with u policies ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000385280000000000000000-Nov-2000000003 Hello, Jonathan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Comments for TSS Yougo to the site and find the device tuned on withou the application entering. The statu of the host PC is checked. The PC is powered with 2300000000V inpu, however, the hard drive and fan LEDs do not light u and there is no activity in the net000000000000000000000000001k cable LEDs. The PC is removed from the cabinet, the cover is ucovered, the connections of the Souce with the disks and motherboard are verified withou sucess, the lanPC does not respond Problem Description by Engineer: Cutomer Fuction/Role: maintenance Device ue: no Eupected and ueupected Resuts: does not tun on uer Impact: does not tun on Patient Impact: does not tun on SE8.2.17 Diagnostic performed by Engineer: Comments for TSS Yougo to the site and find the device tuned on withou the application entering. The statu of the host PC is checked. The PC is powered with 2300000000V inpu, however, the hard drive and fan LEDs do not light u and there is no activity in the net000000000000000000000000001k cable LEDs. The PC is removed from the cabinet, the cover is ucovered, the connections of the Souce with the disks and motherboard are verified withou sucess, the lanPC does not respond ----------------------------------------------------------------------------- | 2023-10-29 01:14 | No
3046 | Yes | Defibrillator blades that are not being recognized by the device. | Defibrillator blades that are not being recognized by the device. | BuQ&R Feedback: No Reason: Corrective Maintenance Trouleshooting Action: OPP-198578 Qute: 00000000113322 ## engenharia.clinica00epaco.org.br Safety Qustions: Case: 0000000022500000000000000008 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Nicole Carmo Ferreira Last udated by: Nicole Carmo Ferreira 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment malfuction or fail at the time of installation? uknown Problem Reported by Cutomer: Equpment Problem: Defibrillator blades that are not being recognized by the device. Priority: Stationary Search for:Run Pablo Indutry:Clinical Engineering Phone: 55970000000013616 Phone: 5598000000008177 ##engenharia.clinica00epaco.org.br uit: SOCIAL CARDBOARD PAPER IND SERVICE Address: Ru Verguiro 0000000000000000Sao Pauo-SP 0000000000000000-900000000 ----------------------------------------------------------------------------- | 2023-10-29 01:29 | Yes
3047 | Yes | Cable drape damaged. System down | Cable drape damaged. System down | 000000000000000000000000001O##: 00000000000000000000000000100000000297983 Niemyjski Mark Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-10-29 02:49 | No
3048 | Yes | Battery Charging port malfunction | Battery Charging port malfunuon | Internal Remarks : ##d 2nd requst to cutomer for SW rev for loaner. Internal Remarks : Interaction Type:New Case Caller: Gary Vasil IP & Issu: 61500000000 Battery Charging port malfuction Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-29 03:29 | No
3049 | Yes | Image disc failure | Image disc failure | 000000000000000000000000001O##: 0000000000000000000000000010000000032520000000000000000-Nov-2000000003 Clement Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: 3rd party Device ue: service Eupected and ueupected Resuts: IPPC failue uer Impact: 000000000000000000000000001kflow delay Patient Impact: none SE 000Parts Consued : 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD 1.00000000000000000000000000000000800000000772161,Spare part set 200000000E100000000ED lamp 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replaced the IP PC and drives. IP PC software was reloaded and system was tested. The system is operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000298316 00000000-Nov-2000000003 Clement Michael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: 3rd Party Vendor Device ue: Repair Eupected and ueupected Resuts: image drives not fuctional uer Impact: delay of ue Patient Impact: none SE000Follow u Requred : Replace IPPC Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Diagnostic performed by Engineer : While trying to only replace the image drives in the IPPC, the OS drive bay stopped fuctioning and the bios got corruted. Ordering new IPPC and OS drive to go withe the new image drives ----------------------------------------------------------------------------- Internal Remarks : FSE is replacing original IPPC with Revised II and had qustion abou different connections. Ref changed from DP to DVI so FSE will need to install a DVI SL transmitter Internal Remarks : Interaction Type: New Case Caller: Mike Dagley IP & Issu: 500000000900000000;Image disc failue. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Create case and qute Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : FSE is replacing the image drives on an IPPC. As soon as he replaced these, the OS disk power buton stopped 000000000000000000000000001king. FSE tried to reseat the drive bay, bu this didn't help. Bios is not recognizing any of the drives, so it seems the drive bay connection is fuly broken. This is an original Nehalem IPPC, so it is qute old. -Approved: IPPC Revised II no HDD: 0000000080000000050000000092 Disk for revised II: 0000000080000000050000000000000000----------------------------------------------------------------------------- | 2023-10-29 04:29 | No
3050 | Yes | 0/u FD0Image disk problem | 0R/u FD0Image disk problem | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Hard disk uLLcannot be euposed uCMG is processing SAP Transfer: Case Transferred to SAP Euternal Remarks: 28.1000000002000000003 1000000007:00000000 uC (86000000006500000000) Service content: Repair image compuer. Service resut: After repair, the device 000000000000000000000000001ks normally SAP Transfer: Case Transferred to SAP Internal Remarks: 28.1000000002000000003 1000000007:00000000 uC (86000000006500000000) "RDI Accuacy: Yes" $Q7: (confirmed reported problem &Image disc ful is report) $Q8: (check IPPC power u, it seems ok, check IPPC harddisc, it seems ok, do database check, no uLL maybe IPPC SW problem) $Q9: (reload IPPC SW, and Swap Image disc, do database receive, fuctiontest is ok ) $Q100000000 (Test and inspection data refer to: PA.&Device retuned to ful fuctionality.) $Q11: Calibrated Tools name:NA T2 Activities:Problem Description: Overview of cutomer problems: The device cannot eupose and store the image.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Frontal IPPC memory 2 has failed; Corruted images received from image detection; Flu storage not possible du to an Image disk problem (>Rel7.2) Remote diagnostics and test resuts: Image storage issus PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Caue 1 Flu storage not possible du to an Image disk problem (>Rel7.2) Soluion 1 Carry ou a Re-create image store via FSF. Soluion 2 If problem still persists replace Imagedisk(s) Caue 2 This uLLcan be caued du to fact the IP-PC (TP) did not start u (properly). Soluion 1 Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Please refer to the Produt Manul of the IP-PC and Host PC (Nehalem) on InCenter for the LED statu details. Soluion 2 If LED statu is Normal then issu is probably net000000000000000000000000001k related. Check the Ethernet connection (TPu100000000- THu13) between the Host PC and the Lateral IP-PC. Soluion 3 If LED statu is not Normal then issu is Boot problem in Lateral IP-PC, check IP-PC power and perform cold reboot of the system and replace Lateral IP-PC (TP) if problem Recommended spare parts uuand designation: Notes on cutomer requrements: Onsite Safety Qustions: Case: 0000000022510000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-29 07:14 | No
3051 | Yes | 0/u FD0Host PC cannot power on | 0R/u FD0Host PC cannot power on | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 00000000.11.2000000003 13:16:58 uC (8600000000600000000)RDI Accuacy: YES $Q7:Device fails to power on. $Q8:The onsite check foud that the host coud not start normally $Q9: The internal board of the host is removed and reinserted onsite. The host starts normally, Scan patients normally $Q100000000 Test and Inspection data refer to :PA,Device retuned to ful fuctionality. $Q11: Calibrated Tools name (N/A )S/N ( N/A) Neut Calibration date (N/A ) T2 Activities:Problem Description: Overview of cutomer problem: After the device is tuned on, all screens are blank.Engineers are requred to be present. Is the faut caued by the equpment itself? Yes Trouleshooting Action: Remote diagnosis and test resuts: It is supected that the Host PC is not powered on. PRS TR statu: No PRS Repair Action: Recommended onsite repair steps: Plu and uplu the Host PC board, clean the PC dut, check whether the BIOS battery in the Host PC is charged, reinstall the Host software, and check whether the Host PC's power suply is normal. Notes on cutomer requrements: Onsite Problem Reported by Cutomer: Can't tun on the compuer SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000225100000000600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 13:16:51 uC (8600000000600000000) Service content: The host cannot be started. The internal board of the host is upluged and inserted on site. The host starts normally and the test fuction is normal. Service resut: The device operates normally. Euternal Remarks: 31.1000000002000000003 00000000:26:15 uC (86000000006252)I approve all of the docuentation in this S000000000000000000000000001and confirm this repair in a safe and complete manner SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 28.1000000002000000003 10000000027:35 uC (8600000000600000000)"SuT. Ready for Clinical ue local time" : " 2000000003-10000000028 16:00000000" Internal Remarks: 31.1000000002000000003 00000000:26:18 uC (86000000006252) I approve all of the docuentation in this S000000000000000000000000001and confirm this repair in a safe and complete manner SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-29 07:54 | No
3052 | Yes | Euposure is not possible unable to do scan | Euposure is not possible unable to do scan | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000157 21-Nov-2000000003 Ramachandran, Sekar The system meets the specification for the performed service and is retuned to ue Resoluion : cleaned and refiued the DIMM in the memory slots and checked. now the System is booting normally and 000000000000000000000000001king withou uLL system handed over for clinical uage. Problem Description by Engineer : Problem description by engineer : Euposue is not possible uable to do scan When was the first occurence? Within... : 300000000Oct-2000000003 12:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Euposue is not possible uable to do scan uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.17.00000000Diagnostic performed by Engineer : on inspection foud the problem with the Memory modue uLLin the IPPC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000500000000Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Euposue is not possible uable to do scan When was the first occurence? Within... : 300000000Oct-2000000003 12:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Euposue is not possible uable to do scan uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.17.00000000Resoluion : Image processing malfuction uLLoccus. Image disk needs replacement. Diagnostic performed by Engineer : Image processing malfuction uLLoccus. Image disk needs replacement. Internal Comments : uLL# and/or description of uLL: Euposue is not possible uable to do scan Repair Action : Image processing malfuction uLLoccus. Image disk needs replacement. Follow u Requred : uLL# and/or description of uLL: Euposue is not possible uable to do scan Repair Action : Image processing malfuction uLLoccus. Image disk needs replacement. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-10-29 15:04 | No
3053 | Yes | 0/u FD0No euposure | 0R/u FD0No euposure | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 00000000.11.2000000003 16:25:28 uC (8600000000900000000) $Q7:FSE confirmed reported problem, the footswitch screw broken and the connector cout't connect $Q8: check the footswitch is ok, wrap the broken screws and replace the screws $Q9: replace the screws make sue the footswitch is ok $Q100000000System is ok. Test and inspection data refer to PA, device retuned to ful fuctionality $Q11: Calibrated Tools name: N/A SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: No euposue SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 00000000.11.2000000003 15:0000000000000000uC (8600000000900000000) Service content: If the fiuing screw of the foot brake is broken, remove the broken wire and replace the fiuing screw. Service resuts: After repair, the equpment retuned to normal operation. Safety Qustions: Case: 0000000022511000000008 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: nannan wang Last udated by: nannan wang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-29 17:34 | No
3054 | Yes | 0/u FD0Unable to start the machine | 0R/u FD0Unable to start the machine | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 00000000225112500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: nannan wang Last udated by: nannan wang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: uable to start the machine SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 00000000:5000000009 uC (8600000000000000001) Service content: IPPC cannot boot, replace IPPC compuer. Service resuts: After repair, the equpment retuned to normal operation. Internal Remarks: 00000000.11.2000000003 00000000:5000000006 uC (8600000000000000001) "RDI Accuacy:N/A" $Q7: (Confirmation the report problem.IPPC coudn't power on.) $Q8: (Onsite check the problem, the IPPC inputing power was normal. the Power of IPPC was defective, repalce the IPPC.) $Q9: (Replace the IPPC.uu0000000080000000012200000000) $Q100000000 (Test and inspection data referto: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name(Mutimeter Flue 287) S/N (22900000000100000000Neut Calibration date (12/00000000000000003) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage:N/A ----------------------------------------------------------------------------- | 2023-10-29 20:04 | No
3055 | Yes | Machine is automatically shutting down | Machine is automatically shutting down | 000000000000000000000000001O##: 00000000000000000000000000100000000382000000002 00000000-Nov-2000000003 Mohammed, Ashfaq The system meets the specification for the performed service and is retuned to ue Resoluion : Monitored system for cases. System is 000000000000000000000000001king with in recommended specifications Problem Description by Engineer : Problem description by engineer : Machine is auomatically shuting down When was the first occurence? Within... : 00000000-Nov-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is auomatically shuting down uer Impact: No uer impact Patient Impact: No patient impact SE 7.2.16.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000035200000000 00000000-Nov-2000000003 Mohammed, Ashfaq The system meets the specification for the performed service and is retuned to ue Follow u Requred : Monitoring after repair Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000001-11 00000000:00000000:00000000 MuTIMETER,2000000003-12-00000000 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000001-11 00000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : Machine is auomatically shuting down When was the first occurence? Within... : 00000000-Nov-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is auomatically shuting down uer Impact: No uer impact Patient Impact: No patient impact SE 7.2.16.00000000Resoluion : Monitored for cases after repair. System is 000000000000000000000000001king normally. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000321000000002 00000000-Nov-2000000003 Mohammed, Ashfaq The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced collimator and completed requred calibration. Tested system sucessfuly after replacement. Monitored for few cases and system is 000000000000000000000000001king within recommended specifications Parts Consued : 1.00000000000000000000000000000000800000000858162,NICOL V3 CV FD 1.00000000000000000000000000000000216700000000585,SIB_u BOARD Follow u Requred : N/A Problem Description by Engineer : Problem description by engineer : Machine is auomatically shuting down When was the first occurence? Within... : 00000000-Nov-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is auomatically shuting down uer Impact: No uer impact Patient Impact: No patient impact SE 7.2.16.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000319837 00000000-Nov-2000000003 Viswanath Buti, Hemanth The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Machine is auomatically shuting down When was the first occurence? Within... : 00000000-Nov-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is auomatically shuting down uer Impact: No uer impact Patient Impact: No patient impact SE 7.2.16.00000000Diagnostic performed by Engineer : Shuters uavailable uLLuer message , uuu cine not happening , foud EPu issu and collimator NICOL issu , need to order part Follow u Requred : Part ordering ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000319337 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Machine is auomatically shuting down Repair Action : 1) System was in paid category. 2) System has mutiple problems - Spoke to FSE Hemanth - and uderstood issu with collimator and euposue. 3) FSE visit requred for futher diagnostics. Diagnostic performed by Engineer : 1) System was in paid category. 2) System has mutiple problems - Spoke to FSE Hemanth - and uderstood issu with collimator and euposue. 3) FSE visit requred for futher diagnostics. Follow u Requred : uLL# and/or description of uLL: Machine is auomatically shuting down Repair Action : 1) System was in paid category. 2) System has mutiple problems - Spoke to FSE Hemanth - and uderstood issu with collimator and euposue. 3) FSE visit requred for futher diagnostics. Technician Name : Vijayakuar, K.S Resoluion : 1) System was in paid category. 2) System has mutiple problems - Spoke to FSE Hemanth - and uderstood issu with collimator and euposue. 3) FSE visit requred for futher diagnostics. Problem Description by Engineer : Problem description by engineer : Machine is auomatically shuting down When was the first occurence? Within... : 00000000-Nov-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is auomatically shuting down uer Impact: No uer impact Patient Impact: No patient impact SE 7.2.16.00000000----------------------------------------------------------------------------- | 2023-10-29 21:04 | No
3056 | Yes | not able to do runs, generator starting up, uray not available tube fault error | not able to do runs, generator starting up, uray not available tube fault error | Internal Remarks : Interaction Type: New Case Caller: Nicolas Hernanadez IP & Issu: not able to do rus, generator starting u, uray not available tue faut uLL Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Qute Neut Step Action: Awaiting Qute Verbal PO Nuber: 85200000000600000000 Verbal PO Approver: Nicolas Hernanadez Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-29 22:29 | No
3057 | Yes | QAASA:0/u FD0host unable boot | QAASA:0R/u FD0host unable boot | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 00000000.11.2000000003 00000000:59:19 uC (860000000060000000000000000 "RDI Accuacy: Yes" $Q7: (Confirmation & Description of problem): Confirmed reported problem. Host PC cannot power on. $Q8: (Trouleshooting & Resuts: 2200000000AC power suply is OK. Host PC defect $Q9: (Repair steps): replace Host PC. Ful software installation. system observation $Q100000000 (Test and Inspection data & Device statu: Test and inspection data refer to: PA. Device retuned to ful fuctionality. $Q11: Calibrated Tools name ( N/A ) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 00000000/ 200000000 2000000003 ^Curent Tue uage: scan second_ 32500000000.25LuProblem Reported by Cutomer: QAASA:00000000R/u uuuFD200000000Euternal Remarks: 00000000.11.2000000003 00000000:59:17 uC (860000000060000000000000000Faut description: HOST PC cannot start normally Service content: Replace HOST PC and completely reinstall the system.Systematic observation. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022511611 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD00000000800000000122391*1 Arranged for 200000000delivery to the venu FSE Chen Shuing -86-13800000000676976 T2 Activities:Problem Description: Overview of cutomer problems: Host uable to boot Detailed description of cutomer problem or faut phenomenon: n/a Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: There is no display on the main control screen, and the uuumodue displays host not connect. Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check the HOST PC startu statu, front panel indicator light, and rear indicator light.If the HOST PC fails, it is recommended to try replacing the bios battery. 2. Check whether the PC has 2200000000 inpu power when it is tuned on.If not, it is recommended to check the power suply of the back panel of M cabinet. Recommended spare parts uuand designation: PPD uu0000000080000000012200000000 u Haswell IP PC no HDD 200000000to Hosp PPD uu0000000080000000000000000321 PD.ASSY.rearpanel boued 200000000to Hosp Notes on cutomer requrements: Please contact the cutomer for on-site repair SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 0000000080000000000000000321 1 Order has been arranged, 0000000080000000012200000000 1 Arranged to go to the hospital 200000000ous a day, FSE Chen Shuing -86-13800000000676976 Euternal Remarks: 29.1000000002000000003 12:11:21 uC (860000000060000000000000000SuT. Ready for Clinical ue local time" : " 2000000003-10000000029 11:00000000 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: C arm movement issu Detailed description of cutomer problem or faut phenomenon: C-arm movement failue Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: There is no display on the main control screen, and the uuumodue displays host not connect. Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check the HOST PC startu statu, front panel indicator light, and rear indicator light.If the HOST PC fails, it is recommended to try replacing the bios battery. 2. Check whether the PC has 2200000000 inpu power when it is tuned on.If not, it is recommended to check the power suply of the back panel of M cabinet. Recommended spare parts uuand designation: PPD uu0000000080000000012200000000 u Haswell IP PC no HDD 200000000to Hosp PPD uu0000000080000000000000000321 PD.ASSY.rearpanel boued 200000000to Hosp Notes on cutomer requrements: Please contact the cutomer for on-site repair ----------------------------------------------------------------------------- | 2023-10-30 06:49 | No
3058 | Yes | the system wont boot | the system wont boot | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000600000000300000000Oct-2000000003 Marshall James The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 Diagnostic performed by Engineer : system not booting. stopping at zero. field service frame000000000000000000000000001k not availible. check foud disk with connector issus. backu disk uavailible. cutomer cleaned. SIB not initializing to allow system to boot. cutomer to order parts. SIB, cable kit and drives. will follow u with cutomer. Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u. uer Impact: no impact Patient Impact: no impact SE 7.2 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Marcu Wyant IP & Issu: 00000000600000000200000000 the system 000000000000000000000000001t boot Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Was the device in Clinical ue at the time of the event? No Preferred Service Time: tomorrow morning Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created case & qute for onsite Neut Step Action: N/A Verbal PO Nuber: 275500000000000000000000000092 Verbal PO Approver: Marcu Wyant Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-30 08:49 | No
3059 | Yes | Patient Complaint that device is getting too hot - Awaiting Product Return 00 | Patient Complaint that device is getting too hot - Awaiting ProduuReturn 0/0 | Test Equpment Info : Nul Device Alarm/Alert as intended? : No Problem Observed : facility has contacted the local lead tech who was informed that the device may have buned a patient on the "R" shouder. uit is being sequstered by the Qulity control person on site. It was reported that the incident occued on Thusday the 26th bu u was not notified util Friday the 27th. Ou Of Bou Failue? : No Labor: Labor ID : L-198630000000000000000ork Performed : At this time I was notified by Christine Deeds (COOO that the device may have buned the patient on the right shouder. The device is being held in sequster by Kim Marie Lopez who is believed to be the qulity control officer. A follow u phone call will occue on Monday. Diagonsis : Not Answered | 2023-10-30 10:44 | Yes
3060 | Yes | stuck in the boot up stage did a hard reset but will not come up | stuck in the boot up stage did a hard reset but will not come up | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000311 300000000Oct-2000000003 Pena, Jessie The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: technician Device ue: FleuVision PC Eupected and ueupected Resuts: FleuVision PC reads Hardrive, it doesn’t read Hardrive uer Impact: FleuVision PC not reading Hardrive Patient Impact: NA SE8.2 Diagnostic performed by Engineer : FleuVision PC is not reading Hardrive, Follow-u 000000000000000000000000001requred to replace FleuVision PC. Follow u Requred : Replace FleuVision PC ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Enriqu IP & Issu: 7311000000007 stuk in the boot u stage did a hard reset bu will not come u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Afterhous Entitlement Options: Entitled Reason For Entitlement Claim N/A Action Taken: created onsite, added po and sent to rp to dispatch asap per cutomer requst Neut Step Action: N/A Verbal PO Nuber: Enriqu Sanchez Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Enriqu Sanchez IP & Issu: 7311000000007 & stuk in the boot u stage did a hard reset bu will not come u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: requsted qute Neut Step Action: Awaiting Qute Verbal PO Nuber: Enriqu Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-30 12:19 | No
3061 | Yes | Monitor blank | Monitor blank | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000016000000001-Oct-2000000003 Madison James The system meets the specification for the performed service and is retuned to ue Resoluion : Duing trouleshooting process, it was foud that the harddisk power is available after manul started. Auomatic start of harddrive is not possible, this symptom lesd to the possibilites of CMOS battery coud be drained to nul. After measuement, it was foud the battery uuuity is empty. This means no need to replace battery nor the PC, since the rootcaue is the cmos battery. Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Engineer Device ue: Cathlab FD200000000 Eupected and ueupected Resuts: System failed to start initialization uer Impact: NA Patient Impact: NA SE 8.1.17 Resoluion : Replacement of CMOS battery, udate configuation on BIOS, and perform start u. System start u normally, test uray passed. System left in ful fuctionality. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000900000000Madison James The system does not meet the specification for the performed service and is not retuned to ue Resoluion : replace Harddrive and HostPC Follow u Requred : replacement of PC and harddisk Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Engineer Device ue: Cathlab FD200000000 Eupected and ueupected Resuts: System failed to start initialization uer Impact: NA Patient Impact: NA SE 8.1.17 Diagnostic performed by Engineer : gude Hospital technician to open M cabinet and check Harddisk indicator, whether its runing or no. foud that the HDD indicator is not lit. this means PC cannot access HDD it self, while the power indicator of Host PC is lit. this means PC is tuned on. ----------------------------------------------------------------------------- | 2023-10-30 12:29 | No
3062 | Yes | 0/u FD0/image space low | 0R/u FD0/image space low | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 31.1000000002000000003 1000000005:23 uC (8600000000900000000) Service content: 1.Image cannot be saved 2. Replace the image disk, the problem is solved, and the machine retuns to normal Service resuts: The device powers on normally, perspective and euposue are normal, and the image displays normally. Safety Qustions: Case: 00000000225118100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/u uuuFD200000000Memory is ful SAP Transfer: Case Transferred to SAP Neut Step Action: PPD Hard disk spare part 00000000800000000500000000000000002 Arranged to go to the hospital 00000000, FSE Cao Peng-86-1810000000062800000000 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that image storage space is low and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Free disk space is too low. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Euecue Database Consistency Test and Repair, delete the Repaired patient, and see if the space is restored. 2. If there is no effect, euecue Recreate Image Store. 3. If there is no effect, try Bypass hard disk, and then reinstall the IPPC software. 00000000If there is no effect, ue PC. Diagnostic Tool performs diagnostic tests and replaces corresponding spare parts 5. Check machine uLLreports and perform faut repairs Recommended spare part uuand name: " 0000000080000000050000000092 IP PC Revised_II no HDD &0000000080000000050000000000000000Hard disk spare part " Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Internal Remarks: 31.1000000002000000003 1000000005:25 uC (8600000000900000000) $Q1: (Y) $Q2: (N) $Q3: (NA) $Q00000000(NA) $Q5: (NA) $Q6:（Image cannot be saved ) $Q7: (Confirmed reported problem, Image cannot be saved) $Q8: (image disk damaged, cannot repair) $Q9: (replaced the 0000000080000000050000000000000000Harddisk spare part, tun normal) $Q100000000 (Test and Inspection data refer to PA. Device retuned to ful fuctionality.) ----------------------------------------------------------------------------- | 2023-10-30 12:34 | No
3063 | Yes | Unable to charge | Unable to charge | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor Safety Qustions: Case: 0000000022511817 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: weiyu Last udated by: wei yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A OneEMS internal remarks: uable to charge, battery and device statu need to be checked on site.Please transfer to buiness for follow-u ----------------------------------------------------------------------------- | 2023-10-30 12:44 | No
3064 | Yes | C arm and Table movement not happening | C arm and Table movement not happening | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000329 00000000-Nov-2000000003 Rarish, PP The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : C arm and Table movement not happening When was the first occurence? Within... : 300000000Oct-2000000003 12:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: C arm and Table movement not happening uer Impact: No uer impcat Patient Impact: No patient impcat uu00000000Diagnostic performed by Engineer : Checked the log and foud AMC Drive failed . Tools & Equpment : MuTIMETER,2000000003-12-00000000 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000800000000990000000051,AMC Triple Servo Drive HP-LP-LP Resoluion : Replaced part with new one . Performed requred calibration. Checked all the fuctions. Handed over the machine in good 000000000000000000000000001king condition ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000526 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : C arm and Table movement not happening When was the first occurence? Within... : 300000000Oct-2000000003 12:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: C arm and Table movement not happening uer Impact: No uer impcat Patient Impact: No patient impcat uu00000000Diagnostic performed by Engineer : Ethercat topology test uLL are there in the logs. It is indicating one of the AMC drive in frontal stand gone defect. Follow u Requred : uLL# and/or description of uLL: C arm and Table movement not happening Repair Action : Ethercat topology test uLL are there in the logs. It is indicating one of the AMC drive in frontal stand gone defect. Technician Name : Vijayakuar, K.S Resoluion : Ethercat topology test uLL are there in the logs. It is indicating one of the AMC drive in frontal stand gone defect. Internal Comments : uLL# and/or description of uLL: C arm and Table movement not happening Repair Action : Ethercat topology test uLL are there in the logs. It is indicating one of the AMC drive in frontal stand gone defect. ----------------------------------------------------------------------------- | 2023-10-30 13:04 | No
3065 | Yes | In ur mode the message appears: CO0ervice required and in service mode:CO0ervice required.The CO0odule does not | In ur mode the message appears: CO0service required and in service mode:CO0service required.The CO0module does not | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000662 00000000-Nov-2000000003 RON FRISSEN The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: no Device ue: No. Eupected and ueupected Resuts: No. uer Impact: No. Patient Impact: no SE uknown Resoluion: CO2 modue defective, replaced. Fuctional check and safety test. Took defective CO2 modue for futher investigation. Water Filter Assembly cannot be bent tightly becaue otherwise an occluion will occu. In a short occluion the CO2 modue will not become defective, bu does mean a fuctional check. Parts Consued: 1.0000000000000000000000000000000056000000001711,Intrepid CO2 Modue ASSY Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000005-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 20000000000000000000000003 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003000000005900000000ENNIE SuNEN The system meets the specifications for the performed service and is retuned to ue Follow u Requred: Please schedue Ron Frissen on 3 November on site to check and possibly repair the intrepid. Please also send the following part to the hospital for the attn. of Rob ten Haaf 1 time 0000000056000000001711 Intrepid CO2 Modue ASSY Patrick already coordinated this via e-mail with Rob ten Haaf (see feeds in the case). Problem Description by Engineer: Cutomer Fuction/Role: no Device ue: No. Eupected and ueupected Resuts: No. uer Impact: No. Patient Impact: no SE uknown Diagnostic performed by Engineer: The CO2 modue is probably defective. Someone will be sent on site for the repair. ----------------------------------------------------------------------------- | 2023-10-30 13:54 | No
3066 | Yes | A message is displayed that the shutter cannot be ud. | A message is displayed that the shutter cannot be ud. | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000819 300000000Oct-2000000003 Mori, Hinata The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: A message is displayed that the shuter cannot be ued. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: A message is displayed that the shuter cannot be ued. uer Impact: No effect Patient Impact: No effect SE R2.2.6 Resoluion: When I visited the site, I confirmed that the problem was the same as it was at the time the problem occured, and that the collimeter was not 000000000000000000000000001king. The problem disappeared after the restart, and I confirmed from the log at the time of the problem that the collimeter was temporarily uuable. I opened the equpment cover, checked the internal wiring and checked to see if the connector was loose, etc., and checked the operation. I confirmed that it can be ued withou a problem. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000792 Saka, Shuji The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: A message is displayed that the shuter cannot be ued. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: A message is displayed that the shuter cannot be ued. uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: RC Response: Remote not connected. Visit and investigate. Internal Comments: [Suject] 00000000DS:SS:u Shiiba: A message is displayed that the shuter cannot be ued. [Internal Comments] 00000000800000000858162 NICOL V3 CV FD Osaka [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Yoko Mori] Mori, Hinata - FSE1 Visit date and time 2000000003/1000000003000000009:00000000 - FSE1 000000000000000000000000001k time 2 [Euamination category] Normal euamination Follow u Requred: ⋆ ----------------------------------------------------------------------------- | 2023-10-30 14:19 | No
3067 | Yes | Stem failure with patients lying under anesthesia | Stem failure with patients lying under anesthesia | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000 November 2, 2000000003 Groenenberg, Frank The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: RR Ms. Rosenstengel Phone +00000000691799659 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Resoluion: The reported faut is confirmed. The following actions were performed for uLLcorrection: Defective detector sensor replaced, adjutment performed, info discusion with Medical Technology. Parts Consued: 1.00000000000000000000000000000000230000000023000000005,SENSOR FD200000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000779 Hellenbrandt, Helmu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: uLL# and/or description of uLL According to log, problems with the detector body gurd: WARNING: DETECTOR bodygurd active According to the system: BG calibrated OK plate[id pu_00000000] offset= 73 BG calibrated OK plate[id pu_00000000] offset= 65 BG calibrated OK plate[id pu_00000000] offset= 53 BG calibrated OK plate[id pu_00000000] offset= 68 Replace detector body gurd Repair Action: Replace detector body gurd Problem Description by Engineer: Problem description by engineer: RR Ms. Rosenstengel Phone +00000000691799659 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: Body gurd problems Euamination was not yet started. Bodygurd adjutment performed with uer - then the euamination coud be started. Delay 35 minues Follow-u requred: uLL# and/or description of uLL According to log, problems with the detector body gurd: WARNING: DETECTOR bodygurd active According to the system: BG calibrated OK plate[id pu_00000000] offset= 73 BG calibrated OK plate[id pu_00000000] offset= 65 BG calibrated OK plate[id pu_00000000] offset= 53 BG calibrated OK plate[id pu_00000000] offset= 68 Replace detector body gurd Repair Action: Replace detector body gurd Technician name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Cutomer comment: Ms. Rosenstengel asked abou the processing statu and asks for call back. / Tel.: +00000000691799659 Cutomer comment: ## Arranged appointment with Ms. Rosenstengel for 11/2 after 10000000000000000/100000000300000000am ## ----------------------------------------------------------------------------- | 2023-10-30 14:54 | No
3068 | Yes | uLLmessage: Complaining about bodyguard function and so they did a restart and then the screen went black. The... | uLLmessage: Complaining about bodyguard funuon and so they did a restart and then the screen went black. The... | 000000000000000000000000001O##: 000000000000000000000000001000000003000000008100000000ALAR JIM The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow-u 000000000000000000000000001k schedued Follow-u Requred: Host PC does not have any power to the HDD / 000000000000000000000000001't start. Either a bad PC or HDD in host PC. Order a new HOST PC and HDD for the cutomer. They can try to ghost back the system to get it to start 000000000000000000000000001king again. The system code is 7220000000000000000Diagnostic performed by Engineer: Host PC does not have any power to the HDD / 000000000000000000000000001't start. Either a bad PC or HDD in host PC. Order a new HOST PC and HDD for the cutomer. They can try to ghost back the system to get it to start 000000000000000000000000001king again. The system code is 7220000000000000000Problem Description by Engineer: Cutomer Fuction/Role: MT Device ue: Host PC 000000000000000000000000001't start, no activity on HDD, no green light/power Eupected and ueupected Resuts: Can't ue system uer Impact: NA Patient Impact: NA SE8.1.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003000000001900000000Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- | 2023-10-30 15:19 | No
3069 | Yes | System does not start, PC remains at stand still during start-up | System does not start, PC remains at stand still during start-up | 000000000000000000000000001O##: 000000000000000000000000001000000003000000002100000000ct 31, 2000000003 Schmitz, Ingo The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MTA Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE 8.1 Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000006-15 00000000:00000000:00000000 MuTIMETER, 200000000000000006-1000000000000000000000000:00000000 Parts Consued: 1.00000000000000000000000000000000800000000000000007813,GEO Q35 IPC COA 1.00000000000000000000000000000000800000000000000002323,FRuSpare Fue Bou M-Cab Resoluion: The reported faut is confirmed. The following actions were performed for uLLcorrection: GEO IPC replaced and installed/configued, host PC restored from backu/ghost and comprehensive fuction check performed. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003000000001300000000Oct 300000000 2000000003 Schmitz, Ingo The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: The system cannot be started. The R cabinet is withou power suply. MPDufues F2 and F12 triggered. Follow-u requred: Replace GEO IPC Problem Description by Engineer: Cutomer Fuction/Role: MTA Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE 8.1 Diagnostic performed by Engineer: - M cabinet checked - MPDuchecked - R cabinet checked - GEO IPC checked The trouleshooting revealed that the power suply of the GEO IPC is defective and leads to the MPDufues triggering F2 and F12. Geo IPC mut be replaced. A follow-u appointment with spare part is requred. Tools & Equpment: MuTIMETER, 200000000000000006-1000000000000000000000000:00000000 ----------------------------------------------------------------------------- Cutomer comment: u Schmitz: Follow-u service on 31.100000000 9:00000000 am + part Cutomer comment: # u uer has INFO abou service today approu. 8:300000000am. # Cutomer comment: u uer will call and requst a call-back at tel. 0000000000000000512000000001772 Will a technician come directly today? Internal Remarks: Via u Schmitz / PH ----------------------------------------------------------------------------- | 2023-10-30 15:19 | No
3070 | Yes | Request for corrective intervention. | Request for correu intervention. | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000000000002-Nov-2000000003 Samule, Franchetti The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Replaced fleuvision PC and reloaded software of the entire system. Fuctional tests were carried ou with a positive oucome. The system is 000000000000000000000000001king for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000718000000002,FleuViewing PC 00000000 W7 + W100000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000312000000005 00000000-Nov-2000000003 Samule, Franchetti The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Performed diagnosis: fauty power suply fleuvision PC. The part was ordered. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003000000000000000000000000Samule, Villa The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: necessary visiton-site with part 00000000800000000718000000002 - FleuViewing PC 00000000 W7 + W1000000001 FSE 00000000Problem Description by Engineer: Cutomer Fuction/Role: u Device ue: no patient involved Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE n.a. Diagnostic performed by Engineer: as requsted by the colleagu the following part is ordered 00000000800000000718000000002 - FleuViewing PC 00000000 W7 + W100000000to replace PC ----------------------------------------------------------------------------- Internal Remarks: Order part nuber: 00000000800000000718000000002. ----------------------------------------------------------------------------- | 2023-10-30 15:54 | No
3071 | Yes | Remote Alert: RMT - ur guidance: Image storage is not possible becau of an image disk problem. | Remote Alert: RMT - ur guidance: Image storage is not possible becau of an image disk problem. | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000585 Adrián Sancho López The system meets the specification for the performed service and is retuned to ue Internal Comments : uLL# and/or description of uLL: N/A Repair Action : Attended to on case 000000002200000000586. Resoluion : Attended to on case 000000002200000000586. Diagnostic performed by Engineer : Attended to on case 000000002200000000586. Problem Description by Engineer : Problem description by engineer : Attended to on case 000000002200000000586. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicina Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No impact uer Impact: No impact Patient Impact: Sin impacto SE 1.2.5.00000000Reguatory Qustions : N/A ----------------------------------------------------------------------------- | 2023-10-30 16:24 | No
3072 | Yes | unable to boot up. | unable to boot up. | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000055 00000000-Nov-2000000003 Mahmou, Mari The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE RIS and PACS configuation requred, Remote connectivity is not 000000000000000000000000001king Diagnostic performed by Engineer : - RIS and PACS need to be configued. - Remote connection need to be configued. Resoluion : - RIS and PACS have been configued. - Remote connectivity has been established. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000060000000000000000-Nov-2000000003 Dharmaraj Thaiappan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uable to boot u. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: None Patient Impact: system uable to ue SE 8.2.3000000005 Diagnostic performed by Engineer : Foud Host PC with Hard drive defective and replaced the same. Performed software installation and carried ou all the requred calibration. Handed over the system for clinical uage. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.00000000000000000000000000000000800000000122391,u Haswell Monoplane Host PC no HDD Tools & Equpment : DOSE-KVP-LIGHT METER,2000000008-0000000030000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000002-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003300000000000000003-Nov-2000000003 Mahmou, Mari The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE System stops at 00000000:00000000 and doesn't log in Diagnostic performed by Engineer : done by another fse ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000 300000000Oct-2000000003 Aneesh Nair The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace HOST PC. Please send Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000Qty 00000000 from local stock along with the ordered PC. Diagnostic performed by Engineer : Cutomer informed that the system is not booting u. Diagnosed the system and foud that Host PC is defective. need to order and replace the PC , the PC will be be shipped from SPS NL system is down. Problem Description by Engineer : Problem description by engineer : uable to boot u. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: None Patient Impact: system uable to ue SE 8.2.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000191 Roble Ramanan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : uable to boot u. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: None Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer : uable to boot u. Internal Comments : uLL# and/or description of uLL: 00CC Move from local stock u Haswell Monoplane Host PC no HDD 00000000800000000122391 Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000Qty 00000000 Repair Action : Supect issu with Host PC software. reload backu Follow u Requred : uLL# and/or description of uLL: 00CC Move from local stock u Haswell Monoplane Host PC no HDD 00000000800000000122391 Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000Qty 00000000 Repair Action : Supect issu with Host PC software. reload backu Technician Name : Roble Ramanan ----------------------------------------------------------------------------- | 2023-10-30 16:24 | No
3073 | Yes | Suainability Issue_Failed to deliver shock | Suainability Issue_Failed to deliver shock |  | 2023-10-30 16:42 | Yes
3074 | Yes | PSA_Legal_uCapnoTrak ETCO0ccessories - with 0iece coniguration_Enhmnt | PSA_Legal_uCapnoTrak ETCO0accessories - with 0piece coniguration_Enhmnt |  | 2023-10-30 17:20 | No
3075 | Yes | uappeared to double count the HR | uappeared to double count the HR | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000925 Marc Chambers The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Prof Richard Lyon MBE - Consutant in Emergency Medicine Device ue: In ue Eupected and ueupected Resuts: uappeared to doule cout the HR, despite it being a good trace. Clinically the patient had a puse of approu 12000000001300000000 The patient was on a traua sats probe - even with this the sats trace wasn’t good enouh to get the HR off and NIBP 000000000000000000000000001ldn’t pick u uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : RDT Buhave requsted that the end uer [rovide device logs ASAP so that the details can be reviewed in ful. ----------------------------------------------------------------------------- Internal Remarks : RDT Produt Specialist awaiting logs from end uer. Once received, and action plan on how to progress received, then futher 000000000000000000000000001ks can be planned if requred. Internal Remarks : Event Date - 21:25 200000000h October 2000000003 Clinical Qustions: 1. What was the medical condition of the patient? - . Major traua - supected internal bleeding. Pedestrian hit by car. 2. What was the AGE of the patient? - 28 Male 3. Was a Tempu LS defibrillator ued together with the Tempu Pro at the time of the event? - Tempu Pro Only 00000000How did the event impact patient care (e.g. no impact, delay of life-saving therapy, inability to diagnose, wrong diagnosis leading to inappropriate treatment) - Difficuty in monitoring patient in flight. Difficut in making decision on blood transfuion. Difficuty diagnosing bleeding 5. How was the reported event managed (e.g., another monitor was ued)? - Clinical jugment uing patient’s radial puse 6. Did delay a care with a deterioration in the patient's state of health occu becaue of the reported event? - No ----------------------------------------------------------------------------- | 2023-10-30 17:29 | No
3076 | Yes | Does not recognize battery | Does not recognize battery |  | 2023-10-30 17:46 | No
3077 | Yes | Device continued operating or became inoperative at the time of phenomenon: | Device continued operating or became inoperative at the time of phenomenon: |  | 2023-10-30 18:10 | No
3078 | Yes | 00Y0osition the patient is on is wrong and cant change to fluoro | 00CDY0position the patient is on is wrong and cant change to fluoro | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000118 300000000Oct-2000000003 Stewart, Aleu The system meets the specification for the performed service and is retuned to ue Resoluion : Adjuted Pivot Potentiometer voltage in the base of the AD7 table. Performed necessary cals and retuned to service. Problem Description by Engineer : Problem description by engineer : RMT - POS: SAM_TblPivot Not Calibrated 300000000Oct-2000000003 00000000:58:00000000 RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport 300000000Oct-2000000003 00000000:57:55 When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: in clinical ue Eupected and ueupected Resuts: Eupected image to be correct orientation, however image is sideways uer Impact: uable to ue system as image is side to side instead of top to bottom Patient Impact: Patient was moved to another room SE 2.2.6.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000065 Van Pelt Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - POS: SAM_TblPivot Not Calibrated 300000000Oct-2000000003 00000000:58:00000000 RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport 300000000Oct-2000000003 00000000:57:55 When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: in clinical ue Eupected and ueupected Resuts: Eupected image to be correct orientation, however image is sideways uer Impact: uable to ue system as image is side to side instead of top to bottom Patient Impact: Patient was moved to another room SE 2.2.6.00000000Diagnostic performed by Engineer : Confirmed uers have rebooted system, problem persists after reboot. Contacted Clinical Applications Suport to discus issu. Reviewed log files and foud uLL. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - POS: SAM_TblPivot Not Calibrated: 1. The Pivot movement of the Patient Suport is not calibrated. Root Caue for RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport: 1. u ray PC Malfuction. 2. Grabber card malfuction of the Fleuvision PC. 3. Defective DVI cable, Converter Repair Action : Service actions for RMT - POS: SAM_TblPivot Not Calibrated: 1.1. Carry ou Pivot movement calibration of the Patient Suport via PSC. Service actions for RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport: 1.1. Check the additional information in the uLL abou channel no. from which image frames are missing. If image frames reported from channel 11 to15 most likely euternal Souce image uray PC is malfuctioning. 2.1. Check the Grabber card statu in the fleuvision PC. Load fleuvision pc software. 3.1. Bad DVI cable, converter of the respective channel leads to missing image frame. replace defective cable or converter Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - POS: SAM_TblPivot Not Calibrated: 1. The Pivot movement of the Patient Suport is not calibrated. Root Caue for RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport: 1. u ray PC Malfuction. 2. Grabber card malfuction of the Fleuvision PC. 3. Defective DVI cable, Converter Repair Action : Service actions for RMT - POS: SAM_TblPivot Not Calibrated: 1.1. Carry ou Pivot movement calibration of the Patient Suport via PSC. Service actions for RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport: 1.1. Check the additional information in the uLL abou channel no. from which image frames are missing. If image frames reported from channel 11 to15 most likely euternal Souce image uray PC is malfuctioning. 2.1. Check the Grabber card statu in the fleuvision PC. Load fleuvision pc software. 3.1. Bad DVI cable, converter of the respective channel leads to missing image frame. replace defective cable or converter Technician Name : Van Pelt Daniel ----------------------------------------------------------------------------- Internal Remarks : Consuted with Dan trying to determine caue of rotated image. Not an applications issu. Onsite service call was dispatched. Internal Remarks : Interaction Type: New Case Caller: JAMES MILLER IP & Issu: position the patient is on is wrong and cant change to fluro Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-30 19:29 | No
3079 | Yes | Touchscreen not responding properly - Sent 0 GFE 00 | Touchscreen not responding properly - Sent 0d GFE 0/0 | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Parts ID - Replacement uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 00000000561539000000001 - RP-Touh Screen,V600000000- 990000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states touhscreen isn't responding properly and is requsting part nuber and price for replacement touhscreen. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Laith Qaqish IP & Issu: 77865500000000/ Parts ID - replacement Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-30 19:34 | No
3080 | Yes | SA:0/u FD0Machine cannot start | SA:0R/u FD0Machine cannot start | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Dear Engineer DuQunwei has been contacted 13720000000000000000100000000 doctor Teacher QuThe First Affiliated Hospital of ui'an Jiaotong uiversity School of Medicine 56962225 The machine cannot start, secuity issu Y N Best Regards jiwenjin SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 3000000001000000002000000003 00000000:10000000000000000 uC (86000000000000000013) Service content: blu screen at boot, software failue, software reinstallation Service resuts: After repair, the equpment retuned to normal operation Problem Reported by Cutomer: Machine cannot start Safety Qustions: Case: 0000000022512231 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: nannan wang Last udated by: nannan wang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 3000000001000000002000000003 00000000:20000000000000000uC (86000000000000000013) $Q7:FSE confirmed reported problem system can't boot u $Q8:Visul checked the uLLinforamtion; juged the software was broken $Q9: Reinstall the software $Q100000000 system t is ok. Test and inspection data refer to PA, device retuned to ful fuctionality $Q11: Calibrated Tools name: N/A SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-30 19:39 | No
3081 | Yes | Defibrillator reports that battery test did not pass during functional check. | Defibrillator reports that battery test did not pass during funuonal check. | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000259 100000000ov-2000000003 CHIARA BERTACCINI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in monitoring Eupected and ueupected Resuts: battery no signal uer Impact: no Patient Impact: no SE na Parts Consued: 1.0000000000000000000000000000000098000000000000000000000000DFM100000000 Lithiu Ion Battery Resoluion: Performed troule shooting and diagnosis, replaced battery, ran tests, everything is ok. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000500000000STEFANO ZAMBONI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: I spoke with u Alugi, the device fails the fuctional test on the battery test. He requsts a visit from the technician to check the fuctioning. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in monitoring Eupected and ueupected Resuts: battery no signal uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Spoke with u Alugi, the device fails the fuctional test on the battery test. He requsts a visit from the technician to check the fuctioning. ----------------------------------------------------------------------------- | 2023-10-30 19:44 | No
3082 | Yes | SPLIT- CO0oor is broken off- See Reference PR 00 | SPLIT- CO0door is broken off- See Reference PR 000 |  | 2023-10-30 20:26 | No
3083 | Yes | Remote Alert: RMT - POS: uLLon [movement=DetectorRotation] - current error | Remote Alert: RMT - POS: uLLon [movement=DeteurRotation] - current error | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000728 00000000-Nov-2000000003 ROBERTO MuTAS The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: RMT - POS: uLLon [movement=DetectorRotation] - curent uLL How often is issu occuring? : Intermittent Malfuctioning area: Geometry Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: Diagnostic Eupected and ueupected Resuts: RMT - POS; uLLon [movement=DetectorRotation] - curent uLL uer Impact: no Patient Impact: no SE 7.2.16.00000000Resoluion: Performed calibration of curent movements of thettor. Fuctional tests. Everything is reguar. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000671 DAVIDE VECCHI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: RMT - POS: uLLon [movement=DetectorRotation] - curent uLL How often is issu occuring? : Intermittent Malfuctioning area: Geometry Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: Diagnostic Eupected and ueupected Resuts: RMT - POS; uLLon [movement=DetectorRotation] - curent uLL uer Impact: no Patient Impact: no SE 7.2.16.00000000Diagnostic performed by Engineer: following log analysis it was decided to check as per indications on case at the same time as the neut PM Internal Comments: uLL# and/or description of uLL RMT - POS: uLLon [movement=DetectorRotation] - curent uLL Repair Action: Radar Follow u Requred: uLL# and/or description of uLL RMT - POS: uLLon [movement=DetectorRotation] - curent uLL Repair Action: Radar What are the neut Trouleshooting steps: Carry ou checks and calibrations as per instrutions. 1 FSE 1 hou The system is not down. Technician Name: DAVIDE VECCHI ----------------------------------------------------------------------------- | 2023-10-30 20:34 | No
3084 | Yes | Touched position was observed to be out of alignment | Touched position was observed to be out of alignment |  | 2023-10-30 20:53 | No
3085 | Yes | Touch screen is malfunctioning | Touch screen is malfunuoning | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005300000000Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced Display Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king additionally there are streaks in the display itself. to correct the issu we will replace the Display Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000300000000Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king additionally there are streaks in the display itself. to correct the issu we will replace the Display Assembly. Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000621 Serra, Rafael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Device ue:Not in Clinical ue Eupected and ueupected Resuts:uable to ue device uer Impact:No impact Patient Impact:No impact SELatest s/w Problem Description/ueupected Resuts: Touh screen is malfuctioning Diagnostic performed by Engineer : Problem Description: Touh screen is malfuctioning Touh screen is malfuctioning Form ued: Bench Trouleshooting Steps Taken: Touh screen is malfuctioning Cutomer reported touh screen uable to calibrate Will send in to repair and requst loaner Resoluion : Sent to Bench Follow u Requred : Touh screen uable to calibrate ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jeremy West IP & Issu: 91880000000072 Touh screen is malfuctioning Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Requsted loaner was shipped to Cut with FedEu Tracking# 71000000001755832. QuST SO# 20000000093 CASE#: 0000000022516865 PN: 00000000-10000000000000000 SN: 616158 Railings: Both Cutomer SN: 615296 Tu Johnson Couty Ambuance District 500000000 E Youg Ave Warrensbug, MO 6000000003-1200000000 Attn: Jermey West 66000000007000000005735 Based On Sales Orders 20000000093. ----------------------------------------------------------------------------- | 2023-10-30 21:34 | No
3086 | Yes | Ceiling mounted radiation shield cover fall | Ceiling mounted radiation shield cover fall |  | 2023-10-30 21:41 | No
3087 | Yes | Alarms or Messages/ Unit operational, not alarming outside the patient room/central station/Loose Connection | Alarms or Messages/ Unit operational, not alarming outside the patient room/central station/Loose Conneuon | 000000000000000000000000001O##: 000000000000000000000000001000000003000000005500000000utin, Mike The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Problem Description: not hearing alerts at suveillance station. -iu -instruted to investigate speaker connection. -Biomed foud the standoff speaker cable screws onto was uscrewed. once pu back into place soud was now being heard. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Monitor Patients Eupected and ueupected Resuts: Monitor Patients or Not Monitor Patients uer Impact: None Patient Impact: None SE N/A ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jessica FRENCH IP & Issu: " 100000000212500000000This device is OPERATIONAL, bu it's NOT alarming ouside the patient room/central station" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: PCCI Action Taken: I created a case, CSP, R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-30 22:09 | No
3088 | Yes | Speakers not connecting - Bench repair/ audible sound found at Bench Foulink-s/nu0ther HW /Replaced parts. | Speakers not conneung - Bench repair/ audible sound found at Bench Foulink-s/nu00other HW /Replaced parts. | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000700000000 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Speakers not 000000000000000000000000001king. Diagnostic performed by Engineer : Problem Description: Speakers not 000000000000000000000000001king. Speakers are not 000000000000000000000000001king. Form ued: Bench Trouleshooting Steps Taken: NA Resoluion : Sent to bench. Follow u Requred : Speakers are not 000000000000000000000000001king. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Rob Lu561-00000000-8226 IP & Issu: u156M000000001 Bench Repair: speakers arent 000000000000000000000000001king not connecting Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: created remote 00000000000000000000000000100000000300000000700000000 dispatched to HC-NA-Tier2-PMCC-Cut, Cut-Tech-Tele NGS Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-30 22:14 | No
3089 | Yes | Delay in recognizing PEA arrest on the monitor. No Malfunction-Alarm not acknowledged. Provided info. | Delay in recognizing PEA arrest on the monitor. No Malfunuon-Alarm not acknowledged. Provided info. |  | 2023-10-30 22:54 | Yes
3090 | Yes | Enhancement Request _IGT-S (uand u table mattress does not physically stick to the table | Enhancement Request _IGT-S (uand u table mattress does not physically stick to the table |  | 2023-10-30 23:41 | No
3091 | Yes | Machine and prouimal pressure sensors failed (EC 00 | Machine and prouimal pressure sensors failed (EC 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000315 Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Biomed is reporting replacing the the flow sensor assembly and then uit is now displaying uLL1000000007. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. The cutomer has advised closing the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: Biomed is reporting replacing the the flow sensor assembly and then uit is now displaying uLL1000000007. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 1000000003000000002000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Cutomer advised to remove all power and reconnect the DA-MC cable. Biomed is reporting attempting reseating the cable a few times bu the uLLhas not been cleared. Diagnostic code 1000000007 = Machine and prouimal pressue sensors failed. 1. Replace the DA PCBA to MC PCBA cable. 2. Replace the DA PCBA. 3. Replace the MC PCBA. 00000000Replace the PM PCBA. 5. Replace the CPuPCBA. Advised Biomed to order and replace the data acqusition to motor controller cable. Part # 00000000561536131 RP-Cable, DA to MC PCBA, 2nd Gen, V600000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Wibert 313-576-10000000000000000u 65800000000 IP & Issu: 000000002200000000322 - Pu O2 sensor in trying to do leak test and needs help to determine issu. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-31 00:04 | No
3092 | Yes | 00N0-ray not possible. Tube problem. Contact tech support. | 00CDN0u-ray not possible. Tube problem. Contautech support. | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000590000000031-Oct-2000000003 Holder Antwan The system meets the specification for the performed service and is retuned to ue Resoluion : System 000000000000000000000000001ld not make u-ray. Checked LEDS of the IPS and other parts and there were none. The IPS Power Suply wasn't sending 200000000olts ou to the other parts of the system. Tested the system by making euposues and checking LEDS statu. LEDs were lit and the system didn't fail. Parts Consued : 1.00000000000000000000000000000000210000000020000000000000000,IPS Certeray R5 Problem Description by Engineer : Problem description by engineer : The cutomer is uable to enable the u-ray. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: Room is down Patient Impact: Patient had to be moved to a different room. uu00000000Tools & Equpment : MuTIMETER,2000000005-00000000-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000511 Sharvari Karua The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : The cutomer is uable to enable the u-ray. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: Room is down Patient Impact: Patient had to be moved to a different room. uu00000000Diagnostic performed by Engineer : The tech was uable to enable u-ray. Downloaded the logs and there is a commuication loss between the u-ray PC and generator. Internal Comments : uLL# and/or description of uLL: 200000000SIEC0000000000000000000000001 uer gudance: u-ray not possible. Tue problem. Contact technical suport. The u-Ray PC coud not open or lost connection with the Frontal Certeray generator. Repair Action : 30000000000000000600000000600000000 - uSC Certeray Resoluion : Coud not resolve the issu and requres onsite suport. Follow u Requred : uLL# and/or description of uLL: 200000000SIEC0000000000000000000000001 uer gudance: u-ray not possible. Tue problem. Contact technical suport. The u-Ray PC coud not open or lost connection with the Frontal Certeray generator. Repair Action : 30000000000000000600000000600000000 - uSC Certeray What are the neut Trouleshooting steps : 1) Check the power to the generator. 2) Check 200000000to uSC. If present, uSC might be defective. Technician Name : Sharvari Karua ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Holder Antwan with trouleshooting no comms to uSC from Sute PC. FSE will have uSC for trouleshooting and also sugested FSE check 200000000C coming from IPS as possible issu. 00000000210000000020000000000000000 - IPS Certeray R5 ----------------------------------------------------------------------------- | 2023-10-31 00:09 | No
3093 | Yes | 00T0-ray not available | 00CDT0-ray not available | 000000000000000000000000001O##: 0000000000000000000000000010000000032000000009 00000000-Nov-2000000003 Brown Mike The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000230000000020000000000000000IBRE CABLE FOR RTO OPT. Tools & Equpment : DOSE-KVP-LIGHT METER,2000000005-10000000031 00000000:00000000:00000000 SCOPEMETER,200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-10000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000000000000031 00000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : u-ray not available When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Reboot does not resolve uer Impact: uable to continu with clinical applications Patient Impact: delay in proceedue uu00000000Resoluion : Puled new fiber cable into stand and throuh cable trouh. System u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000031600000000 00000000-Nov-2000000003 Brown Mike The system does not meet the specification for the performed service and is not retuned to ue Parts Consued : 1.000000000000000000000000000000008000000000000000000000000FDCSIB 1.0000000000000000000000000000000080000000000000000900000000,PDuDC Modue Tools & Equpment : SCOPEMETER,200000000000000000000000031 00000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : u-ray not available When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Reboot does not resolve uer Impact: uable to continu with clinical applications Patient Impact: delay in proceedue uu00000000Diagnostic performed by Engineer : Installed SIB and Power suply. Foud issu was still present. Checked leds on FD and confirmed fiber link not present. Checked fiber and only one link was operational. Ordered new fiber cable Follow u Requred : Install new fiber for detector ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000828 31-Oct-2000000003 Brown Mike The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : u-ray not available When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Reboot does not resolve uer Impact: uable to continu with clinical applications Patient Impact: delay in proceedue uu00000000Diagnostic performed by Engineer : Foud system 000000000000000000000000001ld not make u-ray, futher investigation of logging showed the collimator and FD were both showing uLL. Checked the FDSIB (MDY) and foud troule lights on, not passing post. Ordered FSSIB for neut day. Follow u Requred : Install new part neut day Tools & Equpment : SCOPEMETER,200000000000000000000000031 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000579 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : u-ray not available When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Reboot does not resolve uer Impact: uable to continu with clinical applications Patient Impact: delay in proceedue uu00000000Diagnostic performed by Engineer : John says there is no u-ray ready signal. Rebooting doesn't resolve. Advised that RSW is reporting defective fan and power tray. Called and left messages to cell regarding F1 in the power tray as possible caue of failue. Internal Comments : uLL# and/or description of uLL: Pattern Name: RMT - PWR: Defective Fan and Power Tray Root Caue 1: Mutiple parts are failing simutaneouly: SIBBOu, u's, collimator, FDC and detector. Most likely the Fan and Power tray is defective. 200000000c is missing. Service Action 1: Replace Fan and Power Tray (NY-FPT) located in the M-cabinet. Repair Action : Inspect logs and rsw to reccomend to local fse Follow u Requred : uLL# and/or description of uLL: Pattern Name: RMT - PWR: Defective Fan and Power Tray Root Caue 1: Mutiple parts are failing simutaneouly: SIBBOu, u's, collimator, FDC and detector. Most likely the Fan and Power tray is defective. 200000000c is missing. Service Action 1: Replace Fan and Power Tray (NY-FPT) located in the M-cabinet. Repair Action : Inspect logs and rsw to reccomend to local fse What are the neut Trouleshooting steps : Check f1 on fan/power tray. Check fues for NY5,16, 9. Reccomended 0000000080000000000000000900000000 - PDuFanTray and DCPS Perform FCO72200000000529 while on site as only takes minues. Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : FDCSIB uLL on boot-u lead to redownload attempt of software to FDCSIB. Reinstall system device failed to FDCSIB. Approved for replacement 000000008000000000000000000000000- FDCSIB Internal Remarks : FSE installed a new FDCSIB and is getting the same issu with the flat detector failings its BIST and the uLLlog has collimator uLL. Had FSE look at the LED's on the detector and the D2 Auoa link is blinking, which means it is not synchronized. FSE shined phone light throuh the cable and was able to see light on the red fiber bu not the green fiber, so it seems the cable might be bad. FSE will ru a new fiber across the floor and see if this corrects anything. -Approved: Fiber cable for TTO OPT: 00000000230000000020000000000000000---------------------------------------------------------------------------- | 2023-10-31 00:54 | No
3094 | Yes | M0A - u0- Failed to apply shock in the operational test | M00A - u00 - Failed to apply shock in the operational test |  | 2023-10-31 01:25 | No
3095 | Yes | Grounding test failing on the O0itting retention plate | Grounding test failing on the O0fitting retention plate | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000989 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: duing PM Eupected and ueupected Resuts: grouding test failing on the o2 fitting retention plate uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order the o2 fitting retention plate and the screws, and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and cannot get the groud to pass at the o2 fitting retention plate. Part #00000000561535731, plate, o2 fitting retention, list price $79.00000000 Part #00000000561500000000651, screws 100000000pack, list price $163.00000000 ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Arthu Lacy - phone 619 700000000000000006 IP & Issu: 7000000000000000078 - Cutomer requsting suport - PARTS ID : plate, O2 fitting retention Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-31 01:39 | No
3096 | Yes | Tidal Volume Issue - in progress / onsite unscheduled | Tidal Volume Issue - in progress / onsite unscheduled | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000762 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : while on patient vent not suplying requsted volues. Cutomer requets onsite repair. Perform ful PVT to diagnose. Possible flow issu. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient Eupected and ueupected Resuts: not reaching set volues uer Impact: swapped vent Patient Impact: vent was swapped with no issu SE uknown Diagnostic performed by Engineer : Cutomer requts onsite suport to checkou the vent and repair as needed. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Michael Gladfelter +1(585) 723-7892 IP & Issu: "100000000000000006957 not achieving at title volmes" Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Created R000000000000000000000000001 roued to Technical Suport Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-10-31 01:59 | No
3097 | Yes | Audio-No Sound / No Audio found at Bench- Foulink s/n u00 Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at Bench- Foulink s/n u000 / Other HW-Speaker-HW / Replaced Parts |  | 2023-10-31 03:17 | No
3098 | Yes | Audio-No Sound / No Audio found at bench-Foulink s/n u0/ Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at bench-Foulink s/n u000/ Other HW-Speaker-HW / Replaced Parts |  | 2023-10-31 03:19 | No
3099 | Yes | Audio-no sound/No audio found at Bench-Foulink s/n u0other HW-Speaker-HW/Replaced. | Audio-no sound/No audio found at Bench-Foulink s/n u00/other HW-Speaker-HW/Replaced. |  | 2023-10-31 03:21 | No
3100 | Yes | Audio-no sound/No audio found at Bench- Foulink-s/n u00ther HW-Speaker-HW/Replaced Parts | Audio-no sound/No audio found at Bench- Foulink-s/n u000/other HW-Speaker-HW/Replaced Parts |  | 2023-10-31 03:23 | No
3101 | Yes | A touched position was observed to be out of alignment | A touched position was observed to be out of alignment |  | 2023-10-31 05:09 | No
3102 | Yes | SA:0/u FD0System software runs slowly | SA:0R/u FD0System software runs slowly | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 00000000:26:29 uC (8600000000000000007)[Reload the host and IPPC system, the machine 000000000000000000000000001ks normally] Service resuts (requred): [After this service, the equpment will retun to normal ue] Other important matters: [N/A] SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 0000000080000000050000000092 *1, spare parts are ou of stock, have been ordered, please dispatch FSE Liao Min-86-18600000000316500000000T2 Activities:Repair Action: Alternative nuber confirmation: 0000000080000000012200000000 u Haswell IP PC no HDD T2 Activities:Repair Action: PPD/ 0000000080000000050000000092 IP PC Revised_II no HDD T2 Activities:Problem Description: Overview of cutomer problems: The device software rus slowly and freezes, and the software needs to be reinstalled.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: The device software rus slowly and freezes, and the software needs to be reinstalled.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Plu and uplu the Host PC host board, clean the host dut, replace the host BIOS battery, and reinstall the host software Notes on cutomer requrements: Onsite Safety Qustions: Case: 00000000225127200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 00000000:26:32 uC (8600000000000000007) RDI Accuacy: "Yes" $Q7: (Confirmation & Description of problem) : confirmed, system hug u intermoittent duing power on or operation $Q8: (Trouleshooting & Resuts): check system logfile and hospital information, might system and IPPC SW issu, reload system whole SW, re-load IPPC SW $Q9: (Repair steps):restore system and reload IPPC, system 000000000000000000000000001k normal $Q100000000 (Test and Inspection data & Device statu) test and inspection data refer to: PA, Visul check: pass, system retuned to fulfuction $Q11: Calibrated Tools name S/N Neut Calibration date: NA ----------------------------------------------------------------------------- | 2023-10-31 06:04 | No
3103 | Yes | 0 - u0- Operation check failure | 00- u000- Operation check failure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Qulity inspection failed Safety Qustions: Case: 00000000225127600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-31 06:24 | No
3104 | Yes | 0/uLuCentron/System cannot bootup normally | 0R/uLuCentron/System cannot bootup normally | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 3000000001000000002000000003 11:00000000:00000000uC (86000000008200000000 Service content: The system cannot be tuned on. Plu and uplu the board and install the software. Service resuts: After repair, it was confirmed that all fuctions were operating normally and the equpment retuned to normal operation. "SuT. Ready for Clinical ue local time" : " 2000000003-10000000030000000010000000000000000 Neut Step Action: PPD spare parts: 00000000800000000122391*1, timeliness: 12H shipped FSE: Yu Jiangmin-86-18630000000000000000181 T2 Activities:Problem Description: Overview of cutomer problems: After the device is tuned on, the content on the data screen goes to the ref screen, and other screens are black.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: Host PC display failue.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Plu and uplu the graphics card of the Host PC, clean the dut inside the host, reinstall the host software, and replace the host. Recommended spare parts uuand designation: " 00000000800000000122391 u Haswell Monoplane Host PC no HDD" Notes on cutomer requrements: Onsite Problem Reported by Cutomer: uable to boot Safety Qustions: Case: 0000000022512771 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 3000000001000000002000000003 11:00000000:500000000uC (86000000008200000000 "RDI Accuacy:Yes" $Q7: (Confirmed reported problem, the system did not bootu) $Q8: (the startu program display on reference monitor, Other monitors show a black screen, maybe the software corruted or graphics card on host pc issu make the startu program failed) $Q9: (re-plu the graphics card on host pc and resore the image) $Q100000000 (Test and inspection data refer to PA, performing the cardiac scanning, perfectedimages display on monitors, Device retuned to ful fuctionality) $Q11:NA For Mainland China CV / uR, CT engineers only.Please fill in: ^Curent Tue installed date: ^Curent Tue uage: scan second_ ----------------------------------------------------------------------------- | 2023-10-31 06:29 | No
3105 | Yes | Unable to discharge. | Unable to discharge. | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ASP jinuu follow u Neut Step Action: Dear, GuQiliang, 173 176000000006538 Safety Qustions: Case: 0000000022513000000006 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Shanghai Changhai Hospital No discharge Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-31 08:14 | No
3106 | Yes | QA 000000- Device showed malfunction during self-test | QA 0000000 - Device showed malfunuon during self-test |  | 2023-10-31 09:07 | No
3107 | Yes | QA 000000 DEVICE CANNOT TURN ON | QA 00000000 - DEVICE CANNOT TURN ON |  | 2023-10-31 09:14 | No
3108 | Yes | Device can't power on | Device can't power on | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Dear, ZhouYonguiang, 13776600000000592 Problem Reported by Cutomer: JiangsuMidea Clean Electric Co., Ltd. Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: ASP Honglejia follow u Safety Qustions: Case: 00000000225133500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-31 09:29 | No
3109 | Yes | Direct line, no sound on the monitor, pas de sonnerie avec le console, ICU | Direuline, no sound on the monitor, pas de sonnerie avec le console, ICU | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000600000000Rioseco Inst. The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Clinical Eupected and ueupected Resuts: No soud on monitor. uer Impact: Same Patient Impact: n/a SE uk Diagnostic performed by Engineer: Afterhous callback. Cutomer reported no soud on the monitor. Biomed requsting onsite suport. Dispatched FSE as requsted. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Follow u Requred: No auible alarm, need additional trouleshooting. ----------------------------------------------------------------------------- Neut Step Action: Hi Sonaly, Kamel only needs a remote 000000000000000000000000001 so he can docuent. Pu it as a P3. Kamel will make the necessary to find the root caue. Have a nice day/ Have a G'day! ----------------------------------------------------------------------------- | 2023-10-31 09:49 | No
3110 | Yes | Vent cart not stable | Vent cart not stable | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : CART KEEPS ROCKING Ou Of Bou Failue? : No Labor: Labor ID : L-19860000000000000000000000000000000000000000001k Performed : REPAIRED THE VENT CART Diagonsis : TIFGTHEN ALL SCREWS | 2023-10-31 11:06 | No
3111 | Yes | uray Cannot Euposure | uray Cannot Euposure | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000973 00000000-Nov-2000000003 SILA TuRuA The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad Tech Device ue: in ue Eupected and ueupected Resuts: uray Cannot Euposue uer Impact: uray Cannot Euposue Patient Impact: not impact. SE R2 Resoluion : uray Cannot Euposue (intermittence) caue collimator defect replace new collimator check and test system ok. Parts Consued : 1.00000000000000000000000000000000800000000858162,NICOL V3 CV FD ----------------------------------------------------------------------------- | 2023-10-31 11:44 | No
3112 | Yes | Breathing Circuit is blocked, GFE00OV0 Response received/IS complete | Breathing 0cuit is blocked, GFE0 0NOV00Response received/IS complete | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 00000000.11.2000000003 00000000:00000000:59 uC (86000000001000000005) *Department information: [Department: Equpment Department Contact: Teacher 00 13521675291]* $ASP engineer: [Qiangyi liqiuan] $Q1: (No) $Q2 : (No) $Q3: (No) $Q00000000(No) $Q7: (Confirmed reported problem & The uLLreports that the breathing circut is blocked) $Q8: (Restart the machine, check the device, The uLLreports that the breathing circut is blocked, After replacing the tubine, the machine is 000000000000000000000000001king normally, From this it is juged that there is a problem with the tubine) $Q9:（Replace the tubine） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned to ful fuctionally.) $Q11:Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration date (00000000/11/20000000000000000Calibrated Tools name (TSI000000001H) S/N (000000001100000000000000002) Neut Calibration date ( 10000000000000000/20000000000000000Problem Reported by Cutomer: ASP Qiangyi shi gong 000000003521675291 Call for qutation, mailed Contacted Su Bo 18513000000001300000000for qutation SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022513798 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhang yiheng Last udated by: zhang yiheng 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 00000000.11.2000000003 00000000:00000000:00000000 uC (86000000001000000005) $ASP engineer: [Qiangyi liqiuan] $Q1: (No) $Q2: (No) $Q3: (No) $Q00000000(No) $Q7: (Confirmed reported problem & The uLLreports that the breathing circut is blocked) $Q8: (Restart the machine, check the device, The uLLreports that the breathing circut is blocked, After replacing the tubine, the machine is 000000000000000000000000001king normally, From this it is juged that there is a problem with the tubine) $Q9: (Replace the tubine) $Q100000000 (Test and inspection data refer to :PA checklist.Device retuned to ful fuctionally.) $Q11:Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration date (00000000/11/20000000000000000 Calibrated Tools name (TSI000000001H) S/N (000000001100000000000000002) Neut Calibration date (10000000000000000/20000000000000000----------------------------------------------------------------------------- | 2023-10-31 11:44 | No
3113 | Yes | Geometry restart do not change uLL | Geometry restart do not change uLL | 000000000000000000000000001O##: 00000000000000000000000000100000000310000000090000000031-Oct-2000000003 Kim, Sanggyeom The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Operator Device ue: uing geometry Eupected and ueupected Resuts: geometry problem uer Impact: N/A Patient Impact: N/A SE R8.1.00000000 Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 Resoluion : Checked the log. Checked the geoi pc commuication uLL Reseated the geoi pc cable. Tested the fuction. The fuction was sucessfu. ----------------------------------------------------------------------------- | 2023-10-31 12:04 | No
3114 | Yes | SPLIT_Burning smell coming from uYy monitoring rack. Network-Other-HW. Replaced or upgraded part. | SPLIT_Burning smell coming from uYy monitoring rack. Network-Other-HW. Replaced or upgraded part. | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : Received call from Haverhill 00000000telemetry floor nuse manager, Lisa. Buning smell coming from u monitoring rack and no patients showing u on the central. Staff called the fire department once they smelled smoke. Fire dept pinpointed the smell coming from the patient monitoring rack uing their heat sensor. Travel to Holy Family Haverhill. 00000000central has no patient rhythms and central appears to be in local mode. Trouleshot issu, was able to get the central connected back to the database server. Patient rhythms are still not coming u on central. Noticed some lights are off on the component below the switch, the POE (power over ethernet). Reached ou to Tom Smith and trouleshot. The POE rus the access points and looks like it has buned ou which also caued the buning smell. upluged the POE at requst of nuse manager to prevent any futher buning smell. Tom said to pu in a service call to u Patient Monitoring and when he's dispatched for the call, he will order a replacement compatible switch/POE for overnight delivery. Informed nuse manager, Lisa, and Renee Deyarmin of what appeared to have happened and that they'll be withou telemetry util that POE is replaced. Pu service call into u Patient Monitoring, case# - 0000000022512513. Equpment will have to be entitled before call is dispatched. Ou Of Bou Failue? : N/A Labor: Labor ID : L-198661200000000ork Performed : Received call from Haverhill 00000000telemetry floor nuse manager, Lisa. Buning smell coming from u monitoring rack and no patients showing u on the central. Staff called the fire department once they smelled smoke. Fire dept pinpointed the smell coming from the patient monitoring rack uing their heat sensor. Travel to Holy Family Haverhill. 00000000central has no patient rhythms and central appears to be in local mode. Trouleshot issu, was able to get the central connected back to the database server. Patient rhythms are still not coming u on central. Noticed some lights are off on the component below the switch, the POE (power over ethernet). Reached ou to Tom Smith and trouleshot. The POE rus the access points and looks like it has buned ou which also caued the buning smell. upluged the POE at requst of nuse manager to prevent any futher buning smell. Tom said to pu in a service call to u Patient Monitoring and when he's dispatched for the call, he will order a replacement compatible switch/POE for overnight delivery. Informed nuse manager, Lisa, and Renee Deyarmin of what appeared to have happened and that they'll be withou telemetry util that POE is replaced. Pu service call into u Patient Monitoring, case# - 0000000022512513. Equpment will have to be entitled before call is dispatched. Travel home. Diagonsis : Not Answered | 2023-10-31 12:20 | No
3115 | Yes | M0A - u00 Unable to boot up with only battery | u Unable to boot up with only battery | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022510000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: ASP changqingkaihe follow u Neut Step Action: Received Engineer Wang uingfeng 15127267100000000 Problem Reported by Cutomer: 00000000G/HeartStart uL Defibrillator/Monitor Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-31 12:54 | No
3116 | Yes | No voice prompts | No voice prompts | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002251000000001 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Cutomer’s euplanation of the problem ### No voice prompts, other devices are normal The cutomer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### After upluging and inserting the battery to initialize the machine, the faut persists. Repair Action: T2 Recommended Field Diagnostic Steps### Recommend on-site inspection T2 Recommended spare part uuwith description### na Cutomer demand### na Test/Inspection: na Test Equpment: na SAP Transfer: Case Transferred to SAP Neut Step Action: 1392800000000000000007 Call to inqure abou the progress of the repair application, and have ##d ASP Honglejia Neut Step Action: The equpment is being ordered. There is no reply for the time being when asking the ordering team abou the ordering period on the 2nd. I have jut ##d lui again to confirm that the ordering period is in place. Please be informed. Thank you Problem Reported by Cutomer: Zhonghai-Flur Heavy Indutry Co., Ltd., Jinwan District, Zhuai City, Gungdong Province ----------------------------------------------------------------------------- | 2023-10-31 13:19 | No
3117 | Yes | u0Control panel arm is broken off./ *There are NO case lines and there is NO escalation for this case in SF | u0Control panel arm is broken off./ *There are NO case lines and there is NO escalation for this case in SF | 000000000000000000000000001O##: 0000000000000000000000000010000000031000000006000000003-Nov-2000000003 Ammar Huejeh The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Engineer Device ue: demo Eupected and ueupected Resuts: mechanical issu uer Impact: after Patient Impact: after SE 9.000000005 Control panel arm broke off. Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000003-12-27 00000000:00000000:00000000 Parts Consued: 1.000000000000000000000000000000005618821600000000ssy, Artic Arm, Control Panel Resoluion: Replaced the broken Control Panel Arm with a new one for the system Affinti 300000000 s/n u123E0000000013-NZE2869. All necessary 000000000000000000000000001k has been done according to the factory specifications, the system is 000000000000000000000000001king well. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000310000000078 BAS VERSTEDEN The system does not meet the specifications for the performed service and is not retuned to ue Follow u Requred: Please schedue Ammar to replace the control panel arm. - Requred time: 00000000ous. - Order: 000000005618821600000000rticuating Arm Assembly, Control Panel Problem Description by Engineer: Cutomer Fuction/Role: Engineer Device ue: demo Eupected and ueupected Resuts: mechanical issu uer Impact: after Patient Impact: after SE after Control panel arm broke off. Diagnostic performed by Engineer: Control panel arm needs to be replaced. ----------------------------------------------------------------------------- | 2023-10-31 13:29 | No
3118 | Yes | Request for supply of upatient cable. | Request for supply of upatient cable. | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000566 FALASCA, IVO The system meets the specification for the performed service and is retuned to ue. Follow u Requred: If part cannot be ordered via SPS, proceed differently Make NEMO withou collection at the same time. Tu Osp. Vito Fazzi P.zza Filippo Muatore 73100000000 Lecce (LE) Department: Cardiology 3rd floor - DEA Ac. u Lugi Cavallo/Anna Stifani Tel.000000003266100000000 Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: patient monitoring Eupected and ueupected Resuts: ucable is 000000000000000000000000001n ou and does not 000000000000000000000000001k anymore uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Nuse u. Scollato confirms the need to replace the ucable as it is not 000000000000000000000000001king. Accepts the shipment of the new cable. ----------------------------------------------------------------------------- | 2023-10-31 13:39 | No
3119 | Yes | C-arc cannot be rotated anymore | C-arc cannot be rotated anymore | 000000000000000000000000001O##: 00000000000000000000000000100000000310000000037 November 5, 2000000003 Wilksch, Tim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Called Ms. Pahlke throuh the control center How often is issu occuring? : Continuuly Malfuction area: Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000Resoluion: Bodygurd calibrated, then everything was OK again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000863 Hellenbrandt, Helmu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: Called Ms. Pahlke throuh the control center How often is issu occuring? : Continuuly Malfuction area: Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000Diagnostic performed by Engineer: Problems with geometry Internal Comments: uLL# and/or description of uLL According to the log, the following messages: [Movement=SHIFTDETECTORFRONTAL] uer gudance: The collision sensor is dirty. Perform movements at you own risk. And Application uLL Collision detected [auis=FrtLong] SYMPTOM 1: Stand longituinal movement is restricted or not available at all. CAuE 1: Application uLL Collision detected [auis=FrtLong] SOLuION 1: Read from the logline the auis statu (operation or stand-by) and the position where the collision occured. If the movement hits a mechanical obstrution, this is normal behavior. Otherwise, ru a fuctional test on this specific movement. Possible corrective actions: • Perform curent calibration of movement • Perform force sensor calibration • Clean mechanics or cabling. Repair Action: Trouleshooting on site: Body gurd and longituinal C-arm movement Follow-u requred: uLL# and/or description of uLL According to the log, the following messages: [Movement=SHIFTDETECTORFRONTAL] uer gudance: The collision sensor is dirty. Perform movements at you own risk. And Application uLL Collision detected [auis=FrtLong] SYMPTOM 1: Stand longituinal movement is restricted or not available at all. CAuE 1: Application uLL Collision detected [auis=FrtLong] SOLuION 1: Read from the logline the auis statu (operation or stand-by) and the position where the collision occured. If the movement hits a mechanical obstrution, this is normal behavior. Otherwise, ru a fuctional test on this specific movement. Possible corrective actions: • Perform curent calibration of movement • Perform force sensor calibration • Clean mechanics or cabling. Repair Action: Trouleshooting on site: Body gurd and longituinal C-arm movement Technician name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Internal Remarks: Case was already opened Cutomer comment: u Leberl +0000000015000000000000000031 will call and asks for prompt service. If possible, please already tomorrow. Requst for appointment consutation: Dr. Jacob +000000001500000000761900000000contact. Cutomer comment: No one is reporting uder any of the DWlen. Today no service can be schedued in the holiday area, we schedue for tomorrow and arrange taui delivery. Internal Remarks: Call-ou attempt at the specified nuber - withou sucess /rings throuh Cutomer comment: u Richter has appointment info 00000000.11. from approu. 00000000:300000000AM. Internal Remarks: Ms. Eitz +0000000000000000 5000000000000000031 asks where the technician is. Only reached T. Wilksch on bou. Internal Remarks: Dear Team, Order has been released as requsted. ETA: 00000000/11/2000000003 0000000000000000AM Thank You Best regards & Stay Safe Làszlà Papp Senior Cutomer Service Coordinator Buapest Command Center DHL Suply Chain Internal Remarks: Please ask the cutomer tomorrow morning whether we can get to the system and then inform u Wilksch. ----------------------------------------------------------------------------- | 2023-10-31 13:44 | No
3120 | Yes | M0A - NA - umonitoring lead wire is damaged SPLIT: (See reference PR# 00 | M00A - NA - umonitoring lead wire is damaged SPLIT: (See reference PR# 00) |  | 2023-10-31 14:13 | No
3121 | Yes | System did not go on. | System did not go on. | 000000000000000000000000001O##: 00000000000000000000000000100000000326787 November 9, 2000000003 Adov, Aleuander The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Called Gabriela Kleist Phone +0000000000000000931791 The uLLmessage is confirmed. How often is issu occuring? : Intermittent Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 000Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000009-1000000000000000000000000:00000000 Parts Consued: 1.0000000000000000000000000000000080000000050000000012,Host PC Biplane Revised_II no HDD COA 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD Resoluion: Reported problem: RAM memory in the host PC and image memory lat PC fauty. Both PCs host and image lat. replaced, configuation installed and fuction test performed. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003111700000000November 1, 2000000003 Adov, Aleuander The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: Called Gabriela Kleist Phone +0000000000000000931791 The uLLmessage is confirmed. How often is issu occuring? : Intermittent Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: Reported problem: System does not boot. Trouleshooting performed, system restarted, uLLdoes not occu. uLLmemory read, at the reported time the system was properly powered u, bu no switch-off is visible. According to the cutomer, there were power ouages. Futher uLL in the log, 000000000000000000000000001king memory in the host PC and image memory lat. PC are fauty. Informed the cutomer, follow-u service will be schedued for the replacement of the t000000000000000000000000001PCs. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Follow-u requred: Schedue follow-u service 3 hous (not after 1:00000000 pm, cutomer does not keep the appointments arranged!) and the following spare parts: 0000000080000000050000000022 ClarityIQ ll IP PC 0000000080000000050000000012 Host PC biplane version ll ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000031100000000 Hellenbrandt, Helmu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: uLL# and/or description of uLL Log file withou findings According to the description from Ms. Kleist, the compuer got stuk in bios. Check compuer and Bios battery Repair Action: Check host compuer and Bios battery Follow-u requred: uLL# and/or description of uLL Log file withou findings According to the description from Ms. Kleist, the compuer got stuk in bios. Check compuer and Bios battery Repair Action: Check host compuer and Bios battery Technician name: Hellenbrandt, Helmu Problem Description by Engineer: Problem description by engineer: Called Gabriela Kleist Phone +0000000000000000931791 The uLLmessage is confirmed. How often is issu occuring? : Intermittent Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: System coud not be started After several restarts then everything 000000000000000000000000001ked ----------------------------------------------------------------------------- Cutomer comment: ### Appointment set with Ms. Kleist for 00000000.11. 12:300000000pm. Cutomer comment: ## Arranged appointment with Ms. Buter for 11/9 after 10000000000000000 am. ----------------------------------------------------------------------------- | 2023-10-31 14:29 | No
3122 | Yes | uA0Geometry partly available | uA0 Geometry partly available | 000000000000000000000000001O##: 000000000000000000000000001000000003100000000200000000Oct 31, 2000000003 NICOLAS PEENE The system meets the specifications for the performed service and is retuned to ue Diagnostic performed by Engineer: Visully checked the Geo PC. The front is doing strange things. PC needs to be replaced Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE 8.2 Resoluion: The Geo PC does not 000000000000000000000000001k anymore. Replaced the PC with a new one. After eutensive testing everything 000000000000000000000000001ks normally again. Parts Consued: 1.00000000000000000000000000000000800000000300000000532,GEO PC Q35 CAN card only, COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000310000000000000000PATRICK ATTN. The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : Geometry partly available Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer: Log files are indicating a problem with the Geo IPC Internal Comments: uLL# and/or description of uLL: Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. When problem is persistent it coud be caued du to a defective control net000000000000000000000000001k switch (TS). Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. Bad cabling and/or connections. Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. Issu with GeoIPC OS/AS software. It coud be du to a not correct shudown before, Windows OS can be corruted hence it needs more time in the neut start u. If this happens only sporadically, please ignore it. Else see other possible caues & soluions. Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. Root Caue for RMT - POS: uable to commuicate with GeoIPC. No movements are available (reboot done/needed)..: 1. Please ignore this alert, if it only occus sporadically. Treat it then as a 'one time only'. Root Caue for RMT - POS: uable to commuicate with GeoIPC. No movements are available (reboot done/needed)..: 1. GeoIPC is not booting u at all or fast enouh. As a resut the Host PC is uable to commuicate with the Geo IPC within a specified time. Root Caue for RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement.: 1. Motorized movement uavailable duing Beam Longituinal Frontal movement. If this only occus sporadically and cutomer is not complaining please ignore, else see soluions below. Repair Action: Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check control net000000000000000000000000001k switch (TS) and its power suply (adapter). Replace when necessary. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check the cable RC:u25 to SCC1:u22. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. 1.2. Tighten all the connectors of GeoIPC and SCC1 uit. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. If SW is really corruted please reload GeoIPC software uing uB destroy stick. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start-u of the system. If so, replace GEO IPC. Service actions for RMT - POS: uable to commuicate with GeoIPC. No movements are available (reboot done/needed)..: 1.1. Perform a reboot of the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other Soluion here below. Service actions for RMT - POS: uable to commuicate with GeoIPC. No movements are available (reboot done/needed)..: 1.1. Verify the behavior of GEOIPC at first boot u of the system in the morning (at cold start-u): If inpu power (2300000000) is present bu the GEOIPC does not power u at all, it is caued by a defective internal power suply of GEO IPC itself. 1.2. If GeoIPC does power u fine bu Host PC is uable to commuicate with it in time, it is caued by a defective hard disk. (corruted OS/AS) Then replace the GeoIPC. 1.3. Check whether the cable pins between GeoIPC RC-u15 and IPC adaptor SCC1-u22 are maybe too short. Advise: tighten the RC-u15 and SCC1-u12 connectors or replace the cable between RC-u15 and SCC1-u12. 1.00000000It coud be caued by a defective control net000000000000000000000000001k switch (TS) in the R cabinet or its Power Suply. Service actions for RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement.: 1.1. Clean / degrease the longituinal tracks/rails. 1.2. Check condition of motor ruber wheel. 1.3. Check and/or adjut the tension of the spring. 1.00000000Carry ou the Beam longituinal curent calibration via FSF. 1.5. If problem persists replace the longituinal potmeter SPL R1 1.6. For Clea: Replace Longituinal Drive Assy (SPL uit) 1.7. For Poly G2: If problem still persists replace SPC2 (Eutension board) and/or SPC3 (LuV00000000Resoluion: Geo PC needs to be replaced ----------------------------------------------------------------------------- | 2023-10-31 14:34 | No
3123 | Yes | The live monitor turned off. | The live monitor turned off. |  | 2023-10-31 14:49 | No
3124 | Yes | Email 000 sometimes doesn't u-ray | Email 0000sometimes doesn't u-ray | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000000000003 00000000-Nov-2000000003 Tiago Aguar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Tec. Device ue: yes Eupected and ueupected Resuts: uer Impact: no Patient Impact: no SE Resoluion: Replacement of the MIuuit in the generator and the control board of the cooling uit. Fuctional tests. Operational equpment. Parts Consued: 1.00000000000000000000000000000000210000000015552,PCP Control Circut Cu3100000000 1.00000000000000000000000000000000210000000019971,Mains Interface uit (MIu Certeray Tools & Equpment : MuTIMETER, 2000000003-12-13 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-31 14:59 | No
3125 | Yes | The Defib test fail during OP check. | The Defib test fail during OP check. |  | 2023-10-31 15:04 | No
3126 | Yes | The battery does not work / The defibrillator does not turn on with only the battery. - | The battery does not work / The defibrillator does not turn on with only the battery. - | 000000000000000000000000001O##: 00000000000000000000000000100000000381865 Andrea Ardighieri The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: low battery uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: ou of suport letter was sent to Eng. Matteo Narciso. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000167 FRANCESCO GIuFRIDA The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Low battery we suply in NEMO withou collection 1u 0000000056000000005911 - uuuKit - Lithiu Ion Battery If it is not available, try SIDEM channel. To be shipped. S. Marco Hospital 1 Pignatelli Street Grottaglie Taranto PPI Department Att. Coordinator Ciracì 339/0000000085600000000Problem Description by Engineer: Cutomer Fuction/Role: ppi department Device ue: mutiP monitor Eupected and ueupected Resuts: low battery uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Low battery we suply in NEMO withou collection 1u 0000000056000000005911 - uuuKit - Lithiu Ion Battery ----------------------------------------------------------------------------- Internal Remarks: As per maintenance contract in progress, maintenance technical intervention is requred to be To be carried ou ugently, following a failue on the equpment: DEFIBRILLATOR BIPHASIC MANuL, by u MEDICAL SYSTEMS, uY M 3535 A HEARTSTART uuu u0000000050000000092, inventory 0000000028700000000, located at department 118, Of the presidio P.O. SAN MARCO - GROTTAGLIE, Via G. Pignatelli Grottaglie, that reported Following faut: Battery is not 000000000000000000000000001king and replacement of malfuctioning M3539A power suply. Futher clarifications can be requsted from the department contact: Ciracì 339/0000000085600000000call Global service center: 000000009 9560000000000000000 - e-mail sic.taranto00cspa.it The intervention will have to be carried ou according to the contract in place. When the intervention has been carried ou, this company will have to commuicate also to the address ##areatecnica.ingclinica00sl.taranto.it the resoluion of the faut or, in case of delay Intervention, the reasons that prevented its resoluion, indicating the eupected time for The completed 000000000000000000000000001k. It will be u to this company, moreover, to make sue that in every act in the suject of this order (reports Of 000000000000000000000000001k, invoice, etc.), the requst nuber is indicated precisely or a copy is attached. Of the same order. In the absence of what is mentioned above, it will not be possible, for this Administration, Proceed with the liqudation of the invoices relative to the maintenance contract and/or the order. TECHNICAL MANAGEMENT AREA S.S.D. Clinical Engineering And of the Buiness Informative Services ----------------------------------------------------------------------------- | 2023-10-31 15:24 | No
3127 | Yes | M0A / u00 No test load (Device is EOL and will not be repaired) | M00A / u000/ No test load (Device is EOL and will not be repaired) | 000000000000000000000000001O##: 00000000000000000000000000100000000381892 Andrea Ardighieri The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: low battery uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: ou of suport letter was sent to Eng. Matteo Narciso. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000218 FRANCESCO GIuFRIDA The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Fictitiou 500000000om load and battery to be replaced Nemo withou collection. 1u 989800000000100000000831 - HeartStart 500000000ohm Test Load 1u 0000000056000000005911 - uuuKit - Lithiu Ion Battery If the parts are not available, try SIDEN channel. To be shipped to S. Marco Hospital 1 Pignatelli Street Grottaglie Taranto PPI Department Att. Coordinator Ciracì 339/0000000085600000000Diagnostic performed by Engineer: Fictitiou 500000000om load and battery to be replaced Problem Description by Engineer: Cutomer Fuction/Role: PPI department Device ue: defibbrillator Eupected and ueupected Resuts: 500000000hom load is fauty uer Impact: no Patient Impact: no SE na ----------------------------------------------------------------------------- | 2023-10-31 15:34 | No
3128 | Yes | It isn't alarming when it crosses a threshold and it rebooted itself yesterday. | It isn't alarming when it crosses a threshold and it rebooted itself yesterday. | 000000000000000000000000001O##: 00000000000000000000000000100000000331912 00000000-Nov-2000000003 Simon Goldberger The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Resoluion : Passed cutomer to Lynn Bowers as this 000000000000000000000000001ld be a clincal issu with alarms. All ok now, closing case. Problem Description by Engineer : Cutomer Fuction/Role: Biomedical engineer Device ue: In ue Eupected and ueupected Resuts: Cutomer reporting that uit alarms are not alarming when reaching thresholds uer Impact: N\A Patient Impact: N\A SEN\A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000313222 00000000-Nov-2000000003 Andrew Thomas The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomedical engineer Device ue: In ue Eupected and ueupected Resuts: Cutomer reporting uit restarted and want to know why. uer Impact: No central monitoring Patient Impact: No Central monitoring SEN\A Resoluion : Remove logs from device reported to have restarted rouhly 2am 29/10000000023. Logs attached to 000000000000000000000000001 Follow u Requred : u to report to cutomer why central station restarted at rouhly 2am 29/10000000023. Logs attached to 000000000000000000000000001----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000200000000Simon Goldberger The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomedical engineer Device ue: In ue Eupected and ueupected Resuts: Cutomer reporting uit restarted and want to know why. uer Impact: No central monitoring Patient Impact: No Central monitoring SEN\A Diagnostic performed by Engineer : Sending engineer to gather logs from the uit and investigate why it rebooted. Follow u Requred : Please go to site and pul the logs for the past week from the uit and see if youcan see why it rebooted on Suday 29/100000000Resoluion : Asked Lynn Bowers to speak to the cutomer re the alarm limits which she is doing. Andrew from huo going to site to collect logs and investigate. ----------------------------------------------------------------------------- Neut Step Action : T/C Medical engineering contact uavailable spoke to Richard advised contact not in ##sent to contact to arrange engineer visit. ----------------------------------------------------------------------------- | 2023-10-31 15:39 | No
3129 | Yes | Machine not getting charged and showing message as "service required" | Machine not getting charged and showing message as "service required" | 000000000000000000000000001O##: 00000000000000000000000000100000000310000000023 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Diagnostic performed by Engineer : Defective battery Part No.989800000000200000000781 S.No. C210000000000000000 Dt.2000000001-11-25 to be replaced with same part. Requsted channel to sumit qute for the same. Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A ----------------------------------------------------------------------------- | 2023-10-31 15:59 | No
3130 | Yes | Getting no u-rays, fluoroscopy failed. | Getting no u-rays, fluoroscopy failed. | 000000000000000000000000001O##: 00000000000000000000000000100000000336600000000 16-Nov-2000000003 Brendan Kiely The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Parts Consued : 1.0000000000000000000000000000000022000000000000000072,H.V. transformer 65/800000000100000000 KW 1 tue Resoluion : New HT tank installed to solve no u-ray issu. Job complete. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: Getting no u-rays, fluroscopy failed. uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000336600000000 15-Nov-2000000003 Maslobojevas, Sergeju The system meets the specification for the performed service and is retuned to ue Resoluion : Second man Problem Description by Engineer : Cutomer Fuction/Role: RAD Device ue: uknown Eupected and ueupected Resuts: broke uer Impact: low Patient Impact: low SE Reguatory Qustions : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000336600000000 00000000-Nov-2000000003 Maslobojevas, Sergeju The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Diagnostic performed by Engineer : Second man Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: Getting no u-rays, fluroscopy failed. uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000336537 12-Nov-2000000003 Brendan Kiely The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: Getting no u-rays, fluroscopy failed. uer Impact: Patient Impact: SE Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000007-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000007-31 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000210000000016291,Converter 8E Velara 1.000000000000000000000000989000000000000000085100000000,u 200000000 000000007 ROT-GS 10000000000000000iagnostic performed by Engineer : Tue installation carried ou. The system is still not produing u-ray. HT tank ordered for tomorrow. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000329613 00000000-Nov-2000000003 Nigel Huphries The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: Getting no u-rays, fluroscopy failed. uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Futher 000000000000000000000000001k carried ou: replaced kV-mA board and Cue board on supicion of commuication issus bu this has made no improvement. Faut now supected as being the u-ray tue, so this will be ordered for neut visit. Follow u Requred : 2nd man for lifting Follow u Requred : 2nd man for lifting Follow u Requred : 2nd day of tue replacement Follow u Requred : 1. Fit original converter Q (anode) in 2Q (cathode) position and attempt tue adaptation. If this fails, then: 2. Replace tue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032165000000002-Nov-2000000003 Nigel Huphries The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: Getting no u-rays, fluroscopy failed. uer Impact: Patient Impact: SE Parts Consued : 1.000000000000000000000000000000002100000000000000002,Safety resistor 00000000, 300000000133mm Follow u Requred : Replace kV/mA control uit and/or cue. Adapt tue and test system. Diagnostic performed by Engineer : HT converter has been replaced bu the faut remains. uLLlogging is slightly different to yesterday so futher checks carried ou on u-ray control uits. Tue filaments are confirmed good. Despite generator control uit (Cue) and kV/mA control are passing their self-tests, these uits are now supect so will be ordered for tomorrow. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000312000000005 00000000-Nov-2000000003 Nigel Huphries The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: Getting no u-rays, fluroscopy failed. uer Impact: Patient Impact: SE Tools & Equpment : DOSE-KVP-LIGHT METER,2000000008-0000000030000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-0000000030000000000000000:00000000:00000000 Follow u Requred : Replace converter EQ (left side). Adapt tue and test. Coincidentally, one of the safety resistors also needs to be replaced. Diagnostic performed by Engineer : Faut diagnosis has revealed a defect in the HT converter part of the u-ray generator. Replacement HT converter and safety resistor are on order and retun visit requsted. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000311196 David Couens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to carry ou tue conditioning and tue adaptation Diagnostic performed by Engineer : RSW states that there is asymmetrical kV detected. To carry ou tue conditioning and adaptation to see if this resolves the issu. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: Getting no u-rays, fluroscopy failed. uer Impact: Patient Impact: SE Reguatory Qustions : N/A ----------------------------------------------------------------------------- Cutomer comment : Lorenza Clarke called to ask where the parts were and what is happening with the engineers. I euplained the parts are schedued to be delivered by tomorrow am and 2 engineers are schedued for tomorrow and the neut day as well (00000000 and 00000000/11) Cutomer comment : From: George, Anna <Anna.Georgeucom> Sent: Tusday, October 31, 2000000003 11:00000000 AM To: P.Frollo00helondonclinic.co.u; R.Solomon00helondonclinic.co.u Cc: PH.LTP (Fuctional Accout) <ph.ltpucom> Suject: uGENT u Health Systems: The London Clinic - Service Dates Good morning, Please review the proposed service dates below and confirm whether they are sutable at you earliest convenience. These services are being completed ouside of the grace period. Please sign and retun the attached letter to acknowledge this. Technical ID Equpment Location Contract Period Case Statu Visit 1 GB100000000600000000000000001-000000001 u FD200000000Angio Sute 7th Floor 25.0000000023 - 3000000000000000020000000012172100000000 TBC 11.00000000.200000000855000000006 GB100000000600000000000000005-000000001 u uuuFD200000000OR Table 00000000 Floor Hybrid lab 25.0000000023 - 3000000000000000020000000012161800000000TBC 00000000.00000000.200000000000000002300000000 Have a lovely day. Kind regards, Anna George Long-term Planner Health Systems u&I u uI Main Cutomer Line: 00000000000000000000000060000000086 Northern Ireland Line: 00000000000000000000000060000000000000000Repulic of Ireland Line: +3531800000000832300000000 u ELECTRONICS u LIMITED Ascent 1, Aerospace Bouevard, Farnborouh, Gu00000000uW Registered in England No. 00000000897, ## Anna.Georgeucom Connect with u Cutomer comment : Left voic##for Richard to advise FSE attending tomorrow ----------------------------------------------------------------------------- | 2023-10-31 16:19 | No
3131 | Yes | Operation check takes more time than normal (Software is reloaded). | Operation check takes more time than normal (Software is reloaded). | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000895 Madhan, kuar The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : 000000000000000000000000001king fine after sw reload. Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A ----------------------------------------------------------------------------- | 2023-10-31 16:24 | No
3132 | Yes | PSA_PSE_Legal Issue_Missing Asystole Alarm= - (CHR//RMR//TI child record) | PSA_PSE_Legal Issue_Missing Asystole Alarm= - (CHR//Ru/TI child record) |  | 2023-10-31 16:44 | Yes
3133 | Yes | [Q]Drive issue. | [Q]Drive issue. | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000300000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- Internal Remarks : New Case Solomon Tekabe - 80000000070000000086 IP & Issu: 68000000002316/Drive issu. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Claims Contract (with reason) Reason for Entitlement Claim: N/A Action Taken: Created case, summited cu claims Neut Step Action: Awaiting Entitlement ----------------------------------------------------------------------------- | 2023-10-31 17:04 | No
3134 | Yes | BZS0: Wireless footwitch defect | BZS00: Wireless footwitch defect | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000JEROME CHARLES, The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Biotech Device ue: No. Eupected and ueupected Resuts: No. uer Impact: No. Patient Impact: no SEuknown Diagnostic performed by Engineer: A qute requst was done. ----------------------------------------------------------------------------- | 2023-10-31 17:29 | No
3135 | Yes | Remote Alert: RMT: Host PC POST failed - memory 3 | Remote Alert: RMT: Host PC POST failed - memory 3 | 000000000000000000000000001O##: 00000000000000000000000000100000000332185 200000000Nov-2000000003 Skala John The system meets the specification for the performed service and is retuned to ue Resoluion : RSE M2M call host PC memory uLL. Problem Description by Engineer : Cutomer Fuction/Role: M2M call Device ue: Remote service team foud uLLon Host PC. Eupected and ueupected Resuts: Replace host PC. uer Impact: No impact for cutomer Patient Impact: No patient harmed SE 000Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000066 00000000-Nov-2000000003 Skala John The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Need to order Host PC. Problem Description by Engineer : Cutomer Fuction/Role: M2M call Device ue: Remote service team foud uLLon Host PC. Eupected and ueupected Resuts: Replace host PC. uer Impact: No impact for cutomer Patient Impact: No patient harmed SE 000Diagnostic performed by Engineer : M2M call foud uLLon Host PC. ----------------------------------------------------------------------------- Internal Remarks : Performed Lic swap sent to FSE ----------------------------------------------------------------------------- | 2023-10-31 17:29 | No
3136 | Yes | 0 - u00 Operation check failed | 00- u000- Operation check failed | 000000000000000000000000001O##: 000000000000000000000000001000000003100000000900000000adhan, kuar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : Reported issu not foud now. All testes in operational check passed. ----------------------------------------------------------------------------- | 2023-10-31 17:44 | No
3137 | Yes | [Q] Replacement of the euposition release. | [Q] Replacement of the euposition release. | 000000000000000000000000001O##: 000000000000000000000000001000000003000000000000000031-Oct-2000000003 Marek Jeske The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: uer Device ue: yes Eupected and ueupected Resuts: cosmetic issu - cable insuation broken, handswitch replaced uer Impact: no Patient Impact: no SE2.23 Resoluion: Problem reported with arm rest - cracked - replaced. A problem with the buton insuation was reported - replacement. The fauts have been fiued. The equpment was pu into clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000759000000003,Arm suport board 1.00000000000000000000000000000000220000000010000000083,Handswitch 9-Pin D-Su (f) white ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000023 Marek Jeske The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: uer Device ue: yes Eupected and ueupected Resuts: mechanical problem with euposue switch cable - to be replaced uer Impact: no Patient Impact: no SE Resoluion: An issu has been reported with the secondary insuation of the euposue cable cracking. We will replace the trigger. ----------------------------------------------------------------------------- | 2023-10-31 17:49 | No
3138 | Yes | 0/u FD0Intermittent startup failed | 0R/u FD0Intermittent startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 00000000.11.2000000003 00000000:12:35 uC (86000000008900000000) "RDI Accuacy: yes" $Q7: (Confirmed reported problem, sometimes, there is nothing showing on data monitor after start system.No uLLdisplayed on data monitor.) $Q8: (check the power suply of host pc, we find it is ok.Visul check, we find the light of disk is not on. So it is the reason) $Q9: (replace a motherboard electronics and directly connect hard disk to the motherboard, the system retuned to normal.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (NA) S/N (N/A) Neut Calibration date (N/A) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tueinstalled date: NA ^Curent Tue uage: NA SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: uable to boot Euternal Remarks: 00000000.11.2000000003 00000000:12:32 uC (86000000008900000000) *Department information: [Department: Equpment Department Contact Person: Li Changyou Contact nuber: 1393700000000896]* Service Content: [After powering on, the monitor screen goes black and there is no signal. After repair, the equpment retuned to normal. ] Service resuts: [After repair, the equpment retuned to normal operation. ] Other important matters: [none] Safety Qustions: Case: 0000000022515778 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the host PC cannot be started auomatically intermittently, and the Host PC needs to be tuned on manully, and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Intermittent failue to auomatically boot, the Host PC needs to be powered on manully Remote diagnostics and test resuts: NA PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check whether the Host PC power suply still has power after the machine is powered off. 2. Check whether the BIOS settings are normal and whether the BIOS battery voltage is normal. 3. Check whether the PC power suply voltage is too high. If it is too high, please adjut it. 00000000Youcan try to ue the IPPC power modue. Carry ou euchange 5. If there is no effect, ue PC Diagnostic Tool for diagnostic testing 6. Check the machine uLLreport and perform faut repair. If it cannot be repaired, replace the corresponding spare parts. Recommended spare part uuand name: 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: LiuGungyun 15333710000000086 ----------------------------------------------------------------------------- | 2023-10-31 17:49 | No
3139 | Yes | Patient u - Single chirp equipment alarm for pads possibly delayed therapy | Patient u - Single chirp equipment alarm for pads possibly delayed therapy |  | 2023-10-31 17:49 | Yes
3140 | Yes | 00T0annot get the system to come up, message limited or no connectivity | 00CDT0annot get the system to come up, message limited or no conneuty | 000000000000000000000000001O##: 00000000000000000000000000100000000313000000003 00000000-Nov-2000000003 Cathey Eric The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : cannot get the system to come u, message limited or no connectivity When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: cannot get the system to come u, message limited or no connectivity uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : system ready for ue Diagnostic performed by Engineer : replaced cmos battery in host and set time in bios checked operation ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000312968 Ketcherside Jeff The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : cannot get the system to come u, message limited or no connectivity When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: cannot get the system to come u, message limited or no connectivity uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : cutomer states cannot get the system to come u, message limited or no connectivity. tried to reboot several times bu still an issu Internal Comments : uLL# and/or description of uLL: uable to get the logfile in the systems curent state Repair Action : 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: uable to get the logfile in the systems curent state Repair Action : 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Ketcherside Jeff ----------------------------------------------------------------------------- Internal Remarks : 00000000225228300000000SE Cathey is reporting that the system will not boot and the system does not have Fleuvision. FSE has requsted HOST PC with 000system. Approve: 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part ----------------------------------------------------------------------------- | 2023-10-31 18:14 | No
3141 | Yes | The arm didn't operate_SPLIT | The arm didn't operate_SPLIT |  | 2023-10-31 18:45 | No
3142 | Yes | There was misalignment in the touch position of touchscreen. | There was misalignment in the touch position of touchscreen. |  | 2023-10-31 19:03 | No
3143 | Yes | ucable not functioning | ucable not funuoning | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000088 Gnanasekar, D The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Problem foud utruk cable, defective cable pictue attached, ordering replacement. Follow u Requred : Problem foud utruk cable, defective cable pictue attached, ordering replacement. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : ucable not fuctioning ----------------------------------------------------------------------------- | 2023-10-31 19:04 | No
3144 | Yes | BZA0 New wireless foot pedal | BZA00 New wireless foot pedal | 000000000000000000000000001O##: 00000000000000000000000000100000000311179 PATRICK ATTN. The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer: At the requst of the cutomer requsted qute for new foot pedals. ----------------------------------------------------------------------------- | 2023-10-31 19:14 | No
3145 | Yes | The unit is not giving pressure. (V0blower has no output.) | The unit is not giving pressure. (V0blower has no output.) | 000000000000000000000000001O##: 00000000000000000000000000100000000313600000000 00000000-Nov-2000000003 Galan, Alberto The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561512761,RP-Blower Assy,V600000000Diagnostic performed by Engineer : On 11/00000000/2000000003, A.G- FSE completed the installation of a Blower Assembly and conduted a ful PVT following the manuactuer's gudelines. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000007-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000007-31 00000000:00000000:00000000 MuTIMETER,200000000000000007-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000uLLto operate as eupected. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000blower has no oupu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000031300000000 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Problem Description: V600000000blower has no oupu. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000031/2000000003 Trouleshooting Steps Taken: Cutomer is requsting onsite evalution of V600000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000uLLto operate as eupected. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000blower has no oupu. Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve. Follow u Requred : V600000000blower has no oupu. Part # 00000000561512761 RP-Blower Assy, V600000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kirk Delos-Reyes IP & Issu: 5660000000057 - The uit is not giving pressue. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-31 19:44 | No
3146 | Yes | Remote Alert: RMT - IPU: Fluoro storage not possible due to an Image disk problem (>Rel0 | Remote Alert: RMT - IPU: Fluoro storage not possible due to an Image disk problem (>Rel00 | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000000000007 00000000-Nov-2000000003 David Masip Castilla The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Resoluion : Se recrea IPPC y se verifica fucionamiento del sistema Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- | 2023-10-31 19:54 | No
3147 | Yes | geometry movements partially available, toptop free floating | geometry movements partially available, toptop free floating | 000000000000000000000000001O##: 0000000000000000000000000010000000031583000000001-Nov-2000000003 Rockenstire, Michael The system meets the specification for the performed service and is retuned to ue Resoluion : replace net000000000000000000000000001k switch in geometry cabinet Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: system down Eupected and ueupected Resuts: geometry lost uLL uer Impact: coudn't ue system Patient Impact: can't ue room SE uknown ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000031372000000003-Nov-2000000003 Rowles Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: On boot u geo not 000000000000000000000000001king uLL uer Impact: had to cancel and reschedue this case and others for today and tomorrow. Patient Impact: had to cancel case and reschedue. SE Parts Consued : 1.00000000000000000000000000000000800000000237000000001,REPL. KIT ETHERNET SWITCH CRCB Resoluion : Replaced net000000000000000000000000001king switch. After tun on the geo uLL went away and everything was 000000000000000000000000001king fine. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000313321 31-Oct-2000000003 Rowles Michael The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : confirmed the uLLmessage. checked the GEO rack and foud the switch has all the lights on which only the 3 cable slots pluged in shoud have lights on. Also foud no link lights on the GEO pc net000000000000000000000000001k card. One of those is cauing the issu. I'm ordering a net000000000000000000000000001k switch and a geo PC. Follow u Requred : replace net000000000000000000000000001k switch and maybe the GEO PC Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: On boot u geo not 000000000000000000000000001king uLL uer Impact: had to cancel and reschedue this case and others for today and tomorrow. Patient Impact: had to cancel case and reschedue. SE ----------------------------------------------------------------------------- Internal Remarks : Assisted Mike Rowels with TS Switch issus not 000000000000000000000000001king correctly and also lab has GEO related problems. Approved replacement of TS switch and GEO IPC for trouleshooting. FSE will attempt to euchange switch and see if system will start 000000000000000000000000001king correctly again and have the GEO IPC as backu. Internal Remarks : Contacted FSE to discus issu and FSE states he is onsite and to assign him the 000000000000000000000000001O#for right now. Remote diagnostics will not add valu to service event since FSE is already at the system. Onsite service dispatched as requsted by FSE. ----------------------------------------------------------------------------- | 2023-10-31 20:04 | No
3148 | Yes | 0B (CBIT) Check Vent: Prouimal Pressure Sensor Auto Zero Failed, No GFE required/IS complete | 00 (CBIT) Check Vent: Prouimal Pressure Sensor Auto Zero Failed, No GFE required/IS complete | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000018000000003-Nov-2000000003 Nuvo, Joseph The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: n/a Device ue: not in ue Eupected and ueupected Resuts: repaired uer Impact: none Patient Impact: none SE 3.200000000Tools & Equpment : FLOW METER,200000000000000006-30000000000000000:00000000:00000000 MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000006-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 TORQu WRENCH,200000000000000006-30000000000000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000561512561,RP-PCBA, Data Acqusition, V600000000Diagnostic performed by Engineer : DRPTA, no uLL after repair ----------------------------------------------------------------------------- | 2023-10-31 20:04 | No
3149 | Yes | QA000000he table could not move, and the table, control panel and tube could not be ud | QA00000000_the table could not move, and the table, control panel and tube could not be ud |  | 2023-10-31 20:07 | No
3150 | Yes | During the operating test, uLLmessage: "multifunction electrode type unknown". | During the operating test, uLLmessage: "multifunuon eleuode type unknown". | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000229 Thierry LE GONIDEC The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Resoluion: The cutomer is not affected by the FSN since her paddles have code date 00000000/2000000002 and that the FSN concerns the paddles manuactued before Auut euactly. u Bodivit from AVF went to the cutomer's on 11/8/2000000003 in order to find a soluion. The defective paddle will be recovered by Franck Petit. Diagnostic performed by Engineer: Duing the operating test, uLLmessage: "mutifuction electrode type uknown". Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown ----------------------------------------------------------------------------- Internal Remarks: Cutomer callback for follow-u. Teams in Le Gonidec T. ----------------------------------------------------------------------------- | 2023-10-31 20:14 | No
3151 | Yes | High pressure alarm | High pressure alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000518 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: High pressue alarm. uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states device was tuned into biomed shop with high pressue alarm. cutomer powered uit on in service mode. uable to find any uLLin event log. cutomer powered on device in normal operational mode, foud that High inspiratory pressue alarm is set to only 6. cutomer adjuted alarm setting and alarm went away. cutomer states he will inform staff abou alarm setting and call back if futher assistance is needed. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gerald Sanders 000000005589300000000 IP & Issu: 100000000000000001618/High pressue alarm. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-31 20:54 | No
3152 | Yes | uray stopped during a case | uray stopped during a case | 000000000000000000000000001O##: 0000000000000000000000000010000000031000000002 00000000-Nov-2000000003 Autin Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uray stopped duing a case - uray acqusition is not available - happened three times When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System to uray/System keeps stopping uraying uer Impact: Had to reboot the system three times Patient Impact: Delay in procedue of approu. 200000000minues for three reboots. Procedue was completed SE 1.2.5.00000000Resoluion : Replaced sute PC and reloaded software Parts Consued : 1.00000000000000000000000000000000800000000717981,Sute PC W7 + W100000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000313528 Vilkaukas Viktorija The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uray stopped duing a case - uray acqusition is not available - happened three times When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System to uray/System keeps stopping uraying uer Impact: Had to reboot the system three times Patient Impact: Delay in procedue of approu. 200000000minues for three reboots. Procedue was completed SE 1.2.5.00000000Diagnostic performed by Engineer : Joyce is reporting that this morning the system has stopped uraying three times and system had to be rebooted. She said this has happened to them before bu has been several months ago. The have received a message: uray acqusition is not available. System rebooted three times and the procedue was completed. Created a logfile, Not able to find the same message as reported, bu before every reboot there is a message that a processes has crashed appears to be from Sute PC. They are done with cases for today and neut case is on Thusday 11/2/2000000003. Internal Comments : uLL# and/or description of uLL: 200000000SIEC000000001000000002 SessionManager: Stopping process failed, now killing it 200000000SIEC000000001000000003 SessionManager: Process crashed Repair Action : Joyce is reporting that this morning the system has stopped uraying three times and system had to be rebooted. She said this has happened to them before bu has been several months ago. The have received a message: uray acqusition is not available. System rebooted three times and the procedue was completed. Created a logfile, Not able to find the same message as reported, bu before every reboot there is a message that a processes has crashed appears to be from Sute PC. They are done with cases for today and neut case is on Thusday 11/2/2000000003. Please consut an NSS. It is not clear where the issu is and appears to be software related. Follow u Requred : uLL# and/or description of uLL: 200000000SIEC000000001000000002 SessionManager: Stopping process failed, now killing it 200000000SIEC000000001000000003 SessionManager: Process crashed Repair Action : Joyce is reporting that this morning the system has stopped uraying three times and system had to be rebooted. She said this has happened to them before bu has been several months ago. The have received a message: uray acqusition is not available. System rebooted three times and the procedue was completed. Created a logfile, Not able to find the same message as reported, bu before every reboot there is a message that a processes has crashed appears to be from Sute PC. They are done with cases for today and neut case is on Thusday 11/2/2000000003. Please consut an NSS. It is not clear where the issu is and appears to be software related. Technician Name : Vilkaukas Viktorija ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Autin Ryan with trouleshooting sute PC process crashes. Approved TPR for new Sute PC and FSE will reload SW ----------------------------------------------------------------------------- | 2023-10-31 20:54 | No
3153 | Yes | M0A / u00/ Device uLL therapy knob failure. Device is no longer charging. | M00A / u000/ Device uLL therapy knob failure. Device is no longer charging. | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : Battery fuly charged Ou Of Bou Failue? : No Labor: Labor ID : L-19865939 ; L-198679700000000000000000000000000000000001k Performed : Cleared TK uLLand performed 2 operational checks both passed. Verified charger and battery stau as in spec. ; Spoke with staff. Device was upluged from AC and battery was low. Souds like the uit shu off mid OPS check becaue of low battery. Battery charged and device passes OPS check. Retuned to service. Diagonsis : Cleared TK failue and replaced battery as cauionary measue ; Cleared TK failue and replaced battery as cauionary measue | 2023-10-31 21:01 | No
3154 | Yes | System is not booting up to login screen, screen ju displays the uYy logo | System is not booting up to login screen, screen ju displays the uYy logo | 000000000000000000000000001O##: 00000000000000000000000000100000000317233 00000000-Nov-2000000003 Cortez, Joseph The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Loaded ghost backu and new license file to host pc. Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: no patient involved Eupected and ueupected Resuts: system will boot u / system will not boot u uer Impact: cannot ue system Patient Impact: no patient involved SE uknown ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000 31-Oct-2000000003 Cortez, Joseph The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Diagnosed the host PC. Host PC is bad. Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: no patient involved Eupected and ueupected Resuts: system boots u / system will not boot u uer Impact: cannot ue system Patient Impact: no patient involved SE uknown ----------------------------------------------------------------------------- Internal Remarks : spoke with FSE Joseph Cortez and he can take this service call today at 12PM. Waiting on qute and PO for onsite 000000000000000000000000001O#to be released. Internal Remarks : Assisted FSE Cortez, Joseph with license swap for HOST PC euchange. New MAC 00000000C0000000081B00000000 ----------------------------------------------------------------------------- | 2023-10-31 21:09 | No
3155 | Yes | Touchscreen malfunctioned | Touchscreen malfunuoned | 000000000000000000000000001O##: 00000000000000000000000000100000000313136 Seva, Romain The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Requst for parts uder the contract. uer Impact: No. Patient Impact: No. SE / Diagnostic performed by Engineer: The touh screen is defective. It does not 000000000000000000000000001k anymore. There are impacts and it does not respond anymore. Requsts a new touh panel for replacement. Resoluion: Sending the touh panel uder the contract. Delivery to the biomedical department. Follow u Requred: Sending parts uder the contract. Delivery to the biomedical department. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000310000000000000000 Consued parts: 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000----------------------------------------------------------------------------- | 2023-10-31 21:09 | No
3156 | Yes | Screen froze | Screen froze |  | 2023-10-31 21:36 | No
3157 | Yes | 0 (CBIT) Check Vent: Battery Failed - Pending quote for repair 00 | 00 (CBIT) Check Vent: Battery Failed - Pending quote for repair 0/0 |  | 2023-10-31 22:01 | No
3158 | Yes | Equipment Showing Pipe Failure | Equipment Showing Pipe Failure | BuQ&R Feedback: No SAP Cutomer Symptom Code: ACQP Cutomer Symptom Code: Imaging/Patient Recording/Acqusition Problem Reason: Corrective Maintenance Engineer Failue Code: Electrical/Electronic Failue Repair Activity Code: Instruted Cutomer SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Problem description by engineer : u-ray failue Trouleshooting Action: What steps are taken duing the system diagnosis : 100000000300000000- Thiago (eng) informs that the equpment is presenting an uLLin the u-ray: Low load flu flavor selected. Tue rotor problem; Cutomer-oriented save a log; Analyzed log and verified ADufailue: MIu Fue trip on rail voltage oupu to ADu Verified that the ADuRotation State LED is off; NOT OK Tuned off the entire power panel and the F1 switch of the GPDu Disconnected and reconnected the ADucables; OK Reconnected the Cucable; OK Measued the 00000000 phases of the power panel (00000000V); OK Cuoil level checked; OK Connected the power frame and GPDu Tun on the equpment; Verified that the ADus Rotation State LED was GREEN and measued at 200000000; OK Image acqusition tests, no more failues; OK NOTE: Huidity was high 77% in the technical room. Asking the cutomer to repair so that it is a mauimu of 600000000; Equpment cleared for euaminations uder observation; CONTACT: Thiago 62 99638000000009 uLL# and/or description of uLL: uGC: MIuFue trip on rail voltage oupu to ADu- Frontal Low load flu flavor selected. Tue rotor problem Repair Action: Thiago aware of the closue of the call. 000000000000000000000000001king equpment; OK "Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable" Test/Inspection: OK Test Equpment: N/A OneEMS internal remarks: Problem description by engineer : u-ray failue What steps are taken duing the system diagnosis : 100000000300000000- Thiago (eng) informs that the equpment is presenting an uLLin the u-ray: Low load flu flavor selected. Tue rotor problem; Cutomer-oriented save a log; Analyzed log and verified ADufailue: MIu Fue trip on rail voltage oupu to ADu Verified that the ADuRotation State LED is off; NOT OK Tuned off the entire power panel and the F1 switch of the GPDu Disconnected and reconnected the ADucables; OK Reconnected the Cucable; OK Measued the 00000000 phases of the power panel (00000000V); OK Cuoil level checked; OK Connected the power frame and GPDu Tun on the equpment; Verified that the ADus Rotation State LED was GREEN and measued at 200000000; OK Image acqusition tests, no more failues; OK NOTE: Huidity was high 77% in the technical room. Asking the cutomer to repair so that it is a mauimu of 600000000; Equpment cleared for euaminations uder observation; CONTACT: Thiago 62 99638000000009 uLL# and/or description of uLL: uGC: MIuFue trip on rail voltage oupu to ADu- Frontal Low load flu flavor selected. Tue rotor problem Repair Action "Euternal" : Thiago aware of the closue of the call. 000000000000000000000000001king equpment; OK "Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable" Test/Inspection: OK Test Equpment: N/A Eutra Internal comment field : CSR André Confalonieri Problem Reported by Cutomer: Equpment Presenting Pipe Failue Name: Thiago Taveira Indutry: Clinical Engineering Phone ¹: 62 9 9638000000009 Phone²: ## thiagotaveira00rbisengenhariaclinica.com.br INSTALLED IN: INNOVATION MANAGEMENT ASSOCIATION AND Avenida Anhangura 1000000007 Goiânia-GO 7000000003-336 BR Safety Qustions: Case: 00000000225171600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Aline Dias Last udated by: Aline Dias 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A T2 Activities:Problem Description: Problem description by engineer : u-ray failue Trouleshooting Action: What steps are taken duing the system diagnosis : DAY 10000000031 - Equpment did not present any more failues. uLL# and/or description of uLL: uGC: MIuFue trip on rail voltage oupu to ADu- Frontal Repair Action: Thiago aware of the closue of the call. 000000000000000000000000001king equpment; OK "Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable" Test/Inspection: OK Test Equpment: N/A OneEMS internal remarks: Problem description by engineer : u-ray failue What steps are taken duing the system diagnosis : DAY 10000000031 - Equpment did not present any more failues. uLL# and/or description of uLL: uGC: MIuFue trip on rail voltage oupu to ADu- Frontal Repair Action "Euternal" : Thiago aware of the closue of the call. 000000000000000000000000001king equpment; OK "Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable" Test/Inspection: OK Test Equpment: N/A Eutra Internal comment field : CSR André Confalonieri ----------------------------------------------------------------------------- | 2023-10-31 22:19 | No
3159 | Yes | uCVu-uLL0(SW issue) / WED 0EC0-MIR report open | uCVu-uLL0 (SW issue) / WED 0DEC00MIR report open | 000000000000000000000000001O##: 000000000000000000000000001000000003000000002600000000ovember 7, 2000000003 Good, Kevin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Med Tech Device ue: Cardio Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE 9.000000005 Resoluion: Cutomer reports the uLL200000000.VFPY. K00000000.56 - SSD replaced - Software reloaded Parts Consued: 1.00000000000000000000000000000000561810000000085,ASSY, HARD DRIVE, SSD w/ LABEL ----------------------------------------------------------------------------- | 2023-10-31 22:44 | No
3160 | Yes | Audio & Sounds/Loudspeaker failure/Speaker | Audio & Sounds/Loudspeaker failure/Speaker | 000000000000000000000000001O##: 00000000000000000000000000100000000328237 November 100000000 2000000003 Alejandro Rodrigo Mauier Córdoba The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We tried to speak with nuse Joaquna Ortega 9158680000000000000000eut 00000000182 / 915290000000082 / 915290000000039 bu we were not able to. Device ue: uknown Eupected and ueupected Resuts: Louspeaker failue, the alarms do not 000000000000000000000000001k. uer Impact: uknown Patient Impact: uknown SE Not apply. Parts Consued: 1.00000000000000000000000000000000261000000009000000001,IV-MP5 CBL Speaker Assembly 1.00000000000000000000000000000000261000000009331,IV-MP5 MECHASY Assy Quckmout Tools & Equpment: PATIENT SIMuATOR, 200000000000000001-18 00000000:00000000:00000000 Resoluion: Internal speaker is replaced. Assembly anchoring has been replaced. Verification of correct operation ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000031000000000000000000000000-Nov-2000000003 Alejandro Rodrigo Mauier Córdoba The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We tried to speak with nuse Joaquna Ortega 9158680000000000000000eut 00000000182 / 915290000000082 / 915290000000039 bu we were not able to. Device ue: uknown Eupected and ueupected Resuts: Louspeaker failue, the alarms do not 000000000000000000000000001k. uer Impact: uknown Patient Impact: uknown SE Not apply. Resoluion: Small Part Kits are requsted. Several euternal parts are damaged. Follow u Requred: Material ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000031000000006 Rau Cordero Calderón The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Tried to speak with nuse Joaquna Ortega 9158680000000000000000eut 00000000182 / 915290000000082 / 915290000000039 bu we were not able to. Description of the failue: There is a failue in the speaker. the alarms do not 000000000000000000000000001k. Assign fSE. Read INOP Message Speaker Malfuction is displayed page 137 Service Gude Release N.00000000000000005600000000281 Order part: 00000000261000000009000000001 - IV-MP5 CBL Speaker Assembly uSemi-basement, Infant Emergency, Observation Area D. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Tried to speak with nuse Joaquna Ortega 9158680000000000000000eut 00000000182 / 915290000000082 / 915290000000039 and we were not able to Device ue: uknown Eupected and ueupected Resuts: Louspeaker failue, the alarms do not 000000000000000000000000001k. uer Impact: uknown Patient Impact: uknown SE Not apply. Diagnostic performed by Engineer: Tried to speak with nuse Joaquna Ortega 9158680000000000000000eut 00000000182 / 915290000000082 / 915290000000039 bu we were not able to. Description of the failue: There is a failue in the speaker. the alarms do not 000000000000000000000000001k. Assign fSE. Read INOP Message Speaker Malfuction is displayed page 137 Service Gude Release N.00000000000000005600000000281 Order part: 00000000261000000009000000001 - IV-MP5 CBL Speaker Assembly uSemi-basement, Infant Emergency, Observation Area D. ----------------------------------------------------------------------------- | 2023-10-31 22:54 | No
3161 | Yes | Speaker malfunction. | Speaker malfunuon. | 000000000000000000000000001O##: 0000000000000000000000000010000000031000000006 Roper, Ken The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Problem Description: Butons are not 000000000000000000000000001king; uit shows speaker malfuction. no tone Butons are not 000000000000000000000000001king; uit shows speaker malfuction. no tone Form ued: Bench Trouleshooting Steps Taken: Butons are not 000000000000000000000000001king; uit shows speaker malfuction. no tone Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Butons are not 000000000000000000000000001king; uit shows speaker malfuction. no tone Diagnostic performed by Engineer : Problem Description: Butons are not 000000000000000000000000001king; uit shows speaker malfuction. no tone Butons are not 000000000000000000000000001king; uit shows speaker malfuction. no tone Form ued: Bench Trouleshooting Steps Taken: Butons are not 000000000000000000000000001king; uit shows speaker malfuction. no tone ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kiersten Osborne IP & Issu: " 68817398 Butons are not 000000000000000000000000001king; uit shows speaker malfuction." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-31 23:04 | No
3162 | Yes | Monitor will not charge | Monitor will not charge | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000661 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged. To correct the issu we will replace the Front Case Assembly. Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 Resoluion : Device was repaired by u technician Chongolnee, Roongruchai with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000315176 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged. To correct the issu we will replace the Front Case Assembly. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair Follow u Requred : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000310000000000000000Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Follow u Requred : Problem Description: tech reports the tempu Pro will not charge, the charging pin broken off , sending to bench; requsting a loaner Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repair Confirmed delivery address ##d bench repair form Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE 7.28 Problem Description/ueupected Resuts: tech reports the tempu Pro will not charge, the charging pin broken off , sending to bench; requsting a loaner Diagnostic performed by Engineer : Problem Description: tech reports the tempu Pro will not charge, the charging pin broken off , sending to bench; requsting a loaner Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form ----------------------------------------------------------------------------- Internal Remarks : Received Teams message on 11/6/23 from RSE - Devon Adams asking that I accept 000000000000000000000000001O#for another Case 000000002255800000000 from quu and call cuback to assist Replied with agreement and accepted 000000000000000000000000001O# --- Later reviewed Case 000000002255800000000, foud cuwanted to send device to the Bench, wanted to receive a shipping label and loaner device, mentioned having not received a shipping label for another device that requred repair, wanted to receive information on when he shoud eupect to receive loaner devices and shipping labels for devices Foud cuwas referring to this case, foud no Bench repair form had been ##d to cu foud shipping label and loaner device had not been requsted Called cuback at 500000000853 1593 to assist, left a voic## ##d Irvine CA Bench forms for both cases to cu sumitted requsts for shipping labels and loaner devices to be sent by Bench to cu Sent ##to cumaking him aware of the above, performed actions --- Labor to be confirmed: 11/6/23 from 00000000M to 0000000000000000PM ----------------------------------------------------------------------------- | 2023-10-31 23:29 | No
3163 | Yes | Geometry movements partially available | Geometry movements partially available |  | 2023-10-31 23:32 | No
3164 | Yes | Equipment failed operational test. | Equipment failed operational test. |  | 2023-10-31 23:33 | No
3165 | Yes | It shows an uLLmessage.. | It shows an uLLmessage.. |  | 2023-11-01 00:02 | No
3166 | Yes | Slow when starting the equipment | Slow when starting the equipment |  | 2023-11-01 00:57 | No
3167 | Yes | M0A - u00 Has defibrillation uLLin operational test. | u Has defibrillation uLLin operational test. | BuQ&A Feedback: No Reason: Corrective Maintenance Trouleshooting Action: Schedued for FSE Gulherme on 21-11 Internal Remarks: 00000000.12.2000000003 00000000:37:17 POLARIS (POLARIS) ================= Gutavo de Assis Soua - 0000000000000000576000000003200000000200000000000000000000000038 =================== ==== GG6: Operational tests approved. GG13: No calibratable tools were ued. Neut Step Action: Analysis carried ou and directed to sending a buget Euternal Remarks: 00000000.12.2000000003 00000000:37:17 POLARIS (POLARIS) ================= Gutavo de Assis Soua - 0000000000000000576000000003200000000200000000000000000000000038 =================== ==== Initial failue: Maintenance uLLReplaced uuuDefibrillator uuuitor series: u00000000563976 Tested display, butons, charge with Qed.6. Operational tests, the k. Note: The insuation ruber for the metal sheets is not 000000000000000000000000001n very muh, it is recommended that the metal sheets of the equpment be replaced. NF 000000001 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022518000000007 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Aline Dias Last udated by: Aline Dias 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT Problem Reported by Cutomer: Displays defibrillation uLLin the operational test. Name: Diego Borges Sector: Clinical Engineering 00¹: 11 367700000000 ## engenharia.cirugico00ospitalsaocamilosp.org.br SOCIEDADE BENEFICENTE SAO CAMILO Avenida Pompeia 1178 Sao Pauo-SP 00000000000000002-000000001 BR SAP Transfer: Case Transferred to SAP Euternal Remarks: 22.11.2000000003 00000000:50000000052 POLARIS (POLARIS) ================= Gulherme Do Rio Costa - 0000000000000000576000000003200000000100000000000000000000000071 =================== ==== Evalution of the u HeartStart uuuDefibrillator 000u00000000563976 was carried ou, Equpment failed to perform the Operational Test and covered the Defibrillation Test. Fuctional test was applied with a Jouimeter and a failue was detected in the uuuitor Assy Part. Replacement of the item above is necessary to resolve the faut presented by the equpment. Equpment NOT OK. Repair Action: schedued for FSE Gutavo on 11/300000000Trouleshooting Action: Qute sent SAP Transfer: Case Transferred to SAP Internal Remarks: 22.11.2000000003 00000000:50000000052 POLARIS (POLARIS) ========== Gulherme Do Rio Costa - 0000000000000000576000000003200000000100000000000000000000000071 =================== ==== Service not yet completed. Tests and Calibrated Tools in PA Tool. SAP Transfer: Case Transferred to SAP Trouleshooting Action: OPP-198623 Qute: 000000001133600000000mail: engenharia.cirugico00ospitalsaocamilosp.org.br ----------------------------------------------------------------------------- | 2023-11-01 01:09 | No
3168 | Yes | Device will not go into standby mode, GFE00OV0 No response/IS complete | Device will not go into standby mode, GFE0 0NOV00No response/IS complete | 000000000000000000000000001O##: 00000000000000000000000000100000000315772 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in patient ue Eupected and ueupected Resuts: uit w/not go into standby uer Impact: none Patient Impact: na SE 2.100000000Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Average Air slpm readou is -000000003. Advised cutomer to perform ful PVT and address any issus foud, supect flow issus. If no issus foud replace the flow sensor assy to resolve the -000000003 reading for the Average Air slpm. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Wilbert 313-576-10000000000000000u65800000000 IP & Issu: 100000000000000007600000000machine will not go into standby preventing therapist from setting u for patients Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: created remote 00000000000000000000000000100000000315772 dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-01 01:14 | No
3169 | Yes | Burning smell and visible burn mark on device, No GFE required/IS complete | Burning smell and visible burn mark on device, No GFE required/IS complete | 000000000000000000000000001O##: 0000000000000000000000000010000000031300000000 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: uit was involved in the recall board was replaced bu has gone bad uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the PM board was replaced per FCO 66 and now there is a buning smell and visible bun mark on device. cutomer states that ASP claimed they were going to take care of it bu hasn't heard from ASP and 000000000000000000000000001ld like to have FSE dispatched to repair device. Confirmed cutomer contact info. Resoluion : Creating on site service at 100000000% discout to have FSE replace PM board Follow u Requred : 100000000% discout. Cutomer claims buning smell coming from device and visible bun mark on board. ASP claimed they were going to take care of it bu hasn't heard back from anyone. requsting FSE on site to repair device uder FCO warranty. ----------------------------------------------------------------------------- Internal Remarks : Caller: Phillip 8500000000879 8298 IP & Issu: 60000000019000000009 - uit was involved in the recall bored was replaced bu has gone bad Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A ----------------------------------------------------------------------------- | 2023-11-01 01:59 | No
3170 | Yes | Fails to pick up a 0ead signal. | Fails to pick up a 0lead signal. | 000000000000000000000000001O##: 00000000000000000000000000100000000316553 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: Sales rep. Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Failed to pick u a 3-lead signal duing PM. Diagnostic performed by Engineer : Problem Description: Failed to pick u a 3-lead signal duing PM. Failed to pick u a 3-lead signal duing PM. Form ued: Trouleshooting/Investigation Object: Tempu LS Not Object: Software Version: NA Date and Time of Issu: 10000000031/2000000003 uer Actions: Where is the Deviation: Tempu LS Trouleshooting Steps Taken: NA Was PRS ued: No PRS Permission: Resoluion : Is Caue Known: No Caue: Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Ordered replacement Tempu LS 000000005600000000181. ----------------------------------------------------------------------------- | 2023-11-01 02:29 | No
3171 | Yes | No fluoro during case | No fluoro during case | 000000000000000000000000001O##: 00000000000000000000000000100000000316859 00000000-Nov-2000000003 Park James The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : no fluro duing case .. reboot no help changed protocol and it 000000000000000000000000001ked When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: no fluro duing case .. reboot no help changed protocol and it 000000000000000000000000001ked uer Impact: change protocol to get fluro Patient Impact: No patient impact reported by cutomer. SE 8.1.3000000005 Parts Consued : 1.00000000000000000000000000000000270000000000000000100000000Footswitch CV 3p 00000000Resoluion : uuubuton the foot switch is intermittently not responding, inspected foot switch and did not see any damage to it. replaced the foot switch and retuned system to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000316568 Clark, Byron The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Coud not resolve the issu and requres onsite suport Problem Description by Engineer : Problem description by engineer : no fluro duing case .. reboot no help changed protocol and it 000000000000000000000000001ked When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: no fluro duing case .. reboot no help changed protocol and it 000000000000000000000000001ked uer Impact: change protocol to get fluro Patient Impact: No patient impact reported by cutomer. SE 8.1.3000000005 Diagnostic performed by Engineer : collected logs created and attached to case,, no uLL seen Internal Comments : uLL# and/or description of uLL: none seen in logs Repair Action : ru fuctional checks and interview uer and try protocol in qustion Follow u Requred : uLL# and/or description of uLL: none seen in logs Repair Action : ru fuctional checks and interview uer and try protocol in qustion Technician Name : Clark, Byron ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: JJ Action: I called FSE-James Park regarding 00000000000000000000000000100000000316859 he accepted it for 11/1/23 from 12pm-1pm with 1hr of travel time each way…Per Director Stew Jones who approved for James Park to go onsite- James is Level 1 trained bu not level 2 trained. When I consider this carefuly, who we have available, the natue of the P00000000all, James' training and the cutomer I think the best decision is to place the call in James' quu. I called and spoke with the cu Michael Kuth abt scheduing. Cu said anytime between 12pm -1pm. I called and left the cu a vm message of the date & time FSE will be onsite tomorrow I dispatched the 000000000000000000000000001per Director requst I pu FADB in the Planning Notes Last udated: 8:00000000PM EDT 10000000031/23 ----------------------------------------------------------------------------- | 2023-11-01 03:14 | No
3172 | Yes | HS0ED is emitting the triple chirp | HS0AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000316975 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they brouh a new battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they brouh a new battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they brouh a new battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 15 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A00000000B-0000000065 Battery Install Before# 12/2000000008 Lot#A00000000200000000-23199 Pads Eup Date: 7/18/2000000008 Pads Lot # ----------------------------------------------------------------------------- | 2023-11-01 03:49 | No
3173 | Yes | Low leak CO0ebreathing risk during PM ST performance test | Low leak CO0rebreathing risk during PM ST performance test | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000700000000Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Resoluion : reviewed test set u and advised to place the Whisper Swivel in line with the test lug. uit now passes all PM testing and is placed back into service. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: PM Service Eupected and ueupected Resuts: duing ST performance test on PM low leak co2 rebreathing risk uer Impact: none Patient Impact: na SE uknonw ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Stephen Marshall IP & Issu: 56631573 duing ST performance test on PM low leak co2 rebreathing risk Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: free remote Reason for Entitlement Claim: N/A Action Taken: created case, R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-01 03:59 | No
3174 | Yes | Alarms or Messages/Speaker malfunction/Speaker | Alarms or Messages/Speaker malfunuon/Speaker | 000000000000000000000000001O##: 00000000000000000000000000100000000317185 Gavin BANH The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : Speaker malfuction. speaker replaced uder warranty. Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Engineer Device ue: Monitoring vitals Eupected and ueupected Resuts: Clinical staff coud not hear alarms uer Impact: uable to hear alarms Patient Impact: No impact to patient SE N.00000000 Diagnostic performed by Engineer : Speaker malfuction. speaker replaced uder warranty. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003171800000000 Parts Consued : 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- | 2023-11-01 05:04 | No
3175 | Yes | 0/FD0Host PC Cannot boot up | 0R/FD0Host PC Cannot boot up | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 00000000.11.2000000003 00000000:17:28 uC (860000000095100000000 "RDI Accuacy: Yes" $Q7: (Confirmed reported problem: System coud not start u normally.) $Q8: (Trouleshooting & Resuts: Checked the Host PC was not powered u.) $Q9: (Repair steps: Started u Host PC by pressing the power buton. Checked uLLlogs, BIOS configuation. Tested the voltage of the Host PC main board battery, was 2.95VDC. Reinstalled the battery and Readjuted the time and date of BIOS.) $Q100000000 (Test and Inspection data:Refer to PA. Performed visul check, it passed. Device statu: Device retun to ful fuctionality.) $Q11: Calibrated Tools name: Refer to PA. Euternal Remarks: 00000000.11.2000000003 00000000:17:25 uC (860000000095100000000 Service content: Confirm that the cutomer reported for repair and coud not start u normally.Remotely gude cutomers to manully press Host After pressing the PC power buton, the device resues ue.Check the BIOS settings and backgroud logs on-site, test the Host PC motherboard battery voltage to be normal, and reset the BIOS time and date. Service resut: The equpment retuns to normal ue. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Neut Step Action: NTD: The cutomer refued to transfer and coudn’t tun on the machine when he came in the morning NNY Problem Reported by Cutomer: 00000000R/FD200000000Can't boot u T2 Activities:Problem Description: Overview of cutomer problem: Hos tPC device cannot be tuned on. Detailed description of cutomer problem or faut phenomenon: Hos tPC device cannot be tuned on. Is it a faut of the device itself? Yes Trouleshooting Action: Key faut codes or phenomena: The Hos tPC device cannot be powered on. Remote diagnosis and test resuts: The Hos tPC device cannot be powered on. Check that the power suply is normal. It is supected that the hostpc is fauty. PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check the HostPC power suply, 2. Reformat the HostPC OS disk and reinstall the system, 3. Replace the host PC. Recommended spare parts uuand name: " 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD" Cutomer demand Remarks: Onsite Safety Qustions: Case: 0000000022519000000009 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-01 05:04 | No
3176 | Yes | 0S:SS: Nothing appears on the screen and the equipment does not start - SPLIT | 0DS:SS: Nothing appears on the screen and the equipment does not start - SPLIT | 000000000000000000000000001O##: 00000000000000000000000000100000000317200000000Saka, Shuji The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Nothing appears on the screen and the equpment does not start. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: Nothing is displayed on the screen, and the equpment does not start. uer Impact: No effect Patient Impact: No effect SE 000Resoluion: RC Response: The power off buton does not 000000000000000000000000001k. It was possible to tun the circut breaker off/on and then re-lauch it. I checked the log. There was no log record for the time in qustion, so I think power was not being suplied to the main PC. It seems this is the first time this has had a problem, so I promised to respond if the problem recus, and I completed the response for today. Internal Comments: [Suject] 00000000CS:SS:u Tokuaga: Nothing appears on the screen and the equpment does not start. [Internal Comments] I reported this to FSE Nomua. 00000000230000000016000000007 MPD CONTROL uIT Fuuka 00000000800000000868300000000MPDureplacement assembly Tokyo [Euamination category] Other: ----------------------------------------------------------------------------- | 2023-11-01 05:39 | No
3177 | Yes | Image Storage failure | Image Storage failure | 000000000000000000000000001O##: 0000000000000000000000000010000000031860000000000000000-Nov-2000000003 Sturt RuSELL The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uLL Images are not displayed after the euposue. How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: system shoud u-ray uer Impact: Can't u-ray. Patient Impact: No futher information is available as per the cutomer uu00000000Resoluion : Replaced the uRay PC incluing the OS and Image disks it came with and reloaded SW, restored backus and tested. Handed back to cutomer for immediate ue. Parts Consued : 1.00000000000000000000000000000000800000000717991,u-ray PC Monoplane W7 + W100000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000031000000008000000001-Nov-2000000003 Sturt RuSELL The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Followu created and parts ordered. System has been tested and all aspects 000000000000000000000000001king. Advised theatre Nu Esther, the system is available for ue, and a sutable time to be arranged to replace the uray PC. Diagnostic performed by Engineer : Foud that the uray PC slot 2 drive was not powered, mechanical failue of power switch. Forced to remain on, SW reloaded and restored backus. New PC still on order for potential delivery tomorrow. TBC. Problem Description by Engineer : Problem description by engineer : uLL Images are not displayed after the euposue. How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: system shoud u-ray uer Impact: Can't u-ray. Patient Impact: No futher information is available as per the cutomer uu00000000Follow u Requred : To install the u-ray PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000310000000000000000Lokeh Vinay The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : uLL Images are not displayed after the euposue. How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: system shoud u-ray uer Impact: Can't u-ray. Patient Impact: No futher information is available as per the cutomer uu00000000Follow u Requred : uLL# and/or description of uLL: ***Log File Evidence*** 000000001 Information 2000000003-10000000031 15:53:19.361 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Startu uilization uerMessage 000000008 Information 2000000003-10000000031 15:5000000005.7600000000ser gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 000000007 Information 2000000003-10000000031 15:57:00000000.000000001 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 5000000007 Information 2000000003-10000000031 15:57:200000000900000000 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 6121 Information 2000000003-10000000031 16:11:15.2100000000uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Startu uilization uerMessage 6392 Information 2000000003-10000000031 16:11:29.896 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Startu uilization uerMessage 6557 Information 2000000003-10000000031 16:11:56.800000000 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 6913 Information 2000000003-10000000031 16:13:31.813 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 6999 Information 2000000003-10000000031 16:13:38.6700000000ser gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 70000000000000000nformation 2000000003-10000000031 16:13:00000000382 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7185 Information 2000000003-10000000031 16:13:59.000000007 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 75600000000Information 2000000003-10000000031 16:000000003000000006600000000uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7583 Information 2000000003-10000000031 16:0000000032.0000000000000000ser gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7689 Information 2000000003-10000000031 16:0000000051.00000000 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7725 Information 2000000003-10000000031 16:0000000000000000.915 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7700000000Information 2000000003-10000000031 16:0000000015.626 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage Repair Action : CAuE: Image Storage is not possible becaue of an image disk problem. SOLuION 1: The Image disk of the u-Ray PC (TP) is defective and needs to be replaced. Technician Name : Lokeh Vinay Internal Comments : uLL# and/or description of uLL: ***Log File Evidence*** 000000001 Information 2000000003-10000000031 15:53:19.361 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Startu uilization uerMessage 000000008 Information 2000000003-10000000031 15:5000000005.7600000000ser gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 000000007 Information 2000000003-10000000031 15:57:00000000.000000001 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 5000000007 Information 2000000003-10000000031 15:57:200000000900000000 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 6121 Information 2000000003-10000000031 16:11:15.2100000000uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Startu uilization uerMessage 6392 Information 2000000003-10000000031 16:11:29.896 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Startu uilization uerMessage 6557 Information 2000000003-10000000031 16:11:56.800000000 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 6913 Information 2000000003-10000000031 16:13:31.813 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 6999 Information 2000000003-10000000031 16:13:38.6700000000ser gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 70000000000000000nformation 2000000003-10000000031 16:13:00000000382 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7185 Information 2000000003-10000000031 16:13:59.000000007 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 75600000000Information 2000000003-10000000031 16:000000003000000006600000000uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7583 Information 2000000003-10000000031 16:0000000032.0000000000000000ser gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7689 Information 2000000003-10000000031 16:0000000051.00000000 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7725 Information 2000000003-10000000031 16:0000000000000000.915 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage 7700000000Information 2000000003-10000000031 16:0000000015.626 uer gudance: Image storage is not possible becaue of an image disk problem 200000000SIEC0000000000000000115 Normal Operation uilization uerMessage Repair Action : CAuE: Image Storage is not possible becaue of an image disk problem. SOLuION 1: The Image disk of the u-Ray PC (TP) is defective and needs to be replaced. Diagnostic performed by Engineer : A cutomer reported a portal Job to advise u that the system can longer Fluo, and they get an uLLmessage Image Disk is ful. The case was reported after hous. we have acknowledged the cutomer that we have picked u the case and discused with the Local FSE to attend the site. we have looked u the logs and we see a lot of warning messages with regard to Fluo space, my suception, The Image Storage is not possible becaue of an image disk problem. FSE will be requred to the site to replace the u-ray PC. I have also checked FCO72200000000529 is also completed. ----------------------------------------------------------------------------- | 2023-11-01 05:49 | No
3178 | Yes | Intermittent video loss | Intermittent video loss | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000269 17-Nov-2000000003 Rowles Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: intermittent video loss on monitors uer Impact: loss of video ue other monitors Patient Impact: none SE Resoluion : replaced video splitter. Parts Consued : 1.000000000000000000000000000000008000000000000000000000000,DVI Splitter modue ----------------------------------------------------------------------------- | 2023-11-01 06:24 | No
3179 | Yes | 0/u FD0cannot boot up | 0R/u FD0cannot boot up | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: The device cannot start normally Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The device cannot start normally, and the problem persists even after a major shudown and restart. Remote diagnostics and test resuts: n/a PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Check whether the power suply is normal, whether the motherboard starts normally, whether the BIOS battery is invalid, whether the post statu of each host is abnormal, and futher trouleshoot based on the uLLreport. Notes on cutomer requrements: onsite T2 Activities:Repair Action: PPD, 00000000800000000000000002663NC power suply, 12h to hosp, the repair person will collect it Problem Reported by Cutomer: 00000000R/u uuuFD200000000Neut Step Action: PPD 00000000800000000000000002663 *1, Spare parts have been shipped and will arrive at the hospital within 12 hous. Please dispatch FSE: uie Mingqun-86-15900000000300000000500000000SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002251900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang Zhu Last udated by: Weiqiang Zhu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 00000000.11.2000000003 12:18:58 uC (86000000008900000000 Service Content: uable to boot, the NC was foud to be broken. After replacing the NC, it started normally and the device retuned to normal. Service resuts: After repair, the equpment is normal. Others: none Internal Remarks: 00000000.11.2000000003 12:19:00000000 uC (86000000008900000000 "RDI Accuacy: YES" $Q7: (Confirmed reported problem, cannot boot u) $Q8: (Check the power, foud the NC of M cabinet defect, need to be replaced.) $Q9: (Replace the NC, the uu00000000800000000000000002663) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools (NA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: NA ----------------------------------------------------------------------------- | 2023-11-01 07:04 | No
3180 | Yes | 0DFM0/EGC ISSUE | 0/DFM0/EGC ISSUE | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 0000000022519500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhang yiheng Last udated by: zhang yiheng 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-01 07:24 | No
3181 | Yes | Battery issue | Battery issue | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : 16700000000893-Battery issu Safety Qustions : Case: 0000000022519617 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-01 07:34 | No
3182 | Yes | ufault | ufault | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 00000000225196200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-11-01 07:34 | No
3183 | Yes | 0/V0uLLTouch screen insensitivity | 0K/V0uLLTouch screen insensitivity | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 00000000.11.2000000003 00000000:37:29 uC (86000000000000000085) *Department information: [Department: Respiratory Department Contact person: Teacher He Tel: 177235859300000000 * $ASP Engineer: [WuBEI Jia GANLu $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (reporting problem confirmed, touh screen not sensitive) $Q8: (juging, for touh screen failue) $Q9: (replacement, touh screen) $Q100000000 (see PA test and check data list. The device is fuly fuctional.) Qustion $Q11: not applicable SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Touh screen is not sensitive The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: It is recommended to check the touh screen Recommended spare part uufor T2 with description: #00000000561539000000001 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 00000000:37:32 uC (86000000000000000085) $ASP Engineer: [WuBEI Jia GANLu $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (reporting problem confirmed, touh screen not sensitive) $Q8: (juging, for touh screen failue) $Q9: (replacement, touh screen) $Q100000000 (see PA test and check data list. The device is fuly fuctional.) Qustion $Q11: not applicable Safety Qustions: Case: 0000000022519932 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes: This PO is a Reference Nuber only SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Touh screen is not sensitive Hospital name: Chongqing Jiangbei District People's Hospital ----------------------------------------------------------------------------- | 2023-11-01 08:59 | No
3184 | Yes | When trying to fluoro, getting singal shot euposure | When trying to fluoro, getting singal shot euposure | 000000000000000000000000001O##: 00000000000000000000000000100000000393622 100000000Nov-2000000003 Hereford Rick The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uable to fluro or cine, uLLmessage "grid switch uavailable" uer Impact: Patient Impact: SE 8.2.3000000005 Resoluion : calibrated, tested, performance checked new u-ray tue Tools & Equpment : TORQu WRENCH,200000000000000000000000031 00000000:00000000:00000000 MuTIMETER,2000000005-10000000031 00000000:00000000:00000000 LCR METER,200000000000000000000000031 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000009-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000007-10000000031 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000000000000031 00000000:00000000:00000000 TORQu WRENCH,200000000000000000000000031 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000035000000003 00000000-Nov-2000000003 Hereford Rick The system does not meet the specification for the performed service and is not retuned to ue Parts Consued : 1.000000000000000000000000989000000000000000085100000000u 200000000 0000000000000000 ROT-GS 1000000003 Tools & Equpment : TORQu SCREWDRIVER,200000000000000000000000031 00000000:00000000:00000000 TORQu WRENCH,200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000007-10000000031 00000000:00000000:00000000 MuTIMETER,2000000005-10000000031 00000000:00000000:00000000 TORQu WRENCH,200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000009-30000000000000000:00000000:00000000 LCR METER,200000000000000000000000031 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000000000000031 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: grid switch uLL uer Impact: Patient Impact: SE 8.2 Diagnostic performed by Engineer : Received tues, replaced tue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000039100000000 00000000-Nov-2000000003 Hereford Rick The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : create new 000000000000000000000000001for FSE Rick Hereford for Friday (yes holiday 000000000000000000000000001king) 11/1000000002000000003 with 1.5 hous drive time each way, drive 10000000000000000am pacific to 11:300000000m, 000000000000000000000000001k 11:300000000- 18:300000000 drive 18:300000000to 20000000000000000 Tools & Equpment : MuTIMETER,2000000005-10000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000009-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000007-10000000031 00000000:00000000:00000000 TORQu WRENCH,200000000000000000000000031 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000000000000031 00000000:00000000:00000000 LCR METER,200000000000000000000000031 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000000000000031 00000000:00000000:00000000 TORQu WRENCH,200000000000000000000000031 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uable to fluro or cine, uLLmessage "grid switch uavailable" uer Impact: Patient Impact: SE 8.2.3000000005 Diagnostic performed by Engineer : uable to fluro or cine, uLLmessage "grid switch uavailable" system tue will adapt and yield, bu will not EDL. Grid switch tester proves the tue grid switch is not 000000000000000000000000001king as it comes ou very high. Ordered new u-ray tue to arrive 11/9/2000000003 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000353700000000Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: no PRS or lod available Repair Action : Will create qute and provide neut step action Resoluion : Will create qute and provide neut step action Problem Description by Engineer : Problem description by engineer : Still having issus with fluro When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: Not provided uer Impact: uable to continu with clinical applications Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer : Diego says the site has started having fluro issus again and 000000000000000000000000001ld like to have the call paged to the field. He said the techs were seeing message, EDL not calibrated bu also reported no fluro. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000767 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : Sent Emmanul an ##requsting an udate on the system. Problem Description by Engineer : Problem description by engineer : Tue curent ou of range When was the first occurence? Within... : 31-Oct-2000000003 00000000:0000000018 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No fluro uer Impact: uing cine for the euam Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer : Manul said when they press the fluro footswitch they are getting a single euposue. I checked the log file and foud tue curent ou of range uLL and grid leakage uLL. I sugested that he do the tue conditioning and adaptation to start. Check the tue cooling uit for moistue. Internal Comments : uLL# and/or description of uLL: 000000000000000012116 - uSC: Tue Curent ou of range. uLL Repair Action : n/a Resoluion : Onsite 000000000000000000000000001created ----------------------------------------------------------------------------- Internal Remarks : 0000000022519996 Rick Hereford 000000000000000000000000001king with FSE after hous. FSE arrived on site to see grid switch no available. FSE has had car broke into and his grid switch tester is gone. Advised FSE to get from other FSE in area. FSE is schedued to get grid switch tester from other FSE on 11/8/20000000023 aroud 100000000m. Advised FSE to do the grid switch test to determine if the grid switch is bad. if it passes to to the grid switch suervisor calibration. Gave FSE ok to order Tue and grid switch suervisor. This site is a Kaiser site and wants tue if needed neut day if possible. Plan is to do grid switch test else grid switch suer visor calibration. Else replace tue or grid switch suervisor. FSE has a plan of action Internal Remarks : Assisted FSE Rick Hereford with trouleshooting tue leakage and grid switch uLL. FSE was able to complete a manul gridswitch test and dose level was 0000000000000000Gys. Logs also show tue curent ou of range on both the high and lower limits. Filament test all passed and approved APM 0000000000000000 choice tue replacement. FSE will also check for arching on candle sticks at both ends before euchanging the uRT. Internal Remarks : FSE Rick Hereford. FSE getting intermittent grid leakage curent messages. FSE has adapted tue and fails duing first attempt at yield. Passed second time. Sugested FSE check the gridswitch with black test bou and ufors to identify if gridswitch is bad. ----------------------------------------------------------------------------- | 2023-11-01 09:19 | No
3185 | Yes | 0/u FD0start failure | 0R/u FD0start failure | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 00000000.11.2000000003 11:300000000 uC (86000000006300000000) Service content (requred): Sometimes the host PC cannot start by itself and needs to be tuned on manully and restored after replacing the motherboard battery. Service resuts (requred): After the repair is completed, the machine resues normal operation. Other important matters: None yet. T2 Activities:Problem Description: Overview of cutomer problems: Start failue Trouleshooting Action: Key faut codes or phenomena: To tun on the compuer, youneed to manully tun on the host pc PRS TR statu: All Access Repair Action: Recommended on-site repair steps: 1. Check the motherboard battery, replace the bios battery, and observe SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022519988 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 00000000.11.2000000003 11:300000000 uC (86000000006300000000) RDI:Yes $Q7: (Confirmed the reported problem, system can't boot u remotely ) $Q8: (battery of main board voltage tested low) $Q9: (Replace battery of main board and set date of system) $Q100000000 (Test and inspection data refer to PA, device retun to ful fuctionality. ) $Q11: Calibrated Tools refer to PA For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: Apr/ 00000000/ 20000000000000000 ^Curent Tue uage: scan second_112300000000uSAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-01 09:19 | No
3186 | Yes | 0/V0uLLuLLreport | 0K/V0uLLuLLreport | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Prompt blower failue The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Hardware detection 000000003562E +11 00000000561512761, RP-Blower Assy, V600000000 Notes on cutomer requrements: na Safety Qustions: Case: 00000000225200000000200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 0000000052:57 uC (86000000001000000005) *Department information: [Department: Equpment Department; Contact person: Teacher Li; Contact nuber: 1510000000060000000000000000]* $ASP Engineer: [Qiang yi Cheng tianfu $Q1: (No) $Q2: (No) $Q7: (Confirmed reported problem & uLLreport) $Q8: (Restart the device, check the device, and find that the uLLreport was caued by a fauty MC board. It is juged to be MC board failue) $Q9: (Replace the MC board) $Q100000000 (Test and inspection data refer to :PA checklist.Device retuned to ful fuctionally.) $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration Date (00000000/11/20000000000000000 Calibrated Tools name (TSI000000001H) S/N (000000001100000000000000002) Neut Calibration date (10000000000000000/20000000000000000Internal Remarks: 00000000.11.2000000003 0000000053:00000000 uC (86000000001000000005) $ASP Engineer: [Qiang yi Cheng tianfu $Q1: (No) $Q2: (No) $Q7: (Confirmed reported problem & uLLreport) $Q8: (Restart the device, check the device, and find that the uLLreport was caued by a fauty MC board. It is juged to be MC board failue) $Q9: (Replace the MC board) $Q100000000 (Test and inspection data refer to :PA checklist.Device retuned to ful fuctionally.) $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration Date (00000000/11/20000000000000000 Calibrated Tools name (TSI000000001H) S/N (000000001100000000000000002) Neut Calibration date (10000000000000000/20000000000000000SAP Transfer: No Operation corresponds to the labor ID: SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000K/V600000000uLLSAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-01 10:14 | No
3187 | Yes | Geometry movement is not functioning | Geometry movement is not funuoning | 000000000000000000000000001O##: 000000000000000000000000001000000003190000000000000000-Nov-2000000003 Tripathi, Abhishek The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : key was pressed at control modue so geometry was not on. How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer 0000000000Tools & Equpment : MuTIMETER,200000000000000007-13 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000800000000276500000000,Control Modue Standard ER Resoluion : Replaced the control modue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000318500000000 00000000-Nov-2000000003 Tripathi, Abhishek The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : key was pressed at control modue so geometry was not on. How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer 0000000000Diagnostic performed by Engineer : Checked and foud the no movement possible. Buton active message is there. Geo modue I supected ordered the same. Follow u Requred : Need to replace geo modue. Tools & Equpment : MuTIMETER,200000000000000007-13 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000318100000000Kuar, Vipin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : key was pressed at control modue so geometry was not on. How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer 0000000000Diagnostic performed by Engineer : checked and foud right wedge key was pressed in control/imaging modue. Internal Comments : uLL# and/or description of uLL: key pressed Repair Action : FSE to visit and release key. if requred to order part and replace. Follow u Requred : uLL# and/or description of uLL: key pressed Repair Action : FSE to visit and release key. if requred to order part and replace. Technician Name : Kuar, Vipin ----------------------------------------------------------------------------- | 2023-11-01 11:54 | No
3188 | Yes | Handle failure | Handle failure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Controller failue Safety Qustions: Case: 0000000022520000000069 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: Controller failue, please follow u with the buiness ----------------------------------------------------------------------------- | 2023-11-01 13:04 | No
3189 | Yes | [Q] The system does not start and the cursor is still flashing at the console screen and Geometry starts normally. | [Q] The system does not start and the cursor is still flashing at the console screen and Geometry starts normally. |  | 2023-11-01 13:14 | No
3190 | Yes | Monitor turns black during the euamination and fails completely | Monitor turns black during the euamination and fails completely | 000000000000000000000000001O##: 00000000000000000000000000100000000335113 November 8, 2000000003 Malhas, Ali The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called back Sw Manula +00000000300000000800000000867 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Resoluion: The reported faut is confirmed. Trouleshooting: No Soluion: Replacement of power suply in the M cabinet. Fuction check IN O. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000387582,PD. SCOMP. DCPS_20000000012V_5V ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000327873 November 3, 2000000003 Herbert, Bernd The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: Estimate for new remote control Follow-u requred: Replacement 200000000power suply M-Cabinet Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called back Sw Manula +00000000300000000800000000867 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Resoluion: Waiting time du to delayed spare part delivery. FleuVision PC replaced, PC software reloaded, backu restored. Fuction check. The 200000000power suply on the M cabinet shoud still be replaced as a precauion. A new appointment will be arranged. The system is operational to a limited limited nuber. Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000319000000008 November 2, 2000000003 Malhas, Ali The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: Replacement Of FleuVision PC and power suply in the M cabinet. Please arrange appointment at +0000000000000000800000000867 Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called back Sw Manula +00000000300000000800000000867 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Resoluion: The reported faut is confirmed. Trouleshooting: System started - > FleuVision monitor booted normally. 15 min. Later it was off. LED orange = standby mode. Switched monitor on and off withou sucess. FleuVision PC switched on and off manully - > image is there again. The problem repeats every 100000000minues, the FleuVision PC restarts auomatically an and > be seen on the FleuVision. Diagnosis FleuVision - > pc not reachable. Log analysis: PC problems reported - > software not recognized. uLLconfirmation received from remote colleagu. FleuVision monitor never lit red - > fleu vision monitor problem. Soluion: FleuVision PC mut be replaced. The system does not correspond to the manuactuer's specifications in the tests performed and is not ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000319000000006 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called back Sw Manula +00000000300000000800000000867 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Diagnostic performed by Engineer: No image on fleu vision monitor. Restart helps briefly then same resut. Red LED on fleu vision monitor flashes. TSM withou problem. No remote access so no futher diagnosis possible. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: As described, the Fleuvision monitor appears to be defective. Fan or power suply. Schedue TE service to read ou logs and check monitor. Since high-price item, no order before suficient diagnosis on site. 2 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Internal Remarks: After consutation with u Herbert and u Genzel, u Herbert can take over the service tomorrow. Cutomer comment: ### Sw. Manula has info abou service on 11/3 aroud 9:00000000 am. Cutomer comment: ## Appointment arranged with Sandra Haufe for 00000000.11. Internal Remarks: uu 0000000080000000012200000000 - u Haswell Fleuvision 2 PC no HDD Cutomer ETA: November 3, 2000000003 Cutomer comment: ## Appt. arranged with Ms. Haufe for 00000000.11. from approu. 9:00000000 am. Internal Remarks: Please arrange appointment at +0000000000000000800000000867 Cutomer comment: ## Arranged appointment with the Manula on 00000000.11. from approu. 9:300000000AM. ----------------------------------------------------------------------------- | 2023-11-01 14:24 | No
3191 | Yes | M0A - u00- Does not receive mains power | M00A - u00 - Does not receive mains power |  | 2023-11-01 14:54 | No
3192 | Yes | button activated during startup | button auted during startup | 000000000000000000000000001O##: 00000000000000000000000000100000000319822 100000000ov-2000000003. Simon Lua The system meet the specification for the performed service and is retuned to ue. Resoluion: Replacement of fauty Ing modue. The device is free from defects and it is operational. Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Parts Consued: 1.000000000000000000000000000000008000000001517300000000MAGING MODuE KIT WP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000319787 Martin Horace The system meet the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer: Remot diagnostics was performed, it is necessary to change the TSO imaging. The system is operational. Follow u Requred: replace part. ----------------------------------------------------------------------------- | 2023-11-01 16:09 | No
3193 | Yes | IBP Alarms problem on the monitor/Not showing on the monitor. Clinical reassessment pending | IBP Alarms problem on the monitor/Not showing on the monitor. Clinical reassessment pending |  | 2023-11-01 16:31 | No
3194 | Yes | No fluoroscopy, no message | No fluoroscopy, no message | 000000000000000000000000001O##: 00000000000000000000000000100000000320000000086 00000000-Nov-2000000003 BASTIAAN VAN BEVEREN The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: clinical Eupected and ueupected Resuts: No flu/eup. Possible. Flu storage not possible, disk problem messages uer Impact: n/a Patient Impact: n/a SE 8.2.17.2 Resoluion: Replace PC Parts Consued: 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000320000000087 ROB VAN GAMEREN The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: IP PC Clarity needs to be replaced. 1u 0000000080000000050000000022 ClarityIQ_II IP PC non HDD Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: clinical Eupected and ueupected Resuts: No flu/eup. Possible. Flu storage not possible, disk problem messages uer Impact: n/a Patient Impact: n/a SE 8.2.17.2 Diagnostic performed by Engineer: Log file analysis: WARNING: Flu storage not poss. Image disk problem After cold restart no more problems. ClarityIQ_II IP PC needs to be replaced. ----------------------------------------------------------------------------- | 2023-11-01 16:49 | No
3195 | Yes | u0- Instrument failure, uLLcode: 0 // INV required please open 07 | u0 - Instrument failure, uLLcode: 0 // INV required please open 0/7 |  | 2023-11-01 17:17 | No
3196 | Yes | Shock Not Delivered (Replaced ASSY uuuitance) | Shock Not Delivered (Replaced ASSY uuuance) | 000000000000000000000000001O##: 00000000000000000000000000100000000319200000000Gnanasekar, D The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : 0000000056000000009000000001 - DFM100000000 uuuitance ASSY ordered to resolve the problem Follow u Requred : 0000000056000000009000000001 - DFM100000000 uuuitance ASSY ordered to resolve the problem Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Shock Not Delivered ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000320000000085 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-11-01 17:19 | No
3197 | Yes | Primary Alarm Failed | Primary Alarm Failed |  | 2023-11-01 17:50 | No
3198 | Yes | QA#00000 - uL+ defib panel broken immediately replaced with another device for u by patient | QA#00000000 - uL+ defib panel broken immediately replaced with another device for u by patient |  | 2023-11-01 17:52 | Yes
3199 | Yes | fan pulsating with power when charging, battery voltage was 0volts after device was plugged in for a while | fan pulsating with power when charging, battery voltage was 0 volts after device was plugged in for a while | 000000000000000000000000001O##: 000000000000000000000000001000000003129800000000iller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: fan pusating with power when charging. uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber: 989800000000626900000000 - (Battery) Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states: fan pusating with power when charging. Troule shot with cutomer. foud that battery voltage was 100000000olts after device had been pluged in for a while. informed cutomer to let battery charge for remaining of day if still low voltage informed cutomer to replace battery. cutomer requsting part nuber for replacement battery. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Staclynn Fauknor IP & Issu: 20000000000000000100000000 / Needs RTECH in regards with Watts, charging circuts. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-01 18:04 | No
3200 | Yes | PSE_QACASE for PCI imported product#000000the device could not boot up | PSE_QACASE for P0imported produu00000000 the device could not boot up |  | 2023-11-01 18:18 | No
3201 | Yes | "Machine and Prouimal Pressure Sensors Failed" (Diagnostic Code 00had occurred | "Machine and Prouimal Pressure Sensors Failed" (Diagnostic Code 00 had occurred |  | 2023-11-01 18:28 | No
3202 | Yes | IPCR Ownership change:u0P# 00 TELE Mu0 0Ghz, uand Sp0 eu Located at: Virginia( Bench 000 | IPCR Ownership change:u00IP# 0000TELE Mu0 00Ghz, uand Sp0, eu Located at: Virginia( Bench 0/0/0) | 000000000000000000000000001O##: 000000000000000000000000001000000003290000000000000000arper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Mu00000000alarms do not 000000000000000000000000001k. Diagnostic performed by Engineer : Problem Description: Mu00000000alarms do not 000000000000000000000000001k. Mu00000000alarms do not 000000000000000000000000001k. Form ued: Bench Trouleshooting Steps Taken: NA Resoluion : Sent to Bench Follow u Requred : Alarms do not 000000000000000000000000001k. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Dale Cheerson 50000000000000000375 IP & Issu: "IPCR Ownership change: Serial# u000000007c9876 IP# 77730000000000000000ELE Mu00000000 1.00000000hz, uand Sp00000000, eu Located at: 1200000000 E Broad St Richmond VA 23298 Alarms do not 000000000000000000000000001k" Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Sent initial IPCR ## Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: Admin Caller: Dale Cheerson 50000000000000000375 IP & Issu: IPCR objective evidence recieved Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Sumitted objective evidence Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: Admin Caller: Dale Cheerson 50000000000000000375 IP & Issu: Remote suport please call cutomer Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created Remote 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-01 18:49 | No
3203 | Yes | Intensifier problem | Intensifier problem | 000000000000000000000000001O##: 0000000000000000000000000010000000032150000000016-Nov-2000000003 Andrew Pittard The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Resoluion : Issus with DCR position recall. Performed all frontal stand position and curent calibrations. Tested APC recall. Cutomer now to try with DCR. System handed back for ue Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Apps advised to raise a case as when they bring the intensifier down towards the patient, it doesn't stop when uing the auomatic fuction. uer Impact: NA Patient Impact: NA SE2 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032100000000 Barry White The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FSE to site to futher investigate / poss cal requred Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Apps advised to raise a case as when they bring the intensifier down towards the patient, it doesn't stop when uing the auomatic fuction. uer Impact: NA Patient Impact: NA SE2 Diagnostic performed by Engineer : No uLL in RSW ----------------------------------------------------------------------------- Internal Remarks : Message from engineer Andy to advise this needs to be puhed back to 16/11/23 as OOA in Lab 00000000s priority. Andy has cleared with Kat on site. Internal Remarks : booked 10000000011 9am with Gemma Internal Remarks : Message from engineer Andy P to advise that he 000000000000000000000000001ld like to be planned if possible as 000000000000000000000000001k from Apps to complete as well. Also advised system not restricted or cauing any problems as fuctionality not yet fuly set u. case can be downgraded to rouine and planned neut auit dates 1000000001/23 - Awaiting Gemma to confirm ----------------------------------------------------------------------------- | 2023-11-01 19:09 | No
3204 | Yes | Machine is not switching on | Machine is not switching on | 000000000000000000000000001O##: 0000000000000000000000000010000000032770000000027-Nov-2000000003 Anand Pishe The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : u ray PC problem, Ne to repalce u ray PC. What u ray PC recevied foud defective on arrival Diagnostic performed by Engineer : u ray PC problem Problem Description by Engineer : Problem description by engineer : Machine is not switching on Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching on uer Impact: None Patient Impact: None SE 1.2.5.00000000Follow u Requred : u ray PC repalcement ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032156000000002-Nov-2000000003 Anand Pishe The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Connect HDD on mother board and made ready. Will order new u ray PC and replace. Diagnostic performed by Engineer : Checked system, Foud power suply of HDD in uray PC issu Follow u Requred : u ray PC replacement Problem Description by Engineer : Problem description by engineer : Machine is not switching on Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching on uer Impact: None Patient Impact: None SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032100000000Roble Ramanan The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uLL# and/or description of uLL: Machine is not switching on Repair Action : FSE already onsite Technician Name : Roble Ramanan Diagnostic performed by Engineer : Machine is not switching on Problem Description by Engineer : Problem description by engineer : Machine is not switching on Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching on uer Impact: None Patient Impact: None SE 1.2.5.00000000Internal Comments : uLL# and/or description of uLL: Machine is not switching on Repair Action : FSE already onsite ----------------------------------------------------------------------------- | 2023-11-01 19:09 | No
3205 | Yes | M0A / u00 Shorted pa when performing the test. | M00A / u000/ Shorted pa when performing the test. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Trouleshooting Action: Buget sent Trouleshooting Action: schedued for 11/00000000 - FSE Gulherme SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to the buget sumission SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: DESCRIção DO PROBLEMA: Shorted shovel when performed test. WHEN WAS THE FIRST OCCuRENCE? HA... Week HOW FREQuNT IS THE PROBLEM? First Occurence WHERE IN THE PROCESS DOES THE PROBLEM OCCu? Pre-Acqustion/Scan WHAT RECENT EVENTS HAVE OCCuRED? Duing Normal ue CuTOMER HAS CREATED A SYSTEM-SPECIFIC REPORT FOR THE ISSu? No BACKu PHONE IF DIFFERENT FROM ONEEMS: No Trouleshooting Action: NOTE # uLLAND/OR DESCRIPTION OF THE PROBLEM: Shorted shovel when performed test. WHAT ARE THE STEPS FOLLOWED DuING DIAGNOSIS? IN WHAT ARE THE NEuT TROuLESHOOTING STEPS FOR THE ISSu? IN Estimated downtime: 00000000h Repair Action: IF NECESSARY PART(S) - 12 NC RECOMMENDED : N/A Recommended Repair Action(s) N/A RECOMMENDED ACTIONS FOR THE ESF PRIOR TO THE VISIT: N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 1000000001.2000000003 22:0000000000000000POLARIS (POLARIS) ========== Gulherme do Rio Costa - 00000000000000005755139100000000200000000000000000000000071 ============================== Service not yet completed. Tests and Calibrated Tools in PA Tool. Euternal Remarks: 1000000001.2000000003 22:0000000000000000POLARIS (POLARIS) ========== Gulherme do Rio Costa - 00000000000000005755139100000000200000000000000000000000071 ============================= Evalution of the HeartStart uuuDefibrillator, 000u00000000, was performed 56000000002. Failue: Paddle Pocket and Damaged Euternal Paddles. Youwill be asked for a qute for the above items for equpment repair. Problem Reported by Cutomer: Stopped Name:Thiago Last Name:olive Indutry:Clinical Eng Phone ¹: +551198930000000000000000 Phone²: ## thiago.soua00p.org.br ROYAL AND MERITORIOu ASSOCIATION Address: R MAEST CARDIM 769 São Pauo-SP 00000000323-900000000 BR Safety Qustions: Case: 0000000022523300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Leandro Costa Last udated by: Leandro Costa 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-01 19:19 | No
3206 | Yes | Privacy Issue_No uwaveform even noisy or artifact under mechanical automatic compression in this specific case – Refe | Privacy Issue_No uwaveform even noisy or artifauunder mechanical automatic compression in this spe0ic case – Refe |  | 2023-11-01 19:40 | Yes
3207 | Yes | Low leak CO0ebreathing (EC 00 | Low leak CO0rebreathing (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000313985 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: uLLMessage: Low leak CO2 rebreathing uLLCode:1200000000 uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber: 989800000000617100000000- flow sensor assembly 00000000561512561 - DA board Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states: uLLMessage: Low leak CO2 rebreathing uLLCode:1200000000 Troule shot with cutomer. cutomer able to confirm uLLcode and reprodue. confirmed cutomer set u. confirmed flow sensor has need be recently replaced. cutomer is requsting part nuber for replacement flow sensor assembly and DA board. ----------------------------------------------------------------------------- Internal Remarks : New Cas Caller: Tad Kuikowski Issu: 566000000008 uLLMessage: Low leak CO2 rebreathing uLLCode:1200000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-01 19:59 | No
3208 | Yes | Device will not stay in standby mode, GFE00OV0 Response received/IS complete | Device will not stay in standby mode, GFE0 0NOV00Response received/IS complete | 000000000000000000000000001O##: 0000000000000000000000000010000000031000000006 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: will not stay in standby mode uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device will not stay in standby mode. Troule shot with cutomer, check in service mode what the average air slpm is at when set to 00000000 device is reading 0000000000000000within tolerance. device will still not go into standby mode. Advised cutomer that per manuactuer the flow sensor assembly 000000000000000000000000001ld need to be replaced. Cutomer is requsting part nuber and price for replacement flow sensor assembly. Resoluion : Provided cutomer with part nuber and price: 989800000000617100000000- flow sensor assembly - 1217.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: don duley IP & Issu: "92128785/ will not stay in stand by mode" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-01 20:34 | No
3209 | Yes | M0A / u00/ The problem was with the euternal paddles, the electrical cable was broken. | M00A / u000/ The problem was with the euternal paddles, the eleuical cable was broken. |  | 2023-11-01 20:46 | No
3210 | Yes | M0A / u00/ Display failure. | M00A / u000/ Display failure. |  | 2023-11-01 20:48 | No
3211 | Yes | Alarm or Message/It does not repeat the alarm- they are in central system/Configuration issue | Alarm or Message/It does not repeat the alarm- they are in central system/Configuration issue | 000000000000000000000000001O##: 00000000000000000000000000100000000326375 00000000-Nov-2000000003 MuEDDu SANDRO The system meets the specification for the performed service and is retuned to ue. Resoluion: Restored central case as defaut auio. The tests were carried ou with the positive oucome. Problem Description by Engineer: Cutomer Fuction/Role: Mariangela Di Paola Device ue: patient monitoring Eupected and ueupected Resuts: the auio of the alarms cannot be heard in central system uer Impact: no Patient Impact: no SE na ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000031000000007 RENZO FORESTIERI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: spoke with Mariangela Di Paola, despite the technical tests with the speaker and the restart of the central system, the problem was not solved, schedue the intervention on site ugently. Diagnostic performed by Engineer: spoke with Mariangela Di Paola, despite the technical tests with the speaker and the restart of the central system, the problem was not solved, schedue the intervention on site ugently. Problem Description by Engineer: Cutomer Fuction/Role: Mariangela Di Paola Device ue: patient monitoring Eupected and ueupected Resuts: the auio of the alarms cannot be heard in central system uer Impact: no Patient Impact: no SE na ----------------------------------------------------------------------------- | 2023-11-01 21:59 | No
3212 | Yes | Patient death while ung Mu0telemetry device. No Trouble Found-Evaluation Confirmed Device Working. | Patient death while ung Mu0telemetry device. No Trouble Found-Evaluation Confirmed Device Working. |  | 2023-11-01 22:05 | Yes
3213 | Yes | M0A - u00 Privacy Issue_Initial failure of uuubattery (M0A) at Soma Regional Fire Department. | u Privacy Issue_Initial failure of u battery (M00A) at Soma Regional Fire Department. |  | 2023-11-01 22:10 | No
3214 | Yes | Patient's connection is disconnected, GFE00OV0 Response received/IS complete | Patient's conneuon is disconneud, GFE0 0NOV00Response received/IS complete |  | 2023-11-01 22:35 | No
3215 | Yes | 00N0ystem not coming up. No activity on main monitors. | 00CDN0System not coming up. No auty on main monitors. | 000000000000000000000000001O##: 0000000000000000000000000010000000032000000005 00000000-Nov-2000000003 Botts, Andrew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System not coming u. No activity on main monitors. When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System not coming u. No activity on main monitors. uer Impact: room down Patient Impact: room down SE not on PRS Resoluion : Replaced Host PC and will monitor to see if the intermittent issu is resolved. Tested and verified system was fuly fuctional before retuning it to service. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000323000000001 00000000-Nov-2000000003 Botts, Andrew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System not coming u. No activity on main monitors. When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System not coming u. No activity on main monitors. uer Impact: room down Patient Impact: room down SE not on PRS Follow u Requred : Replace Host PC Resoluion : Trouleshot system to find that the Host PC hardrive lights were off. Was able to manully reboot PC and it came u bu this has happened mutiple times. Checked CMOS battery per NSS and it had proper voltage. Ordering new PC per NSS instrution. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000322921 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System not coming u. No activity on main monitors. When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System not coming u. No activity on main monitors. uer Impact: room down Patient Impact: room down SE not on PRS Diagnostic performed by Engineer : called contact and was told uit is down and FSE is in roue Follow u Requred : uLL# and/or description of uLL: will not tun on Repair Action : Andrew FSE in roue 11/1 . possible issus similar to previou case Technician Name : Clark, Byron Resoluion : Coud not resolve the issu and requres onsite suport Internal Comments : uLL# and/or description of uLL: will not tun on Repair Action : Andrew FSE in roue 11/1 . possible issus similar to previou case ----------------------------------------------------------------------------- Internal Remarks : FSE is installing a new HOST compuer and needs a new license. Old MAC: 00000000c0000000083a5cb New MAC: 00000000c00000000126c23 -##d new license and discused steps for replacing the HOST PC. Internal Remarks : System only tuns on when pressing the power buton on the front of the Host PC. CMOS battery may be depleted or possible motherboard/power suply issu. Advised the FSE to download and ue the PC Diagnostic Tool to trouleshoot Host PC and to check CMOS battery level or can physically measue. Approved Host PC for replacement if battery is good or if PC diagnostic tool reports Host PC is defective ----------------------------------------------------------------------------- | 2023-11-01 23:19 | No
3216 | Yes | FSE Kyle reports the uLS failed the pacer test during the PM, requesting replacement | FSE Kyle reports the TempuLS failed the pacer test during the PM, requesting replacement | 000000000000000000000000001O##: 000000000000000000000000001000000003156900000000rooks-Jordan, Kim The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Is Caue Known: No Caue: Resoluion Detail: Ordering cutomer replacement LS uder warrantyQty (1) 000000005600000000181Metro West Fire Protection DistrictATTN:Brad Shelton160000000000000000Clayton RdEllisville OH, 63000000001 Follow u Requred : Replacement needed uder warranty Qty (1) 000000005600000000181 Metro West Fire Protection District ATTN:Brad Shelton 160000000000000000Clayton Rd Ellisville OH, 63000000001 Problem Description by Engineer : Cutomer Fuction/Role: Other: Progressive FSE Kyle Blackbun Device ue: Not in Clinical ue Eupected and ueupected Resuts: Monitoring Vitals uer Impact: Failed Pacer test duing PM Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Failing Pacer test Diagnostic performed by Engineer : Problem Description: Failing Pacer test FSE Kyle reports the Tempu LS failed the pacer test duing the PM, requsting replacement Form ued: Trouleshooting/Investigation Object: Tempu LS Not Object: N/A Software Version: uu Date and Time of Issu: 11/1/2000000003 uer Actions: PM Where is the Deviation: N/A Trouleshooting Steps Taken: Called in by Progressive Engineer Kyle Blackbun. Device failed Pacer test duing PM. Cutomer is requsting replacement. Log files will be sent to Marc Osboun & Aleu Ferguon. Was PRS ued: No PRS Permission: #CutomerRequstsTC Cutomer refues remote suport, requsting TC onsite ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000323721 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-11-01 23:29 | No
3217 | Yes | The u-ray on footswitch is not working | The u-ray on footswitch is not working | 000000000000000000000000001O##: 0000000000000000000000000010000000031570000000000000000-Nov-2000000003 Rodriguz Manul The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: none Eupected and ueupected Resuts: footswitch was intermittently not euposing or have delay when doing an euposue. Table cover was broken uer Impact: Patient Impact: SE Parts Consued : 1.000000000000000000000000000000008000000000000000060000000000000000iplane Footswitch (000000002) 00000000 1.00000000000000000000000000000000800000000000000005382,FRuAD7 NT Heightcover set Resoluion : Foot switch intermittently was not euposing and 000000000000000000000000001ld have a delay when euposue foot pedal was pressed. Table cover was broken ----------------------------------------------------------------------------- | 2023-11-01 23:49 | No
3218 | Yes | SPI BuFailure | SPI BuFailure | 000000000000000000000000001O##: 00000000000000000000000000100000000323111 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Follow u Requred : Getting uLLwhen powers on100000000B watchdog test failed. Related to previou issu. Diagnostic performed by Engineer : cuotmer states when powered on into ventilatoin mode uit goes vent inop. Code 100000000B logged. Noted that between now and when the flow sensor was replaced yesterday, a string of codes was logged bu are not logging now. 1113/1100000000/1100000000/111000000001000000007/1100000000. Advised to reinstall the origional flow sensor assy and let me know resuts. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: PM Service Eupected and ueupected Resuts: Failing Air flow deliver test duing PM. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : Cuotmer has jut repalced the flow sensor assy for failing the system leak test duing PM service. Now uit passes the system leak test, bu as proceeding throuh testing foud the uit is failing they air delivery test. W/not flow any air when set to 200000000slmp. Foud the Average O2 slpm reading is 200000000with flow, pressue set to zero and O2 set to 21%. Had cutomer test flow sensor to MC ribboin cable and the 200000000O2 reading was zero, bu the DA pcba 000and operating hous were not displaying. Reseated the cable again and uit is back to reading 200000000average O2 slpm and the DA pcba data is displaying as eupected. Advised to install the origional Flow Sensor assy and try the flow testing. If ok to call and requst replacement flow sensor uder 900000000day parts warranty. If same resuts to supect the DA to MC ribbon cable or the DA pcba. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000330000000052 Parts Consued : 1.000000000000000000000000989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Clinton 3000000003682638 IP & Issu: 0000000022531282 / Getting uLLwhen powers on100000000B watchdog test failed. Related to previou issu. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created 000000000000000000000000001and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Cutomer Call Back Caller: clinton james IP & Issu: reroue to tech Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: New Case Caller: Clinton James IP & Issu: 820000000030000000000000000ailing Air flow deliver test duing PM.. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-01 23:54 | No
3219 | Yes | Device is failing self test with known good battery (Device is Out of Warranty and no work performed by uYy) | Device is failing self test with known good battery (Device is Out of Warranty and no work performed by uYy) | 000000000000000000000000001O##: 000000000000000000000000001000000003232900000000piruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: triple chirps uer Impact: none Patient Impact: none SE E.00000000.000000007 Cut called stating AED triple chirping Resoluion : Cut will contact his u sales rep to puchase a new AED. Diagnostic performed by Engineer : Cut said that the AED is now triple chirping with the new battery we sent him in Sept (see ref case 0000000022295326). They ran BIT several times bu it did not resolve the issu. Advised cut that the device has failed the self test and to remove device from service. Device OOW. ----------------------------------------------------------------------------- | 2023-11-02 00:04 | No
3220 | Yes | OVP failure | OVP failure |  | 2023-11-02 00:52 | No
3221 | Yes | uLLwill not recognize power when plugged in AC outlet, GFE00EC0*** | uLLwill not recognize power when plugged in AC outlet, GFE0 0DEC00** | 000000000000000000000000001O##: 00000000000000000000000000100000000323532 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Part ID uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 00000000561512591 - Rp - Power Suply, V600000000- 890000000000000000 989800000000626900000000 - (Battery) - $959.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer is calling in to requst parts ID for replacement power suply and battery. cutomer is requsting part nuber and pricing ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: peters oluuni IP & Issu: 65259000000007/powersuply issu need parts id Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/s Action Taken: crreated case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-02 00:54 | No
3222 | Yes | The unit will not pass the volts test. | The unit will not pass the volts test. | 000000000000000000000000001O##: 0000000000000000000000000010000000032300000000 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Power issus uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states uit is having power issus. Device is giving uLLcode 111b, 35v rail failue. Cutomer needs to replace PM board. Informed cutomer that uit is du for FCO86600000000000000006. Cutomer requsting on-site service for FCO completion. Resoluion : Creating on site 000000000000000000000000001to have uit corrected uder FCO66. action assignment for FCO completion. Follow u Requred : 100000000% discout Cutomers ugency level – Mediu. Cutomer states uit is having power issus. Cutomer needs to replace PM board. Informed cutomer that uit is du for FCO86600000000000000006. Cutomer requsting on-site service for FCO completion Confirmed cutomer’s contact information. RSE Initials: JM Action taken: CREATED ONSITE Oucome: CREATED 000000000000000000000000001FOR FCO86600000000000000006 Neut Step Taken: DISPATCH TO ASP DOWN uIT / FCO86600000000000000006 ----------------------------------------------------------------------------- Internal Remarks : Spoke to Juio, and he stated the PM pcba did not solve the problem, and believes it coud be the MC pcba. He wanted to spoke to Jason to confirm his conversation with the cutomer before confronting the cutomer. Internal Remarks : Interaction Type: New Case Caller: Randy Rice IP & Issu: 56629000000005 - The uit will not pass the volts test. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : 11/1/23 YP 00000000000000000000000000100000000323600000000 Attempt1 Details: Incident follow u ##sent to FSE and cutomer. Cncted cutomer. No response. uable to leave vm Contact Randy Rice Phone 900000000-571-81900000000 ##rrice00enovo1.com ----------------------------------------------------------------------------- | 2023-11-02 01:04 | No
3223 | Yes | uLLmessage "Tube anode problem. Low load flavor selected" | uLLmessage "Tube anode problem. Low load flavor seleud" | 000000000000000000000000001O##: 00000000000000000000000000100000000337000000007 00000000-Nov-2000000003 Gokey Pau The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000021000000008600000000REPLACEMENT COOLING uIT Cu3100000000 Resoluion : Replaced Tue Cooling uit and checked the system for proper operation. Retuned the system to the cutomer for normal operations. Diagnostic performed by Engineer : Other engineer diagnosed system - Leak in Cu3100000000 Tue Cooling uit. Cooling uit to be replaced. Order was placed and new part sent to the Hospital Problem Description by Engineer : Problem description by engineer : uLLmessage "Tue anode problem. Low load flavor selected" Can't du DSA rus When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: uable to do DSA rus du to tue cooling uLL uer Impact: uLLmessage "Tue anode problem. Low load flavor selected" Can't du DSA rus Patient Impact: no patient impact SE 7.2.16.00000000Tools & Equpment : ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000330000000000000000 00000000-Nov-2000000003 Fitchie Victor The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : uLLmessage "Tue anode problem. Low load flavor selected" Can't du DSA rus When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: uable to do DSA rus du to tue cooling uLL uer Impact: uLLmessage "Tue anode problem. Low load flavor selected" Can't du DSA rus Patient Impact: no patient impact SE 7.2.16.00000000Follow u Requred : Install tue chiller uon arrive 11/6 and check for other issus. Diagnostic performed by Engineer : No 000000000000000000000000001k performed as chiller not arriving util 11/6. Chiller oil on site. Parts Consued : 2.000000000000000000000000000000002100000000173,OIL CAN WITH OIL 2,5L ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000323710000000000000000-Nov-2000000003 Fitchie Victor The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace tue cooling uit and oil, check for other uLL or tue problem du to leak. Problem Description by Engineer : Problem description by engineer : uLLmessage "Tue anode problem. Low load flavor selected" Can't du DSA rus When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: uable to do DSA rus du to tue cooling uLL uer Impact: uLLmessage "Tue anode problem. Low load flavor selected" Can't du DSA rus Patient Impact: no patient impact SE 7.2.16.00000000Diagnostic performed by Engineer : uLL logs showed tue cooling uit uLL visully verified tue chiller was empty of oil du to leak at deairiator knob. Spoke with NSS for approval. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000323600000000Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : uLLmessage "Tue anode problem. Low load flavor selected" Can't du DSA rus When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: uable to do DSA rus du to tue cooling uLL uer Impact: uLLmessage "Tue anode problem. Low load flavor selected" Can't du DSA rus Patient Impact: no patient impact SE 7.2.16.00000000Diagnostic performed by Engineer : Looked at diagnostic findings and saw tue cooling uit uLL. System restricted Internal Comments : uLL# and/or description of uLL: RMT - uGEN: Tue Rotor/Anode and Cooling uit problem - Frontal_2, RMT - uGEN: Warning (00000000TJ) TTS: Cooling uit defect - Frontal Repair Action : Frontal - Cooling uit (CLu defect. Tue anode will be stopped (gracefu degradation is enabled). Service Action 1: NOTE: This uLLcan occu if the system is switched Off and back On again in a short time frame. In that case youcan ignore this uLL Else continu with Soluion 2 and futher. Service Action 2: Check if the cooling uit (CLu is powered and switched on. LEDs: H1 (=Cooler On), H2 (=Mains present) and H3 (=Cooling Requst) shoud be lit. Service Action 3: Check whether the Tue cooling uit flud level is perhaps too low. Please fill to level with Shell Diala oil. Service Action 00000000 Look as well to the pressue of the oil circut, if its qute high (higher than 6 Bar) it maybe related to a oil hose blockade. Check visully if oil hoses are perhaps bend too muh. Service Action 5: Check the oil filter of the cooling uit, if necessary clean or replace filter. Service Action 6: Check the cable from cooling uit (CLu to u-Ray generator. Pattern Name: RMT - uGEN: Tue Rotor/Anode and Cooling uit problem - Frontal_2uY Name: fae50000000000000000ca16-cc62-fc31-00000000d556b7fae3 Root Caue 1: Tue Rotor/Anode and Cooling uit problem Service Action 1: If this uLL(mainly) intermittently occus at first system start in the morning, then: Check whether the Tue cooling uit flud level is perhaps too low. Check whether the flow of the tue cooling uit was not stable enouh duing startu. Check the oil filter of the cooling uit, if necessary clean or replace filter. Check whether one of the cooling hoses is not bended too muh so the flow is restricted. If problem still occus, replace cooling uit Service Action 2: Check whether Frontal Tue cooling uit is runing. If not runing, Check the inpu 2300000000ac powersuply to Cu If it is ok, switch off the system and remove WW100000000u2 connector from the Cu Switch ON the Cuuing its service switch. If it switches ON Check the cable between EZ-u513 to to WW100000000u2 (for Velara generator based systems) Service Action 3: If the Cuis not switched on with service switch as mentioned in soluion 2, Install the 2 m long bypass hose shorting the OIL-IN and OIL-Ou terminals of the cooling uit. Bypass the level switch by placing loopback connector at WW100000000S1 connector in case of Cu100000000. Try to switch on the Cu uing the service switch in the Cu If the Cuswitches on in this condition, restore original connections one by one util the problem souce is identified. Note: If the problem is identified to be with Oil hoses, check /clean the oil filter. If the Cuis not switched on with the conditions mentioned, then Cuitself is supected. Service Action 00000000 It coud be caued du to fauty u-ray Tue stator. Check the MRC u-ray stator indutances and resistances at E-u11 : u- V: 225 mH Â±100000000 00000000ohm ±100000000 V - W: 225 mH Â±100000000 00000000ohm±100000000 u- W: 225 mH Â±100000000 00000000hm±100000000 If not Ok then replace u-ray Tue Service Action 5: It coud also be the problem is caued du to u-ray tue stator (mechanical) bearing problems, try new u-ray tue Follow u Requred : uLL# and/or description of uLL: RMT - uGEN: Tue Rotor/Anode and Cooling uit problem - Frontal_2, RMT - uGEN: Warning (00000000TJ) TTS: Cooling uit defect - Frontal Repair Action : Frontal - Cooling uit (CLu defect. Tue anode will be stopped (gracefu degradation is enabled). Service Action 1: NOTE: This uLLcan occu if the system is switched Off and back On again in a short time frame. In that case youcan ignore this uLL Else continu with Soluion 2 and futher. Service Action 2: Check if the cooling uit (CLu is powered and switched on. LEDs: H1 (=Cooler On), H2 (=Mains present) and H3 (=Cooling Requst) shoud be lit. Service Action 3: Check whether the Tue cooling uit flud level is perhaps too low. Please fill to level with Shell Diala oil. Service Action 00000000 Look as well to the pressue of the oil circut, if its qute high (higher than 6 Bar) it maybe related to a oil hose blockade. Check visully if oil hoses are perhaps bend too muh. Service Action 5: Check the oil filter of the cooling uit, if necessary clean or replace filter. Service Action 6: Check the cable from cooling uit (CLu to u-Ray generator. Pattern Name: RMT - uGEN: Tue Rotor/Anode and Cooling uit problem - Frontal_2uY Name: fae50000000000000000ca16-cc62-fc31-00000000d556b7fae3 Root Caue 1: Tue Rotor/Anode and Cooling uit problem Service Action 1: If this uLL(mainly) intermittently occus at first system start in the morning, then: Check whether the Tue cooling uit flud level is perhaps too low. Check whether the flow of the tue cooling uit was not stable enouh duing startu. Check the oil filter of the cooling uit, if necessary clean or replace filter. Check whether one of the cooling hoses is not bended too muh so the flow is restricted. If problem still occus, replace cooling uit Service Action 2: Check whether Frontal Tue cooling uit is runing. If not runing, Check the inpu 2300000000ac powersuply to Cu If it is ok, switch off the system and remove WW100000000u2 connector from the Cu Switch ON the Cuuing its service switch. If it switches ON Check the cable between EZ-u513 to to WW100000000u2 (for Velara generator based systems) Service Action 3: If the Cuis not switched on with service switch as mentioned in soluion 2, Install the 2 m long bypass hose shorting the OIL-IN and OIL-Ou terminals of the cooling uit. Bypass the level switch by placing loopback connector at WW100000000S1 connector in case of Cu100000000. Try to switch on the Cu uing the service switch in the Cu If the Cuswitches on in this condition, restore original connections one by one util the problem souce is identified. Note: If the problem is identified to be with Oil hoses, check /clean the oil filter. If the Cuis not switched on with the conditions mentioned, then Cuitself is supected. Service Action 00000000 It coud be caued du to fauty u-ray Tue stator. Check the MRC u-ray stator indutances and resistances at E-u11 : u- V: 225 mH Â±100000000 00000000ohm ±100000000 V - W: 225 mH Â±100000000 00000000ohm±100000000 u- W: 225 mH Â±100000000 00000000hm±100000000 If not Ok then replace u-ray Tue Service Action 5: It coud also be the problem is caued du to u-ray tue stator (mechanical) bearing problems, try new u-ray tue Technician Name : Eckbold Glenn ----------------------------------------------------------------------------- Internal Remarks : part id for tue chiller and oil ----------------------------------------------------------------------------- | 2023-11-02 01:24 | No
3224 | Yes | Timer stuck at zero. Fails to boot up. | Timer stuck at zero. Fails to boot up. | 000000000000000000000000001O##: 00000000000000000000000000100000000323763 00000000-Nov-2000000003 Williams, Grey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Timer stuk at zero. Fails to boot u. How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Fails to boot u uer Impact: Room down Patient Impact: None. Happened on boot u. SE 8.1.3000000005 Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Resoluion : I replaced the fleuvision PC and reloaded software. The machine can now boot u. I performed fuctional checks and gave the machine back to the cutomer. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000323683 Santos Ken The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Timer stuk at zero. Fails to boot u. How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Fails to boot u uer Impact: Room down Patient Impact: None. Happened on boot u. SE 8.1.3000000005 Diagnostic performed by Engineer : Checked service history in regards to a failed FleuVision 2 Revised. Saw the PC was rebooted bu not replaced. Cutomer tried to reboot the system and the Fleuvision PC, bu not change in system statu. Internal Comments : uLL# and/or description of uLL: Timer coutdown stuk at zero Repair Action : No PRS files available. If onsite and logs show grabber card fails, then replace the Fleuvision 2 Revised PC. Hit Alt F at the star duing boot u. 0000000080000000058600000000 In stock Resoluion : Pending FSE Follow u Requred : uLL# and/or description of uLL: Timer coutdown stuk at zero Repair Action : No PRS files available. If onsite and logs show grabber card fails, then replace the Fleuvision 2 Revised PC. Hit Alt F at the star duing boot u. 0000000080000000058600000000 In stock Technician Name : Santos Ken ----------------------------------------------------------------------------- | 2023-11-02 01:29 | No
3225 | Yes | Power cord fails EST | Power cord fails EST |  | 2023-11-02 02:10 | No
3226 | Yes | Locked C Arc | Locked C Arc | BuQ&A Feedback: No Reason: Corrective Maintenance Repair Action: FSE Adri has already applied the part released 00000000/00000000 at 00000000h Problem Reported by Cutomer: I need to open a call for the angiograph cv0000000082 C-arm stuk Problem was identified duing the procedue Name: Daniel Temoteo Sector: Clinical Engineering 00¹: +5585986000000008800000000 E-mail: daniel.temoteo00ontekliniku.com.br No one was injued uY u uuuFD200000000HOSPITAL MONTE KLINIKu SS LTDA, and is located at Ru Repulica do Libano 700000000Fortaleza-CE 600000000600000000100000000Safety Qustions: Case: 0000000022525300000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Renata Lima Last udated by: Renata Lima 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 22.11.2000000003 23:00000000:57 POLARIS (POLARIS) ================= Adrianu Petru Maria Zijp - 000000000000000057566751000000003000000000000000000000000900000000================== ==== Geometry control often left the EN-MV (Enable Move ment) signal active withou any movement butons activated. When this happens and duing machine startu, after switching on, the motorized movements do not release. And there is the warning "Motorized movements not available" on the monitor. I confirmed the problem duing the u-Ray lights repair visit. The remote service team sent a new control. And when the new geometry control was received, on 11/00000000/2000000003, engineering soon changed it. (This 000000000000000000000000001ked, the software release was correct) Since now, the arc and table movements always come into operation withou problems when the machine is tuned on. With the control assembled correctly and tested for t000000000000000000000000001weeks, I am closing the S000000000000000000000000001Internal Remarks: 22.11.2000000003 23:00000000:58 POLARIS (POLARIS) ================= Adrianu Petru Maria Zijp - 000000000000000057566751000000003000000000000000000000000900000000================== ==== Alignment: Yes Action plan: OK SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Problem description by engineer : buton pressed when tuned on Trouleshooting Action: What steps are taken duing the system diagnosis : checked in the logs and TSO test showed that the joystick does not retun to 00000000after being activated, intermittent. A TSO will be requsted for testing. There is no shock sensor installed uLL# and/or description of uLL: TSO test failed OneEMS internal remarks: Problem description by engineer: buton pressed when tuned on What steps are taken duing the system diagnosis: checked in the logs and TSO test showed that the joystick does not retun to 00000000after activation, intermittent. A TSO will be requsted for testing. There is no shock sensor installed uLL# and/or description of uLL: TSO test failed Repair Action "Internal" : Send PART+FSE+2 hous of machine downtime recommendations to FSE ► Replace TSO; ► test the new TSO, ► Tun it off, tun it on and check if it goes u withou uLLParts requred : 1 Part : 1, uu: 00000000800000000151715, Description : TSO WP Kit, Quntity : 1, Priority : 1 Eutra Internal comment field : RSE Marcel Ogassawara System u 10000000031/2000000003 - 17:00000000rs EDT ----------------------------------------------------------------------------- | 2023-11-02 02:14 | No
3227 | Yes | AED emits triple chirp, "Remove and reinsert battery to test" uLLmessage | AED emits triple chirp, "Remove and reinsert battery to test" uLLmessage | 000000000000000000000000001O##: 0000000000000000000000000010000000032000000002 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below William Baish 300000000uiversity Dr Corpu Christi, Tu 7800000000 Internal Comments : Tracking#785800000000659300000000 DELIVERED Thusday 11/2/23 at 10000000036 AM ##cutomer to confirmed that AED is 000000000000000000000000001king as it shoud Resoluion : Cutomer has confirmed that the AED is 000000000000000000000000001king as it shoud. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping and will not pass the self-test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Cutomer reported the AED kept chirping and will not pass the self-test. Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and will not pass the self-test. Cutomer removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Remove and reinsert battery to test. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Advise cutomer to remove device ou of service. util he received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# B17H-00000000000000009 Battery Install Before# 2000000008/12/31 Lot#00000000886P Pads Eup Date: 00000000/2000000005 Pads Lot # 221000000003-0000000056 Warranty 00000000/9/2000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032000000003 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-11-02 02:14 | No
3228 | Yes | Won't charge battery although it will run on battery power | Won't charge battery although it will run on battery power | 000000000000000000000000001O##: 00000000000000000000000000100000000320000000000000000Harris, Alvin The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Please send a loaner Tempu Pro monitor with SW rev v7.300000000nd featues u NBP, SPO2, temperatue, capnometer, printer to cu Verified cus ALT ship to address Metro West Fire Protection District 160000000000000000Clayton Rd Ellisville MO 63000000001 ATT: Brad Shelton/Case 000000002253200000000 ##d Irvine CA Bench form to cuDiagnostic performed by Engineer : Problem Description: 000000000000000000000000001't charge battery althouh it will ru on battery power 000000000000000000000000001't charge battery althouh it will ru on battery power Form ued: Bench Trouleshooting Steps Taken: Received Teams message from NSS - Travis Haines asking me if I suport the Tempu Pro monitor, informing me that 3rd party FSE - Kyle Blackbun - (00000000) 31000000009800000000 - kyle.blackbun00hsteam.com had called in to their NSS suport line Responded via Teams stating I did suport the produt Travis informed me that 3rd party FSE - Kyle Blackbun - (00000000) 31000000009800000000 - kyle.blackbun00hsteam.com had called in to their NSS suport line Travis euplained that he was having troule rouing Kyle to an RSE, I was the only 1 that had responded, asked if he coud give Kyle my phone nuber or ##address, Kyle needed assistance requsting a replacement device uder warranty Responded asking if the cuwas a u Rep and whether or not the cuhad a case nuber, stating I coud jut call him Travis stated the cuwas a u Rep and had a preventative maintenance case, provided cus first name and phone nuber --- Called cuback at (00000000) 31000000009800000000, reached Kyle Cueuplained that he is onsite performing preventative maintenance for Tempu devices, he performed the service for this monitor yesterday, 10000000031/23, it passed with no issus, bu today, 11/1/23, it was foud to no longer charge its battery Cueuplained that the device will ru on battery power, bu they have to replace the battery with another charged battery becaue the device will not charge the installed battery Euplained to cuthat a loaner coud be requsted and the device 000000000000000000000000001ld need to be sent to ou Bench Travis agreed with this plan, asked that end uer - Brad Shelton - 636 262 3300000000 - bradsh00etrowest-fire.org be listed as the contact Problem Description by Engineer : Cutomer Fuction/Role:Other: 3rd party Field Service Engineer Device ue:Not in Clinical ue Eupected and ueupected Resuts:Will charge battery althouh it will ru on battery power uer Impact:Having to reguarly swap ou depleted batter with a charged battery when uing device on battery power Patient Impact:No impact SEv7.300000000 Problem Description/ueupected Resuts: 000000000000000000000000001't charge battery althouh it will ru on battery power Resoluion : Sent to Bench ----------------------------------------------------------------------------- | 2023-11-02 03:09 | No
3229 | Yes | DPM hardware failure. | DPM hardware failure. | 000000000000000000000000001O##: 0000000000000000000000000010000000032000000005 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Is Caue Known: No Caue: Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.Ordered replacemen Tempu LS 000000005600000000181. Problem Description by Engineer : Cutomer Fuction/Role: Other: Sales rep Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: SE NA Problem Description/ueupected Resuts: DPM HW uLL Happened duing training. NO PATIENT INVOLVED. Diagnostic performed by Engineer : Problem Description: DPM HW uLL Happened duing training. NO PATIENT INVOLVED. DPM HW failue. Happened duing training - NO PATIENT INVOLVED. Form ued: Trouleshooting/Investigation Object: Tempu LS Not Object: Software Version: NA Date and Time of Issu: 11/1/2000000003 uer Actions: Where is the Deviation: Tempu LS Trouleshooting Steps Taken: NA Was PRS ued: No PRS Permission: Follow u Requred : Tu Rio Verde Fire District Attn: Joshu Johnson 25600000000 N. Forest Rd. Rio Verde, AZ 85263 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032000000003 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Gary Vasil IP & Issu: " 91383278 DPM hardware failue; ccb for statu udate for part order" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: transferred call to parts department; Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : This is suposed to be fiued uder FSN. FSCA Ref: SAGQI-37 See attachments to this case. Internal Remarks : ##d cutomer asking for log files. Internal Remarks : Interaction Type: New Case Caller: Gary Vasil IP & Issu: "91383278 DPM hardware failue." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-02 04:19 | No
3230 | Yes | uuuis emitting the triple chirp | uuuis emitting the triple chirp | 000000000000000000000000001O##: 0000000000000000000000000010000000032000000002 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 15 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# B00000000G-000000000000000000000000 Battery Install Before# 6/3000000002000000009 Lot#00000000200000000P Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test.uer Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. ----------------------------------------------------------------------------- | 2023-11-02 04:34 | No
3231 | Yes | At start-up, the blue screen appears and it won't start. | At start-up, the blue screen appears and it won't start. | 000000000000000000000000001O##: 00000000000000000000000000100000000356600000000200000000Nov-2000000003 Hayashi, Hiroto The system meets the specification for the performed service and is retuned to ue Internal Comments: November 200000000h: After checking the sitution, we reported to u Tsusui over the phone that there were no problems and his approval was received. Resoluion: Parts ued Part name: Hard disk spare part Part No: 0000000080000000050000000000000000Serial No: 722219329931 Reguatory Qustions: N/A Problem Description by Engineer: A blu screen appears when starting u and does not start u. Cutomer Fuction/Role: Radiology Technician Device ue: Not ued Eupected and ueupected Resuts: A blu screen appears when starting u and does not start u. uer Impact: Test canceled Patient Impact: Test canceled SE 000Diagnostic performed by Engineer: Data monitor blu screen uLLinvestigation Host PC HDD replacement 000000000000000000000000001k Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003200000000000000002-Nov-2000000003 Hayashi, Hiroto The system meets the specification for the performed service and is retuned to ue Resoluion: After arriving on site, we confirmed that the data monitor was showing a blu screen. Confirmed occurence of uLLcode: 000000000000000000000000000000007B. There is a supicion of HDD failue du to an uLLin the main control PC (HOST PC) of the device. When I tried restarting the device, I noticed a pop-u message indicating that the data file on the HOST PC was corruted. We supected that the HDD was the problem, so we replaced the HDD of the HOST PC and then performed data recovery (Ghost Backu/EPu). Confirm that the data has been retuned and perform an operation check. Patient data has been deleted duing this process. Note that. Resoluion: Parts ued Part name: Hard disk spare part Part No: 0000000080000000050000000000000000Serial No: 722219329931 Reguatory Qustions: N/A Problem Description by Engineer: A blu screen appears when starting u and does not start u. Cutomer Fuction/Role: Radiology Technician Device ue: Not ued Eupected and ueupected Resuts: A blu screen appears when starting u and does not start u. uer Impact: Test canceled Patient Impact: Test canceled SE 000Diagnostic performed by Engineer: Data monitor blu screen uLLinvestigation Host PC HDD replacement 000000000000000000000000001k ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032000000006 Saka, Shuji The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: A blu screen appears when starting u and does not start u. Cutomer Fuction/Role: Radiology Technician Device ue: Not ued Eupected and ueupected Resuts: A blu screen appears when starting u and does not start u. uer Impact: Test canceled Patient Impact: Test canceled SE 000Diagnostic performed by Engineer: RC Contents: BIOS screen display, blu screen duing Windows startu. Even after powering u again, the sitution remains the same. Visit suvey. Internal Comments: [Suject] 00000000DS:SS: Dear Ariga: When I start u, a blu screen appears and it does not start u. [Internal Comments] HostPC Ou of stock in Japan. HostPC orders mut be confirmed on-site before ordering R8.1.17 equpment 00000000800000000200000000611 Host PC Monoplane Revised no HDD R8.1.25 or higher 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD 00000000800000000200000000612 Nehalem Host PC Monoplane Revised no HDD Parts] 0000000080000000050000000000000000 Hard disk spare part QTY:1 ZSID Sendai (u13) [Dispatch] 1 RC -----------------FSE registration information------------- ---- -Visit FSE1 Name [Hiroto Hayashi] Hayashi, Hiroto -FSE1 Visit date and time 2000000003/11/0000000012:300000000-FSE1 000000000000000000000000001king hous 2 [Shipping instrutions] -------------- ---Shipping instrutions #1----------------- -Delivery conditions ASAP (transportation arrangement) -Delivery date ASAP -Delivery destination name Local direct delivery -Recipient name Hiroto Hayashi -0000000000000000006722-3179 -Other Mobile phone call uon arrival -Requster name Toshihito Han [Test category] Other Follow u Requred: * ----------------------------------------------------------------------------- | 2023-11-02 04:54 | No
3232 | Yes | No C arm movements | No C arm movements | 000000000000000000000000001O##: 0000000000000000000000000010000000032000000007 00000000-Nov-2000000003 Caldwell, Pau The system meets the specification for the performed service and is retuned to ue Resoluion : stand long collision Problem Description by Engineer : Problem description by engineer : No C arm movements When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No C arm movements uer Impact: System is down Patient Impact: None/ in progress SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032000000009 Whitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : No C arm movements When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No C arm movements uer Impact: System is down Patient Impact: None/ in progress SE 1.2.5.00000000Diagnostic performed by Engineer : Dan reports that the C arm and table are not moving, no uLL are reported. They are finishing u this case bu 000000000000000000000000001ld like onsite service tonight if possible. Rebooting did not resolve. Saved a LOD and foud the uLL below Internal Comments : uLL# and/or description of uLL: 200000000SPOS6300000000127 Collision detected [auis=FrtLong] 200000000SPOS6300000000127 Collision detected [auis=FrtLat] 200000000SPOS0000000000000000932 Magnu internal post Done/Failed Repair Action : 11_Read from the logline the 11_Read from the logline the SYMPTOM 1: Stand longituinal movement is restricted or not available at all. CAuE 1: Application uLL Collision detected [auis=FrtLong] SOLuION 1: Read from the logline the auis statu (operation or stand-by) and the position where the collision occured. If the movement hits a mechanical obstrution, this is normal behavior. Otherwise, ru a fuctional test on this specific movement. Possible corrective actions: • Perform curent calibration of movement • Perform force sensor calibration • Clean mechanics or cabling. 16_Please verify with the cutomer 16_Please verify with the cutomer SYMPTOM 1: Patient suport movement is restricted or not available at all. CAuE 1: Duing boot u of the system the Power On Selftest failed for the Maqut table. u & Magnu detected Magnu POST uLL(uLL00000000). Most likely table movements are restricted or not possible. SOLuION 1: Please verify with the cutomer if they euuuence any inconvenience, maybe additional info from Maqut is necessary. uLL00000000 = MAGNu_DRIVER_uLLMAGNu_POST_uLL Follow u Requred : uLL# and/or description of uLL: 200000000SPOS6300000000127 Collision detected [auis=FrtLong] 200000000SPOS6300000000127 Collision detected [auis=FrtLat] 200000000SPOS0000000000000000932 Magnu internal post Done/Failed Repair Action : 11_Read from the logline the 11_Read from the logline the SYMPTOM 1: Stand longituinal movement is restricted or not available at all. CAuE 1: Application uLL Collision detected [auis=FrtLong] SOLuION 1: Read from the logline the auis statu (operation or stand-by) and the position where the collision occured. If the movement hits a mechanical obstrution, this is normal behavior. Otherwise, ru a fuctional test on this specific movement. Possible corrective actions: • Perform curent calibration of movement • Perform force sensor calibration • Clean mechanics or cabling. 16_Please verify with the cutomer 16_Please verify with the cutomer SYMPTOM 1: Patient suport movement is restricted or not available at all. CAuE 1: Duing boot u of the system the Power On Selftest failed for the Maqut table. u & Magnu detected Magnu POST uLL(uLL00000000). Most likely table movements are restricted or not possible. SOLuION 1: Please verify with the cutomer if they euuuence any inconvenience, maybe additional info from Maqut is necessary. uLL00000000 = MAGNu_DRIVER_uLLMAGNu_POST_uLL Technician Name : Whitaker, Steven ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Requstor :Pau Caldwell Requst: reschedue Oucome:: reschedued 0000000000000000000000000010000000032000000007 to 11/700PM Internal Remarks : FSE Pau Caldwell. FSE 000000000000000000000000001king on uFleumove. When going longituinal the C will only go part way down length of table, via controls or manully moving. FSE looking for directions to remove cover to get into energy cahin to see if there is binding. ----------------------------------------------------------------------------- | 2023-11-02 05:19 | No
3233 | Yes | Device alarmed due to no Ouygen supply | Device alarmed due to no Ouygen supply | 000000000000000000000000001O##: 00000000000000000000000000100000000317593 Sudhararaj,Vidhya The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Given the history of this device with recuring O2 uavailable issu, cutomer 000000000000000000000000001ld like to send to bench for assessment Provided cutomer with address for bench repair to send it in for assessment and mentioned that the costs for parts may not be covered Problem Description by Engineer : Cutomer Fuction/Role: biomedical engineer Device ue: ou of ue Eupected and ueupected Resuts: V600000000Displaying 'No Ouygen Suply' uer Impact: nil Patient Impact: nil SE na Diagnostic performed by Engineer : reviewed the log and there were coule of O2 uavailable uLLand a O2 device failed uLL Advised cutomer that the issu coud be with flow sensor or pressue sensor and advised to do air and o2 flow accuacy and miu accuacy test on this device and verify if it passes these test. Cutomer reported the device passed all the test ----------------------------------------------------------------------------- Internal Remarks : left message for cutomer ----------------------------------------------------------------------------- | 2023-11-02 06:39 | No
3234 | Yes | QAA0/u FD0can't start system | QAA0R/u FD0can't start system | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: QAA00000000R/u uuuFD100000000 T2 Activities:Problem Description: Overview of cutomer problems: can't start system Trouleshooting Action: Key faut codes or phenomena: 3D cannot be tuned on, lights 3 and 00000000n the host panel are always on. Remote diagnostics and test resuts: "Memory modue detected, bu a memory power failue occured. If t000000000000000000000000001or more memory modues are installed in the system, remove the memory modues, then reinstall one modue and restart the compuer.If the compuer starts normally, continu to install additional memory modues (one at a time) util youlocate the failed modue or reinstall all modues correctly.If only one memory modue is installed, try moving it to a different DIMM connector and then restart the compuer.•Install the same type of memory in you compuer that has been proven to 000000000000000000000000001k properly, if available. 51" PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check the memory, clean the host, and start the test SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 1000000001:16 uC (86000000008398) Service content: The host cannot be started, check the uLLand point to the memory, plu and uplu to clear the memory and then restore Service resut: The device retuns to normal SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 100000000000000009 uC (86000000008398) RDI Accuacy: Yes $Q7: (Confirmed reported problem& bootu uLL) $Q8: (checked indicate: the LED 3and 00000000f diagnostic panel lit) $Q9: (refasten the memory card ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated tools name(NA) S/N (NA) Neut calibration date (NA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: second: N/A SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022526222 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-02 06:44 | No
3235 | Yes | alarm is not going off | alarm is not going off | Internal Remarks : Caller: ruen 9000000007828539 IP & Issu: alarm is not going off Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: created r000000000000000000000000001and roued to rse ----------------------------------------------------------------------------- | 2023-11-02 06:54 | No
3236 | Yes | The equipment remains at 0:0 and does not start | The equipment remains at 0:0 and does not start | 000000000000000000000000001O##: 00000000000000000000000000100000000325000000008 00000000-Nov-2000000003 Kawamua, HidetsuuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The equpment remains at 00000000:00000000 and does not start.There is no grid grid line display on the FleuVision monitor.Tuning the power back on does not improve the sitution. The power was not on for the compuer that controls the FleuVision monitor in cabinet B. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The equpment remains at 00000000:00000000 and does not start.There is no grid grid line display on the FleuVision monitor.Tuning the power back on does not improve the sitution. The power was not on for the compuer that controls the FleuVision monitor in cabinet B. uer Impact: No effect Patient Impact: No effect SE R8.2.3000000005 Resoluion: I confirmed an uLLon the PC that controls the large monitor. I replaced the PC and confirmed normal operation. Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000325000000007 Suui, NobuasuThe system meets the specification for the performed service and is retuned to ue Internal Comments: Simp [Suject] 00000000DS:NS:u Kuo: The equpment remains at 00000000:00000000 and does not start. [Internal Comments] 985 FleuVision Warning 2000000003-11-00000000 00000000:58:30000000000000000 FleuVisionControlService:Observer000000000000000000000000001ker PING(FLEuVISION_PC_CONTROL) is not suceeded after 100000000 seconds. Please check if FLEuVISION_PC_CONTROL is powered on and all (net000000000000000000000000001k) cables are correctly connected. [SEVERITY] Warning [EuCEPTION_DESCRIPTION] FleuVisionControlService:Observer000000000000000000000000001ker PING(FLEuVISION_PC_CONTROL) is not suceeded after 100000000 seconds. Please check if FLEuVISION_PC_CONTROL is powered on and all (net000000000000000000000000001k) cables are correctly connected. 8500000000000000000000000000000000 Startu Service NoDetailInfo [Parts from the equpment] 0000000080000000012200000000:u Haswell Fleuvision 2 PC no HDD QTY:1 ZREP/Traceable Osaka (u15) 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Osaka (u15) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Kawamua, Hidetsuu[Eiot Kawamua] - FSE1 Visit date and time 2000000003/11/0000000011:300000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Eiot Kawamua - 000000000000000000000000006-9200000000 - Others Call mobile uon arrival Requster name NobuasuSuui [Euamination category] Other: Follow u Requred: * Problem Description by Engineer: The equpment remains at 00000000:00000000 and does not start.There is no grid grid line display on the FleuVision monitor.Tuning the power back on does not improve the sitution. The power was not on for the compuer that controls the FleuVision monitor in cabinet B. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The equpment remains at 00000000:00000000 and does not start.There is no grid grid line display on the FleuVision monitor.Tuning the power back on does not improve the sitution. The power was not on for the compuer that controls the FleuVision monitor in cabinet B. uer Impact: No effect Patient Impact: No effect SE R8.2.3000000005 Diagnostic performed by Engineer: RC Response: FSE assignment ----------------------------------------------------------------------------- | 2023-11-02 06:54 | No
3237 | Yes | 0V0uLLAlarm:High blower temperature - GFE 0/7 | 0/V0uLLAlarm:High blower temperature - GFE 00/7 | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Linping District First People's Hospital Blower alarm Internal Remarks: pending timeou: processing 0000000022526323, 000000002252600000000, 000000002252600000000 Safety Qustions: Case: 000000002252600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-11-02 07:49 | No
3238 | Yes | Proactive LS | Proau LS | 000000000000000000000000001O##: 00000000000000000000000000100000000338663 00000000-Nov-2000000003 Stout Chad The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: Diagnostics Eupected and ueupected Resuts: Data monitor blank after boot u and intermittent table movement uLLduing ru off stuy uer Impact: Delayed procedue Patient Impact: Delayed procedue SE 8.2.17 Resoluion : Sucessfuly replaced Fleuvision PC. Downloaded necessary software and performed system fuctionality checks withou issus. Retuned system to cutomer for clinical ue. Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032527000000003-Nov-2000000003 Stout Chad The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: Diagnostics Eupected and ueupected Resuts: Data monitor blank after boot u and intermittent table movement uLLduing ru off stuy uer Impact: Delayed procedue Patient Impact: Delayed procedue SE 8.2.17 Diagnostic performed by Engineer : Replaced table longituinal POT and performed all necessary calibrations. Fleuvision PC didn't test good. Will retun to place Fleuvision PC. System retuned to cutomer will restricted uuuilities. Follow u Requred : Replace Fleuvision PC Parts Consued : 1.00000000000000000000000000000000230000000025982,AD7(NT) LONGITuINAL POTMETER ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003200000000 00000000-Nov-2000000003 Stout Chad The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: Diagnostics Eupected and ueupected Resuts: Data monitor blank after boot u and intermittent table movement uLLduing ru off stuy uer Impact: Delayed procedue Patient Impact: Delayed procedue SE 8.2.17 Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 Follow u Requred : Replace Fleuvision PC and longituinal POT Diagnostic performed by Engineer : Foud loose connector on host PC cauing blank screen on data monitor. Checked uLLlogs. Foud uLL related to table movement mismatch duing movement. Recommended replacing Fleuvision PC and table longituinal POT. ----------------------------------------------------------------------------- Internal Remarks : Table locked u doing a bolu chase ru. Cutomer rebooted a coule times and then the FleuVision PC went ou and system 000000000000000000000000001ld not boot u. The CAT tool keeps crashing, bu I was able to verify with Remote Service 000000000000000000000000001kspot that grabber card uLL are present and the longituinal uLLis a position slip. Sugested for FSE to try recalibrating the AD7 longituinal movement before replacing the POT. -Approved: u Haswell Fleuvision 2PC no HDD: 0000000080000000012200000000 Disk for revised II: 0000000080000000050000000000000000 AD7 Longitudnal Potmeter assy: 00000000230000000025982 ----------------------------------------------------------------------------- | 2023-11-02 08:29 | No
3239 | Yes | 0 - u00- Battery life reduced | 00- u0000- Battery life reduced | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 00000000225266600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: mengdi uu Last udated by: mengdi uu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: FSE uuZiqing MA&TC(185100000000687900000000Call u for a qute to puchase batteries ----------------------------------------------------------------------------- | 2023-11-02 08:59 | No
3240 | Yes | Operation Check failed (Issue resolved after uleadset cleaned | Operation Check failed (Issue resolved after uleadset cleaned | 000000000000000000000000001O##: 00000000000000000000000000100000000318000000006 Nandha, Rahu The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : uleadset cleaned confirmed 000000000000000000000000001king ok. Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: yes Eupected and ueupected Resuts: na uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- | 2023-11-02 09:59 | No
3241 | Yes | Machine not booting | Machine not booting | 000000000000000000000000001O##: 00000000000000000000000000100000000330000000069 00000000-Nov-2000000003 Pandey, Rajesh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Blu dup is coming Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion : Replaced the host pc hdd and restored ghost backu. Foud ok. Handed overt the system in good 000000000000000000000000001king condition Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Tools & Equpment : MuTIMETER,200000000000000002-17 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000325758 00000000-Nov-2000000003 Pandey, Rajesh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Blu dup is coming Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Reseated all reported the problem, reseated all hardware and connections of host pc, foud ok. Observed system in few procedues. Resoluion : Need to replace hdd disk Follow u Requred : Need to replace part Tools & Equpment : MuTIMETER,200000000000000002-17 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000325271 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Blu dup is coming Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Blu dup is coming Internal Comments : uLL# and/or description of uLL: Blu dup is coming Repair Action : Need to replace host PC HDD and reload the ghost Follow u Requred : uLL# and/or description of uLL: Blu dup is coming Repair Action : Need to replace host PC HDD and reload the ghost What are the neut Trouleshooting steps : Need to replace host PC HDD and reload the ghost Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-11-02 11:24 | No
3242 | Yes | Direct to sale: Charge button is not working | Direuto sale: Charge button is not working |  | 2023-11-02 13:04 | Yes
3243 | Yes | AED is emitting triple beeps - log analysis requested | AED is emitting triple beeps - log analysis requested | 000000000000000000000000001O##: 00000000000000000000000000100000000319000000003 Suui, Kosue The system does not meet the specification for the performed service and is not retuned to ue Resoluion: After analyzing the main uit log at the manuactuing factory in the uited States, we foud a record of failue of the high voltage (HV) calibration test. The main uit will need to be replaced du to a diagnostic failue du to an electrical failue in a high voltage element within the main uit, bu after sumitting the estimate, the repair was canceled and the complete set will be retuned to you On November 28, 2000000003, we received you requst to cancel the repair via ## [Simutaneou FCO 000000000000000000000000001k] There are no FCOs that requre 000000000000000000000000001k on this equpment. November 28, 2000000003 Fuua Denshi KitatohokuHanbai We will send the 000000000000000000000000001k report along with the equpment to the person in charge. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit the 000000000000000000000000001k report on paper? <YES>[Final confirmation] Perform operation check after 000000000000000000000000001k based on prescribed instrutions NO Tools & Equpment : MuTIMETER, 200000000000000003-15 00000000:00000000:00000000 Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Fuua Electronic Sales Representative Device ue: Not ued Eupected and ueupected Resuts: 3 beeps uer Impact: None Patient Impact: None SE None Diagnostic performed by Engineer: After receiving the produt, a battery installation test shows that it cannot be ued. To test, please remove the battery and reinsert it.'' Annoucement and beep 3 times The phenomenon was confirmed. ----------------------------------------------------------------------------- Internal Remarks: Faut details Alarm occured 2 to 3 times.Please check. LOT Adut Pad C120000000022-00000000 Battery 98122P ----------------------------------------------------------------------------- | 2023-11-02 14:14 | No
3244 | Yes | 0V0uLLTouch not working | 0/V0uLLTouch not working | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only Safety Qustions: Case: 0000000022527836 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No Problem Reported by Cutomer: Touhing the screen does not 000000000000000000000000001k SAP Transfer: Case Transferred to SAP Internal Remarks: 29.11.2000000003 00000000:0000000056 uC (86000000000000000000000000) $ASP engineer:Yingwei[suuhu$Q1:no $Q2：no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.The machine uer Interface is fauty.$Q8: Download "DRPT" report, uLLCode have not, Shoud be a problem of uer Interface Q9: Replace the uer Interface.Q100000000 Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality.$Q11:Calibrated Tools name (FLuET6500000000 S/N (000000006000000002) Neut Calibration date (200000000000000001/11) Euternal Remarks: 29.11.2000000003 00000000:0000000053 uC (86000000000000000000000000)*Department information: [Department: Medical Materials Suport Division Contact person: uuZhiyang Contact nuber: 1800000000396600000000]* $ASP engineer:Yingwei[suuhu$Q1:no $Q2: no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.The machine uer Interface is fauty.$Q8: Download "DRPT" report， uLLCode have not, Shoud be a problem of uer Interface Q9: Replace the uer Interface.Q100000000Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality.$Q11:Calibrated Tools name (FLuET6500000000 S/N (000000006000000002) Neut Calibration date (200000000000000001/11) ----------------------------------------------------------------------------- | 2023-11-02 14:29 | No
3245 | Yes | 0/u FD0Geometry cannot move | 0L/u FD0Geometry cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000L/u uuuFD200000000 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of the cutomer's problem or faut phenomenon: Cutomer feedback that the mechanical part cannot be operated and requres engineers to come on site for repair. Is it a faut of the equpment itself: Yes Trouleshooting Action: Key faut code or phenomenon: Warning: Geometry Movement Partly Available Remote Diagnosis and test resuts: NA PRS TR Statu: ureachable Repair Action: Sugested on-site maintenance steps: 1. Check if there is an IPC commuication problem 2. Check if there is a problem with PCu3. Check if there is an ENMV high problem 00000000Check the machine uLLand perform faut repairs. If it cannot be repaired, replace the corresponding spare parts. It is recommended that the spare part is uuand its name: 00000000800000000931372 GEO IPC Ivybridge no ZMP with CAN and IO &00000000800000000151715 GEO MODuE WP Kit Cutomer demand notes: Contact the cutomer in time to confirm the Onsite time SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002252800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Case Transferred to SAP Internal Remarks: 25.11.2000000003 15:00000000:00000000uC (860000000065600000000RDI Accuacy?YES " $Q7: (Confirmed the problem, Geometry can't move. ) $Q8: (Reviewed the uLLlog, geometry have no power, DCPS have no oupu, DCPS is defective.) $Q9: (connact buiness order spare parts and inform cutomer need wait for parts and replace parts later. ) $Q100000000 (Test and Inspection data refer to PA & Device retuned to ful fuctionality.) $Q11: Calibrated Tools:(NA). Internal Remarks: 25.11.2000000003 100000000:58 uC (86000000006551) "RDI Accuacy?YES " $Q7: (Confirmed the problem, Geometry can't move. ) $Q8: (Reviewed the uLLlog, geometry have no power, DCPS have no oupu, DCPS is defective.) $Q9: (connact buiness order spare parts and inform cutomer need wait for parts and replace parts later. ) $Q100000000 (Test and Inspection data refer to PA & Device retuned to ful fuctionality.) $Q11: Calibrated Tools:(NA). SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 25.11.2000000003 100000000:56 uC (86000000006551) Service content: The mechanical movement cannot move. Check the log and find that geometry is not started. The DCPS is foud to be broken on site. Wait for spare parts qutation and replace it. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: Euternal Remarks: 25.11.2000000003 15:00000000:00000000uC (860000000065600000000Service content: The mechanical movement cannot move. Checking the log foud that geometry was not started. The DCPS was foud to be broken on site. Waiting for spare parts qutation and replacement. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-02 14:59 | No
3246 | Yes | udoes not turn on | udoes not turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000319631 Buak Yediyildiz The system meets the specification for the performed service and is retuned to ue Resoluion : replaced PDuFanTray and DCPS Diagnostic performed by Engineer : replace PDuFanTray and DCPS Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Power problem sn 288 uSSW 1 2 5 ----------------------------------------------------------------------------- | 2023-11-02 15:44 | No
3247 | Yes | 0/u FD0Poweron failed | 0R/u FD0Poweron failed | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 00000000.11.2000000003 13:26:28 uC (8600000000000000001) $Q7: (Confirmation & Description of problem) Confirmed reported problem, The machine sudenly loses power and cannot start. $Q8: (Trouleshooting & Resuts) On site inspection showed that the MPDucontrol uit was fauty. The MPDucontrol uit was replaced and the faut was eliminated. $Q9: (Repair steps) On site inspection showed that the MPDucontrol uit was fauty. The MPDucontrol uit was replaced and the faut was eliminated. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:N/A; SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 13:26:22 uC (8600000000000000001) "SuT. Ready for Clinical ue local time" : " 2000000003-11-00000000 11:00000000" Service content: The machine sudenly lost power and coud not be started.On-site inspection is MPDu The control uit is fauty. Replace the MPDucontrol uit and the faut is eliminated. Service Resuts: Equpment leaves site fuly fuctional. Other important matters: None yet. T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine did not respond when pressing the power buton and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine does not respond when pressing the power buton Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the MPDupower suply is normal. 2. Check whether the MPD Control board is normal. Check whether the fue on it is damaged. If damaged, replace it. 3. Carry ou faut repair according to FIP. Recommended spare part uuwith name: " 00000000230000000016000000007 MPD CONTROL uIT" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Dispatched for Onsite Suport Neut Step Action: Cheng Jianping-86-18820000000030000000000000000Problem Reported by Cutomer: Circut board failue SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022528656 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: 00000000500000000500000000600000000Teacher uus landline called Neut Step Action: PPD00000000230000000016000000007*1 Arranged 00000000delivery to the venu FSE Cheng Jianping-86-18820000000030000000000000000SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-02 16:34 | No
3248 | Yes | SPLIT - Faulty screen - out of service | SPLIT - Faulty screen - out of service | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000238 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: The screen calibration is completely off and it is also juping between screen displays sporadically. It is also missing a sim card slot cover. - foud duing rouine inspection uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Cover replaced on SD card slots, uit tested OK. Device retuned as the cutomer ugently requres it back, cracked SPO2 connector to be replaced at a later date. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 Follow u Requred : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000335000000001 Robert Jon de Ste Croiu The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: The screen calibration is completely off and it is also juping between screen displays sporadically. It is also missing a sim card slot cover. - foud duing rouine inspection uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : Screen re-calibrated OK. SPO2 connector damage, contacted Brad 00DT to see if a new connector/front enclosue is requred. Follow u Requred : Screen re-calibrated OK. SPO2 connector damage, contacted Brad 00DT to see if a new connector/front enclosue is requred. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000320000000089 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: The screen calibration is completely off and it is also juping between screen displays sporadically. It is also missing a sim card slot cover. - foud duing rouine inspection uer Impact: n/a Patient Impact: n/a SE n/a Reguatory Qustions : N/A Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u Front Enclosue 0000000056000000005000000001 Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u Front Enclosue 0000000056000000005000000001 ----------------------------------------------------------------------------- Cutomer comment : 0000000030000000000000000000000009OIudQAG Internal Remarks : Hi Emily , Please let u know you collection address and a sutable date for you Tempu Pro, Printer, IBP SN 60000000099 Faut : Screen Internal Remarks : Tracking Nuber:1Z3R6000000006692980000000012 COLLECT MONDAY Collection & Retuns address: Devon Air Ambuance Trading Company Limited 5 Sandpiper Cout, Harrington Lane Eueter, Devon, Eu00000000NS Internal Remarks : Good morning Tracy, Many thanks for arranging this collection. I can confirm that the uit was collected from u yesterday. Kind regards, Emily Emily Davies (she/her) Patient Services Administrator Devon Air Ambuance T: 00000000392 00000000666 ----------------------------------------------------------------------------- | 2023-11-02 16:34 | No
3249 | Yes | Speaker defective. No Trouble Found-Evaluation Confirmed Device Working. No fault found. | Speaker defeu. No Trouble Found-Evaluation Confirmed Device Working. No fault found. | 000000000000000000000000001O##: 0000000000000000000000000010000000032700000000 00000000-Nov-2000000003 JEAN-FRANCOIS GEORGES The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: None Patient Impact: None SE Resoluion: No more problem present. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000327222 u Jarrick The system does not meet the specifications for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: None Patient Impact: None SE Diagnostic performed by Engineer: Speaker defective. Follow u Requred: 2 hous onsite on warranty 1* 0000000056000000006971 Louspeaker Assy ----------------------------------------------------------------------------- | 2023-11-02 16:34 | No
3250 | Yes | M0A / u00 Device uLL - | M00A / u000/ Device uLL - | 000000000000000000000000001O##: 00000000000000000000000000100000000326672 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Contacted technician Alba, performed fuctional test, defibrillator/electrode test fails, device is not suported, proceeded with attempt to order condenser. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: defibrillator test failed - electrodes failed uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Foud problem with defibrillator test. The device is not suported. We proceed with the order of the spare part. ----------------------------------------------------------------------------- | 2023-11-02 16:54 | No
3251 | Yes | No uRAY as system was down | No uRAY as system was down | Internal Remarks : Caller: Gerardo Adofina - Web portal case IP & Issu: System is hard down. Please give me a call as soon as possible. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: Requsted Qute Action Taken: Awaiting Qute - Close Pending Entitled Remote 000000000000000000000000001 Neut Step Action: ----------------------------------------------------------------------------- | 2023-11-02 17:24 | No
3252 | Yes | Direct to sale: Therapy delivered failed | Direuto sale: Therapy delivered failed |  | 2023-11-02 18:17 | Yes
3253 | Yes | Ground resistance test failure during PM | Ground resistance test failure during PM | 000000000000000000000000001O##: 000000000000000000000000001000000003280000000000000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer replaced the power cord to solve the issu. All groud test passed sucessfuly. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: duing service Eupected and ueupected Resuts: failing the groud test uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the groud test is failing. At the groud going to the flow sensor assy, is very high. Recommended to try a different power cord and call back with resuts. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kim Bell IP & Issu: 60000000065800000000- Duing the PM, the uLLis not passing the groud resistance test. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-02 18:19 | No
3254 | Yes | AED is emitting triple chirp | AED is emitting triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000328157 Kat Gilliam The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: end uer Device ue: ouside of clinical uage Eupected and ueupected Resuts: Cutomer called in and the AED is triple chirping. When the blu “I” buton is pressed it jut says uLL Tried to do BIT test with a new and old battery and it still triple chirps and says uLL AED is uder IB warranty and to be replaced at zero cost to the cutomer. Device serial# A19D-000000006700000000 uer Impact: none Patient Impact: none SE Internal Comments : Cutomer confirmed via case ##AED is in service and 000000000000000000000000001king proeperly. Diagnostic performed by Engineer : Cutomer called in and the AED is triple chirping. When the blu “I” buton is pressed it jut says uLL Tried to do BIT test with a new and old battery and it still triple chirps and says uLL AED is uder IB warranty and to be replaced at zero cost to the cutomer. Device serial# A19D-000000006700000000nternal Comments : Waiting on order to ship, order nuber 000000000000000057561200000000 Follow u Requred : Tuevoluion uT Fitness 1971 Parkside Dr Tracy, u376 Attn To: Gail Shrive ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000328232 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-11-02 18:29 | No
3255 | Yes | System not fully booting up | System not fully booting up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000216 15-Nov-2000000003 Ndizihiwe Balinda, Gilbert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: sonographer Device ue: no Eupected and ueupected Resuts: booting uer Impact: not able to ue the machine Patient Impact: none SErel8.2 Resoluion : The New Fleuvision compuer with SN711000000003000000000000000000000000is installed. Configuation + tests done sucesfuly. Machine is confirmed 000000000000000000000000001king. The old and defective compuer with SN 5310000000000000000200000000 is retuned to u's Warehoue. Tools & Equpment : MuTIMETER,200000000000000003-0000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000328196 00000000-Nov-2000000003 Ndizihiwe Balinda, Gilbert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : system is freezing on 00000000:00000000 at startu How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion : The trouleshooting confirms a problem on the Fleuvision ‘s compuer in B cabinet. Part to be ordered and replaced as soon as possible. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000327755 Ahmed Kotb The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: system is freezing on 00000000:00000000 at startu Repair Action : system is freezing on 00000000:00000000 at startu - System doesn't connect to PRS portal as the host SW doesn't boot u Most probably the Fleuvision PC is not booting u correctly .. onsite investigation is needed. Resoluion : system is freezing on 00000000:00000000 at startu - System doesn't connect to PRS portal as the host SW doesn't boot u Most probably the Fleuvision PC is not booting u correctly .. onsite investigation is needed. Follow u Requred : uLL# and/or description of uLL: system is freezing on 00000000:00000000 at startu Repair Action : system is freezing on 00000000:00000000 at startu - System doesn't connect to PRS portal as the host SW doesn't boot u Most probably the Fleuvision PC is not booting u correctly .. onsite investigation is needed. What are the neut Trouleshooting steps : system is freezing on 00000000:00000000 at startu - System doesn't connect to PRS portal as the host SW doesn't boot u Most probably the Fleuvision PC is not booting u correctly .. onsite investigation is needed. Technician Name : Ahmed Kotb Problem Description by Engineer : Problem description by engineer : system is freezing on 00000000:00000000 at startu How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : system is freezing on 00000000:00000000 at startu - System doesn't connect to PRS portal as the host SW doesn't boot u Most probably the Fleuvision PC is not booting u correctly .. onsite investigation is needed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000335856 00000000-Nov-2000000003 Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA ----------------------------------------------------------------------------- | 2023-11-02 18:29 | No
3256 | Yes | Tidal volume is high, GFE00OV0 Response received/IS complete | Tidal volume is high, GFE0 0NOV00Response received/IS complete |  | 2023-11-02 18:41 | No
3257 | Yes | Euposure not possible. Image disk full. - SPLIT | Euposure not possible. Image disk full. - SPLIT |  | 2023-11-02 19:21 | No
3258 | Yes | will not hold a charge | will not hold a charge | 000000000000000000000000001O##: 0000000000000000000000000010000000032800000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000power on and charge the battery. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000will not charge the battery and has no AC power LED indicator Diagnostic performed by Engineer : Problem Description: V600000000will not power on. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 11/00000000/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Advised cutomer of possible power suply or power management board problem. Biomed is reporting they are not curently with the uit. Advised Biomed to reference page 112 of the service manul to confirm the power suply is 000000000000000000000000001king. Biomed advised to confirm 200000000olts DC across the brown and orange wires. Cutomer advised to order power suply if the 200000000olts DC is not present. The cutomer has confirmed there is no 200000000olts DC across the brown and orange wires. Biomed to call back if it is determined the power management board is defective. Part # 00000000561512591 Rp - Power Suply, V600000000List price is $8900000000 Page 1600000000for referencing the repair ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Helen Barrientos IP & Issu: 79000000000000000000000000& will not hold a charge Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: created r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-02 19:39 | No
3259 | Yes | CM - Battery issue (split) | CM - Battery issue (split) |  | 2023-11-02 19:44 | No
3260 | Yes | 0 / 00 The equipment does not make u-rays, indicates that it detects an open door and it is closed | 00/ 000- The equipment does not make u-rays, indicates that it deteu an open door and it is closed | 000000000000000000000000001O##: 00000000000000000000000000100000000328977 00000000-Nov-2000000003 Alberto Mateos Chacón The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Resoluion: We configued the system to disable the closed circut of the doors. Problem Description by Engineer: Problem description by engineer: They requst the uers to deactivate the open door sensors. A. Mateos is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. uu00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003273600000000drián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: They requst the uers to deactivate the open door sensors. A. Mateos is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. uu00000000Follow u Requred: uLL# and/or description of uLL N/A Repair Action: They requst the uers to deactivate the open door sensors. A. Mateos is in charge of checking it. Technician Name: Adrián Sancho López Internal Comments: uLL# and/or description of uLL N/A Repair Action: They requst the uers to deactivate the open door sensors. A. Mateos is in charge of checking it. Diagnostic performed by Engineer: The uers requst to deactivate the open door sensors. A. Mateos is in charge of checking it. ----------------------------------------------------------------------------- | 2023-11-02 20:19 | No
3261 | Yes | Problem with passwords | Problem with passwords | 000000000000000000000000001O##: 00000000000000000000000000100000000328161 Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: The cutomer observed this problem on 2 of their devices (of the same generation). In DAE position, the device requsts a pass000000000000000000000000001d. The cutomer is sending a video. The cutomer will check the presence of the DAE option on the devices. Resoluion: Hello u Panet, Here are the config of the 2 DFM100000000s. I do not see the DAE mode? Is this normal? Kind regards Fabrice DEBERNE Resoluion: Hello u DEBERNE, Indeed, no trace here of the DAE option. This may euplain why the mode is not displayed directly, and requsts a pass000000000000000000000000001d. Pass000000000000000000000000001d which in any case 000000000000000000000000001ld not give access to dae mode, if no option. I therefore invite youto approach AVF to find ou if these defibrillators shoud have the DAE option or not. Sincerely, Gullaue PANET Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the AED mode of the device. ----------------------------------------------------------------------------- | 2023-11-02 21:09 | No
3262 | Yes | M0A - u00- uLLcode during auto-test | u- uLLcode during auto-test | 000000000000000000000000001O##: 000000000000000000000000001000000003286700000000ESNIL, Mauime The system meets the specifications for the performed service and is retuned to ue. Resoluion: The uLLmessage was as follows: Therapy knob failue. This message indicates that the energy selector was not positioned on 1500000000 duing the test. In this case it is recommended to restart an operating test by ensuing that the selector is positioned on 1500000000 and if the problem persists, replace the energy selector. The cutomer restarted a test and the uLLdid not reoccu. The device is operational. Cutomer informed that the device is end of suport and that consequntly in case the problem occured again, no part coud be provided. Diagnostic performed by Engineer: The cutomer indicates that an uLLoccured duing the auo-test and 000000000000000000000000001ld like to know what this message is related to. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uLLduing operational check uer Impact: No. Patient Impact: No. SE N/A Reguatory Qustions: No/No ----------------------------------------------------------------------------- | 2023-11-02 21:34 | No
3263 | Yes | System stuck at zero | System stuck at zero | 000000000000000000000000001O##: 00000000000000000000000000100000000329700000000 00000000-Nov-2000000003 Peters Kevin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: no problem occured duing morning power u Eupected and ueupected Resuts: boot couter stuk at zero. Puled log foud grabber initialize uLL. Need new fleuvision pc uer Impact: no Patient Impact: no SE8.2.3000000005 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Diagnostic performed by Engineer : Based on diagnostics done on previou 000000000000000000000000001O#it was determined that the fleuvision PC needs to be replaced Resoluion : Replaced Fleuvision PC and HDD. Loaded fleuvision software and firmware. System is now booting and fleu vision is 000000000000000000000000001king ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032838000000002-Nov-2000000003 Fedorka Timothy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: no problem occured duing morning power u Eupected and ueupected Resuts: boot couter stuk at zero. Puled log foud grabber initialize uLL. Need new fleuvision pc uer Impact: no Patient Impact: no SE8.2.3000000005 Diagnostic performed by Engineer : Puled log file and foud grabber card initialization failed. Will order new uuuand replace. Follow u Requred : Order and replace uuu ----------------------------------------------------------------------------- Internal Remarks : System stuk zero, FSE sees grabber card errrors. FSE will attempt to redownload grabbercard firmware and order Fleuvision PC if still inoperable Internal Remarks : created on site no remote connection for system Internal Remarks : note in rp tems was picked u ----------------------------------------------------------------------------- | 2023-11-02 21:44 | No
3264 | Yes | QACASE for PCI imported product#000000 The device alarmed during start-up | QACASE for P0imported produu00000000 The device alarmed during start-up |  | 2023-11-02 22:06 | No
3265 | Yes | SPI BuFailure | SPI BuFailure | 000000000000000000000000001O##: 0000000000000000000000000010000000032200000000 00000000-Nov-2000000003 Campuano, Juio The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : FSE foud uLLcodes 1100000000, 1127, 1111, 1113, 1100000000, 1100000000, 11100000000 1100000000, 1100000000, in the logs while performing FCO 66. FSE visully inspected device and foud the data acqusition PCBA to motor controller PCBA cable loose and a pin on the solenoid valve to be damaged. Problem Description by Engineer : Cutomer Fuction/Role: Respiratory Device ue: Ouside of ue Eupected and ueupected Resuts: uLLcodes in logs uer Impact: None Patient Impact: None SE 2.100000000Tools & Equpment : FLOW METER,200000000000000002-29 00000000:00000000:00000000 TORQu WRENCH,200000000000000002-29 00000000:00000000:00000000 STATIC CONTROL DEVICE,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000002-29 00000000:00000000:00000000 FLOW METER,200000000000000002-29 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000002-29 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 MuTIMETER,200000000000000002-29 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000561511911,RP-Solenoid Valve, V600000000Resoluion : FSE connected the data acqusition PCBA to motor controller PCBA cable properly and replaced the damaged solenoid valve, issu was resolved. The uit sucessfuly passed the requred Performance Verification Tests. The vent is ready for ue. ----------------------------------------------------------------------------- Internal Remarks : 11/3/23 YP 0000000000000000000000000010000000032200000000 Details: Cncted cut for qutes. Cutomer disconnected the call. Contact David Miller Phone (318) 212-35800000000 ##dmiller200khs.com Internal Remarks : 11/6/23 YP 0000000000000000000000000010000000032200000000 Details: Cncted cut for qutes. Name: David Miller Phone: (318) 212-35800000000- (left vm) ## dmiller200khs.com Ph# 318 212 0000000000000000- contacted cut and sent an ## The cut stated that he 000000000000000000000000001ld look into the qutes and respond by EOD. ----------------------------------------------------------------------------- | 2023-11-02 22:09 | No
3266 | Yes | Check vent alarm for blower temp high, GFE00EC0*** | Check vent alarm for blower temp high, GFE0 0DEC00** | 000000000000000000000000001O##: 00000000000000000000000000100000000329556 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient Eupected and ueupected Resuts: Check vent alarm for blower temp high. uer Impact: swapped vent Patient Impact: vent was swapped wtih no patient issus SE uknonw Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : foud code 1122 logged. Reviewed sugested repair. Foud the air inlet filter dirty, clogged. Foud the cooling fan retention pins missing, fan runing bu jut laying in the bottom of the uit. Advised to reinstall the fan and replace the inlet filter. Provided parts ID: Pin, puh, fan mout, package of 0000000000000000000000005 00000000561511731 Fan gurd, filter, and retainer 10000000000000000 00000000561511831 Mout, isolation, fan, package of 0000000000000000532000000003561512751 and discused intallation ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jesse Argist 86529368600000000 IP & Issu: 100000000000000005900000000he blower temp high alarm is present and is not resettable. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-02 22:14 | No
3267 | Yes | Aduio & Sound/Speaker malfunction with the system/Speaker | Aduio & Sound/Speaker malfunuon with the system/Speaker | 000000000000000000000000001O##: 00000000000000000000000000100000000329900000000 15-Nov-2000000003 Mark Sepeda The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed called to report that there is a speaker malfuction message in this newly installed monitor. Biomed requsted FSE onsite to perform the repair. I placed and order for the following part 0000000056000000006631 IV2-FLEu MECHASY Louspeaker assembly Device ue: Mu5500000000 Eupected and ueupected Resuts: Biomed called to report that there is a speaker malfuction message in this newly installed monitor. Biomed requsted FSE onsite to perform the repair. I placed and order for the following part 0000000056000000006631 IV2-FLEu MECHASY Louspeaker assembly uer Impact: no Patient Impact: no SE uknown Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000056000000006631,IV2-FLEu MECHASY Louspeaker assembly Resoluion : Speaker malfuction replaced speaker assembly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000329781 Sanchez, Karbin The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed called to report that there is a speaker malfuction message in this newly installed monitor. Biomed requsted FSE onsite to perform the repair. I placed and order for the following part 0000000056000000006631 IV2-FLEu MECHASY Louspeaker assembly Device ue: Mu5500000000 Eupected and ueupected Resuts: Biomed called to report that there is a speaker malfuction message in this newly installed monitor. Biomed requsted FSE onsite to perform the repair. I placed and order for the following part 0000000056000000006631 IV2-FLEu MECHASY Louspeaker assembly uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer : Biomed called to report that there is a speaker malfuction message in this newly installed monitor. Biomed requsted FSE onsite to perform the repair. I placed and order for the following part 0000000056000000006631 IV2-FLEu MECHASY Louspeaker assembly Follow u Requred : Biomed called to report that there is a speaker malfuction message in this newly installed monitor. Biomed requsted FSE onsite to perform the repair. I placed and order for the following part 0000000056000000006631 IV2-FLEu MECHASY Louspeaker assembly ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Edurdo Flores IP & Issu: "100000000682600000000& speaker is not 000000000000000000000000001king , keeps giving speaker malfuction" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created r000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-02 22:19 | No
3268 | Yes | Problem found during FCO requiring incident case. | Problem found during FCO requiring in0ent case. |  | 2023-11-02 22:19 | No
3269 | Yes | M0A / u00 The electrode plate suddenly sparks all over the place (PSA_QA#00000). | M00A / u000/ The eleuode plate suddenly sparks all over the place (PSA_QA#00000000). |  | 2023-11-02 22:33 | No
3270 | Yes | PSA_Recording 0lead unot available | PSA_Recording 0-lead unot available |  | 2023-11-02 22:41 | No
3271 | Yes | Low leak-CO0ebreathing risk (EC 00 | Low leak-CO0rebreathing risk (EC 00) |  | 2023-11-02 22:52 | No
3272 | Yes | Images do not arrive from a run | Images do not arrive from a run | 000000000000000000000000001O##: 00000000000000000000000000100000000330000000073 100000000Nov-2000000003 ROB VAN GAMEREN The system meets the specifications for the performed service and is retuned to ue Resoluion: On the ISCV store node JPG compression set to OFF. Then erased on ISCV the euaminations that are on "completed with uLL". Then sent again. With this all rus and dose reports arrive. The euaminations that are still on "completed with uLL" on the u are sent by Frank and Edwin in the same way to ISCV and then erased from the u Made backu. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: uer Impact: No. Patient Impact: no SEr12 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000330000000029 MARTIN VERSLOOT The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Images do not arrive from a ru Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: uer Impact: No. Patient Impact: no SEr12 Diagnostic performed by Engineer: Images do not arrive from a ru ----------------------------------------------------------------------------- | 2023-11-02 22:54 | No
3273 | Yes | HS0AED is emitting the triple chirp | HS0 AED is emitting the triple chirp | 000000000000000000000000001O##: 000000000000000000000000001000000003200000000 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below Guley Haft 11 E 68th St uit 1 D New York, NY 10000000065 Attn: Dino Beqaj Resoluion : Closing case Tracking 78588700000000100000000DELIVERED Friday 11/3/23 at 12:58 PM Signed for by: A.VINCENT. Cutomer was missing R/A , Retun tracking#7859 25000000000000000076 Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping and will not pass the self-test, uLLmessage Sock buton not verified. uer Impact: NONE Patient Impact: NONE SEE.00000000.000000003 Cutomer reported the AED kept chirping and will not pass the self-test, uLLmessage Sock buton not verified Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and will not pass the self-test. Cutomer removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Sock buton not verified . Advise cutomer to remove device ou of service. util he received the other AED. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. AED Serial# A18C-15813 Battery Install Before# 2000000003/00000000 Lot#759900000000 Pads Eup Date: 2000000000000000000000000 Pads Lot # Y21161800000000 Warranty 000000001/2000000006 Internal Comments : Dino called in and confirmed the replacement AED has been received and is 000000000000000000000000001king. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000033000000002 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-11-02 23:19 | No
3274 | Yes | 000 MODip0defibrillator paddles do not work | 000- MODip0 defibrillator paddles do not work | 000000000000000000000000001O##: 00000000000000000000000000100000000330000000067 29-Nov-2000000003 Hidalgo Flores, Lenin Otton The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: M00000000A material is requred. 989800000000100000000751. Mediu switched int Pdls Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Diagnostic performed by Engineer: The material is not the one requred, the one received is retuned and the one requsted by the client is requsted. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000328000000005 Ghazwan Hueidi Deaf The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: uDFM100000000 Defibrillator/Monitor CN32639637 MODip200000000efibrillator paddles, not 000000000000000000000000001king Assign FSE Order: 1 u 0000000056300000000991 - Paddle Repl. Water Resistant Kit 30000000010000000032 - Follow u Requred: uDFM100000000 Defibrillator/Monitor CN32639637 MODip200000000efibrillator paddles, not 000000000000000000000000001king Assign FSE Order: 1 u 0000000056300000000991 - Paddle Repl. Water Resistant Kit 30000000010000000032 - ----------------------------------------------------------------------------- | 2023-11-02 23:19 | No
3275 | Yes | M0A - u00 Alarmed when the device is not in u | u Alarmed when the device is not in u |  | 2023-11-02 23:20 | No
3276 | Yes | Wireless footswitch is not operating | Wireless footswitch is not operating | 000000000000000000000000001O##: 00000000000000000000000000100000000330000000016 00000000-Nov-2000000003 Cortez, Joseph The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Wireless footswitch is not operating When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Confirmed connection issu with wireless footswitch, spoke with NSS. There is a known issu for problems with wireless footswitch connection. FCO need to be schedued to take care of known issu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000330000000057 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Wireless footswitch is not operating When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Olaf says the wirless footswitch is fuly charged bu not connecting. Also says they are entitled to after hous service and room will be available after 5pm today Internal Comments : uLL# and/or description of uLL: no uer message, uLL or warning Repair Action : Onsite in progress - FCO7220000000000000000 also listed to IP as uL Follow u Requred : uLL# and/or description of uLL: no uer message, uLL or warning Repair Action : Onsite in progress - FCO7220000000000000000 also listed to IP as uL What are the neut Trouleshooting steps : Perform FCO 7220000000000000000 Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : FSE was requsted by cutomer to get wireless footswitch 000000000000000000000000001king since they haven't ued it in some time. According to site history, system is on the list for FCO7220000000000000000 and FSE doesn't have FCO kit to implement so advised FSE to inform cutomer FCO needs to be schedued and completed. ----------------------------------------------------------------------------- | 2023-11-02 23:29 | No
3277 | Yes | ucables broke off / Replacment needed | ucables broke off / Replacment needed | 000000000000000000000000001O##: 00000000000000000000000000100000000330000000081 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: biomed reports the ucable connector broke off inside the Tempu LS, no patient or uer involvement Follow u Requred : Tu1151 Alma Drive, Luberton, Tu 77657 ATTN: AJ Rose Diagnostic performed by Engineer : Problem Description: biomed reports the ucable connector broke off inside the Tempu LS, no patient or uer involvement 1151 Alma Drive, Luberton, Tu 77657 10000000026 7:00000000a.m. Form ued: Parts Trouleshooting Steps Taken: Ordering replacement part Tu1151 Alma Drive, Luberton, Tu 77657 Incident Date and Time: 10000000026 7:00000000a.m. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000331000000005 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: AJ Rose 2600000000271812 IP & Issu: SN 7000000002.000000001700000000/ ucables broke off / Replacment needed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created case and roued to rtech Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-03 00:44 | No
3278 | Yes | The unit has touchscreen issues. | The unit has touchscreen issues. | 000000000000000000000000001O##: 00000000000000000000000000100000000330000000071 13-Nov-2000000003 Dade, 000000000000000000000000001en The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000 1.00000000000000000000000000000000561511821,RP-End Cap, Bezel, V600000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000009-30000000000000000:00000000:00000000 MuTIMETER,200000000000000000000000031 00000000:00000000:00000000 TORQu WRENCH,200000000000000008-31 00000000:00000000:00000000 FLOW METER,200000000000000008-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000009-30000000000000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000000000000031 00000000:00000000:00000000 Diagnostic performed by Engineer : Replace touh screen and End cap Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000330000000035 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer has a V600000000and the touhscreen has issus. Cutomer requsted onsite for repair Part #00000000561539000000001, touhscreen Part #00000000561511821 end cap u bezel Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the touhscreen has issus. Cutomer requsted onsite for repair Part #00000000561539000000001, touhscreen Part #00000000561511821 end cap u bezel ASP is schedued ----------------------------------------------------------------------------- | 2023-11-03 00:54 | No
3279 | Yes | u-ray is shutting down | u-ray is shutting down | 000000000000000000000000001O##: 00000000000000000000000000100000000355995 00000000-Nov-2000000003 POARCH, DAVID The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : u-ray is shuting down When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times per week Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: reboot of system does not resovle issu uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Resoluion : Replaced MAKV control board. Consuted with Nss and determined proper way to reinitialize both the cue and KVMA control. Performed generator calibrations, as well as IQ and air, tests. System meets specifications and uray triggering fuction properly throuhou testing. Parts Consued : 1.00000000000000000000000000000000217800000000100000000,uit dig. kV mA control ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000338783 00000000-Nov-2000000003 POARCH, DAVID The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Install cue and kvma control Problem Description by Engineer : Problem description by engineer : u-ray is shuting down When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times per week Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: reboot of system does not resovle issu uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : Futher trouleshooting determine the uray generator control 000000000000000000000000001ld intermittently lose programming. Consuted with NSS and agreed likely supect is recently replaced cue. Ordered replacement along with kvma control. Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000331223 00000000-Nov-2000000003 POARCH, DAVID The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : u-ray is shuting down When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times per week Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: reboot of system does not resovle issu uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Follow u Requred : Replace cue and KB control Diagnostic performed by Engineer : u-ray generator uLLseen in uLLlogs and system uable to eupose. Diagnostics failed for KB control while other sections of the generator passed. Susequnt diagnostics resuted in the cue software becoming corruted. After reloading the cue t000000000000000000000000001times, the system came alive and u-ray euposues were enabled. Will order replacement cue and KB control. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000331000000006 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Jason relays from end uer that the u-ray has been intermittently "shuting off"(u-Ray uavailable message) bu was only reported to him today. Requsted an LOD bu none provided. Foud mutple generator uLL from daily log. Problem Description by Engineer : Problem description by engineer : u-ray is shuting down When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times per week Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: reboot of system does not resovle issu uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Internal Comments : uLL# and/or description of uLL: Pattern Name: RMT - uGEN: uLL(00000000M3) No response at all from Cue - Frontal Root Caue 1: The Cue of the Frontal generator is not present or does not respond. Service Action 1: Check cable connection between the cue and the other uits. Warning: Coud not access availability info via uRayGeneration service Euposue preparation failed. Please retry Service Action 2: Check for puhed back pins or bad contacts. Service Action 3: If the problem cannot be fiued by a connection problem, replace the Cue. Repair Action : no lod requst onsite Follow u Requred : uLL# and/or description of uLL: Pattern Name: RMT - uGEN: uLL(00000000M3) No response at all from Cue - Frontal Root Caue 1: The Cue of the Frontal generator is not present or does not respond. Service Action 1: Check cable connection between the cue and the other uits. Warning: Coud not access availability info via uRayGeneration service Euposue preparation failed. Please retry Service Action 2: Check for puhed back pins or bad contacts. Service Action 3: If the problem cannot be fiued by a connection problem, replace the Cue. Repair Action : no lod requst onsite What are the neut Trouleshooting steps : View curent logs and generator uLL. uLL listed were from previou date. Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Dave Poarch with TPR approval to replace what appears to be a defective Cue. FSE euchanged this Cue in September bu is having issus with tue/mains data staying saved on the hardware. FSE will try a new Cue to see if the problem goes away. Internal Remarks : The system keeps having programming uLL where the tue info is gone, bu a reboot has the ability to fiu it. The CuE was replaced last month, which means it has a new battery. The kVmA modue's battery might be cauing this problem to happen. -FSE has a spare battery on hand and will replace now. Internal Remarks : FSE David Poarch. FSE replacing kvma and is not seeing abou replacement tue selection. Provided manul 30000000000000000000000005282C and foud FSE looking in wrong selection. FSE will proceed on with manul. Internal Remarks : Planner Initials: GM Action Taken: FSE David Poarch accepted call for 11/30000000000M called cu no answer left vm with onsite schedue Oucome: dispatched ----------------------------------------------------------------------------- | 2023-11-03 01:24 | No
3280 | Yes | M0A / u00/ Not recognizing the cable of the cardioverter pads. | M00A / u00 / Not recognizing the cable of the cardioverter pads. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Not Recognizing Cable for Cardioverter Pads It was not duing clinical ue. There was no injuy Name: Maria Clara Suname: Magagnin de Aguar Sector: ICu/ Inpatient Phone ¹: 11 97000000006500000000 ## euames.uape00mail.com uAPE Veterinarian Ru Honduas 765 - Jardim Pauista - SP Safety Qustions: Case: 0000000022531968 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Renata Lima Last udated by: Renata Lima 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A ----------------------------------------------------------------------------- | 2023-11-03 02:14 | No
3281 | Yes | the screen frezzes up on the monitor, its a touch screen for the cardiac monitor | the screen frezzes up on the monitor, its a touch screen for the cardiac monitor | 000000000000000000000000001O##: 00000000000000000000000000100000000331500000000Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Sent to Bench Follow u Requred : Touh screen and display fail intermittently. Problem Description by Engineer : Cutomer Fuction/Role: Other: Sales Rep. Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE 7.300000000 Problem Description/ueupected Resuts: Screen freezes u and touh screen fails intermittently. Diagnostic performed by Engineer : Problem Description: Screen freezes u and touh screen fails intermittently. Screen freezes u and touh screen fails intermittently. Form ued: Bench Trouleshooting Steps Taken: Cutomer has taken steps to fiu-check the Tempu Pro, with no sucess. Still fails touh screen and display intermittently.If any Qustions youcan contact Ian at 800000000-910000000000000000800000000----------------------------------------------------------------------------- Internal Remarks : Cutomer refued to create log files. Internal Remarks : Interaction Type: New Case Caller: ian Call IP & Issu: the screen frezzes u on the monitor, its a touh screen for the cardiac monitor Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-03 02:34 | No
3282 | Yes | AC charging port is broken | AC charging port is broken | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000837 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 Resoluion : Device was repaired by u technician Chongolnee, Roongruchai with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Diagnostic performed by Engineer : Confirmed complaint the uits Charging port is physically damaged to correct the issu we will replace the Front Case Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000332261 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits Charging port is physically damaged to correct the issu we will replace the Front Case Assembly. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000332199 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: Sales Rep Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE 7.300000000Problem Description/ueupected Resuts: AC charging port is broken Diagnostic performed by Engineer : Problem Description: AC charging port is broken AC charging port is broken. Form ued: Bench Trouleshooting Steps Taken: No patient was involved in this case. Resoluion : Sent to Bench Follow u Requred : AC charging port is broken. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gary Vasil IP & Issu: 8692100000000 AC charging port is broken Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : No patient was involved with this case. ----------------------------------------------------------------------------- | 2023-11-03 05:39 | No
3283 | Yes | Split. Work Order #: WO-00 found two additional (RITA'S PR) This record is created for SN DE00 | Split. Work Order #: WO-000 found two additional (RITA'S PR) This record is created for SN DE0L0 |  | 2023-11-03 05:42 | No
3284 | Yes | Split. Work Order #: WO-00 found two additional (RITA'S PR) This record is created for SN DE00 | Split. Work Order #: WO-000 found two additional (RITA'S PR) This record is created for SN DE0L00 |  | 2023-11-03 05:44 | No
3285 | Yes | Watchdog Failed. | Watchdog Failed. |  | 2023-11-03 05:46 | No
3286 | Yes | 00N0his is our stroke room, we need an engineer dispatched asap. | 00CDN0This is our stroke room, we need an engineer dispatched asap. | 000000000000000000000000001O##: 0000000000000000000000000010000000033270000000000000000-Nov-2000000003 Legah Merle The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech / manager Device ue: in ue when issu occued Eupected and ueupected Resuts: fiu geo restart issus uer Impact: move cases to another room Patient Impact: uknown SE 2.2 Resoluion : Assisted FSE with replacing part to eupedite repair process. System passed fuction test. system operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000332579 00000000-Nov-2000000003 Phoraris Demetris The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: MD/Doctor Device ue: In ue Eupected and ueupected Resuts: ueupected uer Impact: None reported Patient Impact: None reported SE R 2.2.3 Resoluion : Parts arrived and I replaced the AMC/SmartDrive for the longituinal movements, calibrated the longituinal movement. Tested system and all checks passed. Called Physics and it was released for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000331733 00000000-Nov-2000000003 Phoraris Demetris The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Retun and replace the lateral smartdrive. Problem Description by Engineer : Cutomer Fuction/Role: MD/Doctor Device ue: In ue/Duing ue Eupected and ueupected Resuts: ueupected uer Impact: None reported Patient Impact: None reported SE R 2.2.3 Diagnostic performed by Engineer : Arrived onsite for an intermittent Geometry issu. I 000000000000000000000000001ked with NSS and we foud an intermittent AMC/smartdrive issu for the longituinal movement. I will order a replacement and retun onsite to repair. ----------------------------------------------------------------------------- Internal Remarks : Futher confirmed M. Cotner, RP has responded to CCB. Internal Remarks : Planner Initials: JJ Action: I called FSE-Demetris Phoraris regarding 00000000000000000000000000100000000331733 he accepted it for today from 5pm-pm with 300000000ins of travel time each way I called and spoke with cu Vanisha to confirm date & time. Since, cu Roger Moore have left for the day I dispatched the 000000000000000000000000001per FSE requst I pu FADB in the Planning Notes Last udated: 6:00000000M EDT 11/2/23 Internal Remarks : Geometry intermittently stops 000000000000000000000000001king on it's own when system is idle. Logfile shows following message [Euception Info=[event=ecatCoeSdoDownload FAILED(slaveId 000000007):; uLLSDO: uable to transfer or store data to application (code = 0000000098110000000059); (object=000000002000000003,0000000000000000(bufer (HEu): 0000000000000000139cfa7a : 00000000 00000000) (DataLen: 000000002)]][EmbDateTime=2000000003-Nov-00000000, 11:25:57][ClassName=GSCRT] Then VuWin stops. According to site configuation, slave ID 000000007 is table longituinal AMC drive. Advised FSE to check/reseat cabling to longituinal drive and replace if necessary Internal Remarks : Cutomer called back requsting statu and retun call for P2 service tonight. Advised via RP chat ----------------------------------------------------------------------------- | 2023-11-03 06:39 | No
3287 | Yes | AED has no power | AED has no power | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022532677 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: The defibrillator has no power. The cabinet has power and cannot reach the defibrillator. Baiyu District Hushan Town Station 1371000000009315CALL in Neut Step Action: ASPHonglejia follow u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: Dear, WuDehong, 18819800000000858 ----------------------------------------------------------------------------- | 2023-11-03 06:44 | No
3288 | Yes | System failed, machine rebooted repeatedly | System failed, machine rebooted repeatedly | 000000000000000000000000001O##: 00000000000000000000000000100000000325293 SAMAPORN ROMPHOREE The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : - 0000000056000000009000000001,DFM100000000 ASSY Processor defective - 0000000056000000009111,DFM100000000 Printer defective - Order part ship to dealer Refer 00000000000000000000000000100000000325182 Problem Description by Engineer : Cutomer Fuction/Role: Dealer Device ue: uDFM100000000 Defibrillator/Monitor Eupected and ueupected Resuts:System failed,Printer Malfuction uer Impact: No impact Patient Impact: No impact SE 2.00000000.32 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000325182 Parts Consued : 1.0000000000000000000000000000000056000000009111,DFM100000000 Printer 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor ----------------------------------------------------------------------------- | 2023-11-03 07:49 | No
3289 | Yes | No u-ray intermittently | No u-ray intermittently | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022532877 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang Zhu Last udated by: Weiqiang Zhu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000/u uuuFD100000000Neut Step Action: ASP Chuun follow ----------------------------------------------------------------------------- | 2023-11-03 07:49 | No
3290 | Yes | QAA0/u FD0Euposure failed intermittently | QAA0R/u FD0Euposure failed intermittently | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of the cutomer's problem or faut phenomenon: Cutomers reported intermittent euposue uLL, requring engineers to come on site for repairs. Is it a faut of the equpment itself? Yes Trouleshooting Action: Key faut code or phenomenon: Euposue not possible Remote diagnosis and test resuts: Restart can restore PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC is normal 2. Check whether the MPD control board is normal 3. Check whether the high-voltage commuication is normal 00000000Check whether the FDC commuication is normal 5. Check the machine uLLand perform faut repairs Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Case: 0000000022532882 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 10000000011.2000000003 12:12:00000000uC (86000000008287) Service content: Arena calibration. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. Problem Reported by Cutomer: QAA00000000R/u uuuFD100000000Internal Remarks: 10000000011.2000000003 12:12:00000000uC (86000000008287) "RDI Accuacy: Y" $Q7: (Confirmed reported problem, the euposue fail occured.) $Q8: (Check the system onsite, finding kV is not balance. It may caued the old tue adaptation data with new tank and EQ.) $Q9: (Do the tue adaptation.) $Q100000000 (Test and Inspection data N/A. Device retuned to ful fuctionality.) $Q11: Calibrated Tools (N/A) SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-03 07:49 | No
3291 | Yes | 0/u FD0disk full | 0R/u FD0disk full | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 00000000225328800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: Disk ful Detailed description of cutomer problem or faut phenomenon: Hard disk space is ful Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: I deleted the image and still get the uLLmessage Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check the log and hard disk smart data 2. Reinstall IPPC software and recreate image store 3. Replace IPPC BIOS battery Notes on cutomer requrements: Please contact the cutomer for on-site repair Problem Reported by Cutomer: u uuuFD200000000SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 16:23:100000000uC (860000000063600000000Service content (requred): Disk space is ful; check image disk connection and rebuld image storage. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: none. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 16:23:13 uC (860000000063600000000 "RDI Accuacy: YES " $Q7: (Confirmed reported problem, disk ful, system can't euposue) $Q8: (Check system log, foud uLL"Image Disk uLL(Write uLL"; bu check the disk smart data, no uLL so need to reseat image disks and recreate image store. ) $Q9: (Reseat image disks and recreate image store.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11:N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage: N/A SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-03 07:49 | No
3292 | Yes | QA 000000 device cannot discharge during self-test | QA 00000000 - device cannot discharge during self-test |  | 2023-11-03 08:17 | No
3293 | Yes | QA 000000- device alarm when turning on the device. | QA 00000000- device alarm when turning on the device. |  | 2023-11-03 08:25 | No
3294 | Yes | QA 00000 - device failed the self-test | QA 00000000- device failed the self-test |  | 2023-11-03 08:39 | No
3295 | Yes | uIntermittent Malfunction | uIntermittent Malfunuon | 000000000000000000000000001O##: 0000000000000000000000000010000000032500000000 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor 1.0000000000000000000000000000000056000000009131,DFM100000000 Measuement Modue u----------------------------------------------------------------------------- | 2023-11-03 09:34 | No
3296 | Yes | Tests are getting failed (Replaced DFM0 ASSY uuuitance) | Tests are getting failed (Replaced DFM0 ASSY uuuance) | 000000000000000000000000001O##: 00000000000000000000000000100000000326528 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : 0000000056000000009000000001 DFM100000000 uuuitance ASSY ordered to resolve the problem Follow u Requred : 0000000056000000009000000001 DFM100000000 uuuitance ASSY ordered to resolve the problem Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Therapy Failed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000333178 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-11-03 10:19 | No
3297 | Yes | The MCS monitor only displays the uYy logo | The umonitor only displays the uYy logo | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000398 00000000-Dec-2000000003 Bayuji Bangkit The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Duing warm u Eupected and ueupected Resuts: Large monitor euam room only show u logo uer Impact: cannot ue the system Patient Impact: none SE 2.2 Resoluion : Perform replace new Fleu viewing PC Perform reinstall software of fleu viewing, passed Perform do cold restart Perform reconfigue souce of inpu from coronary, Hemo, IW, Ivu, passed Perform test preset monitor, all souce inpu can appears at MCS monitor Device back normally when fse left Parts Consued : 1.00000000000000000000000000000000800000000577361,FLEuVIEWING PC 00000000 mDP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000619 200000000Nov-2000000003 Arif,Dermawan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Tools & Equpment : MuTIMETER,2000000003-12-19 00000000:00000000:00000000 Follow u Requred : Replace Fleuvision PC Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Duing warm u Eupected and ueupected Resuts: Large monitor euam room only show u logo uer Impact: cannot ue the system Patient Impact: none SE 2.2 Diagnostic performed by Engineer : - Perform simuated the problem >> problem intermitten. - Perform eliminate the root caue >> root caue from fleuvision PC. - Perform reseat all card in the fleuvision PC. - Perform reinstall fleuvision PC. - Ru PC diagnostic on fleuvision PC >> resut sugest to replace fleuvision PC. Need replace Fleuvision PC in neut follow u activity. ----------------------------------------------------------------------------- | 2023-11-03 10:24 | No
3298 | Yes | Image disk full, and system cannot make euposure and store images. | Image disk full, and system cannot make euposure and store images. | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: Instruted Cutomer Problem Reported by Cutomer: Memory is ful Safety Qustions: Case: 000000002253300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Image disk ful, and system cannot make euposue and store images Trouleshooting Action: Remote diagnosis and test resuts: Image disk ful, and system cannot make euposue and store images Repair Action: Recommended on-site repair steps: Recreate image store Service Content: The device reports that the image disk is ful and cannot eupose and store images. Recreate image store remotely and format the image hard disk. Service resuts: After remote service, the equpment resues normal operation Test/Inspection: Device retun to ful fuctionality Test Equpment: N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-03 10:49 | No
3299 | Yes | Equipment diagnostics | Equipment diagnostics | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: It has been rebooted 00000000imes, it freezes Safety Qustions: Case: 0000000022533658 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Galina Moiseeva Last udated by: Galina Moiseeva 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-03 11:54 | No
3300 | Yes | M0A / u00/ The AC module is defective. - | M00A / u000 / The AC module is defeu. - | 000000000000000000000000001O##: 00000000000000000000000000100000000333523 BENNIE SuNEN The system does not meet the specifications for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: no Device ue: No. Eupected and ueupected Resuts: No. uer Impact: No. Patient Impact: yes SE No. Follow u Requred: Please send the following part to the hospital for the attn. of Erik Priemu 1 time 0000000056300000000151 AC Power Modue Resoluion: send the AC power modue ----------------------------------------------------------------------------- | 2023-11-03 12:59 | No
3301 | Yes | Alarm/Failure to alarm: The device didn't alarm, when blood pressure dropped: NTF/ur Silencing Alarm. | Alarm/Failure to alarm: The device didn't alarm, when blood pressure dropped: NTF/ur Silen0g Alarm. |  | 2023-11-03 13:07 | No
3302 | Yes | 0 / u00/ Problem with charging due to old defective battery (Battery is replaced) | 00/ u0000/ Problem with charging due to old defeu battery (Battery is replaced) | 000000000000000000000000001O##: 000000000000000000000000001000000003335900000000 Problem Description by Engineer : Problem with charging, Diagnostic performed by Engineer : Old battery Resoluion : Euchanged battery ----------------------------------------------------------------------------- | 2023-11-03 13:09 | No
3303 | Yes | internal blades have a broken connector and their u is not possible | internal blades have a broken conneur and their u is not possible | 000000000000000000000000001O##: 0000000000000000000000000010000000032800000000 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator monitor Eupected and ueupected Resuts: Internal paddles are damaged. uer Impact: Impossible to ue internal paddles. Patient Impact: The equpment is ued with limitations. SE Follow u Requred: We contacted the cutomer (Emma eut: 2376), and he told u that the internal paddles have a broken connector and it is not possible to ue them. Material order is handled and the case is transferred onsite so that the field service assigned to the center performs the euchange. - Equpment located in H-E10000000000000000-0000000000000000 ROOM E. CARDIOVASCuAR ICu - Assign FSE Lenin - Order the mentioned material Resoluion: Material order is handled. Diagnostic performed by Engineer: We contacted the cutomer (Emma eut: 2376), and he told u that the internal paddles have a broken connector and it is not possible to ue them. Material order is handled and the case is transferred onsite so that the field service assigned to the center performs the euchange. ----------------------------------------------------------------------------- | 2023-11-03 13:09 | No
3304 | Yes | Inability to start the system | Inability to start the system | 000000000000000000000000001O##: 00000000000000000000000000100000000327626 00000000-Nov-2000000003 Eleuheriadis, Kostas The system meets the specification for the performed service and is retuned to ue Internal Comments: 0000000080000000050000000000000000 (u1) Host PC Monoplane Revised II no HDD, 0000000080000000050000000000000000(u1) Hard disk spare part, D/A: u187, No: 0000000065, Date: 3/11/2000000003 Problem Description by Engineer: Problem description by engineer : Duing morning start-u Inability to start the system in Application. uable to investigate the failue throuh PRS. Order of spare parts Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Resoluion: Inability to start an angiography device and fuction interrutions duing euams from time to time. The system remains in statu: "System Starting, eupected time needed: 00000000:00000000", withou futher progress or displays a blu screen with message: "STOP: 000000000000000000000000000000007A (00000000E2C92C5C,00000000C0000000000000000185,00000000BF8F1EA1,000000003C3E78600000000. Symptoms that occu qute often and ueupectedly, puting the machine ou of order. Following investigation of the failue, it was foud that the uLLis du to Host PC compuer malfuction and the hard drive of the said device was also replaced. Du to incompatibility it was necessary to load the software from scratch in consecuive series of versions SwR8.1.15, SwR8.1.16, SwR8.1.17 and SwR000 Repair 000000000000000000000000001ks were performed in accordance with the instrutions Du00000000638 ReV B of u. The machine is fuctional, and it was made for clinical ue. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000326196 Spiliopouos, Antonis The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: uLL# and/or description of uLL: Duing morning start-u Inability to start the system in Application. Inability to investigate the failue throuh PRS. Spare parts were ordered Repair Action : Duing morning start-u Inability to start the system in Application. Inability to investigate the failue throuh PRS Spare parts were ordered. Resoluion: Duing morning start Inability to start the system in Application. Inability to investigate the failue throuh PRS Spare parts were ordered. Follow u Requred: uLL# and/or description of uLL: Duing morning start-u Inability to start the system in Application. Inability to investigate the failue throuh PRS. Order of spare parts Repair Action : Duing morning start-u Inability to start the system in Application. Inability to investigate the failue throuh PRS. Order of spare parts Technician Name : Spiliopouos, Antonis Problem Description by Engineer: Problem description by engineer : Duing morning start-u Inability to start the system in Application. uable to investigate the failue throuh PRS. Order of spare parts Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: Duing morning start-u Inability to start the system in Application. Inability to investigate the failue throuh PRS Spare parts were ordered. ----------------------------------------------------------------------------- | 2023-11-03 13:19 | No
3305 | Yes | System remains at a stand still shortly before completion of the boot process | System remains at a stand still shortly before completion of the boot process | 000000000000000000000000001O##: 00000000000000000000000000100000000333536 Hellenbrandt, Helmu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: System does not start completely, gets stuk at "System start in 0000000000000000". FleuVision remains dark. FleuVision PC coud be started manully Problem Description by Engineer: Problem description by engineer : RR Zafer Özen Contact Phone +00000000319222000000007 The faut report is confirmed. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Internal Comments: uLL# and/or description of uLL: FleuVisionControlService:Observer000000000000000000000000001ker PING(FLEuVISION_PC_CONTROL) is not suceeded after 100000000 seconds. Please check if FLEuVISION_PC_CONTROL is powered on and all (net000000000000000000000000001k) cables are correctly connected. Repair Action: Euchange FleuVision PC (Repair Action: Euchange FleuVision PC Follow u Requred: uLL# and/or description of uLL: FleuVisionControlService:Observer000000000000000000000000001ker PING(FLEuVISION_PC_CONTROL) is not suceeded after 100000000 seconds. Please check if FLEuVISION_PC_CONTROL is powered on and all (net000000000000000000000000001k) cables are correctly connected. Repair Action: Euchange FleuVision PC (Repair Action: Euchange FleuVision PC Technician Name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000333632 Kousouakis, Vassilios Diagnostic performed by Engineer: N/A Problem Description by Engineer: Problem description by engineer : RR Zafer Özen Contact Phone +00000000319222000000007 The faut report is confirmed. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 ----------------------------------------------------------------------------- Internal Remarks: uu 0000000080000000012200000000 - u Haswell Fleuvision 2 PC no HDD Cutomer ETA: 00000000-NOV-2000000003 Internal Remarks: Abou u Kousouakis He sends Fu swap euisting ET to the new 000000000000000000000000001 new ET is suject to TPR u Wilken knows. Cutomer comment: ### Appointment with u Özen voted for December 12th from 9 a.m. ### Cutomer comment: Rings throuh Cutomer comment: With Sw. Daniela spoke, u Özen is curently standing sterile at the table RR in 300000000minues at the earliest Cutomer comment: u Özen inqures abou an appointment, December 12th. was confirmed ----------------------------------------------------------------------------- | 2023-11-03 13:24 | No
3306 | Yes | Hot Case A:0uDFM0 Defibrillator/Monitor/No shock delivery | Hot Case A:0/Effi0 DFM0 Defibrillator/Monitor/No shock delivery | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022530000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No Problem Reported by Cutomer: Luhe Central Hospital, the discharge machine did not respond for 00000000onsecuive times, and the patient did not respond. Agent Ms. Wang 177968617900000000called to report repair Trouleshooting Action: The DFM100000000 discharged 00000000imes in a row withou any discharge auio reminder, and the patient did not respond, bu the device’s electric shock record was complete, and there were no failed items in the operation inspection.It is recommended to condut electric shock testing and testing equpment ----------------------------------------------------------------------------- | 2023-11-03 13:39 | No
3307 | Yes | Fault: the system is down. | Fault: the system is down. | 000000000000000000000000001O##: 0000000000000000000000000010000000033380000000000000000-Nov-2000000003 FABIO SPERANDIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: THE SYSTEM IS DOWN. IT DOES NOT DISPLAY ANYTHING ON THE LARGE DISPLAY. When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: COuTDOWN 00000000.00000000 STOPPED uer Impact: NO Patient Impact: NO SE 000Resoluion: Performed replacement of the fleuvision PC and reloaded software on it and necessary configuations. The system was left 000000000000000000000000001king. Parts Consued: 1.00000000000000000000000000000000800000000119123, FleuVision 2 PC No HDD, COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000333663 ANDREA COLOMBO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: THE SYSTEM IS DOWN. IT DOES NOT DISPLAY ANYTHING ON THE LARGE DISPLAY. When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: COuTDOWN 00000000.00000000 STOPPED uer Impact: NO Patient Impact: NO SE 000Diagnostic performed by Engineer: reported, system down, coutdown 00000000.00000000 on data monitor, the large screen in euam room does not show video. Intervention will follow, Internal Comments: uLL# and/or description of uLL Coutdown 00000000.00000000 on data monitor, the large screen in euam room does not show video. Remotely reachable system (RSN, bu the port for ftp is down). Repair Action: Following discusion with Fabio Sperandio, organize onsite visit with parts (fleuvision PC & 1hdd) 1 technician 1 day Follow u Requred: uLL# and/or description of uLL Coutdown 00000000.00000000 on data monitor, the large screen in euam room does not show video. Remotely reachable system (RSN, bu the port for ftp is down). Repair Action: Following discusion with Fabio Sperandio, organize onsite visit with parts (fleuvision PC & 1hdd) 1 technician 1 day What are the neut Trouleshooting steps: Following discusion with Fabio Sperandio, organize onsite visit with parts (fleuvision PC & 1hdd) 1 technician 1 day Technician Name: ANDREA COLOMBO ----------------------------------------------------------------------------- | 2023-11-03 14:19 | No
3308 | Yes | Live monitor is not functioning | Live monitor is not funuoning | 000000000000000000000000001O##: 0000000000000000000000000010000000033300000000 Kuar, Vipin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Checked and foud 19" monitor in MCS is not 000000000000000000000000001king. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer 0000000000Diagnostic performed by Engineer : Checked and foud monitor is not getting On. Internal Comments : uLL# and/or description of uLL: No uLL Repair Action : No uLLFollow u Requred : uLL# and/or description of uLL: No uLL Repair Action : No uLL Technician Name : Kuar, Vipin ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000333900000000 Ramachandran, Sekar Problem Description by Engineer : Problem description by engineer : Checked and foud 19" monitor in MCS is not 000000000000000000000000001king. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer 0000000000Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-11-03 14:24 | No
3309 | Yes | 0 failure | 0V failure |  | 2023-11-03 15:10 | No
3310 | Yes | Machine is not switching On | Machine is not switching On |  | 2023-11-03 15:34 | No
3311 | Yes | Euposure problem | Euposure problem | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 00000000.11.2000000003 00000000:57:00000000 uC (8600000000600000000) only on job training Problem Reported by Cutomer: uable to eupose Internal Remarks: 00000000.11.2000000003 17:00000000:500000000uC (8600000000000000001) RDI Accuacy: Yes. $Q7:(Confirm reported problem, System can't euposue. ) $Q8: (Performed visul inspection, System can't euposue. Onsite check, the Velara generator uLLcode was " 00000000HW", it means kV asymmetrical, Read HW Valus, kV anode : 51.76 kV, kV cathode : 3000000002 kV.kV cathode lower than kV anode. 1. Euchange Q1 and Q2, test euposue, it was kV asymmetrical, and the kV cathode lower than kV anode. 2. Open the tue cover, foud that the tue houing was leaking oil.So I replaced the tue, test euposue, it was kV asymmetrical, and the kV cathode lower than kV anode. 3. Euchange the H.V.cable cathode and anode, test euposue, it was kV asymmetrical, and the kV cathode lower than kV anode. 00000000Euchange the kVmA control board, test euposue, it was kV asymmetrical, and the kV cathode lower than kV anode. 5. take HV tank test according to the manul, HV tank defect.) $Q9: (Replace HV tank, and perform relevant calibration and adjutment. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name [PA] Neut Step Action: PPD00000000210000000000000000*1 +00000000217800000000100000000*1 Arrange ugent order if ou of stock FSE Yang Zhiuiong-86-13560000000020000000057 T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the euposue was intermittently uavailable, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: 00000000HS Remote diagnosis and test resuts: FSE has adjuted Q before PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the Q is normal. 2. Check whether the kV mA board is normal. 3. Check whether the Tank is normal. 00000000Check the machine uLLand perform faut repairs. Recommended spare part uuand name: 00000000217800000000100000000 uit dig. kV mA control&00000000210000000000000000 HE Velara 1T Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 17:00000000:00000000uC (8600000000000000001) [Service content] The euposue is interruted, and the test finds that it is a high voltage problem. Replace the high voltage and perform related calibrations. 【Service Resut】The machine operates normally after repair. 【Other matters】None yet SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 00000000.11.2000000003 00000000:57:00000000 uC (8600000000600000000) only on job training Safety Qustions: Case: 000000002253000000007 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: FSE Yang Zhiuiong 13560000000020000000057 Call for repair +ojt chenhaoyuNeut Step Action: The retun visit has failed. An ##has been sent to MM. Hung Chengwei is informed and reported for repair. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-03 15:54 | No
3312 | Yes | C arc no longer moves into the cranial angulated positions | C arc no longer moves into the cranial angulated positions | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000952 23-Nov-2000000003 Kousouakis, Vassilios The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : The reported uLLis confirmed Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Motor Carc defective. Resoluion: Engine replaced. Parts Consued: 1.000000000000000000000000000000002300000000223500000000OTOR C-ARC RED. POLY-G2 CLEA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000100000000Nov-2000000003 Nölkes, Fabian The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : The reported uLLis confirmed Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer: RR FSE Matuina. LOG still shows no uLLmessage. C-Arc engine at the last appointment DEFOA and has not been renewed. MVR High Power Board has been renewed. uLLeuists according to uers continu. C-Arc engine needs to be replaced. Possibly defective Lu boards. Spare parts determined and follow-u order created. Internal Comments: uLL# and/or description of uLL: Root Caue for RMT - POS: BeamPropeller Curent Not Calibrated - Frontal: 1. Frontal channel: Beam propeller curent is not calibrated. Repair Action: Service actions for RMT - POS: BeamPropeller Curent Not Calibrated - Frontal: 1.1. Carry ou beam propeller curent calibration in FSF. Follow u Requred: uLL# and/or description of uLL: Root Caue for RMT - POS: BeamPropeller Curent Not Calibrated - Frontal: 1. Frontal channel: Beam propeller curent is not calibrated. Repair Action: Service actions for RMT - POS: BeamPropeller Curent Not Calibrated - Frontal: 1.1. Carry ou beam propeller curent calibration in FSF. What are the neut trouleshooting steps: HuRY!! Please plan on-site deployment 00000000ous. Renew C-Arc engine. Carry ou adjutments If necessary, replace the Lu board and eutension board (SCP 1; SCP 2) and carry ou all necessary adjutments. Technician Name: Nölkes, Fabian ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 100000000ov-2000000003 Matuina, Ingo The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : RR ue PlettauPhone 0000000000000000000000006955951 The uLLmessage is confirmed. How often is issu occuring? : Intermittent Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Tools & Equpment: MuTIMETER, 2000000003-12-12 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000023000000002300000000,MVR HIGH POWER BOARD Resoluion: The reported faut is confirmed. The following activities were carried ou to correct the uLL -MVR (High Power Board) renewed -All necessary adjutments carried ou The system complies with the manuactuer's specifications in the tests carried ou and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000033000000007 100000000Nov-2000000003 Nölkes, Fabian The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Please reschedue you deployment. Coud not be carried ou du to emergency patient Please skip parts! Problem Description by Engineer: Problem description by engineer : RR ue PlettauPhone 0000000000000000000000006955951 The faut report is confirmed. How often is issu occuring? : Intermittent Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer: Appointment postponed du to emergency patient. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000033000000001 Hellenbrandt, Helmu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : RR ue PlettauPhone 0000000000000000000000006955951 The uLLmessage is confirmed. How often is issu occuring? : Intermittent Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer: The C-ARC movement hangs sporadically and then only rus piecemeal Internal Comments: uLL# and/or description of uLL: soft-collision detected [FREE_TEuT] movement: BeamCArm, override=00000000at position: 15(000000001[cm¦dg]) Time of occurence : 2000000003-11-00000000, 10000000035:21 Repair Action: Check C.Arc movement. Adjut / replace MVR if necessary Follow u Requred: uLL# and/or description of uLL: soft-collision detected [FREE_TEuT] movement: BeamCArm, override=00000000at position: 15(000000001[cm¦dg]) Time of occurence : 2000000003-11-00000000, 10000000035:21 Repair Action: Check C.Arc movement Alignment / Replace MVR if necessary Technician Name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Cutomer comment: ## uPlettauconfirms appointment for Mon. 20000000011. 12 p.m. - 00000000.m. ## Cutomer comment: u Plettaureports a failue, C Bogen no longer rus at all, please ugently follow u +00000000000000006955951 Internal Remarks: After RS ​​with u Nölkes please also order the C-Arc Rotation Motor Assy Clea1/PolyG2 0000000080000000000000000500000000Cutomer comment: ### Appointment with u Plettauvoted for November 100000000 from 9 a.m. ### Cutomer comment: ### Appointment agreed with the uer for November 100000000h from 1 p.m. ### Cutomer comment: u Plettaureports: The same uLLoccus again with the C bow: It no longer moves into the cranially angled positions. Sporadic since yesterday, more frequnt today. The technician was already on site 1-2 days ago. Asks for RR on this. / Tel.: 0000000000000000000000006955951 Internal Remarks: HuRY!! Please plan on-site deployment 00000000ous. Renew C-Arc engine. Carry ou adjutments If necessary, replace the Lu board and eutension board (SCP 1; SCP 2) and carry ou all necessary adjutments. Cutomer comment: u Plettauhas sugested a date: November 200000000h. 9 a.m. - 1 p.m. or 12 p.m. - 00000000.m. He'll clarify things internally and we'll get back to youaroud 00000000.m ----------------------------------------------------------------------------- | 2023-11-03 16:04 | No
3313 | Yes | [Q] Geometric movements unavailable | [Q] Geometric movements unavailable | 000000000000000000000000001O##: 0000000000000000000000000010000000033000000007 100000000Nov-2000000003 Andrew Patterson The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced and tested supect Geo interface bou.System fuctioning correctly Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: Geometric movements uavailable uer Impact: Patient Impact: SE Reguatory Qustions : N/A Parts Consued : 1.00000000000000000000000000000000800000000265772,GEO IO BOu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000033000000005 David Couens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to replace GIB and power suply if requred Diagnostic performed by Engineer : Logging states a faut with the GIB. I will also order the 200000000power suply also Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: Geometric movements uavailable uer Impact: Patient Impact: SE Reguatory Qustions : N/A ----------------------------------------------------------------------------- Cutomer comment : Confirmed planning with Samantha for Friday 9am (10000000011) Cutomer comment : Part on BO, confirmed with cutomer we will be in touh asap to plan visit. RSE confirmed the BO part is most important to fiu faut. Amended to P3 as cutomer confirmed system is still in ue ----------------------------------------------------------------------------- | 2023-11-03 16:19 | No
3314 | Yes | 0 - u000 Low Battery | 00- u0000- Low Battery | 000000000000000000000000001O##: 0000000000000000000000000010000000032900000000 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: nemo withou collection Clinical Laboratory HC S.p.A. c/o San Bortolo di Vicenza Viale Ferdinando Rodolfi, 37 36100000000 Vicenza Technical Bulding Floor 1 Ac. CONCATO TEL. 00000000000000003257 Problem Description by Engineer: Cutomer Fuction/Role: Engineer Clinical Device ue: defibrillator battery test Eupected and ueupected Resuts: low battery uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: battery to be replaced, proceeded with attempt to order. the device is not suported. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000335292 ----------------------------------------------------------------------------- | 2023-11-03 16:39 | No
3315 | Yes | B Tube is failing to eupose | B Tube is failing to eupose | 000000000000000000000000001O##: 0000000000000000000000000010000000033000000008 00000000-Nov-2000000003 David Couens The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: B Tue is failing to eupose uer Impact: NA Patient Impact: NA SE8.1 Resoluion : Replaced both IPPCs and all 6 HDDs. System is 000000000000000000000000001king now and tested by the cutomer Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000033000000006 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to replace IPPC's and disks and configue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: B Tue is failing to eupose uer Impact: NA Patient Impact: NA SE8.1 Diagnostic performed by Engineer : Supect PC issu , A/B IPPC's to be replaced ----------------------------------------------------------------------------- | 2023-11-03 17:04 | No
3316 | Yes | M0A / u00 Faulty battery. | M00A / u000/ Faulty battery. | 000000000000000000000000001O##: 00000000000000000000000000100000000333555 RENZO FORESTIERI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: uon approval of claim and after contractul verification, send part to the address indicated, make nemo withou collection at the same time Tu Make NEMO withou collection at the same time. Tu Osp. S. Maria di Terni 1 Tristano da Joannucio Street 00000000100000000 Terni (TR) c/o Electromedical Laboratory c/o u Danilo Marinelli Cell.300000000810000000000000000 Emanul Pascuci Cell.3356500000000200000000iagnostic performed by Engineer: Clinical Engineer Moira Sensi requsts battery for uuuas it is not 000000000000000000000000001king, after contractul verification, ship part to the address indicated Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Moira Sensi Device ue: uuudefibrillator Eupected and ueupected Resuts: battery is down uer Impact: no Patient Impact: no SE na ----------------------------------------------------------------------------- | 2023-11-03 17:34 | No
3317 | Yes | The tidal volume is too high | The tidal volume is too high |  | 2023-11-03 17:41 | No
3318 | Yes | PSA_A bandbagor FD sterile bag cover was sucked up into the system as the operator was rotating detector to landscape | PSA_A bandbagor FD sterile bag cover was sucked up into the system as the operator was rotating deteur to landscape |  | 2023-11-03 17:47 | No
3319 | Yes | Complete block of machine. | Complete block of machine. | 000000000000000000000000001O##: 00000000000000000000000000100000000335181 00000000-Nov-2000000003 Alessandro, Avitabile The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: Complete block of machine room 7, at the start-u of the machine it appears on the monitor Main description only u logo on black screen. Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: no Patient Impact: no SE 1.2.5.00000000Parts Consued: 1.00000000000000000000000000000000800000000717981, Sute PC W7 + W100000000Resoluion: Foud fauty Sute-PC. Replaced Sute PC, reinstalled OS + application and restored backu and system license. Fuctional tests were carried ou with a positive oucome. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000333891 Samule, Villa The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: Complete block of machine room 7, at the start-u of the machine it appears on the monitor Main description only u logo on black screen. Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: no Patient Impact: no SE 1.2.5.00000000Diagnostic performed by Engineer: it is necessary to replaced PC and reload of SW Internal Comments: uLL# and/or description of uLL The system is not reachable from PRS. An on-site visit is necessary to continu the diagnosis, 1 FSE 8H with the spare part. Repair Action: An on-site visit is necessary in order to continu the diagnosis 1 FSE 8H with part Follow u Requred: uLL# and/or description of uLL The system is not reachable from PRS. An on-site visit is necessary to continu the diagnosis, 1 FSE 8H with the spare part. Repair Action: An on-site visit is necessary in order to continu the diagnosis 1 FSE 8H with part What are the neut Trouleshooting steps: An on-site visit is necessary in order to continu diagnosis 1 FSE 8H with part 1u 00000000800000000717981 - Sute PC W7 + W100000000 Technician Name: Samule, Villa ----------------------------------------------------------------------------- | 2023-11-03 17:54 | No
3320 | Yes | Cuomer Dissatisfaction_Power management board for new bad, FCO after the machine alarm | Cuomer Dissatisfauon_Power management board for new bad, FCO after the machine alarm |  | 2023-11-03 18:07 | No
3321 | Yes | Failed Battery Insertion Test | Failed Battery Insertion Test |  | 2023-11-03 18:09 | No
3322 | Yes | M0A / u00/ The startup alarm could not be eliminated during self-test (QA#000000. | M00A / u000/ The startup alarm could not be eliminated during self-test (QA#00000000). |  | 2023-11-03 18:43 | No
3323 | Yes | It does not pass the tests. | It does not pass the tests. | 000000000000000000000000001O##: 00000000000000000000000000100000000329565 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: dealer Device ue: defibrillator test Eupected and ueupected Resuts: test load in order to carry ou missing test uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: It is necessary to order test load in order to carry ou operation test, it is missing. The device is not suported. We proceed with the attempt to order the part. Follow u Requred: nemo withou collection Barone Romeo Hospital 100000000azzini Street PATTI (ME) At pharmacy AC. CARMELO VERDI 300000000798600000000 ANTONIO TRuLIO 3286512000000007 u Stefano Chiarandà +39300000000566855 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000335893 Parts Consued: 1.000000000000000000000000989800000000100000000831,HeartStart 500000000ohm Test Load 1.000000000000000000000000989800000000100000000981, HeartStart Hands-free Cable ----------------------------------------------------------------------------- Internal Remarks: Attached is pending FDL. Requst for test load cable. Internal Remarks: Hi, Requst to open a call for: Equpment: D-600000000 DEFIBRILLATOR uY: M3535 A HEARTSTARST uuu Manuactuer: u MEDICAL SYSTEMS 000S00000000315600000000 Installed at: ASP5 ME Site: D000000001, PIEMONTE Medical Department: D000000001-3, EMERGENCY Problem: IT DOES NOT PASS THE TESTS. Notes: IT IS NOT POSSIBLE TO CARRY Ou THE FuCTIONAL TEST Du TO LACK OF LOAD TEST Info Smau If the system is not inserted, ue conf. fictitiou: 1000000007171200000000 Reference: Stefano Chiarandà ----------------------------------------------------------------------------- | 2023-11-03 19:39 | No
3324 | Yes | PSA & PSE_Security Issue_The u0ngiography system repeatedly shows overheating of the Ru tube | PSA & PSE_Security Issue_The u0angiography system repeatedly shows overheating of the Ru tube |  | 2023-11-03 19:39 | No
3325 | Yes | Audio &Sounds: Defective speaker - no sound: Speaker/HW: Replaced (00 | Audio &Sounds: Defeu speaker - no sound: Speaker/HW: Replaced (0000 | 000000000000000000000000001O##: 0000000000000000000000000010000000033000000006 Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: parts only uer Impact: uknown Patient Impact: uknown SE uknown Follow u Requred: Biomedical department Resoluion: Parts ordered. Diagnostic performed by Engineer: Ordered parts: A front panel, a screen and a louspeaker. No more soud, no uLLmessage. No patient incident. ----------------------------------------------------------------------------- | 2023-11-03 19:44 | No
3326 | Yes | Monitor is not switching On (Replaced DFM0 ASSY Processor) | Monitor is not switching On (Replaced DFM0 ASSY Processor) | 000000000000000000000000001O##: 0000000000000000000000000010000000033000000009 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : 0000000056000000009000000001 DFM100000000 ASSY Processor ordered to resolve the problem Follow u Requred : 0000000056000000009000000001 DFM100000000 ASSY Processor ordered to resolve the problem Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Monitor is not switching On ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000335985 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor ----------------------------------------------------------------------------- | 2023-11-03 19:54 | No
3327 | Yes | 00he defibrillator one of the paddles does not work. | 000The defibrillator one of the paddles does not work. | 000000000000000000000000001O##: 00000000000000000000000000100000000330000000000000000Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Impossible to load with the paddles. With the buton located on the device, it 000000000000000000000000001ks. The cutomer will do a cross-test with another set of paddles. Sent an e-mail to the cutomer for follow-u. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the charge of the defibrillator. ----------------------------------------------------------------------------- | 2023-11-03 20:04 | No
3328 | Yes | Duplicate of PR00 | Duplicate of PR008 |  | 2023-11-03 20:04 | No
3329 | Yes | 00larm message: Low Leak – CO0ebreathing Risk | 00 Alarm message: Low Leak – CO0Rebreathing Risk |  | 2023-11-03 20:08 | No
3330 | Yes | Total failure, no work possible | Total failure, no work possible | 000000000000000000000000001O##: 00000000000000000000000000100000000336358 November 6, 2000000003 Dögnitz, Matthias The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Called Sw Andrea at +00000000000000001300000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion: Host PC replaced, software loaded, configued, fuction check The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000335826 Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: Called Sw Andrea at +00000000000000001300000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Follow-u requred: uLL# and/or description of uLL The system failed duing the euamination with a bang. And curently remote not reachable. u Schleef checked the data monitor in the control room. The monitor has voltage - the menuis selectable, bu it has no image. Conversion to internal 5V also did not help. Host PC restarted manully - no Bios visible on the data monitor The euamination room is curently not yet accessible Futher trouleshooting on site requred Repair Action: Trouleshooting Technician name: Hellenbrandt, Helmu Internal Comments: uLL# and/or description of uLL The system failed duing the euamination with a bang. And curently remote not reachable. u Schleef checked the data monitor in the control room. The monitor has voltage - the menuis selectable, bu it has no image. Conversion to internal 5V also did not help. Host PC restarted manully - no Bios visible on the data monitor The euamination room is curently not yet accessible Futher trouleshooting on site requred Repair Action: Trouleshooting Diagnostic performed by Engineer: The system failed duing the euamination with a bang. The OP is curently being continud with a mobile C-arm. No problem for the patient. ----------------------------------------------------------------------------- Internal Remarks: Please call uer Ms. Arnold at +0000000000000000135389 Cutomer comment: uder +0000000000000000135389 rings throuh. Cutomer comment: Still did not reach anyone +00000000000000001300000000 +0000000000000000135389 #. ----------------------------------------------------------------------------- | 2023-11-03 20:39 | No
3331 | Yes | Fluoroscopy not possible - not able to set screen - radiation disabled | Fluoroscopy not possible - not able to set screen - radiation disabled | Cutomer comment : Spoke to Tracy from GE and she has said she will contact the cutomer and ask if this job is still needed. Cutomer comment : Called and offered GE dates for visit ----------------------------------------------------------------------------- | 2023-11-03 20:49 | No
3332 | Yes | Alarms or Messages/System in not alerting/ur Error | Alarms or Messages/System in not alerting/ur Error | 000000000000000000000000001O##: 00000000000000000000000000100000000336331 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Clinical Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: tele tech Asley reports there are no alarms; auio or visul at the PIIC, issu started yesterday, also when the patient hits the nuse call on the Mu00000000the alert is not popuating to the PIIC Diagnostic performed by Engineer : Problem Description: tele tech Asley reports there are no alarms; auio or visul at the PIIC, issu started yesterday, also when the patient hits the nuse call on the Mu00000000the alert is not popuating to the PIIC Nuse call buton uMC Dalton 00000000 000000005565Issu started yesterday, no auioHalls Research u7 utra Receiver Form ued: Information Requst Trouleshooting Steps Taken: Spoke with IT rep Dalton bu he is not responsible for the u equpmentAdvised Dalton to speak with the Nuse Suervisor, we need to engage biomedI will call Dalton back in 1 hr for an udate Resoluion : follow-u call Dalton advised the issu was a loose cord connected to the Halls Research u7 The maintenance person re-seated the cable, system is now 000000000000000000000000001king as eupected Issu resolved The system is now 000000000000000000000000001king as eupected ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ashley Riley IP & Issu: System in not alerting/69826000000003 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create Case and R000000000000000000000000001and roued to RSE Internal Remarks : Interaction Type: Cutomer Call Back Caller: Ashley Riley IP & Issu: 69826000000003 System in not alerting Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Cutomer Call Back to report Issu has been resolved. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Follow-u call to Dalton in IT He advised they do have biomed bu he has not reached them yet He euplained they connected a different Mu00000000and the alert did go to the PIIC Advised we still need biomed to validate the issu of no alert souds at the PIIC Offered Onsite service as an option if they can not locate the biomed Dalton will have the biomed call u directly ----------------------------------------------------------------------------- | 2023-11-03 20:59 | No
3333 | Yes | QACASE for PCI imported product#000000During u, the heater heating is not good, and the temperature | QACASE for P0imported produu00000000_During u, the heater heating is not good, and the temperature |  | 2023-11-03 21:03 | No
3334 | Yes | Power failure, GFE00EC0*** | Power failure, GFE0 0DEC00** |  | 2023-11-03 21:11 | No
3335 | Yes | turns on when it's supposed to be off and plugged in | turns on when it's supposed to be off and plugged in | 000000000000000000000000001O##: 00000000000000000000000000100000000336313 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Annul PM service / FCO86600000000000000006 remediation. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in patient ue Eupected and ueupected Resuts: uit powers on by itself / needs PM / FCO86600000000000000006 uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : Batter mfg date is 20000000000000000 Advised to replace the battery per 5 year PM interval and to potentially address the power on issu. Advised they will recive the bench form to send the uit in for PM Service and FCO86600000000000000006 remediation ( PM pcba replacement) Advised to make sue the battery is removed before shipping. ----------------------------------------------------------------------------- | 2023-11-03 21:19 | No
3336 | Yes | The equipment is leaking oil (Split) | The equipment is leaking oil (Split) | 000000000000000000000000001O##: 000000000000000000000000001000000000000000013 November 200000000 2000000003 Vieyra, Jose The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Replace parts. Problem Description by Engineer: Cutomer Fuction/Role: Technician Device ue: yes Eupected and ueupected Resuts: No oil is coming ou uer Impact: yes Patient Impact: no SE 8.1 Diagnostic performed by Engineer: Traces of oil coming from the tue are foud, and geometry control does not 000000000000000000000000001k correctly. Tools & Equpment: MuTIMETER, 200000000000000006-13 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-11-03 21:34 | No
3337 | Yes | Touchscreen does not work. | Touchscreen does not work. | 000000000000000000000000001O##: 00000000000000000000000000100000000336500000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. The cutomer has advised closing the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000is failing the touhscreen calibration. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 11/00000000/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Provided Cutomer the part # for the touhscreen. Part # 00000000561539000000001 RP-Touh Screen, V600000000List price is $9900000000Page 181 for referencing the repair. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000pass touhscreen calibration uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000touhscreen is failing calibration. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Bob Sulivan IP & Issu: 60000000000000000988 - The touhscreen does not 000000000000000000000000001k. Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-03 21:39 | No
3338 | Yes | Unit will not charge | Unit will not charge | 000000000000000000000000001O##: 00000000000000000000000000100000000336793 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged and to correct the issu the Front Case Assembly need s to be replaced. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000336673 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Follow u Requred : Problem Description: biomed reports Tempu Pro will not charge, tried different cables, same uLL sending to bench, loaner needed Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address: 2125 Point Blvd, Sute 200000000, Elgin, IL 60000000023##d bench repair form Diagnostic performed by Engineer : Problem Description: biomed reports Tempu Pro will not charge, tried different cables, same uLL sending to bench, loaner needed Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address: 2125 Point Blvd, Sute 200000000, Elgin, IL 60000000023##d bench repair form Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE 00000000.28 Problem Description/ueupected Resuts: biomed reports Tempu Pro will not charge, tried different cables, same uLL sending to bench, loaner needed ----------------------------------------------------------------------------- Internal Remarks : Please ue the address on the bench form for shipping puposes. ----------------------------------------------------------------------------- | 2023-11-03 21:39 | No
3339 | Yes | Machine shutdown automatically during case | Machine shutdown automatically during case | 000000000000000000000000001O##: 00000000000000000000000000100000000336777 00000000-Nov-2000000003 Guaraj, Vijayan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine shudown auomatically duing case When was the first occurence? Within... : 00000000-Nov-2000000003 21:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Machine shudown auomatically duing case uer Impact: No uer impact Patient Impact: No patient impact SE 1.2.5.00000000Parts Consued : 1.00000000000000000000000000000000800000000560000000051,Service part Review modue Resoluion : Foud that Review Modue Off buton was stuk inside and the system switches off auomatically du to malfuction of the buton so ordered and replaced with new modue checked the system fuctionality all 000000000000000000000000001king good handed over for clinical ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003338300000000Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Machine shudown auomatically duing case When was the first occurence? Within... : 00000000-Nov-2000000003 21:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Machine shudown auomatically duing case uer Impact: No uer impact Patient Impact: No patient impact SE 1.2.5.00000000Diagnostic performed by Engineer : Logs indicate that normal shudown sequnce happened. FSE to ensue that no uer has pressed off buton accidentally or uknowingly. If not happened - then supected off key fuctionality on review modue. Review modue needs replacement. Follow u Requred : uLL# and/or description of uLL: Machine shudown auomatically duing case Repair Action : Logs indicate that normal shudown sequnce happened. FSE to ensue that no uer has pressed off buton accidentally or uknowingly. If not happened - then supected off key fuctionality on review modue. Review modue needs replacement. Technician Name : Vijayakuar, K.S Resoluion : Logs indicate that normal shudown sequnce happened. FSE to ensue that no uer has pressed off buton accidentally or uknowingly. If not happened - then supected off key fuctionality on review modue. Review modue needs replacement. Internal Comments : uLL# and/or description of uLL: Machine shudown auomatically duing case Repair Action : Logs indicate that normal shudown sequnce happened. FSE to ensue that no uer has pressed off buton accidentally or uknowingly. If not happened - then supected off key fuctionality on review modue. Review modue needs replacement. ----------------------------------------------------------------------------- | 2023-11-03 21:44 | No
3340 |  | The equipment does not start - SPLIT | The equipment does not start - SPLIT |  | 2023-11-03 21:55
3341 | Yes | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0Sent 000 | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0 Sent 0/0/03 |  | 2023-11-03 22:06 | Yes
3342 | Yes | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops.GFE0Sent 000 | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops.GFE0 Sent 0/0/03 |  | 2023-11-03 22:07 | Yes
3343 | Yes | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0Sent 000 | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0 Sent 0/0/03 |  | 2023-11-03 22:08 | Yes
3344 | Yes | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0Sent 000 | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0 Sent 0/0/03 |  | 2023-11-03 22:09 | Yes
3345 | Yes | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0Sent 000 | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0 Sent 0/0/03 |  | 2023-11-03 22:10 | Yes
3346 | Yes | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0sent 00 | SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0 sent 0/00 |  | 2023-11-03 22:11 | Yes
3347 | Yes | Repair broken pin of DC Power connector | Repair broken pin of DC Power conneur | 000000000000000000000000001O##: 000000000000000000000000001000000003000000002 Robert Jon de Ste Croiu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Engineer Device ue: not in ue Eupected and ueupected Resuts: Cutomer reporting the DC connection broken on the uit. uer Impact: None Patient Impact: None SEN\A Diagnostic performed by Engineer : No 000000000000000000000000001ks carried ou and parts cancelled/retuned. Reguatory Qustions : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000330000000072 Simon Goldberger The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Engineer Device ue: not in ue Eupected and ueupected Resuts: Cutomer reporting the DC connection broken on the uit. uer Impact: None Patient Impact: None SEN\A Diagnostic performed by Engineer : sending cutomer a qute for a trip to the bench to replace the rear case 000000005600000000000000001 Reguatory Qustions : N/A Follow u Requred : 000000005600000000000000001 ----------------------------------------------------------------------------- Internal Remarks : Hi Joseph , I hope youare well , Can youconfirm that youare arranging shipment of the Tempu device to ou bench in Bromsgrove please Please note, Huo Bench address: Same Buiness Park different bulding Huo Technology Plot 9 Cartwright Road Sauon Buiness Park Stoke Prior B60000000000000000 Internal Remarks : This case was logged against the incorrect 000- no faut with this device and ndver been received at u Bench. u Bench to reject ordred part as uued and then case will be closed. Correct 000logged agisnt case ref 0000000022710000000088. Internal Remarks : Good Afternoon Z277227 / 0000000000000000 Tempu Pro PHC 0000000000000000000 We will shoud already be having this device sent to you I have noted the change of address for the Huo bench bu some things may already be in flight. Joseph Bracey Suply Chain Coordination - Medical Equpment TVS Suply Chain Soluions Ltd Bristol ¦ u M +00000000(000000007867 25000000002 W tvsscs.com Internal Remarks : Hi Joseph , I hope youare well , Can youconfirm that youare arranging shipment of the Tempu device to ou bench in Bromsgrove please Please note, Huo Bench address: Same Buiness Park different bulding Huo Technology Plot 9 Cartwright Road Sauon Buiness Park Stoke Prior B60000000000000000 ----------------------------------------------------------------------------- | 2023-11-03 22:44 | No
3348 | Yes | 00Y0carm stuck in RAO position and will not move reboot no help | 00CDY0carm stuck in RAO position and will not move reboot no help | 000000000000000000000000001O##: 0000000000000000000000000010000000033700000000 00000000Nov-2000000003 Williams, Grey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : carm stuk in RAO position and will not move reboot no help RMT - u: Geometry modue orientation switch problem. 00000000-Nov-2000000003 12:22:59 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: carm stuk in RAO position and will not move reboot no help uer Impact: room down Patient Impact: room down SE 7.2.16.00000000Resoluion : I foud a lot of contrast on the clea railing that made the c arm stuk in one position. I cleaned the contrast off and now the c arm will move. I performed fuctional checks and gave the machine back to the cutomer. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000337285 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : carm stuk in RAO position and will not move reboot no help RMT - u: Geometry modue orientation switch problem. 00000000-Nov-2000000003 12:22:59 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: carm stuk in RAO position and will not move reboot no help uer Impact: room down Patient Impact: room down SE 7.2.16.00000000Diagnostic performed by Engineer : got part from TSO for trouleshooting log show issu with orientation switch need on site to TS motor and connections Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - u: Geometry modue orientation switch problem.: 1. The table side orientation switch valu of one of the connected Control Modues is invalid. ufortuately the logfile does not contain information which modue is affected. Repair Action : Service actions for RMT - u: Geometry modue orientation switch problem.: 1.1. Please (let the cutomer) check whether the table side orientation switch is in a valid position on all Control Modues. (Meanwhile, Doctor Side orientation is ued for the affected Control Modue.) Ru fuctional checks on TSO and check rotation motor and connections If needed GEO MODuE WP Kit 00000000800000000151715 Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - u: Geometry modue orientation switch problem.: 1. The table side orientation switch valu of one of the connected Control Modues is invalid. ufortuately the logfile does not contain information which modue is affected. Repair Action : Service actions for RMT - u: Geometry modue orientation switch problem.: 1.1. Please (let the cutomer) check whether the table side orientation switch is in a valid position on all Control Modues. (Meanwhile, Doctor Side orientation is ued for the affected Control Modue.) Ru fuctional checks on TSO and check rotation motor and connections If needed GEO MODuE WP Kit 00000000800000000151715 Technician Name : Clark, Byron ----------------------------------------------------------------------------- | 2023-11-03 23:19 | No
3349 | Yes | Audio-No Sound / Speaker will not Work-Foulink s/n u0/ Other-Other-Tampered / Other | Audio-No Sound / Speaker will not Work-Foulink s/n u000/ Other-Other-Tampered / Other | 000000000000000000000000001O##: 00000000000000000000000000100000000337688 Al Hammad Lamiae The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Problem Description: Biomed requst mu00000000bench repair- no soud Biomed requst mu00000000bench repair- no soud Form ued: Bench Trouleshooting Steps Taken: cutomer states no soud at all Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: Biomed requst mu00000000bench repair- no soud Diagnostic performed by Engineer : Problem Description: Biomed requst mu00000000bench repair- no soud Biomed requst mu00000000bench repair- no soud Form ued: Bench Trouleshooting Steps Taken: cutomer states no soud at all ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Nicolae Laputea IP & Issu: "72928285 The speaker will not 000000000000000000000000001k." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-04 00:19 | No
3350 | Yes | AED Warranty Replacement | AED Warranty Replacement | 000000000000000000000000001O##: 00000000000000000000000000100000000337500000000Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Device was not in ue Eupected and ueupected Resuts: Cutomer called and stated that the AED is not 000000000000000000000000001king properly. The AED states instrutions for uing the AED pads on a patient when inserted. The cutomer informed that these are brand when inserted. The pads are uder warranty and have been replaced in case 000000002250000000000000000. When the cutomer received the pads replacement, we spoke and foud that the AED still consistent with immediately stating instrutions for placing pads on the patient. The new set of pads have not been opened and the AED is not operating properly. This is an indication that the AED is the issu. The AED is uder warranty and to be replaced at zero cost to the cutomer Device serial# A19a-10000000038 uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer called and stated that the AED is not 000000000000000000000000001king properly. The AED states instrutions for uing the AED pads on a patient when inserted. The cutomer informed that these are brand when inserted. The pads are uder warranty and have been replaced in case 000000002250000000000000000. When the cutomer received the pads replacement, we spoke and foud that the AED still consistent with immediately stating instrutions for placing pads on the patient. The new set of pads have not been opened and the AED is not operating properly. This is an indication that the AED is the issu. The AED is uder warranty and to be replaced at zero cost to the cutomer Device serial# A19a-10000000038 Internal Comments : Waiting for order to ship, order nuber 000000000000000057568668 Follow u Requred : Ship To San Joaqun Regional Transit Dist 2800000000E Myrtle St STOCKTON, u200000000-000000009 Attention: Richard Foster Resoluion : DELIVERED Monday 11/6/23 at 100000000100000000AM Signed for by: T.DOuLAS I spoke with the cutomer and he confirmed that the AED has been received and while speaking with cutomer we uboued the replacement and inserted new pads and the battery. The AED passed the BIT and is in service ready for ue. Internal Comments : AED retun tracking 785936250000000000000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003378900000000Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-11-04 00:24 | No
3351 | Yes | System shut down on its own. Will not power back up. | System shut down on its own. Will not power back up. | 000000000000000000000000001O##: 00000000000000000000000000100000000383000000009 00000000-Nov-2000000003 Midgett Pau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: Cardiovascuar imaging. System was not in ue at he time of the incident. Eupected and ueupected Resuts: Tech herd a noise in the equpment closet and system shu down. Will not tun on. uer Impact: System is down and cannot do cases. Patient Impact: uknown. No Patient was schedued SE uknown Resoluion : Replaced Image Processor PC . Performed requred performance test. All pass and system is ready for ue. Parts Consued : 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD Tools & Equpment : DOSE-KVP-LIGHT METER,20000000000000000000000000000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000338875 00000000-Nov-2000000003 Midgett Pau The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Re-order part du to part being retuned to shipper becaue of improper docuenting by DHL 0000000080000000050000000092 qty 1 Diagnostic performed by Engineer : Parts shipment delay. Spent most of day trying to figue ou shipping uLLand re-ordering part. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: Cardiovascuar imaging. System was not in ue at he time of the incident. Eupected and ueupected Resuts: Tech herd a noise in the equpment closet and system shu down. Will not tun on. uer Impact: System is down and cannot do cases. Patient Impact: uknown. No Patient was schedued SE uknown ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000337900000000 00000000-Nov-2000000003 Midgett Pau The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Checked system statu. Foud L2 circut breaker tripped in M-Cabinet. Also foud fue F-23 blown. I noticed there was no power to the Image Processor Compuer. I tued off the Image Processor Compuer ,and shu the system down. I reset L2 circut breaker and replaced blown 100000000 fue. I tuned the system on and the system booted u to the application . I checked to verify proper power to the Image Processor Compuer, it was good. I tuned on the Image Processor compuer and the L2 circut breaker tripped and fue F-23 is blown again. The Image Processor Compuer is defective. Part on order. Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 Follow u Requred : Order part and ship to site for Satuday November 00000000. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: Cardiovascuar imaging. System was not in ue at he time of the incident. Eupected and ueupected Resuts: Tech herd a noise in the equpment closet and system shu down. Will not tun on. uer Impact: System is down and cannot do cases. Patient Impact: uknown. No Patient was schedued SE uknown ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000337800000000 George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Check inpu power to the system and for breakers and fues to determine caue of issu. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Debriefed caller. He said they were in the room stocking suplies and not touhing the equpment when it sudenly shu off. They get nothing when trying to tun back on. They have already spoken to Pau, their fse and he is aware. ----------------------------------------------------------------------------- Internal Remarks : Part Id for IP_PC , pc is blowing F23 fue when system is powered on Internal Remarks : Pau - FSE ordered IPPC Nehalem which was retuned. FSE was told to order a Haswell IPPC - since system at R000, not allowed to miu PC types. ----------------------------------------------------------------------------- | 2023-11-04 00:24 | No
3352 | Yes | uuuis emitting the triple chirp and the "Remove and reinsert battery to test" message | uuuis emitting the triple chirp and the "Remove and reinsert battery to test" message | 000000000000000000000000001O##: 00000000000000000000000000100000000337800000000Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been triple chirping uer Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Cutomer reported that AED has been triple chirping Diagnostic performed by Engineer : Cutomer reported that AED has been triple chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 13 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# B100000000-00000000798 ----------------------------------------------------------------------------- | 2023-11-04 00:59 | No
3353 | Yes | High blower temp alarm. | High blower temp alarm. | 000000000000000000000000001O##: 00000000000000000000000000100000000331222 16-Nov-2000000003 Rivas, Marvin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: blower temp high uer Impact: uknown Patient Impact: uknown SE 2.300000000Diagnostic performed by Engineer : 11/16/2000000003 - u- replaced fan, blower assembly, and MC PCBA per service gude instrutions. All T&V resuts passed. Device retuned to ue. Parts Consued : 1.0000000000000000000000000000000056151200000000,RP-Fan, Cooling, V6000000001.00000000000000000000000000000000561512761,RP-Blower Assy,V6000000001.00000000000000000000000000000000561537531,RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V68000000001.00000000000000000000000000000000561512751,RP-Mout, Isolation, Fan, Pgk/00000000V6000000001.00000000000000000000000000000000561511731,RP-Pin, Puh, Fan Mout, Pkg/00000000V600000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 FLOW METER,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Adam Franklin 256-737-2516 IP & Issu: SN 200000000000000002791 / Has blower temp high alarm. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created case and Page FSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-04 01:09 | No
3354 | Yes | Primary alarm failed | Primary alarm failed |  | 2023-11-04 01:13 | No
3355 | Yes | display low prioty alarm and pressure at 0call | display low prioty alarm and pressure at 0call | 000000000000000000000000001O##: 00000000000000000000000000100000000517200000000 00000000Dec-2000000003 Rivas, Marvin The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561512191,RP-Gas Delivery System (GDS), V600000000Tools & Equpment : TORQu WRENCH,200000000000000006-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 FLOW METER,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000006-30000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: ASP CSA FSE Device ue: uknown Eupected and ueupected Resuts: 1100000000 uer Impact: unown Patient Impact: uknown SE 2.300000000Diagnostic performed by Engineer : 12/00000000000000003 - u- replaced GDS per service gude instrutions. All tests resuts passed. Device retuned to ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000051720000000016-Nov-2000000003 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : ASP has a V600000000and it is having a Prouimal pressue sensor auozero failed. He verified there is a code 1100000000. ASP jut replaced a flow sensor assy and a DA pcba, bu uit is still failing. He stated that the pin alignment is correct, and will order a GDS. ASP ordered the GDS and is schedued on 12-5-23 Problem Description by Engineer : Cutomer Fuction/Role: ASP Device ue: not in ue Eupected and ueupected Resuts: prouimal pressue sensor auozero failed uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 21-Nov-2000000003 Rivas, Marvin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: post zero cal failed uer Impact: uknown Patient Impact: uknown SE 2.300000000Diagnostic performed by Engineer : 11/21/2000000003 - u- replaced data acqusition PCBA per service gude instrutions. All T&V resuts passed. Device retuned to ue. Parts Consued : 1.00000000000000000000000000000000561512561,RP-PCBA, Data Acqusition, V600000000Tools & Equpment : TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 FLOW METER,200000000000000006-30000000000000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000331312 16-Nov-2000000003 Rivas, Marvin The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: ASP CSA FSE Marvin Rivas Device ue: diagnostics Eupected and ueupected Resuts: uLLcode 1100000000 uer Impact: none Patient Impact: none SE 2.300000000Diagnostic performed by Engineer : 11/16/2000000003 - u- Diagnostics: puled DRPTA log and foud uLLcode 1100000000. DA PCBA replacement requred. Part placed on order. Follow u 000000000000000000000000001k requred to complete repairs. Device not retuned to ue. Follow u Requred : N/A ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Adam 2567372516 IP & Issu: 566300000000/ display low prioty alarm and pressue at 00000000call Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: created Onsite and asvised in RP Chat Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-11-04 01:29 | No
3356 | Yes | No Audio / SN u00peaker/SPLIT 00 | No Audio / SN u000Speaker/SPLIT 000 |  | 2023-11-04 02:09 | No
3357 | Yes | Audio/No Audio / SN u0/Speaker | Audio/No Audio / SN u00Speaker |  | 2023-11-04 02:13 | No
3358 | Yes | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at Bench-Keyrin s/n u000/ Other HW-Speaker-HW / Replaced Parts |  | 2023-11-04 02:18 | No
3359 | Yes | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 Device-Other-HW / Replacement Sent | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00/ Device-Other-HW / Replacement Sent |  | 2023-11-04 02:21 | No
3360 | Yes | HS0ED is emitting the triple chirp (Removing the AED from service) | HS0AED is emitting the triple chirp (Removing the AED from service) | 000000000000000000000000001O##: 00000000000000000000000000100000000338700000000 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 12 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A11H-000000009700000000 Battery Install Before# 2000000008/12/31 Lot#000000008900000000 Pads Eup Date: 20000000000000000000000007 Pads Lot # Y10000000072100000000----------------------------------------------------------------------------- | 2023-11-04 03:04 | No
3361 | Yes | Not charging | Not charging | 000000000000000000000000001O##: 00000000000000000000000000100000000338773 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged. To correct the issu we will replace the Front Case Assembly. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000395 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged. To correct the issu we will replace the Front Case Assembly. Resoluion : Device was repaired by parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. All PM tests passed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000338696 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE 7.300000000 Problem Description/ueupected Resuts: Fails to charge. Diagnostic performed by Engineer : Problem Description: Fails to charge. Tempu Pro is not charging. There was no patient involved in this case. Form ued: Bench Trouleshooting Steps Taken: Tried several different chargers, all fail with this Tempu Pro. Follow u Requred : Tempu Pro fails to charge. Resoluion : Sent to Bench ----------------------------------------------------------------------------- Internal Remarks : There was no patient involved on this case. Internal Remarks : Please ue address on bench form coud not find on servicemau Internal Remarks : New Case Caller: AJ Rose Issu: 97882955 Not charging Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-04 03:19 | No
3362 | Yes | Battery not operational | Battery not operational | Test Equpment Info : Nul Device Alarm/Alert as intended? : No Problem Observed : Failed for: Ru THE OPERATIONAL CHECK. Ou Of Bou Failue? : No Labor: Labor ID : Not Answered 000000000000000000000000001k Performed : Not Answered Diagonsis : Not Answered | 2023-11-04 12:25 | No
3363 | Yes | will not read capnography | will not read capnography | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits Capnometer connection is physically damaged. to correct the issu we will replace the Capnometer Assembly. Resoluion : Device repaired with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000339000000001 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits Capnometer connection is physically damaged. to correct the issu we will replace the Capnometer Assembly. Follow u Requred : 000000005600000000000000001-Capnometer 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000338965 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE 7.300000000Problem Description/ueupected Resuts: Tempu Pro fails to measue capnography. There was no patient involved in this case. Resoluion : Sent to Bench Diagnostic performed by Engineer : Problem Description: Tempu Pro fails to measue capnography. There was no patient involved in this case. Tempu Pro fails to measue capnography. There was no patient involved in this case. Form ued: Bench Trouleshooting Steps Taken: NA Follow u Requred : uable to measue capnography. ----------------------------------------------------------------------------- Internal Remarks : There was no patient involved in this case. Internal Remarks : Interaction Type: New Case Caller: Lee Godbold (817) 00000000-000000003 IP & Issu: 87397355 will not read capnography Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Created R000000000000000000000000001 roued to Technical Suport Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-11-04 12:51 | No
3364 | Yes | AED is emitting the triple chirp and gave the message "Remove and reinsert battery to test" | AED is emitting the triple chirp and gave the message "Remove and reinsert battery to test" | 000000000000000000000000001O##: 00000000000000000000000000100000000338985 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that they install battery 12/2000000001 and AED has been chirping uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that they install battery 12/2000000001 and AED has been chirping Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that they install battery 12/2000000001 and AED has been chirping. Cutomer called distribuor cutomer was told to call for u to trouleshoot the AED. Cutomer had remove battery , Instruted the Cutomer to insert battery, AED went straight into test mode uLLmessage Remove and reinsert battery to test Not ready to ue. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 11 years old, the issu is with the device. Instruting the cutomer they will need to remove from service and not ue the device on anyone AED Serial# A12E-00000000100000000 Battery Install Before# 9/3000000002000000007 Lot#00000000692P ----------------------------------------------------------------------------- | 2023-11-04 12:52 | No
3365 | Yes | Wired Foot pedal not responding | Wired Foot pedal not responding | 000000000000000000000000001O##: 00000000000000000000000000100000000357300000000 00000000-Nov-2000000003 Gerlach, Geoffrey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : When foot pedal pressed system says "uuuis not possible" When was the first occurence? Within... : 00000000-Nov-2000000003 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: uing the one form the control room. Eupected and ueupected Resuts: The FS shoud u- ray. uer Impact: Cant u-ray Patient Impact: No futher information is available as per the cutomer uu00000000Tools & Equpment : MuTIMETER,200000000000000001-23 00000000:00000000:00000000 Resoluion : Adjuted timing of the foot-switch contacts so both activating at the same time (for each pedal). Tested system eutensively and retuned to ful clinical ue. Test and verification was performed on the equpment and docuented in the Test and Verification Report. Device is retuned to ful fuctionality. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000332000000002 Lokeh Vinay The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: ***Log File Evidence*** 600000000Warning 2000000003-11-00000000 16:0000000037.300000000EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 728 Warning 2000000003-11-00000000 16:0000000039.787 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 788 Warning 2000000003-11-00000000 16:0000000000000000893 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 800000000Warning 2000000003-11-00000000 16:0000000000000000773 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 900000000 Warning 2000000003-11-00000000 16:0000000050000000000000000 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 953 Warning 2000000003-11-00000000 16:0000000000000000.922 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 2100000000Warning 2000000003-11-00000000 16:0000000027.831 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 2281 Warning 2000000003-11-00000000 16:0000000058.00000000 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 15000000007 Warning 2000000003-11-00000000 10000000039:16.362 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 21300000000Warning 2000000003-11-00000000 11:00000000:100000000756 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 0000000019 Warning 2000000003-11-00000000 12:23:17.00000000 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 0000000039 Warning 2000000003-11-00000000 12:38:35.782 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent Repair Action : FSE will be requred to the site to replace the FS's possible internal microswitch issu. Diagnostic performed by Engineer : Discused and confirmed the issu with the cutomer. A cutomer reported an issu to advise u that they can't u-ray with the pedal that is attached to the table base. confirmed with the cutomer and from the logs that they have enabled u-ray and also made sue there is no pedal tapping. On checking the logs, I foud that possibly too muh delay in the activation of the 2 internal switches in the foot pedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the footswitch are not switching in the correct sequnce/in time with respect to each other. This is a possible issu with the FS itself. FSE requred to site to replace the FS. Problem Description by Engineer : Problem description by engineer : When foot pedal pressed system says "uuuis not possible" When was the first occurence? Within... : 00000000-Nov-2000000003 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: uing the one form the control room. Eupected and ueupected Resuts: The FS shoud u- ray. uer Impact: Cant u-ray Patient Impact: No futher information is available as per the cutomer uu00000000Follow u Requred : uLL# and/or description of uLL: ***Log File Evidence*** 600000000Warning 2000000003-11-00000000 16:0000000037.300000000EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 728 Warning 2000000003-11-00000000 16:0000000039.787 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 788 Warning 2000000003-11-00000000 16:0000000000000000893 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 800000000Warning 2000000003-11-00000000 16:0000000000000000773 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 900000000 Warning 2000000003-11-00000000 16:0000000050000000000000000 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 953 Warning 2000000003-11-00000000 16:0000000000000000.922 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 2100000000Warning 2000000003-11-00000000 16:0000000027.831 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 2281 Warning 2000000003-11-00000000 16:0000000058.00000000 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 15000000007 Warning 2000000003-11-00000000 10000000039:16.362 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 21300000000Warning 2000000003-11-00000000 11:00000000:100000000756 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 0000000019 Warning 2000000003-11-00000000 12:23:17.00000000 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent 0000000039 Warning 2000000003-11-00000000 12:38:35.782 EN_u inactive and Hand/Footswitch pressed 200000000SDTC0000000032200000000 Normal Operation Service ApplicationEvent Repair Action : FSE will be requred to the site to replace the FS's possible internal microswitch issu. Technician Name : Lokeh Vinay ----------------------------------------------------------------------------- Cutomer comment : 0000000030000000000000000000000001znQ00000000QAE ----------------------------------------------------------------------------- | 2023-11-04 12:56 | No
3366 | Yes | 0/u FD0geometrty fault | 0R/u FD0geometrty fault | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 00000000225390000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Problem description by engineer: Cutomer Problem Overview: The cutomer has contacted FSE and the machinery cannot be moved. Detailed description of cutomer problem or faut phenomenon: All machinery cannot move Whether it is a faut of the equpment itself: Yes Trouleshooting Action: What steps are taken duing the system diagnosis: Key faut codes or phenomena: 600000000uLL2000000003-11-00000000 00000000:50000000059.0000000000000000Programming uLL ueupected MEI Retun Code 000000003000000001000000007 Normal Operation Service NoDetailInfo 600000000uLL2000000003-11-00000000 00000000:50000000059.0000000000000000Programming uLL MCCV internal uLLor parameter uLL 000000003000000001000000007 Normal Operation Service NoDetailInfo 666 uLL2000000003-11-00000000 00000000:50000000059.0000000000000000Programming uLL MCCV call while MCCV not initialized 0000000030000000010000000000000000ormal Operation Service NoDetailInfo 667 uLL2000000003-11-00000000 00000000:50000000059.0000000000000000POST IPC Geometry Failed 0000000000000000000000009300000000Normal Operation Service NoDetailInfo 700000000 uLL2000000003-11-00000000 00000000:51:38.0000000000000000Application uLL POST failed: IPC>, uMP present 000000003000000000000000038 Normal Operation Service NoDetailInfo 700000000 uLL2000000003-11-00000000 00000000:51:38.0000000000000000Application uLL POST failed: IPC>, uMP_sw_version 000000003000000000000000038 Normal Operation Service NoDetailInfo 700000000 uLL2000000003-11-00000000 00000000:51:38.0000000000000000Application uLL POST failed: IPC>, synqnet_statu 000000003000000000000000038 Normal Operation Service NoDetailInfo 732 uLL2000000003-11-00000000 00000000:51:000000000000000000000000Programming uLL MCCV call while MCCV not initialized 0000000030000000010000000000000000ormal Operation Service NoDetailInfo 7300000000rror 2000000003-11-00000000 00000000:51:000000000000000000000000POST IPC Geometry Failed 0000000000000000000000009300000000Normal Operation Service NoDetailInfo 1552 uLL2000000003-11-00000000 00000000:00000000:15.0000000000000000Programming uLL ueupected MEI Retun Code 000000003000000001000000007 Normal Operation Service NoDetailInfo 1553 uLL2000000003-11-00000000 00000000:00000000:15.0000000000000000Programming uLL MCCV internal uLLor parameter uLL 000000003000000001000000007 Normal Operation Service NoDetailInfo 1578 uLL2000000003-11-00000000 00000000:00000000:000000000000000000000000Application uLL POST failed: IPC>, uMP present 000000003000000000000000038 Normal Operation Service NoDetailInfo 15800000000uLL2000000003-11-00000000 00000000:00000000:000000000000000000000000Application uLL POST failed: IPC>, uMP_sw_version 000000003000000000000000038 Normal Operation Service NoDetailInfo 15800000000rror 2000000003-11-00000000 00000000:00000000:000000000000000000000000Application uLL POST failed: IPC>, synqnet_statu 000000003000000000000000038 Normal Operation Service NoDetailInfo 1600000000 uLL2000000003-11-00000000 00000000:00000000:000000000000000000000000Programming uLL MCCV call while MCCV not initialized 0000000030000000010000000000000000ormal Operation Service NoDetailInfo 1611 uLL2000000003-11-00000000 00000000:00000000:000000000000000000000000POST IPC Geometry Failed 0000000000000000000000009300000000Normal Operation Service NoDetailInfo Remote diagnostics and test resuts: IPC issus PRS TR statu: All Access Repair Action: uLL# and/or description of uLL Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. Bad cabling and/or connections. 2. The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. 3. When problem is persistent it coud be caued du to a defective control net000000000000000000000000001k switch (TS). 00000000Issu with GeoIPC OS/AS software. It coud be du to a not correct shudown before, Windows OS can be corruted hence it needs more time in the neut start u. If this happens only sporadically, please ignore it. Else see other possible caues & soluions. Root Caue for RMT - ACQ : Frontal - Geometry indicates uknown SouceImageDistance : ImagingArea is set with MauRFS and MinSID: 1. The Geometry is uable to determine the Souce Skin Distance (SSD). This warning will occu when the system controller is uable to commuicate with GEO. Completed or Recommended Repair Action "Internal" : Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check the cable RC:u25 to SCC1:u22. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. 1.2. Tighten all the connectors of GeoIPC and SCC1 uit. 2.1. Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start u of the system. If so, replace GEO IPC. 3.1. Check control net000000000000000000000000001k switch (TS) and its power suply (adapter). Replace when necessary. 00000000. If SW is really corruted please reload GeoIPC software uing uB destroy stick. Service actions for RMT - ACQ : Frontal - Geometry indicates uknown SouceImageDistance : ImagingArea is set with MauRFS and MinSID: 1.1. Coud be issu with the PCufor the Geometry please check power suply first. Recommended on-site repair steps: Replacing Geo IPC Recommended spare parts uuand designation: PPD uu 00000000800000000931362 GEO IPC Ivybridge with ZMP CAN and IO mechanical motion control host 200000000to Hosp Notes on cutomer requrements: onsite SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 00000000800000000931362 1 Arranged to go to the hospital 200000000ous a day FSE: Yang Yang-86-1820000000000000000000000008 Euternal Remarks: 00000000.11.2000000003 100000000200000000 uC (8600000000600000000 Service content: Intermittent failue of mechanical movement, replacement of IPC host. Service resut: The device operates normally. Internal Remarks: 10000000011.2000000003 00000000:00000000:00000000 uC (8600000000600000000 "RDI Accuacy:Yes" $Q7: (Confirmed the problem, Intermittent mechanical movement abnormalities) $Q8: (onsite checked, foud Intermittent, C-arm cannot move, device needs to be restarted) $Q9: (onsite Replaced the GEO IPC.reload the SW) $Q100000000 ( $Q100000000 (Test and Inspection data refer to PA, Device retuned to ful fuctionality) ) $Q11: (No tools ued, Neut Calibration date ( NA ) Problem Reported by Cutomer: 00000000R/u uuuFD200000000 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-04 13:26 | No
3367 | Yes | 0/u FD0Host PC cannot power on | 0R/u FD0Host PC cannot power on | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 00000000.11.2000000003 0000000015:100000000uC (8600000000000000002) "SuT. Ready for Clinical ue local time" : " 2000000003-11-00000000 13:00000000" Service content: uable to enter the system after booting, on-site inspection that the Host PC does not start, pluging and upluging the Host to clean it After the PC memory modue and related boards are powered on and tested, they retun to normal and are observed for ue. Service resuts: After repair, the machine can be ued normally. Other important matters: None SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problem: Hos tPC device cannot be tuned on. Detailed description of cutomer problem or faut phenomenon: Hos tPC device cannot be tuned on. Is it a faut of the device itself? Yes Trouleshooting Action: Key faut codes or phenomena: The Hos tPC device cannot be powered on. Remote diagnosis and test resuts: The Hos tPC device cannot be powered on. Check that the power suply is normal. It is supected that the hostpc is fauty. PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check the HostPC power suply, 2. Reformat the HostPC OS disk and reinstall the system, 3. Replace the host PC. Recommended spare parts uuand name: " 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD" Cutomer demand Remarks: Onsite Problem Reported by Cutomer: The system cannot be started Safety Qustions: Case: 0000000022539100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 13:0000000011 uC (8600000000000000002) "RDI Accuacy: Yes" $Q1: () $Q2: () $Q3: () $Q00000000() $Q7: (Confirmation & Description of problem) Confirmed reported problem, host can't boot u $Q8: (Trouleshooting & Resuts) onsite check the device, the host PC can't boot u, clean the host PC memory modue, then boot u host PC normal, the device retuns normal. $Q9: (Repair steps) onsite check the device, the host PC can't boot u, clean the host PC memory modue, then boot u host PC normal, the device retuns normal. $Q100000000 (Test and inspection data refer to PA. Device retuned to ful fuctionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( refer to PA ) Neut Calibration date ( refer to PA ) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: N/A ----------------------------------------------------------------------------- | 2023-11-04 13:27 | No
3368 | Yes | SA:0/u FD0L-arm rotation cover broken | SA:0K/u FD0L-arm rotation cover broken | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Overview of cutomer problem: The cover on the L arm fell off. Detailed description of the cutomer problem or faut phenomenon: The cutomer reported that the bukle of the cover was broken and the cover on the L arm fell off. Is it a faut of the equpment itself? Yes Trouleshooting Action: Key faut code or phenomenon: The cover on the L arm comes off Remote diagnosis and test resuts: NA PRS TR statu: All pass Repair Action: Recommended on-site repair steps: Replace the cover that the L arm rotates. Recommended spare part uuand designation: PPD: cover uu 00000000230000000000000000793 L-arm rotation cove Quntity 1 piece Timeliness 200000000to hospital Cutomer demand notes: onsite Safety Qustions: Case: 0000000022539116 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 00000000230000000000000000793 1 Arranged to be at the venu 200000000ous a day uu Yu-86-1330000000010000000011 Problem Reported by Cutomer: Lid is broken FSE uu yu 00000000330000000010000000011 Call for repair SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: If the retun visit fails, MM will be notified via ## SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-04 13:32 | No
3369 | Yes | 0/V0uLLIntermittent black screen | 0K/V0uLLIntermittent black screen | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022539198 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Intermittent black screen The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: NA Repair Action: T2 Recommended on-site repair steps: Recommended trouleshooting screen Recommended spare part uufor T2 with description: #00000000561528891 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 00000000.11.2000000003 00000000:00000000:00000000uC (86000000000000000085) *Department information: [Department: Respiratory Department Contact Person: Teacher He Tel: 177235859300000000 * $ASP Engineer: [WuBEI Jia GANLu $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (reported problem has been confirmed, device intermittent black screen) $Q8: (juging u board failue) $Q9: (replacing u board) $Q100000000 (see PA checklist for test and check data.) . The device has been restored to ful fuction.) $Q11: not applicable SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Chongqing Bishan District Hospital Occasionally black screen FSE Notes: This PO is a Reference Nuber only Internal Remarks: 00000000.11.2000000003 00000000:00000000:52 uC (86000000000000000085) $ASP Engineer: [WuBEI Jia GANLu $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (reported problem has been confirmed, device intermittent black screen) $Q8: (juging u board failue) $Q9: (replacing u board) $Q100000000 (see PA checklist for test and check data.) . The device has been restored to ful fuction.) $Q11: not applicable ----------------------------------------------------------------------------- | 2023-11-04 13:56 | No
3370 | Yes | 000) 0 therapy supply out of range,value = 0 | 000(00) 0V therapy supply out of range,value = 00 | 000000000000000000000000001O##: 000000000000000000000000001000000003327900000000Parts Consued : 1.0000000000000000000000000000000056000000008951,DFM100000000 AC Power Modue ----------------------------------------------------------------------------- | 2023-11-04 13:57 | No
3371 | Yes | Therapy delivery test fail | Therapy delivery test fail | 000000000000000000000000001O##: 00000000000000000000000000100000000339535 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : 0000000056000000009000000001 DFM100000000 uuuitance ASSY ordered to resolve the problem Follow u Requred : 0000000056000000009000000001 DFM100000000 uuuitance ASSY ordered to resolve the problem Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Therapy delivery test fail ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000339572 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-11-04 16:02 | No
3372 | Yes | 0V0uLL sensor fault - pending repair | 0/V0uLL sensor fault - pending repair | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022500000000500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000887125271 Call Flow Sensors Duiangyan People's Hospital of Sichun Province ----------------------------------------------------------------------------- | 2023-11-04 16:06 | No
3373 | Yes | The device is not going through the test. | The device is not going through the test. | 000000000000000000000000001O##: 00000000000000000000000000100000000339600000000 Buak Tezel The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Diagnostic. Eupected and ueupected Resuts: The device does not pass the test. uer Impact: No. Patient Impact: The device coud not be ued. SE 2.00000000.32 Tools & Equpment : MuTIMETER, 200000000000000002-16 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000003-16 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER, 200000000000000000000000000000000 00000000:00000000:00000000 Parts: *1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance Resoluion: The defective Therapy uuuitor part of the device was replaced with a new one. All tests and checks were performed in according to the procedue listed in the service docuent. No problem was seen. The device was delivered to the uer in 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000339587 Buak Tezel The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: The inspection showed that the device did not pass the operation control test, the Therapy uuuitor part was ou of tolerance and was defective. The defective Therapy uuuitor part needs to be replaced with a new one. According to the failue detection, 0000000056000000009000000001 - A NEW ONE SHOuD BE DIRECTED TO ANTALYA HOSPITAL FOR HEALTH. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Diagnostic. Eupected and ueupected Resuts: The device does not pass the test. uer Impact: No. Patient Impact: No. SE 2.00000000.32 Resoluion: It was detected duing the checks that the device coud not pass the operation control test, the Therapy uuuitor part gave valus ou of tolerance and it was defective. The defective Therapy uuuitor part needs to be replaced with a new one. According to the failue detection, 0000000056000000009000000001 - A NEW ONE SHOuD BE DIRECTED TO ANTALYA HOSPITAL FOR HEALTH. ----------------------------------------------------------------------------- | 2023-11-04 16:17 | No
3374 | Yes | The device is not going through the Tess. | The device is not going through the Tess. | 000000000000000000000000001O##: 00000000000000000000000000100000000339628 Kaan Özkan The system meets the specification for the performed service and is retuned to ue Resoluion: The uuuitor was replaced with a new one, the device was tested and checked. It 000000000000000000000000001ks withou any problems. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Diagnostic. Eupected and ueupected Resuts: The Device Does Not Pass the Test. uer Impact: No. Patient Impact: No. SE 2.00000000.32 Diagnostic performed by Engineer: It is not suject to operational control. The uuuitor is fauty. Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance Tools & Equpment : MuTIMETER, 200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR, 200000000000000002-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000002-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000339618 Buak Tezel The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Duing the checks, it was determined that the device did not pass the operational control test, the Therapy uuuitor part gave a valu ou of tolerance and was defective. According to the faut detection, 0000000056000000009000000001 - SHOuD BE OFFERED TO ANTALYA LIFE HOSPITAL. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Diagnostic. Eupected and ueupected Resuts: The Device Does Not Pass the Test. uer Impact: No. Patient Impact: No. SE 2.00000000.32 Resoluion: Duing the checks, it was determined that the device did not pass the operational control test, the Therapy uuuitor part gave a valu ou of tolerance and was defective. According to the faut detection, 0000000056000000009000000001 - SHOuD BE OFFERED TO ANTALYA YAŞAM HOSPITAL. ----------------------------------------------------------------------------- | 2023-11-04 16:26 | No
3375 | No | PR# 00 Case 000s considered a duplicate of PR# 00Case 000 | PR# 000Case 0000is considered a duplicate of PR# 000Case 00003 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000718 Parts Consued: 1.0000000000000000000000000000000056000000009121,ITP CAPNOTRAK PuP ASSY ----------------------------------------------------------------------------- | 2023-11-04 17:21 | No
3376 | Yes | 0/u FD0Low space | 0R/u FD0Low space | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: No Faut Foud: uconfirmed Problem / Failue Problem Reported by Cutomer: Ou of memory Safety Qustions: Case: 000000002250000000038 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Low space Trouleshooting Action: Instrut cutomer to delete the old patients Repair Action: Instrut cutomer to delete the old patients Service content: Cutomers reported insuficient memory space and instruted cutomers to delete patient images to restore space. Service resut: After remote service, the device 000000000000000000000000001ks normally Test/Inspection: Device retun to ful fuctionality Test Equpment: N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-04 17:27 | No
3377 | Yes | Backup alarm is defective | Backup alarm is defeu | 000000000000000000000000001O##: 0000000000000000000000000010000000033000000001 Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: uLLmessage: Check Vent: Emergency alarm defective. Resoluion: Qute for intervention to be done on behalf of the Althea company. 00000000561537200000000Card - CPuPCBA (software 2.300000000: 1 00000000561537531 Card - Motor Controller PCBA (MC): 1 000000005615000000001 Card - Power Management PCBA (PM) 00000000 GENERATION: 1 Follow u Requred: Qute for intervention to be issud on behalf of the Althea company. M1: 2 hous D1 ----------------------------------------------------------------------------- | 2023-11-04 17:31 | No
3378 | Yes | Fluoroscopy cable PIM is not functioning | Fluoroscopy cable PIM is not funuoning | 000000000000000000000000001O##: 0000000000000000000000000010000000033000000003 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Footswitch not 000000000000000000000000001king properly When was the first occurence? Within... : 00000000-Nov-2000000003 1000000000000000000000000 How often is issu occuring? : First occurence Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Footswitch not 000000000000000000000000001king properly uer Impact: No ur imacpt Patient Impact: No patient impact SE 8.2.3000000005 Diagnostic performed by Engineer : Footswitch connector was not proper - FSE to visit and check the connection and reseat the connection properly. Finally need to check footswitch fuctionality. Internal Comments : uLL# and/or description of uLL: Footswitch not 000000000000000000000000001king properly Repair Action : Footswitch connector was not proper - FSE to visit and check the connection and reseat the connection properly. Finally need to check footswitch fuctionality. Resoluion : Footswitch connector was not proper - FSE to visit and check the connection and reseat the connection properly. Finally need to check footswitch fuctionality. Follow u Requred : uLL# and/or description of uLL: Footswitch not 000000000000000000000000001king properly Repair Action : Footswitch connector was not proper - FSE to visit and check the connection and reseat the connection properly. Finally need to check footswitch fuctionality. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-11-04 17:36 | No
3379 | Yes | uLLmessage/ No active alarm on all vital parameters - No Trouble Found - Cau Unknown | uLLmessage/ No au alarm on all vital parameters - No Trouble Found - Cau Unknown | 000000000000000000000000001O##: 0000000000000000000000000010000000033000000006 MESNIL, Mauime The system meets the specifications for the performed service and is retuned to ue. Resoluion: In the logs no uLLis visible today and no trace of alarm deactivation on the scopes is visible. The logs indicate that duing the mentioned period alarms rang on the monitors and were sent to the t000000000000000000000000001central stations of the department. Regarding the reactivation of the alarms, we only see the reactivation of the QT alarms from the central station by one of the nuses and the reactivation of the puse on the monitor of room neu_2 bu no other reactivation is visible on the other monitors. The analysis of the logs does not allow u to conclue on the souce of the problem, a CFF was opened with the qulity so that a futher analysis of the event is done by the factory, a mail will be sent to the cutomer to detail the resuts of this analysis. Diagnostic performed by Engineer: The cutomer indicates that duing the day of Thusday, November 2, a little before 16:300000000 all the alarms on all the monitors of the department were disabled and had to be reactivated manully to get back to normal operation. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Alaaudeactivated on all the monitors. uer Impact: No. Patient Impact: No. SE C.00000000.00000000 Reguatory Qustions: No/No ----------------------------------------------------------------------------- | 2023-11-04 17:37 | No
3380 | Yes | Machine is not booting and not switching on | Machine is not booting and not switching on | 000000000000000000000000001O##: 00000000000000000000000000100000000339900000000 00000000-Nov-2000000003 Lamba, Ajay The system meets the specification for the performed service and is retuned to ue Resoluion : Reload the ghost. Problem solved handed over system in 000000000000000000000000001king condition. Tools & Equpment : MuTIMETER,2000000003-12-27 00000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : Checked with cutomer system struked at 0000000000000000while startu . and application is not starting u Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000339521 Kuar, Atu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Checked with cutomer system struked at 0000000000000000while startu . and application is not starting u Follow u Requred : uLL# and/or description of uLL: uable to check as logs not available on PRS portal Repair Action : FSE to visit and check Technician Name : Kuar, Atu Internal Comments : uLL# and/or description of uLL: uable to check as logs not available on PRS portal Repair Action : FSE to visit and check Problem Description by Engineer : Problem description by engineer : Checked with cutomer system struked at 0000000000000000while startu . and application is not starting u Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000----------------------------------------------------------------------------- | 2023-11-04 18:06 | No
3381 | Yes | ulead malfunction | ulead malfunuon | 000000000000000000000000001O##: 00000000000000000000000000100000000339600000000Gnanasekar, D The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : 989800000000160000000000000000 u5-Lead Grabber Limb, IEC ordered to resolve the problem Follow u Requred : 989800000000160000000000000000 u5-Lead Grabber Limb, IEC ordered to resolve the problem Diagnostic performed by Engineer : ulead malfuction Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000339939 Parts Consued : 1.000000000000000000000000989800000000160000000000000000,u5-Lead Grabber Limb, IEC ----------------------------------------------------------------------------- | 2023-11-04 18:31 | No
3382 | Yes | operational test failed (Replaced DFM0 ASSY uuuitance) | operational test failed (Replaced DFM0 ASSY uuuance) | 000000000000000000000000001O##: 00000000000000000000000000100000000339886 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : 0000000056000000009000000001 DFM100000000 ASSY uuuitance ordered to resolve the problem Follow u Requred : 0000000056000000009000000001 DFM100000000 ASSY uuuitance ordered to resolve the problem Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : operational test failed ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000339952 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-11-04 18:41 | No
3383 | Yes | the device not booting-split | the device not booting-split | 000000000000000000000000001O##: 00000000000000000000000000100000000335615 13-Nov-2000000003 ue SERRA The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: No Eupected and ueupected Resuts: No video on TSM uer Impact: NO Patient Impact: NO SE000Diagnostic performed by Engineer: Change in TSM duing Monday's intervention. Parts Consued: 1.0000000000000000000000000000000080000000000000000563,Proscribe to TSM repl.kit Control Room ----------------------------------------------------------------------------- Internal Remarks: Requst intervention report from uSERRA dated 00000000/11/2000000003 (passage from uSERRA dated 10000000011/2000000003) ----------------------------------------------------------------------------- | 2023-11-04 18:51 | No
3384 | Yes | uPro ¦ trunk cable failing | TempuPro ¦ trunk cable failing | 000000000000000000000000001O##: 00000000000000000000000000100000000336311 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: biomed advised the truk cable for the Tempu Pro is not 000000000000000000000000001king, requsting a replacement cable Diagnostic performed by Engineer : Problem Description: biomed advised the truk cable for the Tempu Pro is not 000000000000000000000000001king, requsting a replacement cable Form ued: Parts Trouleshooting Steps Taken: Contacted the Sales Rep Shane GregoryPer the repair strategy for the Tempu Pro, the Sales rep is to order all accessories ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Chris 732998500000000 IP & Issu: Tempu - 2 truk cables failing on him - SN 610000000000000000 Needing replacement Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-04 20:16 | No
3385 | Yes | uPro ¦ trunk cable failing | TempuPro ¦ trunk cable failing | 000000000000000000000000001O##: 000000000000000000000000001000000003363300000000Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Problem Description: biomed reports the truk cable for the Tempu Pro is not 000000000000000000000000001king, requsting replacement Form ued: Information Requst Trouleshooting Steps Taken: Advised, Sales Rep, Shane Gregory will order any and all accessories for the Tempu Pro##d the Sales rep and advised him to order replacement cables for the cutomerIssu resolved Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: biomed reports the truk cable for the Tempu Pro is not 000000000000000000000000001king, requsting replacement ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Chris 732998500000000 IP & Issu: Tempu - 2 truk cables failing on him - SN 60000000000000000- Needing replacement. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created case and Page RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-04 20:37 | No
3386 | Yes | M0A / u00/ Does not pass the operational test with the euternal blades. | M00A / u000/ Does not pass the operational test with the euternal blades. |  | 2023-11-04 20:56 | No
3387 | Yes | M0A / u0 / Does not pass the pacemaker test. | M00A / u00 / Does not pass the pacemaker test. |  | 2023-11-04 20:57 | No
3388 | Yes | System blocks after a few seconds uLL"red". | System blocks after a few seconds uLL"red". | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000261 21-Nov-2000000003 GENNARI, MARIANO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Follow u Requred: Cool uit replacement/repair requred. Follow ##for replacement code Diagnostic performed by Engineer: System stopped du to a leak in the u-ray tue cooling system. Cool uit replacement requred. Estimate for replacement will follow. System stopped Tools & Equpment : MuTIMETER, 200000000000000005-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-11-04 21:16 | No
3389 | Yes | primary alarm failure | primary alarm failure | Internal Remarks : Interaction Type: New Case Caller: bob cambpbell for kelly colwell IP & Issu: "921293700000000 primary alarm failue" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-04 22:46 | No
3390 | Yes | 0/u FD0no uray | 0R/u FD0no uray | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 00000000.11.2000000003 00000000:10000000000000000uC (8600000000000000003) $Q7: (Confirmation & Description of problem) Confirmed reported problem: 000000000000000000000000001ld not eupouue $Q8: (Trouleshooting & Resuts) uLLmessage: check the system log, the FDC commuication with HOSTPC was not ready, the swicth was defcted $Q9: ( Repair steps)Replace the withc and make FDC test was successed.$Q100000000 Test and inspection :device has retuned to ful fuctionality Refer to PA$Q11:Calibrated Tools:Refer to PA T2 Activities:Problem Description: Overview of cutomer problems: Cutomer has contacted FSE,no uray Detailed description of cutomer problem or faut phenomenon: uable to eupose Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The switch is damaged Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Replace switch Recommended spare parts uuand designation: PPD uu00000000800000000239552 switch 200000000to Hosp Problem Reported by Cutomer: uable to eupose Safety Qustions: Case: 000000002250000000061 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD00000000800000000239552*1, has been shipped, please send 000000000000000000000000001k FSE: Tang Yi-86-18677920000000096 Euternal Remarks: 00000000.11.2000000003 00000000:55:18 uC (8600000000000000003) Check the system log and find that one interface of the switch is fauty. Replace the switch and perform corresponding tests. All fuctions of the device retun to normal SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-05 10:41 | No
3391 | Yes | System not booting | System not booting | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000090000000000000000-Nov-2000000003 Ramesh, S The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE 8.2.17 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Resoluion : Cutomer complaint regarding system not switching on , reached site & diagnosed & foud Host PC is fauty so the same ordered & foud part is not available in coutry stock so the same replaced from another site and made system was fuctional for time period ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003000000000000000000000000-Nov-2000000003 Ramesh, S The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer complaint regarding system not switching on , reached site & observed & foud Host PC is fauty so the same ordered & foud part is not available in coutry stock so the same replaced from another local site & made system fuctional Problem Description by Engineer : Problem description by engineer : Host PC not booting Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000Kuar, Vipin The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Host PC not booting Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided Follow u Requred : uLL# and/or description of uLL: No uLL Repair Action : Host PC not booting so ordered the same. Technician Name : Kuar, Vipin Internal Comments : uLL# and/or description of uLL: No uLL Repair Action : Host PC not booting so ordered the same. Diagnostic performed by Engineer : Checked and foud Host PC is not switching On ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000061 00000000-Nov-2000000003 FSE_Naseem J, Abdu Kader The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Tools & Equpment : MuTIMETER,200000000000000007-25 00000000:00000000:00000000 Diagnostic performed by Engineer : Checked the system and foud Host PC not booting. Checked the POST LEDs and it’s not glowing. Removed and seated all cables and boards bu same problem persist. Need to replace Host PC Problem Description by Engineer : Cutomer Fuction/Role: Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE 8.2.17 Follow u Requred : Need to replace Host PC ----------------------------------------------------------------------------- | 2023-11-05 13:06 | No
3392 | Yes | 0uDFM0 Defibrillator/Monitor/Electrode self check alarm | 0/Effi0 DFM0 Defibrillator/Monitor/Eleuode self check alarm | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002250000000083 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: uiyuyang Last udated by: uiyuyang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 153565770000000000000000Call u for repair It said that there was a call before, bu now the electrode pad is self-checking and alarming, and remote gudance is needed. Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-11-05 15:11 | No
3393 | Yes | System is not working at all | System is not working at all | 000000000000000000000000001O##: 000000000000000000000000001000000003000000000000000013-Nov-2000000003 AuAbdelrazek The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - SIB: SIBBOu/board failed. How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 000Resoluion : Replaced the SIB BOu modue in the M cabinet and reloaded the its own software.. now the new SIB BOu is 000000000000000000000000001king normally and the system is fuly fuctioning as well system is retun to medical ue as cutomer satisfied. Parts Consued : 1.00000000000000000000000000000000800000000000000002855,SIBBOu uIT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000Amjad Aburqou The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - SIB: SIBBOu/board failed. How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 000Diagnostic performed by Engineer : SIBBOu/board failed. and it need to be replaced. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - SIB: SIBBOu/board failed.: 1. SIBBOu / SIBBoard failed Repair Action : Service actions for RMT - SIB: SIBBOu/board failed.: 1.1. 1.Check the LED statu(Power and statu) of SIB 1.2. 2. InspectÂ POST resuts of SIB. 1.3. 3. Euecue BIST of System Interface uit in FSF. 1.00000000 00000000Perform loopback test in FSF. (For Release <7.6, CF-SIB is orderable separately). 1.5. 5. In the FSF, perform Fuctional test in Diagnostics /System Interface. Replace the defective uit Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - SIB: SIBBOu/board failed.: 1. SIBBOu / SIBBoard failed Repair Action : Service actions for RMT - SIB: SIBBOu/board failed.: 1.1. 1.Check the LED statu(Power and statu) of SIB 1.2. 2. InspectÂ POST resuts of SIB. 1.3. 3. Euecue BIST of System Interface uit in FSF. 1.00000000 00000000Perform loopback test in FSF. (For Release <7.6, CF-SIB is orderable separately). 1.5. 5. In the FSF, perform Fuctional test in Diagnostics /System Interface. Replace the defective uit Technician Name : Amjad Aburqou ----------------------------------------------------------------------------- | 2023-11-05 17:16 | No
3394 | Yes | Speaker Mallfuction | Speaker Mallfuuon | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000051 00000000-Nov-2000000003 Francis Olibrice The system meets the specification for the performed service and is retuned to ue Resoluion : CuOMER REPORTED SPEAKER MALFuCTION INOP AND THAT CENTRAL STAITION ALARMS WERE INTERMITTENT ADVISED CuTOMER THAT THE INOP WAS Du TO THE TELEMETRY BOu AND THAT Mu00000000CARRIES THAT INOP CuTOMER WAS ADAMANT THAT OVERVIEW IN TRAuA00000000WASNT ALARMING VERIFED THAT SPEAKER IS SECuED AND CONNECTION IS NOT LOOSE CONFIRMED THAT SPEAKER IS FuLY OPERATIONAL VERIFED THE ALARMS SETTINGS TESTED THE ALARM AND WAS ABLE TO CONFIRM THAT OVERVIEW IS FuLY OPERATIONAL Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: CuTOMER REPORTED THAT ALARM 000000000000000000000000001LD ONLY ALARM SOMETIMES AND WAS INTERMITTENT Bu AFTER TEST AND VERIFICATION NO ISSu FOuD WITH ALARMS uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: Sporadically -some sectors per the nuse show an INOP speaker malfuction- monitoring with tele Mu00000000. The INOP has cleared by restarting the Mu00000000in qustion today, only happening with one sector-no- no soud per biomed-euternal speaker. Biomed is no longer at the site. The INOP cleared so far. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000300000000Pirela, Manul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: Sporadically -some sectors per the nuse show an INOP speaker malfuction- monitoring with tele Mu00000000. The INOP has cleared by restarting the Mu00000000in qustion today, only happening with one sector-no- no soud per biomed-euternal speaker. Biomed is no longer at the site. The INOP cleared so far. Follow u Requred : Problem Description: Sporadically -some sectors per the nuse show an INOP speaker malfuction- monitoring with tele Mu00000000. The INOP has cleared by restarting the Mu00000000in qustion today, only happening with one sector-no- no soud per biomed-euternal speaker. Biomed is no longer at the site. The INOP cleared so far. Sporadically -some sectors per the nuse show an INOP speaker malfuction- monitoring with tele Mu00000000. The INOP has cleared by restarting the Mu00000000in qustion today, only happening with one sector-no- no soud per biomed-euternal speaker. Biomed is no longer at the site. The INOP cleared so far. Form ued: Trouleshooting/Investigation Object: Traua 00000000 Not Object: Software Version: uuDate and Time of Issu: 11/5/2000000003 uer Actions: Where is the Deviation: Traua 00000000 Trouleshooting Steps Taken: Then biomed cleared the INOP by resetting the Mu00000000 not able to replicate at this time, he is not in front of the system-he also restarted the PC and did not have a spare speaker-it coud also be a soud driver issu- advised to let me know if he wants me to dispatch the case as critical today- he stated he Was PRS ued: No PRS Permission: Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Prevention: Diagnostic performed by Engineer : Problem Description: Sporadically -some sectors per the nuse show an INOP speaker malfuction- monitoring with tele Mu00000000. The INOP has cleared by restarting the Mu00000000in qustion today, only happening with one sector-no- no soud per biomed-euternal speaker. Biomed is no longer at the site. The INOP cleared so far. Sporadically -some sectors per the nuse show an INOP speaker malfuction- monitoring with tele Mu00000000. The INOP has cleared by restarting the Mu00000000in qustion today, only happening with one sector-no- no soud per biomed-euternal speaker. Biomed is no longer at the site. The INOP cleared so far. Form ued: Trouleshooting/Investigation Object: Traua 00000000 Not Object: Software Version: uuDate and Time of Issu: 11/5/2000000003 uer Actions: Where is the Deviation: Traua 00000000 Trouleshooting Steps Taken: Then biomed cleared the INOP by resetting the Mu00000000 not able to replicate at this time, he is not in front of the system-he also restarted the PC and did not have a spare speaker-it coud also be a soud driver issu- advised to let me know if he wants me to dispatch the case as critical today- he stated he Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: smith pierre jean IP & Issu: 78167155/ Speaker Mallfution Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: NA Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Create R000000000000000000000000001and roued to RSE via RP Chat Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-11-05 21:36 | No
3395 | Yes | WHEN PERFORMING TESTS AND IT DOES NOT PASS THE uYy LOGO, THIS IS RANDOM | WHEN PERFORMING TESTS AND IT DOES NOT PASS THE uYy LOGO, THIS IS RANDOM | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000Nieves, Carlos The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: uLLreported with equpment SW, tried several times to call the contact withou sucess. When was the first occurence? Within... : 00000000-Nov-2000000003 1000000000000000000000000 How often is issu occuring? : First occurence Malfuction area: Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostics Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 7.2.16.00000000Diagnostic performed by Engineer: An uLLis reported with the equpment's SOFTWARE, we tried several times to call the contact withou sucess. Follow u Requred: uLL# and/or description of uLL uLLreported with equpment SW, tried several times to call the contact withou sucess. Repair Action: uLLreported with equpment SW, tried several times to call the contact withou sucess. What are the neut Trouleshooting steps: Equpment and log file are being accessed from the beginning. Check the system's power suplies and cards System HD check-u Software reload Technician Name: Nieves, Carlos Internal Comments: uLL# and/or description of uLL uLLreported with equpment SW, tried several times to call the contact withou sucess. Repair Action: uLLreported with equpment SW, tried several times to call the contact withou sucess. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000355399 00000000-Nov-2000000003 Mauimo, Migul The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: uLLreported with equpment SW, tried several times to call the contact withou sucess. When was the first occurence? Within... : 00000000-Nov-2000000003 1000000000000000000000000 How often is issu occuring? : First occurence Malfuction area: Imaging-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostics Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 7.2.16.00000000Resoluion: The equpment was foud to be 000000000000000000000000001king withou the failue reported in the log file, it does not report initialization uLL. Operation tests are satisfactory. ----------------------------------------------------------------------------- | 2023-11-05 23:56 | No
3396 | Yes | Patient u - Message stated shock was not necessary, but the device was still analyzing - log analysis requested | Patient u - Message stated shock was not necessary, but the device was still analyzing - log analysis requested | Internal Remarks: Failue content 2 Even after the "No shock necessary" message was sent, it was analyzed, bu since then the message has not been sent. Please check the log. LOT Adut pad: 210000000012-00000000500000000---------------------------------------------------------------------------- | 2023-11-06 06:11 | Yes
3397 | Yes | Fleuvision display problem | Fleusion display problem | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000083 00000000-Nov-2000000003 seokhyeon jeong The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Fleuvision display problem uer Impact: no Patient Impact: no SE R1 Resoluion : Checked fleuviewing pc statu. -> fleuvision pc boot fail Checked pc power line & power suply. -> mainboard defect Replaced fleuviewing pc Installed pc software & Changed license. Restored system settings. Performed system fuction tests. Parts Consued : 1.00000000000000000000000000000000800000000660000000013,FleuViewing PC 00000000 Tools & Equpment : MuTIMETER,200000000000000001-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000086 00000000-Nov-2000000003 KS, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Fleuvision display problem uer Impact: no Patient Impact: no SE R1 Resoluion : - checked fleuvision pc - Identified fleuvision problem - replaced fleuvision pc - changed license - changed mac address - Restored backu file - checked normal operation Tools & Equpment : MuTIMETER,200000000000000001-18 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000073 H.I ha The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FLEuVIEWING PC 00000000 change Follow u Requred : FLEuVIEWING PC 00000000 change Follow u Requred : FLEuVIEWING PC 00000000 change Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Fleuvision display problem uer Impact: no Patient Impact: no SE R1 Resoluion : FLEuVIEWING PC 00000000 change ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000019 Heeju Byu Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Fleuvision display problem uer Impact: no Patient Impact: no SE R1 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-11-06 06:16 | No
3398 | Yes | Euposure not possible due to oil leakage | Euposure not possible due to oil leakage | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 11.11.2000000003 16:31:23 uC (86000000000000000000000000 Service content: The machine cannot be euposed. The inspection foud that there was oil leakage at the cooling uit end, cauing the cooling uit to not 000000000000000000000000001k properly. After reconnecting the oil leakage point, it 000000000000000000000000001ked normally. Service resuts: After repair, the equpment retuned to normal operation. Problem Reported by Cutomer: The coolant is gone and the tue cannot be euposed. T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported oil leakage from the oil cooler, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: oil cooling oil leakage Remote diagnostics and test resuts: NA PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check whether there is oil leakage at the oil pipe connection and try to re-fiu it. 2. Check the machine uLLand perform faut repairs. Recommended spare parts uuand designation: NA Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Internal Remarks: 11.11.2000000003 16:31:26 uC (86000000000000000000000000 RDI Accuacy: YES $Q7:(Comfirmed reported problem.The device can't euposue.) $Q8:(Check and find it is the cooling uit oil rope leakage oil, need to reconnect the oil rope.) $Q9:(Reconnect the oil rope, and add the oil to the cooling uit.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: N/A. Safety Qustions: Case: 0000000022500000000100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-06 06:21 | No
3399 | Yes | Start failed | Start failed | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 10000000011.2000000003 00000000:0000000023 uC (8600000000900000000 000000000000000000000000001k content: After booting, it prompts that the line cannot be disconnected. After trouleshooting, it is foud that the IPPC host cannot start normally and all the indicators are not on. Replacement of a new IPPC can eupose the service normally. Resut: After replacing the IPPC, the device resues normal ue. Other matters: None Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Euternal Remarks: 00000000.11.2000000003 00000000:15:00000000 uC (8600000000900000000 000000000000000000000000001k content: After booting, it was prompted that it coud not go online. After investigation, it was foud that the IPPC host coud not start normally and all the indicator lights were not on. Buiness follow-u was requred to replace the new IPPC. Trouleshooting service resuts: Buiness follow-u is requred to puchase IPPC spare parts and continu the maintenance task after waiting for spare parts. Other matters: None Safety Qustions: Case: 000000002250000000033 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Internal Remarks: 10000000011.2000000003 00000000:0000000000000000uC (8600000000900000000 RDI Accuacy: Yes $Q7: (Confirmed reported problem.start u system euposue indicate u-ray cannot possible and monitor in MCS no signal ) $Q8: (check the fsf log, then check the IPPC 000000000000000000000000001king statu, that shows IPPC failue and no power to bootu& check ippc suply 2200000000 ok .need to replace IPPC to fiu the problem ) $Q9: (replaced the IPPC, and download software ;check the fuction ok) $ Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools N/A SAP Transfer: Case Transferred to SAP Neut Step Action: 139000000005000000003 called to remind the order, and the buiness has been notified for follow-u ----------------------------------------------------------------------------- | 2023-11-06 07:12 | No
3400 | Yes | During a 0 the arm did not rotate. | During a 0, the arm did not rotate. | 000000000000000000000000001O##: 000000000000000000000000001000000003000000007000000006-Nov-2000000003 Iizua, Tatsua The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Duing a 3D 3D, the arm did not rotate. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Duing a 3D 3D, the arm did not rotate. uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Resoluion: [000000000000000000000000001k content details] Based on the event log details, I juged that duing a 3DRA DRA, the C-arm propeller was rotating and something was being collided with or something was abou to collide (soft collision) and the equpment recognized it and the 3D 3D image did not start normally. After I replaced the position sensor that detects the position of the C-arm propeller, I adjuted the C-arm propeller rotation position and the drive curent.I checked the operation of the 3D 3D and confirmed that there is no problem. Resoluion: [Parts ued] Part nuber: 00000000230000000017113 /Part name POTENTIOMETER ASSY, ROTATION DRIVE /Serial No.: NA /Nuber of uits: 1 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000037 Saka, Shuji The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Duing a 3D 3D, the arm did not rotate. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Duing a 3D 3D, the arm did not rotate. uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Diagnostic performed by Engineer: RC Response: I checked the log. "Safepath policed by the profilemovement itself" was being ued. Visit and investigate. Internal Comments: [Suject] 00000000DM:SS:u Hirono: Duing a 3D 3D, the arm did not rotate. [Internal Comments] Safepath violated by the profilemovement itself soft-collision detected [FREE_TEuT] movement: BeamPropeller, override=00000000at position: -1200000000(000000001[cm¦dg]) Time of occurence : 2000000003-11-00000000, 2:17:19 The uLLis different, bu 00000000230000000017000000007: Rotation drive H.D. was replaced. Previou uLL occurence: 2000000003-100000000100000000 3:17:58 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] uLLreport [movement=BeamPropeller] Imin : -000000001 mA Imau : 2355 mA Iact : -1000000003 mA Iset : 1900000000mA Vact : 500000000 aus Vset : 00000000 aus Pact : -300000000auPset : -313 auState : MotorisedMoving Note : au= 000000001deg. or mm uLLs) foud : curent uLL [Parts from the equpment] 00000000230000000017113:potentiometer assy, rotation drive QTY:1 ZRFA Tokyo (u11) 0000000023000000002300000000:CLEA EuTENSION BOARD 1 QTY:1 ZRFA Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Tatsua Iizua] Iizua, Tatsua - FSE1 Visit date and time 2000000003/11/0000000017:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions CPu(no warehoue remove/transport arrangements) - Delivery Mut arrived by 13:00000000 on November 6, 2000000003. - Delivery address name FSE will remove it from the Tokyo warehoue. - Received person name Tatsua Iizua - 0000000000000000002268-2351 - Others Requster name Toshihito Saka [Euamination category] Normal euamination Follow u Requred: ⋆ ----------------------------------------------------------------------------- | 2023-11-06 08:16 | No
3401 | Yes | The equipment also did not start last Friday. | The equipment also did not start last Friday. | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000012 Suui, NobuasuThe system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: The equpment also did not start last Friday.Over the weekend, the equpment was left runing. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The equpment also did not start last Friday.Over the weekend, the equpment was left runing. uer Impact: No effect Patient Impact: No effect SE R000Internal Comments: [Suject] 00000000DS:NS:u Tabata: The equpment also did not start last Friday. [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name --None-- - FSE1 Visit date and time 2000000003/11/00000000--None-- - FSE1 000000000000000000000000001k time 2 [Euamination category] Normal euamination Resoluion: RC Response: Case:0000000022000000005900000000ill be responded to tomorrow ----------------------------------------------------------------------------- | 2023-11-06 08:51 | No
3402 | Yes | uDFM0 Defibrillator/Monitor/connects to the ulead wire and reports an uLLduring power-on selt-test | Effi0 DFM0 Defibrillator/Monitor/conneu to the ulead wire and reports an uLLduring power-on selt-test | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: Instruted Cutomer Problem Reported by Cutomer: 00000000N/uDFM100000000 Defibrillator/Monitor/ T2 Activities:Problem Description: Cutomer’s euplanation of the problem ### self-test identification The cutomer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### "Key uLLstatements and codes### The device connects to the ulead wire and reports an uLLduing power-on self-test.After instruting the cutomer to remove the u the manul self-test is normal.Restarting the device 000000000000000000000000001ks fine." Repair Action: T2 Recommended Field Diagnostic Steps### After removing the euternal u the device self-test is normal. T2 Recommended spare part uuwith description### na Cutomer demand### na Test/Inspection: na Test Equpment: na Safety Qustions: Case: 0000000022500000000600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: weiyu Last udated by: wei yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-06 09:01 | No
3403 | Yes | The equipment remains at 0:0 and does not start | The equipment remains at 0:0 and does not start | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000081 Suui, NobuasuThe system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000CS:NS:u Morita: The equpment remains at 00000000:00000000 and does not start. [Internal Comments] Simp 1223 FleuVision Warning 2000000003-11-00000000 00000000:300000000000000000000000000000000FleuVisionControlService:Observer000000000000000000000000001ker PING(FLEuVISION_PC_CONTROL) is not suceeded after 100000000 seconds. Please check if FLEuVISION_PC_CONTROL is powered on and all (net000000000000000000000000001k) cables are correctly connected. [SEVERITY] Warning [EuCEPTION_DESCRIPTION] FleuVisionControlService:Observer000000000000000000000000001ker PING(FLEuVISION_PC_CONTROL) is not suceeded after 100000000 seconds. Please check if FLEuVISION_PC_CONTROL is powered on and all (net000000000000000000000000001k) cables are correctly connected. 8500000000000000000000000000000000 Startu Service NoDetailInfo 0000000020000000000000000000000001 FRuFleu-PC HD (withou snapshot) [Euamination category] Normal euamination Problem Description by Engineer: The equpment remains at 00000000:00000000 and does not start.Tuning the power back on does not improve the sitution.On the FleuVision monitor, the layou gridline and images from euternal inpu, suh as PACS, are displayed. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The equpment remains at 00000000:00000000 and does not start.Tuning the power back on does not improve the sitution.On the FleuVision monitor, the layou gridline and images from euternal inpu, suh as PACS, are displayed. uer Impact: No effect Patient Impact: No effect SE R8.1.3000000005 Resoluion: RC Response: He tuned the circut breaker in the M cabinet off/on and then tuned the power off/on (the second time), and it started normally. At that time, the display on the data monitor screen had a blu tinge, so I had him correct the color settings in the settings menu I remotely diagnosed it and confirmed that commuication was not established between the compuer that controls FleuVision and the host compuer duing startu. ----------------------------------------------------------------------------- | 2023-11-06 09:16 | No
3404 | Yes | Response received/QA000000 Device powered off | Response received/QA00000000 Device powered off |  | 2023-11-06 11:19 | No
3405 | Yes | A full hard disk is being loaded. | A full hard disk is being loaded. | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000099 00000000-Nov-2000000003 Meira, Fernando The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: RMT - POS: uable to commuicate with GeoIPC 00000000-Nov-2000000003 00000000:0000000015 Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Geometry does not 000000000000000000000000001k. uer Impact: The equpment cannot be ued normally. Patient Impact: No futher information is available as per the cutomer. SE 000Resoluion: Software reinstallation in GeoPc. Correct operation tests. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000083 Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: RMT - POS: uable to commuicate with GeoIPC 00000000-Nov-2000000003 00000000:0000000015 Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Geometry does not 000000000000000000000000001k. uer Impact: The equpment cannot be ued normally. Patient Impact: No futher information is available as per the cutomer. SE 000Diagnostic performed by Engineer: The equpment does not move and a ful disk message appears. Checked the system event log, foud commuication uLL with the geometry control PC (Geo IPC), it might be necessary to replace it, need to go and check it. It is also recommended to delete the image disks completely from service ("Image Database Consistency test", "Recreate Image Store"). Internal Comments: uLL# and/or description of uLL Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. When problem is persistent it coud be caued du to a defective control net000000000000000000000000001k switch (TS). Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. Bad cabling and/or connections. Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. Issu with GeoIPC OS/AS software. It coud be du to a not correct shudown before, Windows OS can be corruted hence it needs more time in the neut start u. If this happens only sporadically, please ignore it. Else see other possible caues & soluions. Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. Repair Action: Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check control net000000000000000000000000001k switch (TS) and its power suply (adapter). Replace when necessary. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check the RC:u25 to SCC1:u22 cable. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. 1.2. Tighten all the connectors of GeoIPC and SCC1 uit. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. If SW is really corruted please reload GeoIPC software uing uB destroy stick. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start u of the system. If so, replace GEO IPC. Follow u Requred: uLL# and/or description of uLL Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. When problem is persistent it coud be caued du to a defective control net000000000000000000000000001k switch (TS). Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. Bad cabling and/or connections. Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. Issu with GeoIPC OS/AS software. It coud be du to a not correct shudown before, Windows OS can be corruted hence it needs more time in the neut start u. If this happens only sporadically, please ignore it. Else see other possible caues & soluions. Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. Repair Action: Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check control net000000000000000000000000001k switch (TS) and its power suply (adapter). Replace when necessary. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check the RC:u25 to SCC1:u22 cable. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. 1.2. Tighten all the connectors of GeoIPC and SCC1 uit. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. If SW is really corruted please reload GeoIPC software uing uB destroy stick. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start u of the system. If so, replace GEO IPC. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-11-06 13:37 | No
3406 | Yes | System failure. Images do not save | System failure. Images do not save | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000027 November 100000000 2000000003 Krub, Horst The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: System had to be shu down 00000000to ru. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Resoluion: Replacement bios battery and change hard disk in the image compuer. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000079 Baugarten, Rüdiger The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: System had to be shu down 00000000to ru. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: Cutomer is to be Log analysis Internal Comments: uLL# and/or description of uLL IPPC is not recognized; Enable Movement uLL Repair Action: Te with part Follow-u requred: uLL# and/or description of uLL IPPC is not recognized; Enable Movement uLL Repair Action: Te with part Technician name: Baugarten, Rüdiger ----------------------------------------------------------------------------- Cutomer comment: ## Service on 00000000.11. 9:00000000 am u Baugarten: The information is known to the I will call Horst Internal Remarks: Via FSE Krub The part cannot be foud, bu the system is runing again. Please schedue again tomorrow afternoon. ----------------------------------------------------------------------------- | 2023-11-06 15:01 | No
3407 | Yes | Does not boot completely | Does not boot completely | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000091 November 7, 2000000003 Nölkes, Fabian The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.00000000000000000000000000000000800000000200000000611,Nehalem Host PC Monoplane Rev_III NO_HDD Resoluion: The reported uLLis confirmed. Host PC starts withou hard drives. PC replaced. License adjuted. System checked. The system corresponds to the manuactuer's specifications and is ready for clinical ue. Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. u Schroer is on the way. ++0000000031180000000031561 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.25.1 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000100000000obert Kern The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. u Schroer is on the way. ++0000000031180000000031561 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.25.1 Diagnostic performed by Engineer: According to the cutomer, no display on monitors. Net000000000000000000000000001k uLLmessage on TSM. Cutomer replaced data monitor, no change. Cutomer checked 5 V. Shoud be OK. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: Schedue technical service for trouleshooting on site, 2 hous of 000000000000000000000000001k. Take host PC and implementer along. Replacement of what is defective. Technician name: Robert Kern ----------------------------------------------------------------------------- | 2023-11-06 16:22 | No
3408 | No | Broken cuff | Broken cuff | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000065 Andrea Ardighieri The system meets the specification for the performed service and is retuned to ue. Resoluion: The System does not show malfuctions, it retuns accordingly to clinical ue with the delivery of the material of the accessory/consuable (ecvo spo2 pressue cuf) In case of futher problem, as agreed with the department u Lenti a new report will be opened. Problem Description by Engineer: Cutomer Fuction/Role: department Device ue: defibrillator test Eupected and ueupected Resuts: broken cuf and sp2 cable uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Spoke with department contact Lenti, as agreed we proceed with the shipment to the department Follow u Requred: nemo withou collection S. S. ANNuZIATA HOSPITAL 1 BRuO STREET 7000000003 TARANTO FIRST AID DEPARTMENT COORDINATION DEPARTMENT AC. LENTI 000000009000000005191-000000009- 000000005000000005 ----------------------------------------------------------------------------- | 2023-11-06 16:27 | No
3409 | Yes | Machine is not booting | Machine is not booting | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000300000000Madhan, kuar The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : To replace 0000000056000000009000000001 DFM100000000 SOM PCA ASSY Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001 DFM100000000 SOM PCA ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000066 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY SOM ----------------------------------------------------------------------------- | 2023-11-06 16:36 | No
3410 | Yes | uLLdescription: System does not receive power. Patient is on the system. | uLLdescription: System does not receive power. Patient is on the system. | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000092 November 7, 2000000003 Beer, Dietmar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called u Asrawi. +0000000000000000267900000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 7.2.16.00000000Resoluion: Trouleshooting, power-u circut checked, MPD control uit replaced, fuction check, system is ready for clinical service Parts Consued: 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT Tools & Equpment: MuTIMETER, 200000000000000005-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000063 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called u Asrawi. +0000000000000000267900000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 7.2.16.00000000Diagnostic performed by Engineer: Phases lamps light u. Bu no ED is on in the PC bou. System does not react to the switch-on command. Fues F100000000F11 and F 23 all OK. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: TE service to check power suply PC bou and MPDucontrol uit. Replacement of the defective part. 3 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Cutomer comment: ## Appt. arranged with u Al-Asrawi for 00000000.11. from approu. 9:00000000 am. Cutomer comment: Sent an ##to the cutomer. Cutomer comment: u Muz Al-Asrawi +0000000000000000267900000000asks for call back regarding scheduing. Cutomer 000000000000000000000000001ld like to know whether appointments for tomorrow are to be cancelled. Cutomer comment: Confirmed the kVa by ##----------------------------------------------------------------------------- | 2023-11-06 16:41 | No
3411 | Yes | Black screen appears during startup | Black screen appears during startup |  | 2023-11-06 16:57 | No
3412 | Yes | uLLmessage: Memory is full | uLLmessage: Memory is full | 000000000000000000000000001O##: 000000000000000000000000001000000003000000006 November 7, 2000000003 Fueder, Bernhard The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Memory is ful Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Resoluion: Trouleshooting; IP-PC for lateral level replaced and reprogrammed. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Performed check of the electrical systems according to DIN EN 62353 and DGu reguation 3. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000012200000000,u Haswell Clarity PC no HDD Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000000000000000000000000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000077 Baugarten, Rüdiger The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: uLL# and/or description of uLL IPPC supiciou Repair Action: Te with part Follow-u requred: uLL# and/or description of uLL IPPC supiciou Repair Action: Te with part Technician name: Baugarten, Rüdiger Problem Description by Engineer: Problem description by engineer: Memory is ful Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: Called cutomer, log file cannot be downloaded ----------------------------------------------------------------------------- Cutomer comment: ## Service with the uer for 00000000.00000000. from approu. 100000000300000000am. ----------------------------------------------------------------------------- | 2023-11-06 17:31 | No
3413 | Yes | "OVP Circuit Failed" alarm occurred and it could not be reset | "OVP 0cuit Failed" alarm occurred and it could not be reset. - Pending repair 0/8 |  | 2023-11-06 18:10 | No
3414 | Yes | the end ur is unable to delete the patient data from the system | the end ur is unable to delete the patient data from the system | 000000000000000000000000001O##: 0000000000000000000000000010000000056600000000 00000000-Dec-2000000003 Dharmaraj Thaiappan The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Image -re-creation was performed. Tested and handed over the system for clinical uage. Problem Description by Engineer : Problem description by engineer : RMT - IPu Frontal IPPC failed to boot u 29-Nov-2000000003 00000000:00000000:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: RMT - IPu Frontal IPPC failed to boot u 29-Nov-2000000003 00000000;00000000;00000000uer Impact: None Patient Impact: None SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000056600000000 00000000-Nov-2000000003 Roble Ramanan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - IPu Frontal IPPC failed to boot u 29-Nov-2000000003 00000000:00000000:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: RMT - IPu Frontal IPPC failed to boot u 29-Nov-2000000003 00000000;00000000;00000000uer Impact: None Patient Impact: None SE 000Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Frontal IPPC failed to boot u: 1. The Host PC (TH) is uable to commuicate with the Frontal IPPC (TP). Either the IPPC did not boot u du to lack of (Mains) power, a defective BIOS battery or the OS/AS did not boot u in-time. Repair Action : Service actions for RMT - IPu Frontal IPPC failed to boot u: 1.1. Check the back of the IPPC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs in the Trouleshooting manul of the IPPC on InCenter. 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related. Then euamine the ethernet wiring (THu12 - TPu100000000 between Host PC and IPPC and perform Cold reboot of the system. 1.3. If LED statu is NOT Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IPPC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct. 1.5. A defective BIOS battery coud be a reason why the IPPC does not boot. Then replace the Frontal IPPC (TP). Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - IPu Frontal IPPC failed to boot u: 1. The Host PC (TH) is uable to commuicate with the Frontal IPPC (TP). Either the IPPC did not boot u du to lack of (Mains) power, a defective BIOS battery or the OS/AS did not boot u in-time. Repair Action : Service actions for RMT - IPu Frontal IPPC failed to boot u: 1.1. Check the back of the IPPC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs in the Trouleshooting manul of the IPPC on InCenter. 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related. Then euamine the ethernet wiring (THu12 - TPu100000000 between Host PC and IPPC and perform Cold reboot of the system. 1.3. If LED statu is NOT Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IPPC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct. 1.5. A defective BIOS battery coud be a reason why the IPPC does not boot. Then replace the Frontal IPPC (TP). Technician Name : Roble Ramanan Diagnostic performed by Engineer : RMT - IPu Frontal IPPC failed to boot u 29-Nov-2000000003 00000000:00000000:00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000061 00000000-Nov-2000000003 Aneesh Nair The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 16-Oct-2000000003 00000000:21:26 RMT - IPu Free disk space is too low. 16-Oct-2000000003 00000000:21:26 RMT - IPu Frontal - Image Disk uLLIP-PC Image disk 1 cannot be accessed 16-Oct-2000000003 00000000:20000000028 Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: the end uer is uable to delete the patient data from the system uer Impact: Intermittently system uable to ue Patient Impact: None SE 000Diagnostic performed by Engineer : Cutomer informed that they are uable to delete the patient from system. checked the system and foud that the even thouh the patients being deleted from the system , the space is showing same. recreated the image disks bu foud issus with the performance of IPPC and its disks. since the parts are not covered uder the contract,a qute need to sent and approved. once the part are approved and ordered ; a new case will be opened for the part replacement. system is in restricted mode bu in 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000037 Roble Ramanan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 16-Oct-2000000003 00000000:21:26 RMT - IPu Free disk space is too low. 16-Oct-2000000003 00000000:21:26 RMT - IPu Frontal - Image Disk uLLIP-PC Image disk 1 cannot be accessed 16-Oct-2000000003 00000000:20000000028 Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: the end uer is uable to delete the patient data from the system uer Impact: None Patient Impact: None SE 000Diagnostic performed by Engineer : the end uer is uable to delete the patient data from the system Internal Comments : uLL# and/or description of uLL: 00cc move from local stock 0000000080000000050000000000000000HARD DISK SPARE PART Qty 00000000. Repair Action : Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu Frontal - Image Disk uLLIP-PC Image disk 1 cannot be accessed: 1.1. Check the cable between the disk and the motherboard for connection issus. 1.2. Image disk 1 is defect. 1.3. Replace the complete IP-PC (possible defect disk bay board). Service actions for RMT - IPu Frontal - Image Disk uLLIP-PC Image disk 1 cannot be accessed: 1.1. Please check the brand and type of disk. Replace with thee correct type of harddisk as advised via E-SPF. Follow u Requred : uLL# and/or description of uLL: 00cc move from local stock 0000000080000000050000000000000000HARD DISK SPARE PART Qty 00000000. Repair Action : Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu Frontal - Image Disk uLLIP-PC Image disk 1 cannot be accessed: 1.1. Check the cable between the disk and the motherboard for connection issus. 1.2. Image disk 1 is defect. 1.3. Replace the complete IP-PC (possible defect disk bay board). Service actions for RMT - IPu Frontal - Image Disk uLLIP-PC Image disk 1 cannot be accessed: 1.1. Please check the brand and type of disk. Replace with thee correct type of harddisk as advised via E-SPF. Technician Name : Roble Ramanan ----------------------------------------------------------------------------- | 2023-11-06 18:21 | No
3415 | Yes | the big screen not working | the big screen not working | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000081 00000000-Nov-2000000003 Mohamed Karawala The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : the big screen not 000000000000000000000000001king Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Resoluion : -Sucessfuly, corrective maintenance visit performed with reported problem that, the fleuvision screen isn't booting. -Performed software for the fleuvison PC. -Tested the system and it is back to clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000700000000Amjad Aburqou The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : the big screen not 000000000000000000000000001king Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Diagnostic performed by Engineer : the big screen not 000000000000000000000000001king, system not connected to PRS need to plan onsite visit. Internal Comments : uLL# and/or description of uLL: the big screen not 000000000000000000000000001king, system not connected to PRS need to plan onsite visit. Repair Action : the big screen not 000000000000000000000000001king, system not connected to PRS need to plan onsite visit. Follow u Requred : uLL# and/or description of uLL: the big screen not 000000000000000000000000001king, system not connected to PRS need to plan onsite visit. Repair Action : the big screen not 000000000000000000000000001king, system not connected to PRS need to plan onsite visit. Technician Name : Amjad Aburqou ----------------------------------------------------------------------------- | 2023-11-06 18:47 | No
3416 | Yes | power supply fail | 00 power supply fail |  | 2023-11-06 18:53 | No
3417 | Yes | SPLIT - Charging connector and foot are out of work. | SPLIT - Charging conneur and foot are out of work. | 000000000000000000000000001O##: 00000000000000000000000000100000000352256 00000000-Dec-2000000003 Franck PETIT The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: no Patient Impact: no SE / The holding bar was broken and the device no longer charged du to the connector defective power suply Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000005-0000000000000000:00000000:00000000 PATIENT SIMuATOR, 200000000000000003-00000000 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000056000000005111,RDT Tempu Pro Rear Cover w/Printer Kit 1.0000000000000000000000000000000056000000005000000001,RDT Front Enclosue kit (for 00000000-100000000u-R & Resoluion: Replacement by Romain SEVA of the front panel and the rear hul. Performance test: OK carried ou by Franck PETIT The device can be pu back into service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000096 Seva, Roman The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: The device will be retuned to Suesnes for repair. - Replacement of the front panel for the charging connector. - Replacement of the rear panel for the stand (no spare part for the stand only - it is inclued in the rear panel). Follow u Requred: Qute for intervention. MO= 00000000M= M1 No travel to be charged becaue the device will be sent to Suesnes. Franck Petit suport needs to be planned. PLEASE NOTE for parts, when the qute is accepted, they shoud be sent to the following address: u France COMMERCIAL 33 Ru de Verdu Liv Suport MA&TC - 00000000 floor - Franck PETIT 0000000000000000700000000873 921500000000SuESNES France Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: no Patient Impact: no SE / Diagnostic performed by Engineer: The charging connector and the foot are ou of order. Reguatory Qustions: N/A ----------------------------------------------------------------------------- | 2023-11-06 19:12 | No
3418 | Yes | Speaker defect - Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part | Speaker defeu- Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part | 000000000000000000000000001O##: 00000000000000000000000000100000000352371 Gu DE GREVE The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: CASE Device ue: Eupected and ueupected Resuts: uer Impact: None Patient Impact: SE Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000001-17 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000056000000006971,Mu_75000000008500000000& AD_75/85 Louspeaker Assy Resoluion: Replacement of the speaker. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000352297 Gu DE GREVE The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: uder contract, do 000000000000000000000000001bench repair 2 hous for GDG for Friday, November 100000000+ order 1 * 0000000056000000006971 Louspeaker Assy PuO Jette Problem Description by Engineer: Cutomer Fuction/Role: CASE Device ue: Eupected and ueupected Resuts: uer Impact: None Patient Impact: SE Resoluion: uder contract, do 000000000000000000000000001bench repair 2 hous for GDG for Friday, November 100000000+ order 1 * 0000000056000000006971 Louspeaker Assy PuO Jette ----------------------------------------------------------------------------- | 2023-11-06 19:41 | No
3419 | Yes | geometry not avaiable | geometry not avaiable | 000000000000000000000000001O##: 0000000000000000000000000010000000035600000000 00000000-Nov-2000000003 Rodriguz Antonio The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : spoke to Todd the mcs does not move or the table .message on system When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: system did not start as eupected uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion : Picked u part and shipped ou. System is fuctional ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000353598 00000000-Nov-2000000003 Rodriguz Antonio The system meets the specification for the performed service and is retuned to ue Follow u Requred : Retun part Parts Consued : 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V Problem Description by Engineer : Problem description by engineer : spoke to Todd the mcs does not move or the table .message on system When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: system did not start as eupected uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion : Replaced the 200000000power suply. The system is now operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003520000000000000000-Nov-2000000003 Rodriguz Antonio The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace power suply Problem Description by Engineer : Problem description by engineer : spoke to Todd the mcs does not move or the table .message on system When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: system did not start as eupected uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Foud that 200000000power suply was fauty Tools & Equpment : MuTIMETER,2000000005-00000000-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000352382 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : spoke to Todd the mcs does not move or the table .message on system When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: system did not start as eupected uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : speaking to Todd the system was on this morning and he performed hard down and restart ,still getting message on system . recomend move patient Todd saved lod for service . rest of system ,screens uray appears to be u runing . no movements of stand or table . Resoluion : Coud not resolve the issu and requres onsite suport Internal Comments : uLL# and/or description of uLL: Geometry movements not avaiable Repair Action : uY Name: 39af52ae-8ae6-ceeb-12c0000000000000000d6a89182dc Root Caue 1: Geo IPC can card malfuction. Service Action 1: Reseat the can card in the Geo IPC andÂ perform download board software. Service Action 2: Replace with compatible Geo IPC. Produt Name GEO IPC Ivybridge no ZMP with CAN and IO Produt Code 00000000800000000931372 PD.SCOMP.DCPS_20000000012V_5V Material Nuber 00000000800000000387582 lod at case level Follow u Requred : uLL# and/or description of uLL: Geometry movements not avaiable Repair Action : uY Name: 39af52ae-8ae6-ceeb-12c0000000000000000d6a89182dc Root Caue 1: Geo IPC can card malfuction. Service Action 1: Reseat the can card in the Geo IPC andÂ perform download board software. Service Action 2: Replace with compatible Geo IPC. Produt Name GEO IPC Ivybridge no ZMP with CAN and IO Produt Code 00000000800000000931372 PD.SCOMP.DCPS_20000000012V_5V Material Nuber 00000000800000000387582 lod at case level What are the neut Trouleshooting steps : check startu of geo pc and power suply .200000000olt low voltage ,reset can card in geo pc Technician Name : Sulivan Scott ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Rodriguz Antonio with TPR approval for defective DCPS in R-cabinet. 2300000000AC was at the inpu for the power suply. ----------------------------------------------------------------------------- | 2023-11-06 19:47 | No
3420 | Yes | SULTANAH AMINAH HOSPITAL ¦ u FD00¦ ISSUE ¦ Problem system cannot boot up | SULTANAH AMINAH HOSPITAL ¦ u FD00¦ ISSUE ¦ Problem system cannot boot up | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000586 00000000-Nov-2000000003 MOHD ASRAF BIN AB RAHIM The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Device ue: Angiogram Eupected and ueupected Resuts: system cannot boot u. uer Impact: No Impact Patient Impact: No Impact SE R8.2 Resoluion : Confirm reported problems; system cannot boot u. New Fleuvision PC arrived, replace the part accordingly. Perform system fuctional test; all passed. system retun to uer fuly fuctional. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-11-06 19:51 | No
3421 | Yes | : It does not work, notice becau the disk is full, images are deleted and the same defect continues | : It does not work, notice becau the disk is full, images are deleted and the same defeucontinues | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000088 Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: It seems that the access to the image disks is failing, E. Alonso is in charge of more tests. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Images cannot be acqured. uer Impact: Images cannot be acqured. Patient Impact: No impact SE 000Diagnostic performed by Engineer: It seems that the access to the image disks is failing, E. Alonso is in charge of more tests. Internal Comments: uLL# and/or description of uLL N/A Repair Action: It seems that the access to the image disks is failing, E. Alonso is in charge of more tests. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: It seems that the access to the image disks is failing, E. Alonso is in charge of more tests. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000352631 00000000-Nov-2000000003 Alonso Lindell, Eric The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: It seems that the access to the image disks is failing, E. Alonso is in charge of more tests. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Images cannot be acqured. uer Impact: Images cannot be acqured. Patient Impact: No impact SE 000Resoluion: Operation of the IPPC disks is checked, all the stuies are deleted from the disks and the equpment is 000000000000000000000000001king correctly. ----------------------------------------------------------------------------- | 2023-11-06 20:16 | No
3422 | Yes | Device restarted due to uLL | Device restarted due to uLL |  | 2023-11-06 20:48 | No
3423 | Yes | Geometry does not work - repeatedly. | Geometry does not work - repeatedly. | 000000000000000000000000001O##: 00000000000000000000000000100000000353000000006 100000000Nov-2000000003. Miroslav Horáček The system meet the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: NA. Device ue: NA. Eupected and ueupected Resuts: NA uer Impact: NA. Patient Impact: NA. SE NA Diagnostic performed by Engineer: Geometry is blocked repeatedly. Resoluion: Replacement of the Sute PC and FDCSIB control compuer was made. Testing - the device is fuly fuctional, repair did not affect constrution and fuctionality of the device. Parts Consued: 1.000000000000000000000000000000008000000000000000000000000FDCSIB. 1.0000000000000000000000000000000080000000051700000000,SuTE PC mDP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000012 00000000-Nov-2000000003. Jan Odstrčil The system meet the specification for the performed service and is retuned to ue. Follow u Requred: Replacement of the SutePC. Resoluion: Replacement of the Sute PC was made, software was reinstalled, settings backus were restored. Replacement of FDCSIB controller, backu restored. Fuctionality was checked. The system is fuctional withou the defect being detected. Repair did not affect constrution and fuctionality of the device. The system brouht into operation fuctional. Problem Description by Engineer: Cutomer Fuction/Role: NA. Device ue: NA. Eupected and ueupected Resuts: NA uer Impact: NA. Patient Impact: NA. SE NA Geometry is blocked repeatedly. ----------------------------------------------------------------------------- | 2023-11-06 20:52 | No
3424 | Yes | uLLstopped working | uLLstopped working |  | 2023-11-06 21:04 | No
3425 | Yes | Warning light is faulty - Sent 0 GFE 00 | Warning light is faulty - Sent 0d GFE 0/0 |  | 2023-11-06 21:35 | No
3426 | Yes | Battery does not charge. It doesn't work. | Battery does not charge. It doesn't work. | 000000000000000000000000001O##: 00000000000000000000000000100000000353537 29-Nov-2000000003 Medina, Ismael The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Diagnostic performed by Engineer: Does not tun on. Resoluion: Necessary to replace power suply and battery Follow u Requred: Continu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000300000000500000000Ghazwan Hueidi Deaf The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Follow u Requred: uDFM100000000 Defibrillator/Monitor CN32668000000002 Does not charge battery Does not 000000000000000000000000001k Assign FSE Check log and battery uLLIf youneed material, order (via CFC/Delaer): 1 u 989800000000190000000071 - Rechargeable Li-ion Battery Pack uIn the Electromedicine 000000000000000000000000001kshop - 1113300000000- Resoluion: uDFM100000000 Defibrillator/Monitor CN32668000000002 Does not charge battery Does not 000000000000000000000000001k Assign FSE Check log and battery uLLIf youneed material, order (via CFC/Delaer): 1 u 989800000000190000000071 - Rechargeable Li-ion Battery Pack uIn the 000000000000000000000000001kshop of electromedicine - 1113300000000- Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A ----------------------------------------------------------------------------- | 2023-11-06 21:46 | No
3427 | Yes | 0 - u0 - Failure on the defibrillation | 00- u00 - Failure on the defibrillation | 000000000000000000000000001O##: 0000000000000000000000000010000000035300000000 Gullaue PANET The system does not meet the specifications for the performed service and is not retuned to ue. Resoluion: Hello, We tested with a new paddle and it persists. Kind regards Juien Duarquz Resoluion: u DuARQuZ, Thank youfor you sending. Coud youredo an operating test, with another set of paddles if available, otherwise with a test resistance if available? This will allow to confirm if the device is permanently fauty. Several messages are repeated in the logs: (7:30000000000000000 Pacer RFu HV uuuitor Energy LOW (89.85u) This message concerns the condenser with a measued valu < 90000000000000000. This caues problems for the device. In this case, the replacement of the condenser is recommended. (7:38) Defib Test Failue This message concerns the THERAPY management board which is defective and prevents the correct operation of the device. If you tests confirm again the presence of these uLL, and a failue of the defibrillation test, then the condenser and/or therapy board are fauty. ufortuately, becaue of the obsolescence of the device (12/31/2000000002), we will not be able to repair it. Indeed, to date, no part is available anymore. Sincerely, Gullaue PANET Resoluion: Hello, Here are the logs as requsted previouly. The operating tests were done with the paddle and it freezes at the defibrillation step. Good reception, Kind regards Juien Duarquz Resoluion: Hello u DuARQuZ, I will contact youfuther to the opening of the case on the uL+defibrillator. Coud youtell me which step the operating test does not pass? 000000000000000000000000001ld youhave a copy of this test to send me? How is this test done? With the paddles or with the charge resistance? Coud youalso forward me the log boues of the device? To access service mode: 1 Tun the therapy buton on monitor. 2 Press the menuselection buton to display the main menu 3 Navigate to Other and press the menuselection buton. 00000000n the Other menu select Service/Maintenance. The message is displayed: Euit the clinical mode. 5 Navigate util Yes and press the Selection Menubuton to confirm the message Euit Clinical Mode? Youare prompted to enter a pass000000000000000000000000001d. 6 Enter the pass000000000000000000000000001d (27689) by scrolling throuh the list util the desired nuber is highlighted. 7 Press the MenuSelection buton to enter each digit. 8 Select Completed when youentered all the figues. 9 Press the MenuSelection buton to display the main menuof the service mode: To consut the uLLlogs: 1 Select Hardware uLLLog or Software uLLLog in the main service menuand press the menuselection key. 2 Press again the MenuSelection buton to print, euport or delete the log. In the couse of waiting, Sincerely, Gullaue PANET Resoluion: u DuARQuZ, Thank youfor this additional test. The device is therefore permanently defective, ufortuately. Do youwant an obsolescence letter to indicate that the device cannot be repaired? Sincerely, Gullaue PANET Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the operating test of the defibrillator. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Reguatory Qustions: No/No ----------------------------------------------------------------------------- Internal Remarks: The cutomer called again and wants to be contacted as soon as possible. ----------------------------------------------------------------------------- | 2023-11-06 21:47 | No
3428 | Yes | QACASE 000000interface has been in the alarm, cannot be adjued alarm - Final 07 | QACASE 00000000 interface has been in the alarm, cannot be adjued alarm - Final 0/7 |  | 2023-11-06 22:00 | No
3429 | Yes | [Q] Problems with saving the patient data | [Q] Problems with saving the patient data | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000November 13, 2000000003 Maris, Ezra The system meets the specification for the performed service and is retuned to ue Resoluion: When sending the patient data, the Alua cannot send the Strutue Dose Report since the RDSR node is not reachable on the hospital side. On the Alua, the RDSR node is set u correctly. Patients cannot be closed auomatically by the Alua since the Alua only closes them when it has sent the Strutue Dose Report. As mentioned, this is not possible. Becaue of this, patients get stuk in the list and at some point the Alua complains abou this (too little memory). In-houe IT will now contact the euternal IT company Deadalu and commission them to see what has to be changed with the bulding so that the RDSR node 000000000000000000000000001ks. Problem Description by Engineer: Problem description by engineer: RMT - IPu Euposue not possible. Image disk ful 00000000-Nov-2000000003 1000000000000000036 RMT - IPu Free disk space is too low. November 6, -2000000003 1000000000000000059 u Kowski is on the way. +000000005100000000726000000006 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000087 Robert Kern The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: RMT - IPu Euposue not possible. Image disk ful 00000000-Nov-2000000003 1000000000000000036 RMT - IPu Free disk space is too low. November 6, -2000000003 1000000000000000059 u Kowski is on the way. +000000005100000000726000000006 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: Tried to access the system remotely to repair the database. Port is not released. Euplained to cutomer how he can delete manully. Internal Comments: uLL# and/or description of uLL Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu Free disk space is too low: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Repair Action: Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Free disk space is too low: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Follow-u requred: uLL# and/or description of uLL Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu Free disk space is too low: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Repair Action: Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Free disk space is too low: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. What are the neut trouleshooting steps: TE service to check and repair database if necessary, recreate and decision abou ugrade. 2 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Cutomer comment: u Kowski will clarify in advance in the bulding with the colleagu, he asks for call back on 00000000.11. aroud noon. Cutomer comment: ### Appointment arranged with Ms. Fischer for 13.11. 13:00000000 hrs. ----------------------------------------------------------------------------- | 2023-11-06 22:01 | No
3430 | Yes | M0A / u00 uLLmessage on the test. - | M00A / u000/ uLLmessage on the test. - | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000300000000ssad MAJEED The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Recommendation: 989800000000100000000800000000 0000000056000000003851 Requst for qute for suply of parts sent to the qutation department. Follow u Requred: Qute Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer wants to receive a qute to order the maintenance kit of the CO2 modue of the uuudefibrillator. ----------------------------------------------------------------------------- | 2023-11-06 22:27 | No
3431 | Yes | Geometry Unavailable | Geometry Unavailable | 000000000000000000000000001O##: 0000000000000000000000000010000000035520000000000000000-Nov-2000000003 WEINGART, SASHA The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Geo uavailable on morning start u. Reboot usucessfu. RMT - POS: No Lu nodes detected, check GEOCAN connections 00000000-Nov-2000000003 00000000:38:200000000 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Failue of geometry to boot u properly uer Impact: Room fails to boot u Patient Impact: None SE 8.2.3000000005 Resoluion : Installed new low voltage power suply and verified the voltage oupu was correct. Tested, verified normal operation; system ready for ue. Parts Consued : 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000035000000001 00000000-Nov-2000000003 WEINGART, SASHA The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Geo uavailable on morning start u. Reboot usucessfu. RMT - POS: No Lu nodes detected, check GEOCAN connections 00000000-Nov-2000000003 00000000:38:200000000 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Failue of geometry to boot u properly uer Impact: Room fails to boot u Patient Impact: None SE 8.2.3000000005 Diagnostic performed by Engineer : Started troule shooting and foud the low voltage power suply for the R cabinet was dead, 2300000000AC going in bu no DC voltage coming ou. Ordered replacement part will retun tomorrow morning to complete repair. System down in meantime. Follow u Requred : Replace the low voltage power suply Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000353689 Santos Ken The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Geo uavailable on morning start u. Reboot usucessfu. RMT - POS: No Lu nodes detected, check GEOCAN connections 00000000-Nov-2000000003 00000000:38:200000000 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Failue of geometry to boot u properly uer Impact: Room fails to boot u Patient Impact: None SE 8.2.3000000005 Diagnostic performed by Engineer : uLLlogs show hospital mains issus late last week. uS is connected to this system. This morning at 0000000000000000 system geometry fails to boot u despite several attempts to reboot. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - POS: No Lu nodes detected, check GEOCAN connections: 1. No Lu nodes were detected duing start-u. Repair Action : Service actions for RMT - POS: No Lu nodes detected, check GEOCAN connections: 1.1. Check SPP power suply. 1.2. Check GEO can connections. Resoluion : Sending to the field for onsite Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - POS: No Lu nodes detected, check GEOCAN connections: 1. No Lu nodes were detected duing start-u. Repair Action : Service actions for RMT - POS: No Lu nodes detected, check GEOCAN connections: 1.1. Check SPP power suply. 1.2. Check GEO can connections. Technician Name : Santos Ken ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action Taken: Follow u activity for FSE Sasha Weingart for 11/700PM Oucome: dispatched Internal Remarks : Planner Initials: GM Action Taken: FSE Sasha Weingart accepted call for 11/600PM Oucome: dispatched ----------------------------------------------------------------------------- | 2023-11-06 23:02 | No
3432 | Yes | Battery failure | Battery failure | 000000000000000000000000001O##: 0000000000000000000000000010000000035000000000000000000000000ov-2000000003 DAMIANI, MAuIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Following ordinary maintenance, foud that battery is flat and needs to be replaced uer Impact: no Patient Impact: no SE na Resoluion: Replaced battery. The charge test was carried ou with the positive oucome. Parts Consued: 1.000000000000000000000000989800000000626900000000 Assy Battery Li-Ion, 11ah, V600000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000059 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: u Arianna Camponogara reports that following ordinary maintenance the battery was foud to be flat and therefore needs to be replaced. Order battery and schedue technician for the replacement. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Following ordinary maintenance, foud that battery is flat and needs to be replaced uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u. Arianna Camponogara reports that following ordinary maintenance the battery was foud flat and therefore needs to be replaced. ----------------------------------------------------------------------------- | 2023-11-06 23:07 | No
3433 | Yes | ImagePC defective with the result that main fus trigger and system does not work | ImagePC defeu with the result that main fus trigger and system does not work | 000000000000000000000000001O##: 00000000000000000000000000100000000353000000001 November 7, 2000000003 Schut, Andreas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: u Device ue: u Eupected and ueupected Resuts: does not start u uer Impact: high Patient impact: low SE 8.1.17 Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000022.ClarityIQ_II IP PC no HDD Resoluion: ImagePC replaced and installed, fuction test via service software OK. ----------------------------------------------------------------------------- Internal Remarks: Since the order was barely on the cu-off, asked SPS Tier1 to initiate taui requst. ----------------------------------------------------------------------------- | 2023-11-06 23:17 | No
3434 | Yes | ucable is broken off into the port | ucable is broken off into the port | 000000000000000000000000001O##: 0000000000000000000000000010000000035000000007 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: ucable broken off in uconnector. Diagnostic performed by Engineer : Problem Description: ucable broken off in uconnector. ucable is broken off inside the connector. Form ued: Trouleshooting/Investigation Object: Tempu LS Not Object: Software Version: NA Date and Time of Issu: 11/6/2000000003 uer Actions: Testing. Where is the Deviation: Tempu LS Trouleshooting Steps Taken: NA Was PRS ued: No PRS Permission: Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.Ordered replacement Tempu LS 000000005600000000181. Follow u Requred : Tu A.J. Rose 1151 Alma Dr. Luberton, Tu 77657 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000035000000005 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Cutomer comment : No patient was involved in this case. Happened duing testing. Internal Remarks : New Case Caller: AJ ROSE IP & Issu: ucable is broken off into the port;10000000030000000039 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created and roued r000000000000000000000000001to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : ##d cutomer, asked for log files. ----------------------------------------------------------------------------- | 2023-11-06 23:51 | No
3435 | Yes | image disk error | image disk error | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000mjad Aburqou The system meets the specification for the performed service and is retuned to ue Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Free disk space is too low.: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Root Caue for RMT - IPu Image Disk write problem. Images possibly lost.: 1. Image Disk write problem. Images possibly lost. Root Caue for RMT - IPu Fluro storage not possible du to an Image disk problem (>Rel7.2): 1. Flu storage not possible du to an Image disk problem (>Rel7.2) Repair Action : Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu Image Disk write problem. Images possibly lost.: 1.1. Replace image disk. Service actions for RMT - IPu Fluro storage not possible du to an Image disk problem (>Rel7.2): 1.1. Carry ou a Re-create image store via FSF. 1.2. If problem still persists replace Imagedisk(s) Problem Description by Engineer : Problem description by engineer : RMT - IPu Free disk space is too low. 00000000-Nov-2000000003 00000000:17:16 RMT - IPu Image Disk write problem. Images possibly lost. 00000000-Nov-2000000003 10000000056:39 RMT - IPu Fluro storage not possible du to an Image disk problem (>Rel7.2) 00000000-Nov-2000000003 12:00000000100000000 Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 000Resoluion : Access to the system remotely and reload the IPPC OS and SW, recreate image store to hard disks, after that system retun 000000000000000000000000001king fine withou any problem. Diagnostic performed by Engineer : Image Disk write problem. ----------------------------------------------------------------------------- | 2023-11-06 23:51 | No
3436 | Yes | This device is DOWN becau its CHARGING PORT is NOT working | This device is DOWN becau its CHARGING PORT is NOT working | 000000000000000000000000001O##: 0000000000000000000000000010000000035000000008 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Resoluion : Device was repaired by u technician Chongolnee, Roongruchai with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged and to correct the issu we will need to replace the Front Case Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000035000000002 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: biomed reports the charging port is not 000000000000000000000000001king on the device, sending to bench Diagnostic performed by Engineer : Problem Description: biomed reports the charging port is not 000000000000000000000000001king on the device, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form Follow u Requred : Problem Description: biomed reports the charging port is not 000000000000000000000000001king on the device, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repair Confirmed delivery address ##d bench repair form CuT NEEDS A SHIPPING LABEL ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Christopher BARON IP & Issu: " 86950000000000000000his device is DOWN becaue its CHARGING PORT is NOT 000000000000000000000000001king" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: PCCI Action Taken: I created a case, CSP, R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-07 00:11 | No
3437 | Yes | Pressure Sensor Autozero Failed | Pressure Sensor Autozero Failed |  | 2023-11-07 00:27 | No
3438 | Yes | SBC uLLCode | SBC uLLCode | 000000000000000000000000001O##: 00000000000000000000000000100000000350000000000000000Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Calibration only, no parts requred 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits SD card is corruted and the SD card will be suplied by the Bu Device will only requre Calibration for NBP, CO2 and Touh Screen. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000350000000000000000Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE 00000000.300000000roblem Description/ueupected Resuts: tech reports an uLLdisplayed on the Tempu Pro - “SBC dll not loaded”, uer was forced to shu it down the uit, sending to bench Diagnostic performed by Engineer : Problem Description: tech reports an uLLdisplayed on the Tempu Pro - “SBC dll not loaded”, uer was forced to shu it down the uit, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair formCuT NEEDS A LOANER Follow u Requred : Problem Description: tech reports an uLLdisplayed on the Tempu Pro - “SBC dll not loaded”, uer was forced to shu it down the uit, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repair Confirmed delivery address ##d bench repair form CuT NEEDS A LOANER ----------------------------------------------------------------------------- Cutomer comment : 00000000300000000000000000000000021LW8QAM File uloaded by CS Portal uer : File Name: Tempu Image.jpeg. Comments: File(s) uloaded duing case creation Internal Remarks : 11/13/2000000003 Requsted Loaner has been shipped to Cut with FedEu Tracking# 716991710000000081. ----------------------------------------------------------------------------- | 2023-11-07 00:51 | No
3439 | Yes | M0A / u00/ Cannot pass the self-test (QA#00000. | M00A / u000 / Cannot pass the self-test (QA#0000000. |  | 2023-11-07 00:56 | No
3440 | Yes | Touchscreen not responsive in bottom right corner | Touchscreen not responsive in bottom right corner | 000000000000000000000000001O##: 0000000000000000000000000010000000035000000008 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:touhscreen not responsive in bottom right corner uer Impact: not uable needs repair Patient Impact: none SEun Resoluion : Provided part nubers for touhscreen and lcd, advised to check lcd cable connections when replacing the touhscreen and replace lcd or lcd cable if flickering still occus, cutomer acknowledged and will order parts, issu resolved- V600000000touhscreen pn 00000000561539000000001 V600000000lcd 2nd gen pn 00000000561528891 Material Nuber 0000000056153600000000 - RP-Cable, LCD, NEC, 2nd Gen, V600000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states duing set u the touhscreen is not responsive in the bottom right corner of display bu touhscreen cal passes, no patient involvement, also there is a flickering at the top of alarm message banner, requsts part nuber to correct issu. ----------------------------------------------------------------------------- Internal Remarks : neut Step Action: New Case Caller: Victoria McKowsky 230000000070000000012 IP & Issu: 100000000295500000000/Part ID. Touhscreen is not 000000000000000000000000001king. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-07 01:11 | No
3441 | Yes | Alarm or Messages/System did not alarm when it was in u on a patient/ur Error | Alarm or Messages/System did not alarm when it was in u on a patient/ur Error | 000000000000000000000000001O##: 0000000000000000000000000010000000035510000000000000000-Nov-2000000003 Aasim Ali The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Mu800000000 alarms. uer Impact: No impact Patient Impact: No impact SE N.00000000.00000000 Problem Description/ueupected Resuts: Mu800000000 and u2 did not alarm Diagnostic performed by Engineer : Went onsite and collected logs from PIC iu. Performed test and verification with CS Phipps, Suan. Resoluion : Collected Logs and uloaded to notes and attachment ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003500000000 Edmond Omar The system meets the specification for the performed service and is retuned to ue Resoluion : Problem Description: Mu800000000 and u2 did not alarm This device is OPERATIONAL, bu it did NOT alarm when it was in ue on a patient. Patient died. Form ued: Trouleshooting/Investigation Object: An Mu800000000 and u2 Not Object: Other Mu800000000 and u2's. Software Version: N.00000000.00000000 Date and Time of Issu: 11/000000002000000003 uer Actions: uu Where is the Deviation: Cutomer claims device did not alarm. Patient died. Trouleshooting Steps Taken: Biomed tested the Mu800000000 on a simuator and all vitals alarmed.Claim was that apnea did not alarm. Was PRS ued: No PRS Permission: Follow u Requred : Location Name Wellstar Kennestone Hospital Street 677 Chuch St NE City MARIETTA State GA Zip 3000000006000000001100000000Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Mu800000000 alarms. uer Impact: No impact Patient Impact: No impact SE N.00000000.00000000 Problem Description/ueupected Resuts: Mu800000000 and u2 did not alarm Diagnostic performed by Engineer : Problem Description: Mu800000000 and u2 did not alarm This device is OPERATIONAL, bu it did NOT alarm when it was in ue on a patient. Patient died. Form ued: Trouleshooting/Investigation Object: An Mu800000000 and u2 Not Object: Other Mu800000000 and u2's. Software Version: N.00000000.00000000 Date and Time of Issu: 11/000000002000000003 uer Actions: uu Where is the Deviation: Cutomer claims device did not alarm. Patient died. Trouleshooting Steps Taken: Biomed tested the Mu800000000 on a simuator and all vitals alarmed.Claim was that apnea did not alarm. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Roderick TAGGuG IP & Issu: " 58318000000006 This device is OPERATIONAL, bu it did NOT alarm when it was in ue on a patient" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: I created a case, R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-07 01:31 | Yes
3442 | Yes | Touchscreen not working properly, display alignment issue | Touchscreen not working properly, display alignment issue | 000000000000000000000000001O##: 0000000000000000000000000010000000035000000009 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: touhscreen not 000000000000000000000000001king properly uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states that when attempting to make changes, the touhscreen isn't 000000000000000000000000001king properly. it is slightly off from where youactully touh. cutomer is requsting assistances performing touhscreen calibration. Informed cutomer how to perform touhscreen calibration. also informed cutomer how to confirmed calibration with touhscreen diagnostic mode. Cutomer states he will attempt to perform touhscreen and callback if futher assistance is needed. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: michael tassler IP & Issu: 1000000000000000087/display alignment issu Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: New Case Caller: michael tassler IP & Issu: 69163999/display alignment issu Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-07 01:36 | No
3443 | Yes | standby btn is not active - will not go into standby mode | standby btn is not au - will not go into standby mode | 000000000000000000000000001O##: 00000000000000000000000000100000000350000000000000000Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: will not go into standby mode uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 989800000000617100000000- flow sensor assembly - 1217.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states device will not access standby mode. Troule shot with cutomer. had cutomer check in service mode the air flow slpm. cutomer states device is showing 15 with nothing set. Informed cutomer that it shoud only be either -1 or 1 anything else is considered failing and device will need to have flow sensor assembly replaced. cutomer is requsting part nuber and price for replacement flow sensor assembly. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: jamal IP & Issu: 5660000000077 cannot get into standby Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: Created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-11-07 01:41 | No
3444 | Yes | At 00m the entire system went down. Onsite service pending | At 0:0am the entire system went down. Onsite service pending | 000000000000000000000000001O##: 00000000000000000000000000100000000355000000006 00000000-Nov-2000000003 Angel Ramos The system meets the specification for the performed service and is retuned to ue Follow u Requred : Root Caue Analysis with Hospital IT team. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Hosts shoud not lose connection uer Impact: No impact Patient Impact: No impact SE C.00000000.00000000 Problem Description/ueupected Resuts: CuWANTS RCA - Host screens went white - lost connection Resoluion : Follow u visit created, date is set for the 200000000 bu I will contact cutomer and arrange a sooner date. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000350000000000000000rtega, Carmen The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : CuWANTS RCA - Host screens went white - lost connection Diagnostic performed by Engineer : Problem Description: CuWANTS RCA - Host screens went white - lost connection Form ued: Trouleshooting/Investigation Object: All hosts Not Object: Bedsides Software Version: C.00000000.00000000 Date and Time of Issu: 11/6/2000000003 11:00000000AM uer Actions: uu Where is the Deviation: Hospital wide Trouleshooting Steps Taken: Remoted onto primary server. All hosts are fuctional, cutomer wanted to know if why host screens went white. Checked log viewer. Foud some "software udate" message and "time" messages. Informed cutomer that RCA 000000000000000000000000001ld requre on site technician. CuWANTS RCA Was PRS ued: Yes PRS Permission: Biomed - Thomas https://pww.portal.rs.healthcare.u.com/PrsPortal/defaut.aspu?PrsID=282171 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Hosts shoud not lose connection uer Impact: No impact Patient Impact: No impact SE C.00000000.00000000 Problem Description/ueupected Resuts: CuWANTS RCA - Host screens went white - lost connection ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: thomas helms IP & Issu: 87000000007200000000At 11:00000000am the entire system went down/cutomer requsted tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: thomas 682-2800000000756 IP & Issu: case 0000000022558500000000 wants to create onsite 000000000000000000000000001 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: advised onsite 000000000000000000000000001has already been created and a PO was provided with PO# MA000000001-0000000000000000181528 advised "Qute creation in progress" placed note in neut step action Please call Thomas at 682-2800000000756 for scheduing they need service asap.. thanks Tierra V 11/6/23 00000000pm est Neut Step Action: Awaiting Qute Verbal PO Nuber: MA000000001-0000000000000000181528 Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : neut Step Action: Cutomer Call Back Caller: Thomas 682280000000056 IP & Issu: 0000000022558500000000/CCB for scheduing Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Sent AA to RP Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-07 01:41 | No
3445 | Yes | the Patel is sticking | the Patel is sticking | 000000000000000000000000001O##: 000000000000000000000000001000000003000000007 00000000-Nov-2000000003 Mandeville Nathan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: cutomer says that the footswitch is sticking uer Impact: cannot fluro Patient Impact: uknown uuResoluion : Euamined the footswitch for debris and cleaned off the micro switch and pedal. No futher issus according to the cutomer. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jason Serrano IP & Issu: 812925600000000the Patel is sticking Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE/ contacted on call RE --- adc v cut of callback Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-07 01:46 | No
3446 | Yes | Eut. 0 - the system is slow and does not respond - system is slow and its not answering | Eut. 00 - the system is slow and does not respond - system is slow and its not answering | 000000000000000000000000001O##: 000000000000000000000000001000000003500000000000000007-Nov-2000000003 Vijoui, Negar The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Coronary Eupected and ueupected Resuts: system software mut fuction / review is down. uer Impact: No impact. Patient Impact: No impact SE 2.2 Resoluion: Following PC, it is replaced. The software is reinstalled. The problem is solved. Following PC is replaced. Software is reloaded. Problem solved. Parts Consued: 1.0000000000000000000000000000000080000000051700000000, FOLLOWING MDP PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000097 00000000-Nov-2000000003 Vijoui, Negar The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: Following PC replacement Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Coronary Eupected and ueupected Resuts: system software mut fuction / review is down. uer Impact: No impact. Patient Impact: No impact SE 2.2 Diagnostic performed by Engineer: The pc sut software is broken down. The software re-installation did not solve the problem. The Following PC will be replaced. Following PC software is down. Software reload did not solve the issu. Following PC will be replaced. ----------------------------------------------------------------------------- Internal Remarks: Negar - u"review" does not 000000000000000000000000001k, following PC bootu issu, FSE to reload following PC software. Internal Remarks: FSE Negar wants to know if she can receive the part tomorrow. Transferred her to POD Internal Remarks: Negar - requsted new license key for Following PC. Internal Remarks: FSE Negar Vijoui. FSE reloaded the Suter PC. Review monitor has u logo and black screen. Disconnected u300000000rom Sute PC and the display went blank. Video path is good. Reset following PC and nothing coming u. Approved following PC replacement. 0000000080000000051700000000 - FOLLOWING MDP PC ----------------------------------------------------------------------------- | 2023-11-07 02:11 | No
3447 | Yes | Bru head "snapped in 0while bruing. No injury. | 00u Bru head "snapped in 0 while bruing. No injury. | 2000000003-12-00000000 19:28:00000000***Claims*** ##d cu 3rd reminder. Awaiting response.;2000000003-11-22 18:0000000053:***claims*** ##d cu to follow u with safety qustions. Awaiting response.;2000000003-11-00000000 100000000:00000000:***claims*** ##d cu with safety qustions. Awaiting response. | 2023-11-07 04:36 | No
3448 | Yes | Lost power to u-ray system | Lost power to u-ray system | 000000000000000000000000001O##: 0000000000000000000000000010000000035668000000008-Nov-2000000003 Stockton Wayne The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Lost power to u-ray system How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Resoluion : Replaced MPD power suply and tested operations normal. Ran diagnostics on the system passing. Cleaned u supended euams in archive quu. Ran level one diagnostics was passing. Parts Consued : 1.00000000000000000000000000000000230000000028523,PFS272 POWERTRAY FRuTools & Equpment : MuTIMETER,200000000000000001-30000000000000000:00000000:00000000 OSCILLOSCOPE,2000000003-11-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000355870000000000000000-Nov-2000000003 Lim, Chris The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace Power Tray. System Code 722000000006 Problem Description by Engineer : Problem description by engineer : Lost power to u-ray system How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : System Code 722000000006. Performed trouleshooting with FSE Wayne Stockton. Bypassed uS, and gained power back, bu uable to power u the main PC. Power Tray is receiving power. Replace power tray. Tools & Equpment : MuTIMETER,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000381365 00000000-Nov-2000000003 Stockton Wayne The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Lead u-ray tech Roger Device ue: uLroom 1 Eupected and ueupected Resuts: Power u normally uer Impact: Delayed Patient Impact: Delayed SE 7.2.16 Diagnostic performed by Engineer : Helped Chris Lim trouleshoot problems with schematics and testing. Helped place uS in by-pass. Foud no oupu power from MPD fan tray power suply. Recommended replacement. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003558500000000Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Lost power to u-ray system How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : Roger says they heard a "pop" and then a buzing noise in the equpement room. They are curently uable to boot the system and the table is not 000000000000000000000000001king. He also says there might be some power monitoring equpment hooked u to the system at present time. Internal Comments : uLL# and/or description of uLL: no uLLdetermined. Repair Action : created qute with cutomer provided PO Follow u Requred : uLL# and/or description of uLL: no uLLdetermined. Repair Action : created qute with cutomer provided PO What are the neut Trouleshooting steps : Determine power issu and souce of loss Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action Taken: FSE Wayne Stockton accepted call for 11/600PM already spoke to cutomer Oucome: dispatched Internal Remarks : Planner Initials: GM Action Taken: FSE Chris Lim requsted call to be transferred to him for 11/600PM Oucome: dispatched Internal Remarks : FSE Chris Lim. FSE calling to get part nuber for Marathon uS in Teal. Foud a Marathon nuber 000000001177 and sugested to try All Parts for part. ----------------------------------------------------------------------------- | 2023-11-07 04:56 | No
3449 | Yes | u will not start | 0CS:YO000 and will not start | 000000000000000000000000001O##: 00000000000000000000000000100000000355923 00000000-Nov-2000000003 Ishihara, Yuki The system meets the specification for the performed service and is retuned to ue Resoluion: Parts ued: 0000000080000000012200000000 u Haswell Fleuvision 2 PC no HDD 0002230000000080000000091 1 uit Follow u Requred: N/A Follow u Requred: N/A Reguatory Qustions: N/A Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD Resoluion: The system was stuk at 00000000:00000000 and 000000000000000000000000001ld not boot u. As a resut of the investigation, it was confirmed that there was an uLLin the large monitor control PC (Fleuvison PC) in the euamination room. We have replaced the casing of FleuvisonPC. We have reinstalled the OS and software. After the 000000000000000000000000001k was completed, we performed an operation check and foud no abnormalities. Problem Description by Engineer: Does not start at 0000000000000000 Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: Does not start at 0000000000000000 uer Impact: No impact Patient Impact: No impact SE uknow ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000355911 Ohgua, Yui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Does not start at 0000000000000000 Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: Does not start at 0000000000000000 uer Impact: No impact Patient Impact: No impact SE uknow Diagnostic performed by Engineer: RC Contents: Euplanation on identifying the location of the failue We plan to respond on-site. Follow u Requred: * Internal Comments: [Suject] 00000000DS:YO2: u Nakagawa: Does not start at 0000000000000000 [Ou of stock parts] 0000000080000000050000000000000000Hard disk spare part QTY:1[GATP] ZSID Nagoya (u1000000000000000080000000012200000000:u Haswell Fleuvision 2 PC no HDD QTY:1 [GATP] ZREP/Traceable Nagoya (u100000000[Not issud parts] 0000000080000000058600000000:FleuVision-2_Revised PC no HDD QTY:1[GATP] ZREP/Traceable Nagoya (u100000000Reason for not issud:--None-- [Dispatch] 1 RC -----------------FSE registration information----------------- -Visit FSE1 Name Ishihara, Yuki [Yui Ishihara] - FSE1 Visit date and time 2000000003/11/0000000012:00000000 -FSE1 000000000000000000000000001king time 00000000Shipping instrutions] -----------------Shipping instrutions #1---------- ------- -Delivery conditions CPu(no warehoue pick-u/transport arrangements) -Delivery date Mut arrive 2000000003/11/00000000 10000000000000000 -Delivery address: Nagoya warehoue/pick-u by FSE -Recipient name: Yui Ishihara -0000000000000000007596-1752 -Other Mobile phone call uon arrival -Requster name HirotsuuOkua [Inspection category] Other ----------------------------------------------------------------------------- | 2023-11-07 05:21 | No
3450 | Yes | HS0ED is emitting the triple chirp | HS0AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000355937 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: (Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 12 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A11H-00000000173 Battery Install Before# 2000000009/6/300000000 Lot#00000000000000001P Pads Eup Date: 2000000005/00000000 Pads Lot # C10000000052200000000 ----------------------------------------------------------------------------- | 2023-11-07 05:36 | No
3451 | Yes | SA:0/u FD0Carm cannot move | SA:0R/u FD0Carm cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 00000000.11.2000000003 00000000:00000000:32 uC (8600000000600000000) Service content (requred): The cutomer reported that the C-arm coud not move, and the AMC power cable was reprocessed and retuned to normal. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None Internal Remarks: 00000000.11.2000000003 00000000:00000000:35 uC (8600000000600000000) "RDI Accuacy: Yes" $Q7: (Confirmation & Description of problem/C-arm not move.) $Q8: (Onsite reviewed the uLLlog, "Application uLL Motor assembly uLL[auis=Frtprop]", The HV voltage (not logic power) is below the minimu SOFTWARE specified voltage, Check the amc power suply euception.) $Q9: (Reinsert the AMC power cable, then retuned to normal, system OK) $Q100000000 (Test system OK and inspection data refer to: PA. Device retuned to ful fuctionality. ) $Q11:Calibrated Tools:N/A SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: SA:00000000R/u uuuFD200000000 Safety Qustions: Case: 000000002250000000095 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the C-arm cannot move and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: C-arm cannot move Remote diagnosis and test resuts: Restart has no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the L arm Long direction motor is normal. 2. Check whether the EtherCAT commuication is normal. 3. Check whether the AMC is normal. 00000000Check the machine uLLand perform faut repairs. Recommended spare part uuand name: 00000000800000000262732 CLEA2 LONGITuINAL DRIVE ASSY &00000000800000000990000000051 AMC Triple Servo Drive HP-LP-LP" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 13:0000000000000000uC (8600000000600000000)"SuT. Ready for Clinical ue local time" : " 2000000003-11-00000000 15:00000000" SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-07 06:31 | No
3452 | Yes | SELF TEST FAILED | SELF TEST FAILED | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 000000002250000000091 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : CM - SELF TEST FAILED ----------------------------------------------------------------------------- | 2023-11-07 07:01 | No
3453 | Yes | Failing self test | Failing self test | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 0000000022500000000300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable Problem Reported by Cutomer : CM SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-07 07:16 | No
3454 | Yes | SELF TEST FAILED | SELF TEST FAILED | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : CM - SELF TEST FAILED+Battrey Issu Safety Qustions : Case: 000000002250000000039 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-07 07:16 | No
3455 | Yes | Failing self test | Failing self test | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : CM SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 000000002250000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable ----------------------------------------------------------------------------- | 2023-11-07 07:17 | No
3456 | Yes | AED does not pass self-test | AED does not pass self-test | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : CM Safety Qustions : Case: 0000000022500000000600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-07 07:21 | No
3457 | Yes | the 0didn't rotate | 0DM:YO00, the 0 didn't rotate | 000000000000000000000000001O##: 0000000000000000000000000010000000035650000000000000000-Nov-2000000003 Wakayanagi, Koji The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Not ued. Eupected and ueupected Resuts: Cannot do rotational shooting uer Impact: No effect Patient Impact: No effect SER8.1.3000000005 Resoluion: 00000000000000000000000000100000000357396 000000000000000000000000001k report and received the cutomer's signatue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000357396 00000000-Nov-2000000003 Motoyama, Yojiro The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Parts Consued: 1.00000000000000000000000000000000230000000017000000008,ROTATION DRIVE H.D. Resoluion: At RAO 1200000000degrees, there was a log record of the safety mechanism fuctioning and the arm not moving futher. The same log occured yesterday and today, and it did not occu before the part was replaced on October 19, so I juged that it was an initial failue of the part that was replaced on the 19th so I again replaced the motor incluing the position detector. After I replaced it, I adjuted the system and tested the 3D rotation. Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: In ue Eupected and ueupected Resuts: 3D did not rotate. uer Impact: Yes Patient Impact: None SE8.1.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000356660000000000000000-Nov-2000000003 Abe, Tomoki The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Resoluion: I reported this to a different 000000000000000000000000001 Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: In ue Eupected and ueupected Resuts: Cannot rotate uer Impact: No effect Patient Impact: No effect SE8.1.300000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000035600000000 Ohgua, Yui The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Internal Comments: [Suject] 00000000DM:YO2:u Hirono: The 3D 3D did not rotate. [Internal Comments] 2000000003/11/7 10000000033:00000000 soft-collision detected 2000000003/11/7 10000000033:00000000 CLE_PMD.prop:SP violated 2000000003/11/7 10000000033:00000000 Safepath violated by the profilemovement itself [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Wakayanagi, Koji [Koshi Wakayanagi] - FSE1 Visit date and time 2000000003/11/0000000018:00000000 - FSE1 000000000000000000000000001k time 00000000 [Euamination category] Normal euamination Diagnostic performed by Engineer: RC Response: I euplained how to identify the failue. We plan to respond on site. Problem Description by Engineer: The 3D 3D did not rotate. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The 3D 3D did not rotate. uer Impact: No effect Patient Impact: No effect SE uknow ----------------------------------------------------------------------------- | 2023-11-07 07:51 | No
3458 | Yes | The system can’t turn-ON | The system can’t turn-ON | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000 13-Nov-2000000003 CHATCHAI SANTANA The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: Not in clinical ue. Eupected and ueupected Resuts: The system can’t tun-ON uer Impact: No impact. Patient Impact: No impact. SE 1.2.5 Diagnostic performed by Engineer : PDuFanTray and DCPS was defect. Resoluion : Replace PDuFanTray and DCPS. Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS ----------------------------------------------------------------------------- | 2023-11-07 08:21 | No
3459 | Yes | SA:0/u FD00System cannot boot up | SA:0R/u FD00System cannot boot up | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Overview of cutomer problem: The device couts down to 00000000:00000000 and cannot enter the system normally. I restarted it three times, bu it still doesn't 000000000000000000000000001k.Engineers are requred to be present. Is the faut caued by the equpment itself? Yes Trouleshooting Action: Remote diagnosis and test resuts: The device couts down to 00000000:00000000 and cannot enter the system normally. I restarted it three times, bu it still doesn't 000000000000000000000000001k.Requres engineer presence PRS TR Statu: No PRS Repair Action: Recommended on-site repair steps: Check whether IPPC or IPC is not tuned on, check whether the PC's power suply or PC's BIOS battery is discharged Cutomer demand notes: Onsite Problem Reported by Cutomer: SA:00000000R/u uuuFD100000000100000000Large screen display problem Euternal Remarks: 00000000.11.2000000003 00000000:0000000038 uC (8600000000000000001)"SuT. Ready for Clinical ue local time" : " 2000000003-11-00000000 17:300000000 [Service Content] Large-screen PC software crashes, reinstall the system and software and retun to normal. 【Service Resut】The machine operates normally after repair. 【Other matters】None yet Safety Qustions: Case: 0000000022500000000200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-07 08:31 | No
3460 | Yes | The system stopped during a photon. | 0CS:TK:The system stopped during a photon. | 000000000000000000000000001O##: 00000000000000000000000000100000000356000000002 Kamata, Takayui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: The system stopped duing a 3-bee-shot. u/Ms. Sada hiro uhin: 7722 ・Duing the operation of a patient weighing 1600000000g between 9:00000000 and 10000000000000000 yesterday (November 6). Some message was displayed. He 000000000000000000000000001ld like u to check the log. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The system stopped duing a 3-beg. u/Ms. Sada hiro uhin: 7722 ・Duing the operation of a patient weighing 1600000000g between 9:00000000 and 10000000000000000 yesterday (November 6). Some message was displayed. He 000000000000000000000000001ld like u to check the log. uer Impact: No effect Patient Impact: No effect uu00000000Internal Comments: [Suject] 00000000CS:TK:u Sadahiro: A roning stopped in the middle. [Internal Comments] ＝＝＝ 9558 2000000003-11-00000000 00000000:0000000000000000363 Command: StartEuposue [CHANNEL_INFO=u-Ray Channel Biplane] [Info=StartEuposue Euecue() "" -1] Information Normal Operation uEvent Application State Controller Igc 9582 2000000003-11-00000000 00000000:00000000500000000215 uer gudance: Ru aborted: Tue problem [Component=IGC_IGC] [SuComponent=1] [ID=5000000000000000085] [Priority=19] Information Normal Operation uerMessage iEngineCore Iec 9587 2000000003-11-00000000 00000000:00000000500000000327 urayService: Performance analysis [urayChannel=Lateral] [urayMode=Euposue] [RuTag=1330000000000000000000000009661723] [SeriesNuber=6] [RuLength=38] [ImageWidth=968] [ImageHeight=968] [ImageResoluionChanges=00000000 [ImageResoluionRecord=1:968u968] [TestShotsReceived=00000000 [ImagesReceived=38] [SuImagesReceived=38] [ImageLossOccurences=00000000 [StripLossOccurences=00000000 [AverageFrameRateFE=25.0000000000000000300000000ps] [AverageFrameRateIP=200000000989700000000fps] [MauInterImageTimeFE=000000000000000030000000000000000ms] [MauInterImageTimeIP=000000001350000000000000000ms] [AverageProcessingLoad=600000000632191 %] [MinImageLatency=23.3000000000000000000000000ms] [MauImageLatency=27.7380000000000000000ms] [AverageImageLatency=2000000000000000068 ms] [StdevImageLatency=00000000829768 ms] [AverageAdditionalLatency=00000000000000005568 ms] [MauAdditionalLatency=000000001510000000000000000ms] [MauAdditionalLatencyImageNuber=31] [ThresholdAdditionalLatency=100000000.000000000000000000000000 ms] [Dvlp2BuferSize=15728600000000bytes] Information Normal Operation Performance iEngineCore Iec 959000000002000000003-11-00000000 00000000:00000000500000000359 urayViewer: Performance analysis [urayChannel=Lateral] [RuTag=1330000000000000000000000009661723] [SeriesNuber=6] [MinImageLatency=200000000580000000000000000ms] [MauImageLatency=32.1190000000000000000ms] [AverageImageLatency=26.52000000005 ms] [StdevImageLatency=1.872231 ms] Information Normal Operation Performance iEngineCore Iec 9592 2000000003-11-00000000 00000000:0000000050000000000000000 urayService: channel independent irradiation data [PatientId=89cbda8ab800000000aadb00000000a8e85000000006f8975e1c3b21a67f298cb381576ee5d000000001cf] [EuaminationId=1b000000009e81a5687c86f00000000e2811a9bca10000000089f567d2622cfe00000000261c8bca700000000f9] [RuTag=1330000000000000000000000009661723] [PerformingPhysician=000000008fcb6a7b9e588cb0000000068d1153ee6ac1f000000003000000008700000000c00000000fd71600000000c80000000067] [PatientAge=30000000000000000years] [PatientWeight=1500000000175 kg] [ApplicationId=10000000000000000] [ProcedueGrouId=320000000000000000 [Epu=Cardio Clarity PCI(CAG)_New; Left Coronary_copy1; 25fps Mediu Special] [DateTime=11/00000000/2000000003 00000000:0000000000000000 [Duation=1519ms] [IrradiationType=EuposueStationary] [DoseControlMode=Cine] [ImageCout=38] [PatientType=LargeAdut] [Description=LVG 25 fps Mediu] [FluroFlavor=Normal15] [RoadmapPro=false] [TotalDAP=152162.100000000mGy*cm2] [TotalAK=1338.000000003mGy] [TotalFluroTime=391s] [TotalEuposueTime=300000000 [TableHeight=857mm] [TableLateral=-66mm] [TableLongituinal=6500000000m] Information Normal Operation ApplicationEvent iEngineCore Iec 95900000000000000003-11-00000000 00000000:0000000050000000000000000 urayService: channel dependent irradiation data [PatientId=89cbda8ab800000000aadb00000000a8e85000000006f8975e1c3b21a67f298cb381576ee5d000000001cf] [EuaminationId=1b000000009e81a5687c86f00000000e2811a9bca10000000089f567d2622cfe00000000261c8bca700000000f9] [RuTag=1330000000000000000000000009661723] [CHANNEL_INFO=u-Ray Channel Lateral] [DAP=3289.1000000008mGy*cm2] [AK=25.3161mGy] [AnguationAngleStart=0000000000000000degrees] [AnguationAngleEnd=0000000000000000degrees] [RotationAngleStart=0000000000000000degrees] [RotationAngleEnd=0000000000000000degrees] [Detectoru=200000000mm] [DetectorY=200000000mm] [DetectorDiagonal=300000000mm] [ObjectThickness=00000000m] [Cuilter=000000001mm] [AlFilter=1.00000000m] [FocalSpot=000000008mm] [CollimatedHeight=2100000000] [CollimatedWidth=2100000000] [CollimatedArea=000000002] [Peak mA=580000000057mA] [ms=6.22ms] [Rate=25.00000000puse/s] [PuseCout=38] [Peak kV=125.00000000kV] [SID=1115mm] [ShuterLeft=00000000m] [ShuterTop=00000000m] [ShuterRight=00000000m] [ShuterBottom=00000000m] [WedgeAngle1=1800000000egrees] [WedgeDistance1=100000000mm] [WedgeAngle2=1800000000egrees] [WedgeDistance2=100000000mm] [BeamLongituinal=1621mm] [BeamLateral=00000000m] [CArcRotation=00000000egrees] [PropellorRotation=00000000egrees] [ZRotation=00000000egrees] [DetectorRotation=00000000egrees] [ImageWidth=968piuels] [ImageHeight=968piuels] [ValidRectLeft=5piuels] [ValidRectTop=5piuels] [ValidRectRight=9600000000uels] [ValidRectBottom=9600000000uels] Information Normal Operation ApplicationEvent iEngineCore Iec 9595 2000000003-11-00000000 00000000:0000000050000000000000000 uer gudance: Ru aborted: Tue problem [Component=IGC_IGC] [SuComponent=1] [ID=5000000000000000085] [Priority=19] Information Normal Operation uerMessage iEngineCore Iec 9600000000 2000000003-11-00000000 00000000:00000000500000000595 Ru Data [CHANNEL_INFO=u-Ray Channel Lateral][Info=[EndReason=Other][RuLength=2.20000000099990000000072522][NuberOfPuses=38][Focu=Largest][FirstFilamentCurent=5.0000000072937000000001719][LastFilamentCurent=5.0000000072000000000000000085][OldCFactor=1.00000000000000006629390000000017][NewCFactor=1.000000000000000068000000006700000000][TotalFilamentWear=000000002790000000000000000888700000000][DeltaFilamentWear=2.690000000038000000003700000000-00000000]] Information Normal Operation NoDetailInfo uRay_Control uGC 9600000000 2000000003-11-00000000 00000000:00000000500000000596 After ru data [CHANNEL_INFO=u-Ray Channel Lateral] [RuTag=1330000000000000000000000009661723] [EndReason=uGS_RuENDREASON_OTHER] [RuLength=2.2000000000000000000000000 [NrOfPuses=38] [Focu=uGS_FOCu_LARGEST] [FirstFilamentCurent=5.00000000729] [LastFilamentCurent=5.00000000725] [OldCFactor=1.000000000000000066] [NewCFactor=1.000000000000000068] [TotalFilamentWear=0000000027900000000] [DeltaFilamentWear=00000000000000000000000000000000] Information Normal Operation Interface Acqusition Igc From Generator device: PoInt EVR: Gate driver faut ＝＝＝ [Euamination category] Normal euamination Diagnostic performed by Engineer: I checked the log. - 9:000000005000000009:0000000012,9:0000000038 " From Generator device: PoInt EVR: Gate driver faut" occured 3 times. - After each 100000000photos, "Ru aborted: Tue problem" was displayed and the image stopped. - The u-ray conditions were almost Mmau. I think it was a malfuction caued by a high load. I reported the above to u Sada hiro, and he agreed to ue it as-is for now. ----------------------------------------------------------------------------- | 2023-11-07 09:21 | No
3461 | Yes | QA 000000- euternal paddles cannot be ud | QA 0000000 - euternal paddles cannot be ud |  | 2023-11-07 10:00 | No
3462 | Yes | Machine is not booting up while starting up | Machine is not booting up while starting up | 000000000000000000000000001O##: 00000000000000000000000000100000000356789 00000000-Nov-2000000003 Shelar, Ritesh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Host PC is not booting intermittently Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: yes Eupected and ueupected Resuts: eupected uer Impact: No Patient Impact: Not SE 8.2.17.00000000Tools & Equpment : MuTIMETER,200000000000000006-26 00000000:00000000:00000000 Resoluion : Replaced host pc cmos battery. Made system fuctional. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000356700000000 Kuar, Vipin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Host PC is not booting intermittently Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: Not provided SE 8.2.17.00000000Follow u Requred : uLL# and/or description of uLL: No uLL Repair Action : FSE to visit site and check CMOS battery of Host PC. if it is not charged well then replace the same and correct date and time in Host PC. Technician Name : Kuar, Vipin Internal Comments : uLL# and/or description of uLL: No uLL Repair Action : FSE to visit site and check CMOS battery of Host PC. if it is not charged well then replace the same and correct date and time in Host PC. Diagnostic performed by Engineer : Host PC is not switching on intermittently ----------------------------------------------------------------------------- | 2023-11-07 10:11 | No
3463 | Yes | QA 00000- uFailure | QA 0000000 - uFailure |  | 2023-11-07 10:12 | No
3464 | Yes | There was a problem with moving the L arm and C arm. | There was a problem with moving the L arm and C arm. |  | 2023-11-07 10:40 | No
3465 | Yes | QAAOCAP:0/u FD0Footswitch problem | QAAOCAP:0R/u FD0Footswitch problem | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: QAAOCAP:00000000R/u uuuFD100000000T2 Activities:Repair Action: Recommended on-site repair steps: Nemo/euposue foot brake/200000000to hospital Service Content: The euposue foot switch is intermittently malfuctioning and the butons do not reboud well. Please ue the spare parts eupress service to replace the euposue foot switch. SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002250000000018 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-07 10:41 | No
3466 | Yes | Failed Testing | Failed Testing | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002250000000032 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: weiyu Last udated by: wei yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: sa:ocap:00000000r/HeartStart uL Defibrillator/Monitor T2 Activities:Problem Description: Cutomer’s euplanation of the problem ### Failed to pass self-test The cutomer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### "Key uLLstatements and codes### uable to pass self-test" Repair Action: T2 Recommended Field Diagnostic Steps### sugestion T2 Recommended spare part uuwith description### na Cutomer demand### na Test/Inspection: na Test Equpment: na Neut Step Action: Missed retun visit, ##d MM Neut Step Action: Dear, I jut received a requst from a teacher at the First Affiliated Hospital of China Medical uiversity to repair the following equpment. The details are as follows: 1. uY: MP00000000 SN: DE0000000000000000833, Faut description: No uwaveform is oupu, hospital contact teacher: u uiao from the Department of Anesthesia, Tel: 15900000000819000000003 2. uY: MP00000000 SN: DE0000000000000000832, Faut description: The alarm mue buton is aging and the alarm cannot be mued, u uiao, anesthesia department, Tel: 15900000000819000000003 3. uY: M000000005A, SN: u000000000000000069, Faut description: uable to pass the self-test, contact the hospital teacher: u uiao from the Department of Anesthesia, Tel: 15900000000819000000003 Please help directly S000000000000000000000000001to issu 100000000000000000000000000000000001O# to me and issu 200000000000000000000000000000000001O# to ASP Honglejia for follow-u maintenance. Thank you Best Regards QuZongyao QuZongyao SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 00000000:0000000052 uC (860000000000000000100000000 *Department information: [Department: Department of Anesthesia Contact person: Teacher uiao Tel: 15900000000819000000003]* $ASP engineer:[honglejia uushuhu] $Q1: N $Q2: N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem: Operation check failed $Q8: Reformed visul inspection, the cable of the treatment handle is loose and not secuely connected. $Q9: Reconnect the treatment handle and perform an operational check. $Q100000000Test and inspection data refer to:Visul check: Passed.Device retuned to ful fuctionality $Q11: Calibrated Tools: N/A No open Field Change Order foud for this service. Internal Remarks: 00000000.11.2000000003 00000000:0000000055 uC (860000000000000000100000000 $ASP engineer:[honglejia uushuhu] $Q1: N $Q2:N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem: Operation check failed $Q8: Reformed visul inspection, the cable of the treatment handle is loose and not secuely connected. $Q9: Reconnect the treatment handle and perform an operational check. $Q100000000Test and inspection data refer to:Visul check: Passed.Device retuned to ful fuctionality $Q11: Calibrated Tools: N/A ----------------------------------------------------------------------------- | 2023-11-07 10:46 | No
3467 | Yes | Space low - Unable to save images | Space low - Unable to save images | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Engineers need to come on site immediately to clean the disk SAP Transfer: Dispatched for Onsite Suport Neut Step Action: 00000000882110000000000000000ALL IN SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002250000000079 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the image hard drive reported an uLLand requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: WARNING: Write problem. Images possibly lost Remote diagnostics and test resuts: NA PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC starts normally, check whether the statu of the POST LED behind the IPPC is normal, and ue the PC Diagnostic Tool for diagnostic testing. 2. Euecue Recreate Image Store. If it cannot recreate, try to bypass the hard disk bou, or format the hard disk. 3. Check the IPPC and Is the commuication net000000000000000000000000001k cable between the Host PC normal? 00000000Check the machine uLLand perform faut repairs. Recommended spare part uuand name:" 0000000080000000050000000000000000Hard disk spare part &0000000080000000050000000092 IP PC Revised_II no HDD Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Euternal Remarks: 00000000.11.2000000003 00000000:31:00000000uC (8600000000600000000) Service content (requred): Cutomer reported for repair: Insuficient disk space was prompted; the image disk retuned to normal after cleaning. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 00000000:31:00000000uC (8600000000600000000) "RDI Accuacy: Yes" $Q7: (Confirmation & Description of problem/Prompt disk space is insuficient.) $Q8: (Onsite inspection, The available disk space is small, Need to clean disk.) $Q9: (Do the database consistency test and recreate image store, Restart equpment, the problem was solved. Device retun to ful service.) $Q100000000 (Test system OK and inspection data refer to: PA. Device retuned to ful fuctionality. ) $Q11:Calibrated Tools:N/A ----------------------------------------------------------------------------- | 2023-11-07 11:07 | No
3468 | Yes | Tidal volume does not display values | 0/V0uLLTidal volume does not display values | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Tidal volue does not display valu Neut Step Action: Peking uiversity International Hospital Safety Qustions: Case: 0000000022500000000300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: ##d ASP qiangyi for follow u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-11-07 12:17 | No
3469 | Yes | Failed BIT | Failed BIT |  | 2023-11-07 12:56 | No
3470 | Yes | System cannot start up | System cannot start up | 000000000000000000000000001O##: 00000000000000000000000000100000000356795 Saat Muna Berlinang The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: Euposue not possible, reselect application uer Impact: Patient Impact: SE Resoluion : Gude biomed to restart system ----------------------------------------------------------------------------- | 2023-11-07 13:07 | No
3471 | Yes | Screen failure | Screen failure | 000000000000000000000000001O##: 0000000000000000000000000010000000035759000000001-Dec-2000000003 Eu Gok The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The device does not tun on When was the first occurence? within... : 200000000ct-2000000003 00000000:00000000:00000000 How often is issu occuring? : Intermittent Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system shoud boot and be ready uer Impact: system is not ready Patient Impact: no impact SE 1.000000002.1 Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Resoluion: The Host PC, which was previouly foud to be fauty, was replaced with a new one. The system was delivered operational after trials were carried ou. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000357200000000Buak Yediyildiz The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The device does not tun on When was the first occurence? within... : 200000000ct-2000000003 00000000:00000000:00000000 How often is issu occuring? : Intermittent Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system shoud boot and be ready uer Impact: system is not ready Patient Impact: no impact SE 1.000000002.1 Diagnostic performed by Engineer: I met with u Sercan and u Erman. HostPC 000000000000000000000000001't boot. It was opened via HostPC, the system is active. HostPC will change. Internal Comments: uLL# and/or description of uLL HostPC power is not on Repair Action: I talked to u Sercan and u Erman. HostPC 000000000000000000000000001't boot. It was opened via HostPC, the system is active. HostPC will change. Follow u Requred: uLL# and/or description of uLL HostPC power is not on Repair Action: I talked to u Sercan and u Erman. HostPC 000000000000000000000000001't boot. It was opened via HostPC, the system is active. HostPC will change. What are the neut Trouleshooting steps: I met with u Sercan and u Erman. HostPC 000000000000000000000000001't boot. It was opened via HostPC, the system is active. HostPC will change. Technician Name: Buak Yediyildiz ----------------------------------------------------------------------------- | 2023-11-07 13:31 | No
3472 | No | antenna were replaced( Split) | antenna were replaced( Split) |  | 2023-11-07 13:45
3473 | Yes | Order label. | Order label. | 000000000000000000000000001O##: 000000000000000000000000001000000003000000006 ----------------------------------------------------------------------------- | 2023-11-07 14:26 | No
3474 | Yes | FM: Reboot device | FM: Reboot device | 000000000000000000000000001O##: 00000000000000000000000000100000000357939 300000000Nov-2000000003 Biebl, Joseph The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : The uLLmessage is confirmed. no RR Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Resoluion: Evalution of log files. uLLno longer occured. The system meets the manuactuer's specifications in the tests carried ou and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000357800000000 Robert Kern The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : The uLLmessage is confirmed. no RR Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: System has not been online since Juy. Clarification via remote not possible Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow u Requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: TE ue to check the BIOS settings HOST PC. Checking remote connection. AZ 2 h Technician Name: Robert Kern ----------------------------------------------------------------------------- Cutomer comment: Kd. has info Tel: 00000000000000000000000028210000000000000000that the rep. will be done on PM day November 9th ----------------------------------------------------------------------------- | 2023-11-07 14:41 | No
3475 | Yes | MES SR-00- uLLgenerator bu no u-rays possible. In the middle of a primary medical acquisition and only... | MES SR-000- uLLgenerator bu no u-rays possible. In the middle of a primary medical acquisition and only... | 000000000000000000000000001O##: 00000000000000000000000000100000000388500000000 00000000-Nov-2000000003 Andy Carter The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: cardio Device ue: cardio Eupected and ueupected Resuts: no u-ray oupu uer Impact: system down Patient Impact: system down SEna Resoluion : Replaced tue and tested serviceable ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038800000000 00000000-Nov-2000000003 Phil Crouh The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: uLLgenerator buy no u-rays possible. In the middle of a primary medical acqusition and only takes a single shot of the device for it to uLL uer Impact: system down. Patient Impact: none. SE2.2 Resoluion : Tue replaced. System configued and tested. All tests passed. System handed back to cutomer for med physics testing. Parts Consued : 1.000000000000000000000000989000000000000000086752,u 200000000+ 0000000000000000 ROT-GS 1000000003 Tools & Equpment : DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000381821 00000000-Nov-2000000003 Andy Carter The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: uLLgenerator buy no u-rays possible. In the middle of a primary medical acqusition and only takes a single shot of the device for it to uLL uer Impact: NA Patient Impact: NA SE2.2 Diagnostic performed by Engineer : Replaced HiVo and system adapted bu then failed duing tue yield. Follow u Requred : Replace tue Parts Consued : 1.00000000000000000000000000000000210000000011828,HiVo high voltage transformer ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000035799000000007-Nov-2000000003 Andy Carter The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: uLLgenerator buy no u-rays possible. In the middle of a primary medical acqusition and only takes a single shot of the device for it to uLL uer Impact: NA Patient Impact: NA SE2.2 Diagnostic performed by Engineer : performed tue diagnostics. system reporting low curent Follow u Requred : replace supected components ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000357862 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : In case u-ray tue curent Ou of range is too HIGH, please first carry ou u-ray Tue Conditioning and then u-ray tue Adaptation. In case u-ray tue curent Ou of range is too LOW, then it coud be caued du to microleakages in the u-ray tue, a defective filament or short circuting of the filament against the cathode cu. Then carry ou the filament test in PSC: Diagnostics - Generator - Fuctional Test - Filament Test. Depending of the resuts the root caue can be determined: u-ray tue filament, Grid switch, HiVo or PoInt. Service Action 2: Also check HT cables and candles - inspect its pins. If a contact pin is deformed, readjut the slit uing pin gaue tool. 2. Apply silicon paste as per recommendation in "u-Ray Generator - Known Issus - Gude FautFinding u-Ray tues u" docuent. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: uLLgenerator buy no u-rays possible. In the middle of a primary medical acqusition and only takes a single shot of the device for it to uLL uer Impact: NA Patient Impact: NA SE2.2 Diagnostic performed by Engineer : Tue Arcing/Curent uLL ----------------------------------------------------------------------------- | 2023-11-07 14:47 | No
3476 | Yes | Device will not accept charge. | Device will not accept charge. | 000000000000000000000000001O##: 00000000000000000000000000100000000358181 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u Front Enclosue 0000000056000000005000000001 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Medical Electronics Device ue: not in ue Eupected and ueupected Resuts: Device will not accept charge. (Dock charger base terminals will not charge device & side charger port will not charge. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u Front Enclosue 0000000056000000005000000001 ----------------------------------------------------------------------------- Internal Remarks : Tracking Nuber:1Z3R6000000006695619000000003 COLLECT TOMORROW Internal Remarks : Tracking Nuber:1Z3R6000000006692000000000000000055 RECOLLECTION TOMORROW Internal Remarks : Tracking Nuber:1Z3R6000000006692000000000000000055 recollection tomorrow ----------------------------------------------------------------------------- | 2023-11-07 15:27 | No
3477 | Yes | It does not start. | It does not start. | 000000000000000000000000001O##: 000000000000000000000000001000000003000000006 00000000-Nov-2000000003 Luas Santolino Simón The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Phisician Device ue: Hemodynamics Eupected and ueupected Resuts: black screen. No boot u. uer Impact: delay Patient Impact: Na SE 8.2 Resoluion: Corrut software in the Host. Hard disk is replaced and backu copy is reloaded. Adjutments are checked and correct operation of the system is verified. ----------------------------------------------------------------------------- | 2023-11-07 15:41 | No
3478 | Yes | Air cooling fan issue - Complete 09 | Air cooling fan issue - Complete 0/9 | 000000000000000000000000001O##: 00000000000000000000000000100000000358551 November 100000000 2000000003 Alejandro Rodrigo Mauier Córdoba The system meets the specification for the performed service and is retuned to ue. Parts Consued: 2.00000000000000000000000000000000561512751,RP-Mout, Isolation, Fan, Pgk/00000000V600000000 2.00000000000000000000000000000000561511731,RP-Pin, Puh, Fan Mout, Pkg/00000000V600000000 1.00000000000000000000000000000000561511831,RP-Fan Gurd, Filter & Retainer, V600000000Resoluion: Incident duing preventive maintenance. Fan and filter stents are replaced. Material is consued. Problem Description by Engineer: Cutomer Fuction/Role: Field Service Engineer Device ue: Preventive Maintenance Eupected and ueupected Resuts: Air cooling fan issu uer Impact: NO Patient Impact: NO SE 3.200000000Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000353000000006 00000000-Nov-2000000003 Alejandro Rodrigo Mauier Córdoba The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Field Service Engineer Device ue: Preventive Maintenance Eupected and ueupected Resuts: Air cooling fan issu uer Impact: NO Patient Impact: NO SE 3.200000000Resoluion: Loose air inlet fan. Fastenings are replaced and position is adjuted. Preventive maintenance is performed. 00000000000000000000000000100000000192195 Replacement material is requsted in this 000000000000000000000000001 Follow u Requred: Material ----------------------------------------------------------------------------- | 2023-11-07 15:56 | No
3479 | Yes | Battery faiilure | Battery faiilure | 000000000000000000000000001O##: 00000000000000000000000000100000000388175 00000000Dec-2000000003 DAMIANI, MAuIZIO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: CLINICAL ING Device ue: battery test Eupected and ueupected Resuts: battery to be replaced uer Impact: no impact Patient Impact: no impact SE na Resoluion: Battery replaced Parts Consued: 1.000000000000000000000000989800000000626900000000Assy Battery Li-Ion,11ah,V600000000Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000003-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003527700000000Ardighieri, Andrea The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: CLINICAL ING Device ue: battery test Eupected and ueupected Resuts: battery to be replaced uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Battery euhauted following maintenance. Follow u Requred: Battery foud euhauted following maintenance, proceed with ordering replacement part. ----------------------------------------------------------------------------- | 2023-11-07 16:51 | No
3480 | Yes | All centrals, clients and servers lost connection to primary server in Helsingborg and Landskrona /GFE0ent 00 | All centrals, clients and servers lost conneuon to primary server in Helsingborg and Landskrona /GFE0sent 0/0 | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000085 ABDELBARI MOHAMMED The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Biomed informed that several Suveillances lost connection to primary and physio servers intermittently. This occured several times throuhou the day. First we tried to reboot the primary and physio servers, since the utime was very high, this improved the sitution bu issu still persisted. RSE checked the logs and noticed uLL in the database and lengthy qury entries that coud be the caue, We therefore agreed to clear the physio databases and also the internal database on several other iu hosts that were affected. This solved the issu. System has been u and runing several days now withou any issus. Info provided to cutomer staff. We agreed to close the case. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In ue. Eupected and ueupected Resuts: Suveillance lost connection to primary and physio servers intermittently. uer Impact: Suveillance lost connection to primary and physio servers intermittently. Patient Impact: None. SE C.00000000 ----------------------------------------------------------------------------- | 2023-11-07 17:21 | No
3481 | Yes | "Ouygen Not Available" alarm occurred, and a Check Vent alarm occurred | "Ouygen Not Available" alarm occurred, and a Check Vent alarm occurred |  | 2023-11-07 18:01 | No
3482 | Yes | AED beeps continuouy despite battery change | AED beeps continuouy despite battery change | Internal Remarks: I call the cutomer to mention sending her an e-mail for IPCR. She did not observe any new "beep-beeps", so she will see how to follow u. If it happens again, she will answer ou e-mail. Otherwise, she will call u to cancel the case. However, she has the correct ref. =" a17l10000000035 (which is not well localized in any case) ----------------------------------------------------------------------------- | 2023-11-07 18:16 | No
3483 | Yes | Display is getting blurred | Display is getting blurred | 000000000000000000000000001O##: 00000000000000000000000000100000000380000000091 Nandha, Rahu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: yes Eupected and ueupected Resuts: na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Display cable reconnected to resolve. ----------------------------------------------------------------------------- | 2023-11-07 18:31 | No
3484 | Yes | M0A - u00 Could not be turned on normally | u Could not be turned on normally |  | 2023-11-07 18:34 | No
3485 | Yes | Equipment stopped- replaces the battery | Equipment stopped- replaces the battery | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Equpment Stopped Name: Erik Silva Santos Indutry: Clinical Engineering Phone ¹: 11 972000000009197 Phone ²: 11 9000000000000000000000000 ##Engenharia.clinica00mbm.org.br BENEF ISRAELITABRAS HOSPI SOCIETY Address: 627 Albert Einstein Avenu São Pauo-SP 00000000652-0000000000000000 BR Internal Remarks: Technical Validation Internal Remarks: Technical Validation Safety Qustions: Case: 0000000022556500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Leandro Costa Last udated by: Leandro Costa 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment malfuction or fail at the time of installation? uknown ----------------------------------------------------------------------------- | 2023-11-07 18:36 | No
3486 | Yes | system will not boot up | system will not boot up | 000000000000000000000000001O##: 00000000000000000000000000100000000380000000057 16-Nov-2000000003 Rodriguz Lazaro The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System not booting u all the way When was the first occurence? Within... : 00000000-Nov-2000000003 00000000:0000000013 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System not booting u uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Done by FSE remotely Parts Consued : 1.00000000000000000000000000000000800000000200000000612,Nehalem Host PC Monoplane Rev_III NO_HDD Resoluion : Replaced the compuer and ued the previou hard drive on the new compuer. Installed the new appropriate license file. Rebooted mutiple times and tested the system and is operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000379556 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System not booting u all the way When was the first occurence? Within... : 00000000-Nov-2000000003 00000000:0000000013 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System not booting u uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Dwight said the system is not booting u all the way. When they tun on the system the data monitor shows the initial boot process bu then the monitor goes blank. The fleu monitor has the lines on the monitor bu no video and they can hear the brakes clicking on the table. Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Check host pc, ue another monitor to connect directly into host pc to watch boot process. Coud try reloading ghost back u. If needed, 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD and 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Check host pc, ue another monitor to connect directly into host pc to watch boot process. Coud try reloading ghost back u. If needed, 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD and 0000000080000000050000000000000000- Hard disk spare part Technician Name : Coffman, Josh ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: dwight IP & Issu: 520000000026900000000will not boot Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: Created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-11-07 18:37 | No
3487 | Yes | QACASE for PCI imported product#00000 The device can not save the image | QACASE for P0imported produu00000000 The device can not save the image |  | 2023-11-07 18:56 | No
3488 | Yes | Legal issue_BP caud a bicep tendon rupture in May 0. Emailed legal on 0c0 | Legal issue_BP caud a bicep tendon rupture in May 00 Emailed legal on 0ec03 |  | 2023-11-07 19:02 | Yes
3489 | Yes | problem with 0ead | problem with 0lead | 000000000000000000000000001O##: 00000000000000000000000000100000000381200000000Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Lieuenant states the 00000000 lead on the uis not 000000000000000000000000001king, there are no uLLmessages, they jut cannot get the 00000000 lead to pick u anything; needs to send in for bench repair to get the 00000000 lead to 000000000000000000000000001k. Verified shipping address: North Collier Fire Department ATTN: LT. Colton Muray Taylor Rd 600000000 Naples, FL 3000000009-18500000000 u Lt. states they 000000000000000000000000001ld like the device retuned directly to their shop and not the main address on file; please ship to the address listed above. Diagnostic performed by Engineer : Lieuenant states the 00000000 lead on the uis not 000000000000000000000000001king, there are no uLLmessages, they jut cannot get the 00000000 lead to pick u anything; needs to send in for bench repair to get the 00000000 lead to 000000000000000000000000001k. Verified shipping address: North Collier Fire Department ATTN: LT. Colton Muray Taylor Rd 600000000 Naples, FL 3000000009-18500000000 u Lt. states they 000000000000000000000000001ld like the device retuned directly to their shop and not the main address on file; please ship to the address listed above. Problem Description by Engineer : Cutomer Fuction/Role: Lieuenant states the 00000000 lead on the uis not 000000000000000000000000001king, there are no uLLmessages, they jut cannot get the 00000000 lead to pick u anything; needs to send in for bench repair to get the 00000000 lead to 000000000000000000000000001k. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Resoluion : Verified cutomers contact and shipping information and ##d the u bench form. Created an action assignment for the bench to send the cutomer a shipping label. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001--------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: colton IP & Issu: 9170000000075 problem with 00000000ead Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: Created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-11-07 19:26 | No
3490 | Yes | Low ouygen pressure | 0K/V0uLLLow ouygen pressure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022556765 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Low ouygen pressue The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: It is recommended to check the sensor Recommended spare part uufor T2 with description: #989800000000617100000000 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Low ouygen pressue Neut Step Action: ASPhumei FSE Notes: This PO is a Reference Nuber only Neut Step Action: 000000005621126285 call in Neut Step Action: Qingdao Binhai College Internal Remarks: 17.11.2000000003 00000000:29:500000000uC (86000000000000000028) ASP engineer: [Humei Huntong] $Q1:Y $Q2:N $Q3:N $Q00000000 $Q7:Confirmed reported problem, the uLLreport low O2 suply pressue. $Q8:Start the uLL find that the device report low O2 suply pressue.According to service manul, test the central ouygen pressue, find that the central ouygen pressue below normal service pressue.After adjut the central ouygen pressue, the device retun to normal.So,the central ouygen pressue had problem. $Q9:Adjut the central ouygen pressue. $Q100000000Visul check:Passed,the central ouygen pressue test:Passed.Device retuned to ful fuctionality. $Q11: N/A Euternal Remarks: 17.11.2000000003 00000000:29:00000000uC (86000000000000000028) *Department information: [Department: Respiratory Department Contact: Teacher Gun Tel: 1859560000000000000000* ASP engineer: [Humei Huntong] $Q1: Y $Q2:N $Q3:N $Q00000000 $Q7:Confirmed reported problem, the uLLreport low O2 suply pressue. $Q8:Start the uLL find that the device report low O2 suply pressue.According to service manul, test the central ouygen pressue, find that the central ouygen pressue below normal service pressue.After adjut the central ouygen pressue, the device retun to normal.So, the central ouygen pressue had problem. $Q9:Adjut the central ouygen pressue. $Q100000000Visul check:Passed,the central ouygen pressue test:Passed.Device retuned to ful fuctionality. $Q11: N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-07 19:31 | No
3491 | Yes | Air leakage reported | Air leakage reported |  | 2023-11-07 19:41 | Yes
3492 | Yes | uuufunctional check failed | u funuonal check failed | 000000000000000000000000001O##: 00000000000000000000000000100000000395800000000 Ardighieri, Andrea The system meets the specification for the performed service and is retuned to ue Resoluion: Problem resolved by correctly re-runing the fuctional check by the Biamonte technician. Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: defibrillator test Eupected and ueupected Resuts: fuctional verification failed uer Impact: no impact Patient Impact: no impact SE na ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000357612 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: coord Pierpaolo Romanò Device ue: diagnosis Eupected and ueupected Resuts: prob on calibration uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na coord Pierpaolo Romanò, report defibrillator problems Follow u Requred: coordinated by Pierpaolo Romanò, reports problems with the defibrillator, the system is NOT in contract, requres a visit from a technician for complete calibration and susequnt fuctional tests. I open QuTES, then send on site equpped with the necessary tools Diagnostic performed by Engineer: coordinated by Pierpaolo Romanò, reports problems with the defibrillator, as the system is NOT in contract, requres a visit from a technician for complete calibration and susequnt fuctional tests. I open QuTES ----------------------------------------------------------------------------- Internal Remarks: System awaiting change of service provider to u. Iba aware. ----------------------------------------------------------------------------- | 2023-11-07 19:42 | No
3493 | Yes | Does not start | Does not start | 000000000000000000000000001O##: 00000000000000000000000000100000000389129 23-Nov-2000000003 Leuhold Marc The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Diagnostic performed by Engineer: Material retun. Problem Description by Engineer: Cutomer Fuction/Role: MTRA Device ue: yes Eupected and ueupected Resuts: na uer Impact: no Patient Impact: no SE8.1.u ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000035700000000 Leuhold Marc The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: Log files analyzed du to performance issus. Image data transfer stop. Spoken to cutomer, system is curently 000000000000000000000000001king normally. Problem Description by Engineer: Cutomer Fuction/Role: MTRA Device ue: yes Eupected and ueupected Resuts: na uer Impact: no Patient Impact: no SE 8.1.u ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000357752 00000000-Nov-2000000003 Vung, Thanh The system meets the specification for the performed service and is retuned to ue Resoluion: FleuVision PC replaced and fuction OK. System is fuly fuctional. Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: In Eupected and ueupected Resuts: System does not start. Stuk at 00000000:00000000. uer Impact: System down. Patient Impact: No SE NA. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000382100000000 00000000-Nov-2000000003 Leuhold Marc The system meets the specification for the performed service and is retuned to ue Follow u Requred: Material only Problem Description by Engineer: Cutomer Fuction/Role: MTRA Device ue: yes Eupected and ueupected Resuts: na uer Impact: no Patient Impact: no SE8.1.u Diagnostic performed by Engineer: System gets stuk at 00000000:00000000 duing startu and Fleuvision stays dark. Fleuvision PC defective. Trouleshooting, variou fuctional tests carried ou, FrameGrabber cards in the FleuVision PC defective. Fleuvision PC replaced. General fuctional check, system OK. Afterwards problems, system very slow. All movements very slow, no uLLmessage to be seen. After restarting everything is normal, log files checked. Tools & Equpment: MuTIMETER, 200000000000000000000000020000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-11-07 19:42 | No
3494 | Yes | M0A / u0/ The battery does not hold the charge (Replaced uuuKit - Lithium Ion Battery) | M00A / u00/ The battery does not hold the charge (Replaced u Kit - Lithium Ion Battery) | 000000000000000000000000001O##: 00000000000000000000000000100000000389000000002 13-Nov-2000000003 ZAMBONI, STEFANO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Battery replaced, instruent ready for clinical ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000380000000085 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: clinical engineer Gianlua Pasini Device ue: diagnosis Eupected and ueupected Resuts: battery prob uer Impact: not in clinical ue Patient Impact: not in ue on patient SE not heard clinical engineer Gianlua Pasini, reports battery problems Diagnostic performed by Engineer: consuted clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is uder contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager, otherwise we will be forced to write the EOL letter Follow u Requred: heard from clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is in contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager, otherwise we will be forced to send the EOL letter, if we find them they mut be sent NEMO to: OSP BuALINI Viale Giovanni Ghirotti, 286 0000000021 Cesena (FC) c/o Transit Point Ing Clinica c.a. u Marcello Prunoli Tel.00000000000000005200000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000381535 Parts Consued: 1.00000000000000000000000000000000980000000050000000021,uuuKit - Lithiu Ion Battery ----------------------------------------------------------------------------- Internal Remarks: For the attention of: We hereby ask youfor you technical intervention uDER CONTRACT on the equpment indicated below INTERVENTION 000S0000000050000000051 Inv. : 98168 IC CE Class: DEF - DEFIBRILLATOR Manuactuer: PHI - u HEALTHCARE uY: 3H - HEARTSTART uuuM3536A CdC: CE500000000- CE HEALTH EMERGENCY Eup. Warranty: 00000000/00000000/2000000003 uEQuPMENT Machine downtime: NO Ref. Department: 2 tel.: 2 Inv. GAAC: MuB0000000000000000126500000000Intervention ID: 72000000009 Forklift: Faut: BATTERY REPLACEMENT REQuRED AS IT DOES NOT HOLD A CHARGE Inv. Company: DDT Sending: INTERNAL TECHNICIAN TO CONTACT AT 0000000000000000352987: PASINI GIANLuA - u FIS. MEDICAL AND ENGINEERING CLINIC Ou call reference: 2000000003/1000000003/IC/CE dated 00000000/11/2000000003 13:00000000:00000000 In the docuents relating to this requst, always refer to the following nuber: Print Date: CESENA, there 00000000/11/2000000003 13:00000000:28 Technical Assistance uO. FM and IC We remind youthat if the intervention is carried ou in the department, the employment report mut be stamped and signed by the staff of the department itself and delivered to the electronic laboratory of the Medical Physics and Clinical Engineering uit. Please also send any eupense statement by fau, reporting the nuber of 000000000000000000000000001king hous, to the nuber 000000000000000030000000072 or by ##to the address fis.tec.ce00ulromagna.it ----------------------------------------------------------------------------- | 2023-11-07 20:16 | No
3495 | Yes | M0A - u0 - The battery does not hold the charge. | M00A - u00 - The battery does not hold the charge. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000519 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000380000000038 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Eng. Gianlua Pasini Device ue: diagnosis Eupected and ueupected Resuts: prob battery uer Impact: not in clinical ue Patient Impact: not in ue on patient SE heard from clinical engineer Gianlua Pasini, he reported problems on battery Diagnostic performed by Engineer: heard from clinical engineer Gianlua Pasini, he reports problems on battery, system shows u as uder contract, bu the defibrillator has not been suported for a long time, I try to indicate the code of the battery for a local search or with the help of the site manager, otherwise we will be forced to issu the EOL letter, if we find them they need to be sent NEMO. If we cannot find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it Follow u Requred: I spoke to clinical Gianlua Pasini, who reported problems with battery, the system shows u as uder contract, bu the defibrillator has not been suported for a long time, I try to indicate the code of the battery for a local search or with the help of the site manager, otherwise we will be forced to issu the EOL letter, if we find them they need to be sent NEMO. if we cannot find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it NEMO to: BuALINI HOSPITAL Viale Giovanni Ghirotti, 286 0000000021 Cesena (FC) c/o Transit Point Clinical Encephalo u Marcello Prunoli Tel.00000000000000005200000000 ----------------------------------------------------------------------------- Internal Remarks: To the attention of: This is to ask youfor you technical intervention uDER CONTRACT on the equpment below indicated INTERVENTION Nuber: u0000000050000000055 Inv. : 98198 IC CE Class: DEF - DEFIBRILLATOR Manuactuer: PHI - u HEALTHCARE uY: 3H - HEARTSTART uuuM3536A CdC: CE5500000000- SPB uO. S. PIERO HEALTH EMERGENCY Eupires Warranty: 00000000/00000000/2000000003 uEQuPMENT Down machine: NO Ref. Department: 22 tel.: 2 Inv. GAAC: MuB000000000000000012657 ID Intervention: 72000000001 Forlection: Faut: REQuST TO REPLACE BATTERY AS IT DOES NOT HOLD THE CHARGE Inv. Company: DDT Sending: INTERNAL TECHNICIAN TO CONTACT AT 0000000000000000352987: PASINI GIANLuA - u PHYS. MEDICAL AND E.E. CLINIC Ou call reference: 2000000003/1000000001/IC/CE on 00000000/11/2000000003 13:00000000:00000000 In the docuents relative to this requst always refer to the following nuber: Print date: CESENA, there 00000000/11/2000000003 13:00000000:39 Technical Assistance uO. FM and IC Remember that if the intervention is carried ou at the department, the 000000000000000000000000001k report will have to be stamped and signed by the staff of the department and delivered to the electronic laboratory u Medical Physics and Clinical Engineering. Please, also send any sumary of eupenses via fau, reporting the nuber of hous of 000000000000000000000000001k, at nuber 000000000000000030000000072 or via e-mail to fis.tec.ce00ulromagna.it ----------------------------------------------------------------------------- | 2023-11-07 20:37 | No
3496 | Yes | It does not turn on. | It does not turn on. | 000000000000000000000000001O##: 00000000000000000000000000100000000381679 00000000-Nov-2000000003 Cristobal Medina Peña The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The application does not load, it is probably necessary to reload the Software of the SutePC. C. Medina is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment cannot be ued. Patient Impact: No impact SE 1.2.5.00000000Resoluion: Defective u-ray PC. Sw is changed and loaded. It is 000000000000000000000000001king. Parts Consued: 1.00000000000000000000000000000000800000000660000000093,u-ray PC Monoplane Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000008-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000380000000000000000 Adrián Sancho López The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The application does not load, it is probably necessary to reload the Software of the SutePC. C. Medina is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment cannot be ued. Patient Impact: No impact SE 1.2.5.00000000Diagnostic performed by Engineer: The application does not load, it is probably necessary to reload the Software of the SutePC. C. Medina is in charge of checking it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The application does not load, it is probably necessary to reload the Software of the SutePC. C. Medina is in charge of checking it. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: The application does not load, it is probably necessary to reload the Software of the SutePC. C. Medina is in charge of checking it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-11-07 20:41 | No
3497 | Yes | The alarm silence button not func - Cuomer replaced CPU PCBA, Requesting encryption codes | The alarm silence button not func - Cuomer replaced CPU PCBA, Requesting encryption codes | 000000000000000000000000001O##: 000000000000000000000000001000000003815800000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer needs to 000000000000000000000000001k with the hospital IT grou to get tera term u and runing. He already has the encryption codes for the software option. Diagnostic performed by Engineer : Cutomer has a V600000000and they jut replaced the CPupcba and need the software options. He did not recall the original issu with the vent, bu now after replacing the CPupcba, it is operating fine. Informed him to reprogram the serial # and add the power on hous via tera term. He claims he has tera term, bu never ued it yet. Instruted him to check with the hospital IT grou to assist with that. provide him the encryption codes for Avaps, Cfleu, and ramp. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: replaced CPupcba, and needs encryption codes uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Steve Schutz IP & Issu: replaced the main board in uit, and now some of the software is missing/5663000000009 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE ----------------------------------------------------------------------------- | 2023-11-07 20:47 | No
3498 | Yes | WHEN uNG THE DEFIBRILLATOR IT GIVES THE CAPTION "DEFIBRILLATOR DISABLED" | 000WHEN uNG THE DEFIBRILLATOR IT GIVES THE CAPTION "DEFIBRILLATOR DISABLED" | 000000000000000000000000001O##: 00000000000000000000000000100000000352790000000000000000-Nov-2000000003 Telpalo, uises The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: In ue Eupected and ueupected Resuts: uer reported device needed Operational Test uer Impact: Coudn't ue device. Patient Impact: N/A SE uknown Resoluion: uer reported that the equpment got blocked and requsted to do the test, bu it did not let them for a while, soon after it was possible. Checked the uLLlog and detected that the cables were loose when the uer tried the test, performed a check-u and corrected the statu Tools & Equpment: DEFIBRILLATOR ANALYZER, 200000000000000009-20000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-11-07 20:47 | No
3499 | Yes | Device powered down while on patient | Device powered down while on patient | 000000000000000000000000001O##: 0000000000000000000000000010000000039960000000017-Nov-2000000003 Reu, Timothy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: device failed when on patient uer Impact: system restricted Patient Impact: none reported SE 3.00000000 Resoluion : Implemented FCO86600000000000000006 and problem was corrected. Testing and verification was performed on FCO 00000000000000000000000000100000000180000000092. All tests passed. Device is ready for patient ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000352861 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Creating on site service for 100000000% discout to be claimed uder FCO Follow u Requred : 100000000% discout. Device failed while in ue. Powered down will not power on. To be claimed uder FCO66. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: device failed when on patient uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device failed when on patient, no harm reported to patient. Cutomer states device was reported to power down. Cutomer uable to power device on to check event log. Cutomer is requsting on site service to be covered uder FCO66 . ----------------------------------------------------------------------------- Internal Remarks : Contact Name: emsar Method: nsa/##What was discused: cutomer requsted to have fco completed on their uit. they also stated they have a down uit. i went ahead and udated the nsa and udated the contact information. this will pop u on emsars list. emsar will reach ou to the cutomer to schedue this. i also sent emsar an ## incident case- 0000000022557172 fco case 0000000020000000020000000035 Neut Step Action: n/a Internal Remarks : Interaction Type: New Case Caller: Aleu Courouh IP & Issu: 57000000001900000000evice failed when on patient Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: free remote Reason for Entitlement Claim: N/A Action Taken: created case, R000000000000000000000000001Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-07 21:06 | No
3500 | Yes | uuubattery does not hold the charge | u battery does not hold the charge | 000000000000000000000000001O##: 00000000000000000000000000100000000381200000000MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: clinical engineer Gianlua Pasini Device ue: diagnosis Eupected and ueupected Resuts: battery prob uer Impact: not in clinical ue Patient Impact: not in ue on patient SE not heard clinical engineer Gianlua Pasini, reports battery problems Follow u Requred: heard from clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is in contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager otherwise we will be forced to send the EOL letter, if we find them they mut be sent NEMO. if we do not find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it NEMO to: OSP BuALINI Viale Giovanni Ghirotti, 286 0000000021 Cesena (FC) c/o Transit Point Ing Clinica c.a. u Marcello Prunoli Tel.00000000000000005200000000 Diagnostic performed by Engineer: consuted clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is uder contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager, otherwise we will be forced to send the EOL letter, if we find them they mut be sent NEMO. If we can't find the battery, kindly send EOL letter ----------------------------------------------------------------------------- Internal Remarks: We hereby ask youfor you technical intervention IN CONTRACT on the equpment indicated below INTERVENTION 000S0000000050000000057 Inv. : 98193 IC CE Class: DEF - DEFIBRILLATOR Manuactuer: PHI - u HEALTHCARE uY: 3H - HEARTSTART uuuM3536A CdC: CE500000000- CE HEALTHCARE EMERGENCY Eup. Warranty: 00000000/00000000/2000000003 uEQuPMENT Machine downtime: NO Ref. Department: 2 tel.: 2 Inv. GAAC: MuB000000000000000012656 Intervention ID: 72000000008 Forklift: Faut: BATTERY REPLACEMENT REQuRED AS IT DOES NOT HOLD A CHARGE Inv. Company: DDT Sending: INTERNAL TECHNICIAN TO CONTACT AT 0000000000000000352987: PASINI GIANLuA - u FIS. MEDICAL AND ENGINEERING CLINIC Ou call reference: 2000000003/10000000000000000C/CE dated 00000000/11/2000000003 12:59:00000000 In the docuents relating to this requst, always refer to the following nuber: Print Date: CESENA, there 00000000/11/2000000003 12:59:500000000echnical Assistance uO. FM and IC We remind youthat if the intervention is carried ou in the department, the employment report mut be stamped and signed by the staff of the department itself and delivered to the electronic laboratory of the Medical Physics and Clinical Engineering uit. Please also send any eupense statement by fau, reporting the nuber of 000000000000000000000000001king hous, to the nuber 000000000000000030000000072 or by ##to the address fis.tec.ce00ulromagna.it ----------------------------------------------------------------------------- | 2023-11-07 21:11 | No
3501 | Yes | Request for analysis of logs. Desaturation alarms do not sound.= Case in Progress as Dec0 | Request for analysis of logs. Desaturation alarms do not sound.= Case in Progress as Dec0 | 000000000000000000000000001O##: 00000000000000000000000000100000000352938 Seva, Romain The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: No. Patient Impact: No. SE / Diagnostic performed by Engineer: The nuses indicate that some desatuation alarms do not soud. Recommendation: Checked the parameters of the monitor on the desatuation alarms (presence of a delay for the desatuation alarms). If possible, eutraction of the logs for analysis in order to do an analysis and check the behavior of the monitors. Resoluion: The cutomer did not eutract the logs and does not have any euamples for analysis (incident +7 days). The cutomer wants to receive a qute for on-site service so that a technician checks the configuation of the alarms on the central station and the monitors. Possibly also redo a reminder of the behavior of the alarms in the department. Qute for on-site service sent to the cutomer. Follow u Requred: Qute for on-site intervention. MO= 2 hous Type MO= M2 D=D3 The nuses complain that the desatuation alarms do not soud. The cutomer did not eutract the logs and does not have any euamples for analysis (incident +7 days). The cutomer wants to receive a qute for on-site service so that a technician checks the configuation of the alarms on the central station and the monitors. Possibly also redo a reminder of the behavior of the alarms in the department. ----------------------------------------------------------------------------- | 2023-11-07 21:12 | No
3502 | Yes | M0A - u0- The battery does not hold the charge (Replaced uuuKit - Lithium Ion Battery). | M00A - u00- The battery does not hold the charge (Replaced u Kit - Lithium Ion Battery). | 000000000000000000000000001O##: 00000000000000000000000000100000000389000000008 13-Nov-2000000003 ZAMBONI, STEFANO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Battery replaced, instruent ready for clinical ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000381297 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: clinical engineer Gianlua Pasini Device ue: diagnosis Eupected and ueupected Resuts: battery prob uer Impact: not in clinical ue Patient Impact: not in ue on patient SE not heard clinical engineer Gianlua Pasini, reports battery problems Diagnostic performed by Engineer: consuted clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is uder contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager, otherwise we will be forced to send the EOL letter, if we find them they mut be sent NEMO. If we can't find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it Follow u Requred: heard from clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is in contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager otherwise we will be forced to send the EOL letter, if we find them they mut be sent NEMO. if we do not find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it NEMO to: OSP BuALINI Viale Giovanni Ghirotti, 286 0000000021 Cesena (FC) c/o Transit Point Ing Clinica c.a. u Marcello Prunoli Tel.00000000000000005200000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000381985 Parts Consued: 1.00000000000000000000000000000000980000000050000000021,uuuKit - Lithiu Ion Battery ----------------------------------------------------------------------------- Internal Remarks: For the attention of: We hereby ask youfor you technical intervention uDER CONTRACT on the equpment indicated below INTERVENTION 000S0000000050000000052 Inv. : 98188 IC CE Class: DEF - DEFIBRILLATOR Manuactuer: PHI - u HEALTHCARE uY: 3H - HEARTSTART uuuM3536A CdC: CE500000000- CE HEALTH EMERGENCY Eup. Warranty: 00000000/00000000/2000000003 uEQuPMENT Machine downtime: NO Ref. Department: 2 tel.: 2 Inv. GAAC: MuB000000000000000012655 Intervention ID: 72000000007 Forklift: Faut: BATTERY REPLACEMENT REQuRED AS IT DOES NOT HOLD A CHARGE Inv. Company: DDT Sending: INTERNAL TECHNICIAN TO CONTACT AT 0000000000000000352987: PASINI GIANLuA - u FIS. MEDICAL AND ENGINEERING CLINIC Ou call reference: 2000000003/1000000005/IC/CE dated 00000000/11/2000000003 12:58:00000000 In the docuents relating to this requst, always refer to the following nuber: Print Date: CESENA, there 00000000/11/2000000003 12:58:500000000Technical Assistance uO. FM and IC We remind youthat if the intervention is carried ou in the department, the employment report mut be stamped and signed by the staff of the department itself and delivered to the electronic laboratory of the Medical Physics and Clinical Engineering uit. Please also send any eupense statement by fau, reporting the nuber of 000000000000000000000000001king hous, to the nuber 000000000000000030000000072 or by ##to the address fis.tec.ce00ulromagna.it ----------------------------------------------------------------------------- | 2023-11-07 21:21 | No
3503 | Yes | prouimal pressure sensor autozero failed | prouimal pressure sensor autozero failed | 000000000000000000000000001O##: 000000000000000000000000001000000003818800000000erchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer Callback Diagnostic performed by Engineer : Cutomer has a V600000000and the uit has a uLLmessage pressue sensor failed. Cutomer verified there was a code 1100000000 in the significant log. They jut replaced the flow sensor assy, bu the uit now has 62 hous since that part was replaced. Cutomer verified there were many 1100000000 codes in the history. Recommended going throuh this trouleshooting in order 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the machine auo-zero solenoid (SOL 2 and SOL 00000000 3. Replace the DA PCBA. Cutomer called back and stated that some of the pins were ou of the DA pcba. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: prouimal pressue sensor auozero failed uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000382776 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: prouimal pressue sensor auozero failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer called back and needed more information Resoluion : They re-installed the DA pcba back on top of the flow sensor assy to solve the issu. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Callback Caller: " Ashleigh Marcotte" IP & Issu: 898992700000000uLLmessage pressue sensor failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created 2nd R000000000000000000000000001ufer RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: " Ashleigh Marcotte" IP & Issu: 898992700000000uLLmessage pressue sensor failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001ufer RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-07 21:26 | No
3504 | Yes | Stand floating | Stand floating | 000000000000000000000000001O##: 00000000000000000000000000100000000382000000006 00000000-Nov-2000000003 Delahanty Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : spoke to Lindsey the table and stand are free floating How often is issu occuring? : Mutiple times per week Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: system di not power u table and stand uer Impact: uable to move stand Patient Impact: moved patient from table SE 000Resoluion : Cutomer euuuencing start u uLLwith geometry, trouleshot and foud power and control uit cauing intermittent geo uLL on start u, replaced power and control uit and performed all relevant testing. System operating normally and retuned to ue. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003817200000000Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : uLL# and/or description of uLL: no power to stand or table Repair Action : Caue Motorized movement uavailable duing Beam Longituinal Frontal movement. If this only occus sporadically and cutomer is not complaining please ignore, else see soluions below. Soluion 1 Clean / degrease the longituinal tracks/rails. Soluion 2 Check condition of motor ruber wheel. Soluion 3 Check and/or adjut the tension of the spring. Soluion 00000000arry ou the Beam longituinal curent calibration via FSF. Soluion 5 If problem persists replace the longituinal potmeter SPL R1 Soluion 6 For Clea: Replace Longituinal Drive Assy (SPL uit) Soluion 7 For Poly G2: If problem still persist replace SPC2 (Eutension board) and/or SPC3 (LuV00000000Caue 1 Longituinal potmeter is defective Soluion Replace table longituinal potmeter Caue 2 Table longituinal potmeter calibration Soluion Check the cable towards this potmeter and perform the table calibration 00000000230000000025982 longituinal potmeter assy 00000000800000000162371 Motor linear drive 00000000230000000022113 potm.uit longituinal What are the neut Trouleshooting steps : check uLLlogging , perform diagnostics and calibration for longituinal table and Beam longituinal of the stand Technician Name : Sulivan Scott Internal Comments : uLL# and/or description of uLL: no power to stand or table Repair Action : Caue Motorized movement uavailable duing Beam Longituinal Frontal movement. If this only occus sporadically and cutomer is not complaining please ignore, else see soluions below. Soluion 1 Clean / degrease the longituinal tracks/rails. Soluion 2 Check condition of motor ruber wheel. Soluion 3 Check and/or adjut the tension of the spring. Soluion 00000000arry ou the Beam longituinal curent calibration via FSF. Soluion 5 If problem persists replace the longituinal potmeter SPL R1 Soluion 6 For Clea: Replace Longituinal Drive Assy (SPL uit) Soluion 7 For Poly G2: If problem still persist replace SPC2 (Eutension board) and/or SPC3 (LuV00000000Caue 1 Longituinal potmeter is defective Soluion Replace table longituinal potmeter Caue 2 Table longituinal potmeter calibration Soluion Check the cable towards this potmeter and perform the table calibration 00000000230000000025982 longituinal potmeter assy 00000000800000000162371 Motor linear drive 00000000230000000022113 potm.uit longituinal Problem Description by Engineer : Problem description by engineer : spoke to Lindsey the table and stand are free floating How often is issu occuring? : Mutiple times per week Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: system di not power u table and stand uer Impact: uable to move stand Patient Impact: moved patient from table SE 000Diagnostic performed by Engineer : speaking to Lindsey the system table and stand is free floating . She is shuing down and restarting . after restart the table is no longer free floating bu there is no power to the stand .cutomer said this also happened yesterday and they were able to clear with mutiple restarts of system . they have not started case and are moving patient to alternate room system does uray . ----------------------------------------------------------------------------- Internal Remarks : cutomer 000000000000000000000000001ld like on site service ,states they have t and m cap pool for on site coverage ##provided bethanne.mazzola00tvincenthospital.com primary fse ryan states they still need to provide po for on site rse sent qute ----------------------------------------------------------------------------- | 2023-11-07 21:36 | No
3505 | Yes | M0A / u0 / The battery does not hold the charge. | M00A / u00 / The battery does not hold the charge. | 000000000000000000000000001O##: 00000000000000000000000000100000000381337 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Gianlua Pasini Device ue: Diagnosis Eupected and ueupected Resuts: battery problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from clinical Gianlua Pasini, he reports problems on battery. Diagnostic performed by Engineer: heard from clinical engineer Gianlua Pasini, he reports problems on battery, system shows u as uder contract, bu the defibrillator has not been suported for a long time, I try to indicate the code of the battery for a local search or with the help of the site manager, otherwise we will be forced to issu the EOL letter, if we find them they need to be sent NEMO. If we cannot find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it NEMO Follow u Requred: I spoke to clinical Gianlua Pasini, who reported problems with battery, the system shows u as uder contract, bu the defibrillator has not been suported for a long time, I try to indicate the code of the battery for a local search or with the help of the site manager, otherwise we will be forced to issu the EOL letter, if we find them they need to be sent NEMO. If we cannot find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it NEMO to: BuALINI HOSP. Viale Giovanni Ghirotti, 286 0000000021 Cesena (FC) c/o Transit Point Clinical Ac. u Marcello Prunoli Tel.00000000000000005200000000 ----------------------------------------------------------------------------- Internal Remarks: This is to ask youfor you technical intervention uDER CONTRACT on the equpment below indicated INTERVENTION Nuber: u0000000050000000000000000 Inv. : 98178 IC CE Class: DEF - DEFIBRILLATOR Manuactuer: PHI - u HEALTHCARE uY: 3H - HEARTSTART uuuM3536A CdC: CE500000000- CS uO. CESENATICO HEALTH EMERGENCY Scad. Warranty: 00000000/00000000/2000000003 u EQuPMENT Machine downtime: NO Ref. Department: 2 tel.: 2 Inv. GAAC: MuB000000000000000012653 Intervention ID: 72000000006 Forklift: Faut: REQuST TO REPLACE BATTERY AS IT DOES NOT HOLD THE CHARGE Inv. Company: DDT Sending: INTERNAL TECHNICIAN TO CONTACT AT 0000000000000000352987: PASINI GIANLuA - u PHYS. MEDICAL AND E.E. CLINICAL Ou call reference: 2000000003/1000000006/IC/CE on 00000000/11/2000000003 12:57:00000000 In the docuents relative to this requst always refer to the following nuber: Date of print: CESENA, there 00000000/11/2000000003 12:57:57 Technical Assistance uO. FM and IC Remember that if the intervention is carried ou at the department, the 000000000000000000000000001k report will have to be stamped and signed by Staff of the department and delivered to the electronic laboratory u Medical Physics and Clinical Engineering. Please also Send any sumary of eupenses via fau, reporting the nuber of hous of 000000000000000000000000001k, at nuber 000000000000000030000000072 or via e-mail. To the following address: fis.tec.ce00ulromagna.it ----------------------------------------------------------------------------- | 2023-11-07 21:46 | No
3506 | Yes | M0A - u0 - The battery does not hold the charge. | M00A - u00 - The battery does not hold the charge. | 000000000000000000000000001O##: 000000000000000000000000001000000003813800000000MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Gianlua Pasini Device ue: diagnosis Eupected and ueupected Resuts: battery prob uer Impact: not in clinical ue Patient Impact: not in ue on pazinet SE not heard clinical ing Gianlua Pasini, reports battery problems Follow u Requred: heard from clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is in contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager otherwise we will be forced to send the EOL letter, if we find them they mut be sent NEMO. if we do not find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it NEMO to: OSP BuALINI Viale Giovanni Ghirotti, 286 0000000021 Cesena (FC) c/o Transit Point Ing Clinica c.a. u Marcello Prunoli Tel.00000000000000005200000000 Diagnostic performed by Engineer: consuted clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is uder contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager, otherwise we will be forced to send the EOL letter, if we find them they mut be sent NEMO. If we can't find the battery, kindly send EOL letter ----------------------------------------------------------------------------- Internal Remarks: For the attention of: We hereby ask youfor you technical intervention uDER CONTRACT on the equpment indicated below INTERVENTION 000S0000000050000000000000000Inv. : 97618 IC CE Class: DEF - DEFIBRILLATOR Manuactuer: PHI - u HEALTHCARE uY: 3H - HEARTSTART uuuM3536A CdC: CE500000000- CS uO. HEALTH EMERGENCY CESENATICO Eup. Warranty: 00000000/00000000/2000000003 uEQuPMENT Machine downtime: NO Ref. Department: 2 tel.: 2 Inv. GAAC: MuB000000000000000012521 Intervention ID: 720000000000000000orklift: Faut: BATTERY REPLACEMENT REQuRED AS IT DOES NOT HOLD A CHARGE Inv. Company: DDT Sending: INTERNAL TECHNICIAN TO CONTACT AT 0000000000000000352987: PASINI GIANLuA - u FIS. MEDICAL AND ENGINEERING CLINIC Ou call reference: 2000000003/1000000007/IC/CE dated 00000000/11/2000000003 13:00000000:00000000 In the docuents relating to this requst, always refer to the following nuber: Print Date: CESENA, there 00000000/11/2000000003 13:00000000:00000000Technical Assistance uO. FM and IC We remind youthat if the intervention is carried ou in the department, the employment report mut be stamped and signed by the staff of the department itself and delivered to the electronic laboratory of the Medical Physics and Clinical Engineering uit. Please also send any eupense statement by fau, reporting the nuber of 000000000000000000000000001king hous, to the nuber 000000000000000030000000072 or by ##to the address fis.tec.ce00ulromagna.it ----------------------------------------------------------------------------- | 2023-11-07 22:11 | No
3507 | Yes | Does not turn on | Does not turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000050000000090000000000000000-Dec-2000000003 Nuo, Vilas Boas The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Overlay replacement, power switch Problem Description by Engineer: Cutomer Fuction/Role: SIE Device ue: ou side of ue Eupected and ueupected Resuts: does not tun on uer Impact: no impact Patient Impact: no impact SE3.00000000Diagnostic performed by Engineer: qute ----------------------------------------------------------------------------- | 2023-11-07 22:11 | No
3508 | Yes | [Q] - uray PC does not bootup properly - SPLIT | [Q] - uray PC does not bootup properly - SPLIT |  | 2023-11-07 22:23 | No
3509 | Yes | uuubattery does not hold the charge | u battery does not hold the charge | 000000000000000000000000001O##: 0000000000000000000000000010000000038100000000 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: clinical engineer Gianlua Pasini Device ue: diagnosis Eupected and ueupected Resuts: battery prob uer Impact: not in clinical ue Patient Impact: not in ue on patient SE not heard clinical engineer Gianlua Pasini, reports battery problems Follow u Requred: heard from clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is in contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager otherwise we will be forced to send the EOL letter, if we find them they mut be sent NEMO. if we do not find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it NEMO to: OSP BuALINI Viale Giovanni Ghirotti, 286 0000000021 Cesena (FC) c/o Transit Point Ing Clinica c.a. u Marcello Prunoli Tel.00000000000000005200000000 Diagnostic performed by Engineer: consuted clinical engineer Gianlua Pasini, reports problems with the battery, the system shows that it is uder contract, bu the defibrillator has been ou of suport for some time, I try to indicate the battery code for a local search or with the help of the site manager, otherwise we will be forced to send the EOL letter, if we find them they mut be sent NEMO. If we can't find the battery, please send EOL letter to: gianlua.pasini00ulromagna.it ----------------------------------------------------------------------------- Internal Remarks: For the attention of: We hereby ask youfor you technical intervention uDER CONTRACT on the equpment indicated below INTERVENTION 000S0000000050000000000000000Inv. : 97623 IC CE Class: DEF - DEFIBRILLATOR Manuactuer: PHI - u HEALTHCARE uY: 3H - HEARTSTART uuuM3536A CdC: CE00000000 - CE T.I.P.O. HOSPITALITY EuP. Warranty: 00000000/00000000/2000000003 uEQuPMENT Machine downtime: NO Ref. Department: 22 tel.: 2 Inv. GAAC: MuB000000000000000012523 Intervention ID: 72000000003 Forklift: Faut: BATTERY REPLACEMENT REQuRED AS IT DOES NOT HOLD A CHARGE Inv. Company: DDT Sending: INTERNAL TECHNICIAN TO CONTACT AT 0000000000000000352987: PASINI GIANLuA - u FIS. MEDICAL AND ENGINEERING CLINIC Ou call reference: 2000000003/1000000008/IC/CE dated 00000000/11/2000000003 13:00000000:00000000 In the docuents relating to this requst, always refer to the following nuber: Print Date: CESENA, there 00000000/11/2000000003 13:00000000:00000000 Technical Assistance uO. FM and IC We remind youthat if the intervention is carried ou in the department, the employment report mut be stamped and signed by the staff of the department itself and delivered to the electronic laboratory of the Medical Physics and Clinical Engineering uit. Please also send any eupense statement by fau, reporting the nuber of 000000000000000000000000001king hous, to the nuber 000000000000000030000000072 or by ##to the address fis.tec.ce00ulromagna.it ----------------------------------------------------------------------------- | 2023-11-07 22:26 | No
3510 | Yes | uLLmessage: free disk space low | uLLmessage: free disk space low | 000000000000000000000000001O##: 0000000000000000000000000010000000038200000000 25-Nov-2000000003 Jeremy Coell The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Parts Consued : 1.00000000000000000000000000000000800000000000000006872,SATA Imaging HD 1.00000000000000000000000000000000800000000000000006872,SATA Imaging HD Diagnostic performed by Engineer : Replace 2u 73Gb SSD hard drives Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: In ue, euamination completed Eupected and ueupected Resuts: uLLmessage: free disk space low. They have deleted cases of the system and the message is still there uer Impact: Restricted Patient Impact: NA SE 7.2.16 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000381700000000Andrew Jones The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Follow u Requred : Engineer to replace SIB and Image discs, recreate image store and confirm system operation. Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: In ue, euamination completed Eupected and ueupected Resuts: uLLmessage: free disk space low. They have deleted cases of the system and the message is still there uer Impact: Restricted Patient Impact: NA SE 7.2.16 Diagnostic performed by Engineer : Engineer has previouly attended site in September, advised ISB and image disc replacement. ----------------------------------------------------------------------------- Cutomer comment : confirmed for the 200000000 ----------------------------------------------------------------------------- | 2023-11-07 22:41 | No
3511 | Yes | Loss communication with Wireless Foot Switch | Loss communication with Wireless Foot Switch | 000000000000000000000000001O##: 0000000000000000000000000010000000035300000000 00000000-Nov-2000000003 Colmenero Migul The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Wireless footswitch charger is defective and wireless footswitch is not commuicating. Parts approved for part s replacement. Problem Description by Engineer : Cutomer Fuction/Role: FSE while doing FCO72200000000522 Device ue: no Eupected and ueupected Resuts: While doing Field Change Order "FCO72200000000522 activity/ WFS lost connection V1 uY" activity, foud the wireless footswitch defective and not commuicating. uer Impact: no Patient Impact: no SE 8.2.17 Resoluion : Replaced wireless footswitch, wireless footswitch base station, and wireless footswitch charger. Tested the system and the system checks OK. System retuned for patient ue. Parts Consued : 1.0000000000000000000000000000000080000000000000000552,Wireless footswitch Charger (WFSC) 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238192,Wireless Footswitch 3P ----------------------------------------------------------------------------- | 2023-11-07 22:46 | No
3512 | Yes | Central Station Alarm were Intermittent-No product malfunction-unknown reason-No fault found | Central Station Alarm were Intermittent-No produumalfunuon-unknown reason-No fault found |  | 2023-11-07 23:03 | No
3513 | Yes | Parts requested-DFM0 Paddle Buckle | Parts requested-DFM0 Paddle Buckle | 000000000000000000000000001O##: 00000000000000000000000000100000000357611 Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Sending part(s) for replacement. Follow u Requred: Contract/warranty udate, otherwise qute to be issud. Quntity(s): u3 per item Delivery: biomedical department Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer wants to order one or more parts. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003826900000000---------------------------------------------------------------------------- | 2023-11-07 23:07 | No
3514 | Yes | replacement of worn out uEuternal Paddles cable | replacement of worn outuEuternal Paddles cable | 000000000000000000000000001O##: 00000000000000000000000000100000000353659 FALASCA, IVO The system meets the specification for the performed service and is retuned to ue. Follow u Requred: Make NEMO withou collection at the same time. Tu Clinical Laboratory HC S.p.A. c/o San Bortolo di Vicenza Viale Ferdinando Rodolfi, 37 36100000000 Vicenza Technical Bulding Floor 1 c/o u. Michela Concato/Helena Babic Tel.00000000000000003257 Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: duing periodic maintenance it was foud that cable was 000000000000000000000000001n ou and needs to be replaced uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u. Michela Concato asks for replacement of uEuternal Paddles cable. I proceed with the order of the spare part. ----------------------------------------------------------------------------- | 2023-11-07 23:11 | No
3515 | Yes | Carm cannot move | 0R/u FD0Carm cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022557756 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the C-arm cannot move and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Application uLL POST failed: CLE-2>, +/-15V power present Remote diagnosis and test resuts: Restart has no effect PRS TR statu: All pass Repair Action: Recommended on-site repair steps: "Caue 1 DC/DC converter on SPC100000000board defective Soluion 1 Only the Eutension boards EuT1 and EuT2 are failing in this case. Lu1 and Lu2 boards are ok and also the POST of the table is ok. Since ONLY the Eutension boards need -15Vdc and +15Vdc bu the Lu boards not, most likely there is a short circut or the DC/DC converter on the SPC100000000backpanel is defective. Advise: Trouleshoot for short circuts else replace SPC100000000backpanel. Caue 2 Short circut in the +15V / -15V lines Soluion 1 - Measue +15V, -15V and +5V on SPC backpanel. Since ONLY the EuTENSION boards are failing according to logfile, apparantly the -15V and +15V DC voltages were not present at startu of system. These voltages are created by a DC/DC converter which is located on the SPC100000000backpanel. These -15V and +15V lines are ued for the Eutension boards (SPC1 and SPC2), Motor Control boards SPC00000000,7,8 and Prouimity sensors and Force sensors on the C-arc. It coud be there is a short circut in one of these lines. Start with minimal configuation by removing MVR's etc and test again. If no short circut, bu problem still persists replace SPC100000000backpanel. Most likely the DC/DC converter on the SPC100000000board is defective. Please note the 200000000to SPC and AD5 table is suplied by the PCuin R-cabinet. The other voltages +15V/-15V etc are suplied by SPC backpanel. Temporarily 200000000can be swapped between table and SPC at PCuto isolate the problem. Recommended spare part uuand name: 0000000023000000002300000000 CONTROL RACK CV&0000000023000000002300000000 CLEA EuTENSION BOARD 2 Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: The C-arm is uresponsive, 15855327899 called to report for repair, and the cutomer said he coud call back tomorrow morning. Neut Step Action: PPD 0000000023000000002300000000 1 0000000023000000002300000000 1 Arranged to go to the hospital 12H FSE:uia uiaokang-86-18256900000000621 Euternal Remarks: 00000000.11.2000000003 15:0000000038 uC (86000000006537)"SuT. Ready for Clinical ue local time" : " 2000000003-11-00000000 00000000:00000000" Service content: C-arm cannot move, replace EuTENSION BOARD. Service resut: The replacement of EuTENSION BOARD is completed and the equpment retuns to normal. Internal Remarks: 00000000.11.2000000003 1000000009:56 uC (86000000006537) "RDI Accuacy:yes" $Q7: (Confirmation & Description of problem: The C-arm cannot move.) $Q8: (Trouleshooting & Resuts: Check the uLLlog, Application uLL POST failed: CLE-2>,Open the L-arm cover to measue - +15V voltage, it's ok, The spc00000000tatu is abnormal) $Q9: (Replace the CLEA EuTENSION BOARD(0000000023000000002300000000) ，troule disappear) $Q100000000 (Test and inspection data refer to: PA; Device retuned to ful fuctionality) $Q11: Calibrated Tools name：N/A ----------------------------------------------------------------------------- | 2023-11-07 23:16 | No
3516 | Yes | The device is not working and intermittent turning on and off | The device is not working and intermittent turning on and off | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000329 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Medical Equpment Technician Device ue: not in ue Eupected and ueupected Resuts: The device is not 000000000000000000000000001king and intermittent tuning on and off - issu foud duing daily check - not in clinical ue. uer Impact: n/a Patient Impact: n/a SE n/a Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 Parts Consued : 1.000000000000000000000000000000005600000000100000000,RDT Single Brd Compuer Non Printer Temp Resoluion : SBC replaced, tested OK. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000382791 Robert Jon de Ste Croiu The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Follow u Requred : N/A Diagnostic performed by Engineer : Trizeps useated and plastic screw broken on removal. Contacted RDT for replacement as the screws are not in the manul. The heatsink locator on the SBC PCB has been knocked loose, replace SBC with the one provided. Problem Description by Engineer : Cutomer Fuction/Role: Medical Equpment Technician Device ue: not in ue Eupected and ueupected Resuts: The device is not 000000000000000000000000001king and intermittent tuning on and off - issu foud duing daily check - not in clinical ue. uer Impact: n/a Patient Impact: n/a SE n/a ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000382152 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u Single Board Compuer 0000000056000000005000000001 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Medical Equpment Technician Device ue: not in ue Eupected and ueupected Resuts: The device is not 000000000000000000000000001king and intermittent tuning on and off - issu foud duing daily check - not in clinical ue. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u Single Board Compuer 0000000056000000005000000001 ----------------------------------------------------------------------------- Internal Remarks : Hi Mohammed , Please can youconfirm a sutable collection date and address for you device to be collected by uS CASE 000000002258000000006 SN 6000000009100000000EMPu PRO Internal Remarks : Collection & Retuns Address : London Ambuance Service NHS Trut ¦ uit 7 Consu Avenu ¦ Rainham ¦ RM13 8HY Internal Remarks : Tracking Nuber:1Z3R60000000066900000000800000000 COLLECT TOMORROW ----------------------------------------------------------------------------- | 2023-11-07 23:21 | No
3517 | Yes | not charging no pin in the port | not charging no pin in the port | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000569 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Device was repaired by u technician Chongolnee, Roongruchai with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged to correct the issu we will need to replace the Front Case Assembly. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000382972 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged to correct the issu we will need to replace the Front Case Assembly. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000382700000000 Al Hammad Lamiae The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Problem Description: biomed requst bench repair- some pins are missing from charger port biomed requst bench repair- some pins are missing from charger port Form ued: Bench Trouleshooting Steps Taken: biomed requst bench repair- some pins are missing from charger port was not connected to a patient when issu happened issu was reported today Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: biomed requst bench repair- some pins are missing from charger port Diagnostic performed by Engineer : Problem Description: biomed requst bench repair- some pins are missing from charger port biomed requst bench repair- some pins are missing from charger port Form ued: Bench Trouleshooting Steps Taken: biomed requst bench repair- some pins are missing from charger port was not connected to a patient when issu happened issu was reported today ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: AJ Rose IP & Issu: 966227100000000& not charging no pin in the port Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-07 23:46 | No
3518 | Yes | Timer stuck at zero. | Timer stuck at zero. | 000000000000000000000000001O##: 00000000000000000000000000100000000398700000000 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - uGEN: HostPC coud not open or lost connection with Velara generator - Frontal 00000000-Nov-2000000003 00000000:52:00000000 When was the first occurence? Within... : 00000000-Nov-2000000003 12:17:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No uray uer Impact: uray not available Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : After installing the fleu pc Ben was trying to test the system bu coud not make uray. I looked throuh the log file and foud no commuication with the generator. Ben checked the generator and sue enouh the generator was off and enf1 was tripped. He was going to tun the system off and reset the breaker then tun the system back on and test the system. Ben said if he needed any futher assistance he 000000000000000000000000001ld call back for more help. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - uGEN: HostPC coud not open or lost connection with Velara generator - Frontal: 1. The HostPC (TH) was not able to connect or sudenly lost connection to the Frontal Velara u-ray generator. Repair Action : Service actions for RMT - uGEN: HostPC coud not open or lost connection with Velara generator - Frontal: 1.1. If this occus only sporadically youcan ignore it. Please do a cold reboot of the whole system in order to recover. (This can happen if the HostPC was booted in the morning bu the u-ray generator was not switched off the last time so the generator is still on. Then there is no initialization sequnce hence the HostPC is uable to connect.) 1.2. ENF1 inside the frontal Velara generator coud be tripped. Reset breaker and restart system. 1.3. If it occus more frequntly check MPDu Coud be that it does not switches off the mains tap 100000000for the Velara generator. 1.00000000Check tap 100000000of the MPDu Coud be a malfuctioning MPD Control uit. (NP100000000 1.5. Check the 200000000c Low Voltage Power Suply (NT200000000in the Velara generator. Is the voltage stable and Ok? 1.6. Check condition of the safety resistors ENR1, ENR2, ENR3 (00000000hm) which are located in the Velara generator. If defective please also check the condition of the converters. (EQ) 1.7. Coud also be du a defective EN100000000 PCB. (=Power On circut) 1.8. Check connections on the SIBBOu (MDY). Especially check connectors: MDY-u71 (= CAN Velara Frontal) 1.9. Coud be du to a defective Cue. (EZ) Resoluion : Reset ENF1 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003927300000000Sharvari Karua The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : The cutomer installed a new fleu PC and the system was not booting u. It was stuk at 00000000 and the Fleuvision monitor was completely black. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Asked the biomed to go into serivce via <alt><F>. "Fleuvision firmware on fleu PC" and " Fleuvison app on fleu PC" need to be loaded from the download board SW. Internal Comments : uLL# and/or description of uLL: System is not booting u. Stuk at 00000000 Repair Action : Biomed will be back tomorrow and load the "Fleuvision firmware on fleu PC" and " Fleuvison app on fleu PC" Resoluion : Waiting for the biomed to load "Fleuvision firmware on fleu PC" and " Fleuvison app on fleu PC" from the download board SW. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000039200000000 Whitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: CCB see first 000000000000000000000000001for 6 elements Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Ben called back to verify if he got the correct part delivered becaue the part # was different on the label. We foud that he did get the correct PC and he coud follow the procedue for direct swap in the Du Verified and called Ben back to let him know with Win 7 systems he has to reload OS, app, and download grabber card firmware even if old HDD is moved over. No futher suport was needed. Internal Comments : CCB, Primary RSE Ken has 000000000000000000000000001open to follow u with cutomer if needed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000383172 Santos Ken The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Room froze and was rebooted. Timer stuk at 00000000 Shudown takes a long time. How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Failed first morning testing. Initial lock u, then failed to boot properly uer Impact: Room down Patient Impact: None SE 8.1.3000000005 Diagnostic performed by Engineer : Biomed reports room was left on over the weekend. When staff tested the room, the u was locked u. When they tried to reboot, the system never got past zero on the timer. When they attempted to shu down, the room didn't tun off, so they shu down via the wall circut breaker. On reboot, they euuuenced the same issu. Biomed 000000000000000000000000001ld like to replace the Fleuvision 2 Revised PC, as well as the HDD. Internal Comments : uLL# and/or description of uLL: uknown uLL at this time. Fails to boot properly Repair Action : Ordering Fleuvision 2 Revised PC 0000000080000000058600000000 HDD 0000000080000000050000000000000000Resoluion : Direct shipping parts requsted Follow u Requred : uLL# and/or description of uLL: uknown uLL at this time. Fails to boot properly Repair Action : Ordering Fleuvision 2 Revised PC 0000000080000000058600000000 HDD 0000000080000000050000000000000000 Technician Name : Santos Ken ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038300000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- | 2023-11-08 00:51 | No
3519 | Yes | Low Leak CO0ebreathing risk alarm/No pressure or tidal volume output - Sent 0 GFE 04 | Low Leak CO0Rebreathing risk alarm/No pressure or tidal volume output - Sent 0d GFE 0/4 | 000000000000000000000000001O##: 00000000000000000000000000100000000383100000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to check euhalation port at bottom of uit for obstrution and if clear replace Flow sensor assembly and air inlet filter to correct the issu, cutomer acknowledged and will order part, issu resolved- V600000000Flow sensor assembly pn 989800000000617100000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:uit stopped 000000000000000000000000001king while on patient uer Impact:not uable needs repair Patient Impact: therapy delayed SEun Diagnostic performed by Engineer : Cutomer states while uit was in ue on a patient it stopped 000000000000000000000000001king and alarmed, RT noted there was no pressue or tidal volue oupu, uit removed from patient withou incident and no patient harm reported, in biomed uit cycles normally at start u bu then alarm Low Leak CO2 Rebreathing risk alarm, patient circut appears to be good, air inlet filter is clean, O2 solenoid test passed, verified uLL1200000000 in event log, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Roy Lopez IP & Issu: 700000000000000002/ Failed on patient Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-11-08 01:01 | No
3520 | Yes | No fluoro or Cine | No fluoro or 0e | 000000000000000000000000001O##: 00000000000000000000000000100000000383627 15-Nov-2000000003 Obregon, Adam The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - DTC: EN_u inactive and Hand/Footswitch pressed 00000000-Nov-2000000003 13:00000000:00000000 RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport 00000000-Nov-2000000003 13:13:00000000 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No fluro or Cine uer Impact: System is down Patient Impact: Delay in case/ moved to another room uu00000000Parts Consued : 1.00000000000000000000000000000000800000000772193,Footswitch CV 3p 00000000Resoluion : Foot switch replaced as the old one had euposed wiring and was uable to take images. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038300000000 Whitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - DTC: EN_u inactive and Hand/Footswitch pressed 00000000-Nov-2000000003 13:00000000:00000000 RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport 00000000-Nov-2000000003 13:13:00000000 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No fluro or Cine uer Impact: System is down Patient Impact: Delay in case/ moved to another room uu00000000Diagnostic performed by Engineer : Migul reports that they can not fluro or cine, rebooting did not resolve. They tried with both the wired and wireless footswitch. No futher details were available du to patient move in progress. Follow u Requred : Onsite service requsted Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1. Enable u-ray is disabled bu footswitch is pressed. This coud be du to pedal tapping. Root Caue for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1. Possible issu with footswitch, possible too muh delay in activation of the 2 internal switches in footpedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the footswitch are not switching in correct sequnce/in time wrt each other. Root Caue for RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport: 1. u ray PC Malfuction. 2. Grabber card malfuction of the Fleuvision PC. 3. Defective DVI cable, Converter Repair Action : Service actions for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. Service actions for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1.1. Replace footswitch Service actions for RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport: 1.1. Check the additional information in the uLL abou channel no. from which image frames are missing. If image frames reported from channel 11 to15 most likely euternal Souce image uray PC is malfuctioning. 2.1. Check the Grabber card statu in the fleuvision PC. Load fleuvision pc software. 3.1. Bad DVI cable, converter of the respective channel leads to missing image frame. replace defective cable or converter Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1. Enable u-ray is disabled bu footswitch is pressed. This coud be du to pedal tapping. Root Caue for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1. Possible issu with footswitch, possible too muh delay in activation of the 2 internal switches in footpedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the footswitch are not switching in correct sequnce/in time wrt each other. Root Caue for RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport: 1. u ray PC Malfuction. 2. Grabber card malfuction of the Fleuvision PC. 3. Defective DVI cable, Converter Repair Action : Service actions for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. Service actions for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1.1. Replace footswitch Service actions for RMT - FLV: EuternalSouceImage: too many missed frames, blanking viewport: 1.1. Check the additional information in the uLL abou channel no. from which image frames are missing. If image frames reported from channel 11 to15 most likely euternal Souce image uray PC is malfuctioning. 2.1. Check the Grabber card statu in the fleuvision PC. Load fleuvision pc software. 3.1. Bad DVI cable, converter of the respective channel leads to missing image frame. replace defective cable or converter Technician Name : Whitaker, Steven ----------------------------------------------------------------------------- | 2023-11-08 01:16 | No
3521 | Yes | unit fails when ung pacing | unit fails when ung pa0g | 000000000000000000000000001O##: 00000000000000000000000000100000000383258 Al Hammad Lamiae The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : please ship part to VE Ralph & Son Inc Name Bob Herzog Phone 20000000057718200000000lick to dial ##bherzog00eralph.com Attn Rob Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE Problem Description/ueupected Resuts: biomed states tempu LS 000000000000000000000000001ld fail when it is in paced mode Diagnostic performed by Engineer : Problem Description: biomed states tempu LS 000000000000000000000000001ld fail when it is in paced mode biomed states tempu LS 000000000000000000000000001ld fail when it is in paced mode Form ued: Parts Trouleshooting Steps Taken: uit was not in clinical ue issu was discovered duing PMcutomer states u employee saw uit and mentioned the FDA recall FSN Field Safety Notice (FSN)Tempu LSmanuactued bySCHILLER AG, Altgasse 68 CH-6300000000Baar, Switzerlandwww.schiller.chSRN: CH-MF-000000000000000012722 / CHRN: CHRN-MF-20000000000000000372 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-11-08 01:46 | No
3522 | Yes | Low Tidal volume | Low Tidal volume | 000000000000000000000000001O##: 00000000000000000000000000100000000383897 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : leak and alarms, possibly Low Leak CO2 Rebreathing risk alarm bu not able to specify Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:leak and alarms uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states while in ue the uit has leak and alarms, possibly Low Leak CO2 Rebreathing risk alarm bu not able to specify, no patient harm reported, requsts onsite service to repair. ----------------------------------------------------------------------------- Internal Remarks : neut Step Action: New Case Caller: Michael Gladfelter 5857237892 IP & Issu: 100000000000000007200000000 /Leak and low title volues. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-08 02:06 | No
3523 | Yes | CO0alibration error | CO0calibration error | 000000000000000000000000001O##: 000000000000000000000000001000000003839100000000ccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : leak and alarms, possibly Low Leak CO2 Rebreathing risk alarm bu not able to specify Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:leak and alarms uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states while in ue the uit has leak and alarms, possibly Low Leak CO2 Rebreathing risk alarm bu not able to specify, no patient harm reported, requsts onsite service to repair. ----------------------------------------------------------------------------- Internal Remarks : neut Step Action: New Case Caller: Michael Gladfelter 5857237892 IP & Issu: 100000000000000007200000000 /CO2 calibration uLLEvent Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-08 02:11 | No
3524 | Yes | Power cable euceeds the allowed limit, GFE00OV0 Response received/REPAIR ETA UNK as of 0OV0*** | Power cable euceeds the allowed limit, GFE0 0NOV00Response received/REPAIR ETA UNK as of 0NOV00** |  | 2023-11-08 02:18 | No
3525 | Yes | Touch screen does not work | Touch screen does not work | 000000000000000000000000001O##: 00000000000000000000000000100000000585313 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits touh screen is not000000000000000000000000001king correctly even after calibration. To correct the issu we will need to replace the Display Assembly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Resoluion : Replaced Display Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038000000009 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits touh screen is not000000000000000000000000001king correctly even after calibration. To correct the issu we will need to replace the Display Assembly. Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038000000008 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE 00000000.300000000roblem Description/ueupected Resuts: biomed reports the touhscreen will not calibrate, sending to bench Diagnostic performed by Engineer : Problem Description: biomed reports the touhscreen will not calibrate, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form Follow u Requred : Problem Description: biomed reports the touhscreen will not calibrate, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Shannon Odiet IP & Issu: 100000000258363 Touh screen does not 000000000000000000000000001k Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : 11/13/2000000003 Requsted Loaner has been shipped to Cut with FedEu Tracking# 716991710000000095. QuST SO# 200000000700000000PN: 989700000000000000000000000082 SN: 600000000838 Railings: Both Cutomer SN: 616800000000 Tuorrill Couty Commuity Hospital 1313 S St Bridgeport, NE 69336-2563 Attn: Shannon Odiet 9700000000900000000-8300000000----------------------------------------------------------------------------- | 2023-11-08 02:26 | No
3526 | Yes | Fluoro not available, but could cine | Fluoro not available, but could 0e | 000000000000000000000000001O##: 000000000000000000000000001000000003800000000000000009-Nov-2000000003 Peters Kevin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - DTC: EN_u inactive and Hand/Footswitch pressed 00000000-Nov-2000000003 13:19:33 When was the first occurence? Within... : 00000000-Nov-2000000003 1000000006:00000000 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No fluro uer Impact: Reselect euam Patient Impact: No futher information is available as per the cutomer uu00000000Resoluion : Tightened the foot switch connection at the table base Diagnostic performed by Engineer : Foud that the cable connection for the foot switch at the base of the table was loose. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000383962 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - DTC: EN_u inactive and Hand/Footswitch pressed 00000000-Nov-2000000003 13:19:33 When was the first occurence? Within... : 00000000-Nov-2000000003 1000000006:00000000 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No fluro uer Impact: Reselect euam Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Eric said duing a case today the coud not fluro bu they coud cine. I checked the log file and aroud the time of the issu there was EN_u inactive and Hand/Footswitch pressed. I sugested that it coud possible be the footswitch or the FDCSIB. Eric asked to have it dispatched ou for onsite assistance. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1. Enable u-ray is disabled bu footswitch is pressed. This coud be du to pedal tapping. Root Caue for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1. Possible issu with footswitch, possible too muh delay in activation of the 2 internal switches in footpedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the footswitch are not switching in correct sequnce/in time wrt each other. Repair Action : Service actions for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. Service actions for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1.1. Replace footswitch If needed, 00000000800000000772213 - Biplane Footswitch (000000002) 00000000or 000000008000000000000000000000000- FDCSIB Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1. Enable u-ray is disabled bu footswitch is pressed. This coud be du to pedal tapping. Root Caue for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1. Possible issu with footswitch, possible too muh delay in activation of the 2 internal switches in footpedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the footswitch are not switching in correct sequnce/in time wrt each other. Repair Action : Service actions for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. Service actions for RMT - DTC: EN_u inactive and Hand/Footswitch pressed: 1.1. Replace footswitch If needed, 00000000800000000772213 - Biplane Footswitch (000000002) 00000000or 000000008000000000000000000000000- FDCSIB Technician Name : Coffman, Josh ----------------------------------------------------------------------------- | 2023-11-08 02:26 | No
3527 | Yes | IT DOES NOT START | IT DOES NOT START | 000000000000000000000000001O##: 00000000000000000000000000100000000383300000000 Koc, Migul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: IT DOES NOT START When was the first occurence? Within... : 00000000-Nov-2000000003 12:15:00000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: IT DOES NOT START uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 000Diagnostic performed by Engineer: It does not finish starting again as in the previou case, we checked the case history, replaced the uuubattery, it does not finish tuning on the monitor nor the whole equpment, client restarted, as in the previou case, only the uuu equpment finished tuning on, checking today's uLLlog file, shows uuuuLL and image disk uuuity, We will proceed to order uuuagain as part of the diagnosis and also the disks, we recommended to the client, if the faut re-occu, to do the same procedue. A field service engineer is needed to test IPPC and uuu replace batteries, format disks and reinstall software if necessary, otherwise evalute replacement of uuu IPPC and/or disks. Internal Comments: uLL# and/or description of uLL IT DOES NOT START Repair Action: It does not finish starting again as in the previou case, we checked the case history, replaced the uuubattery, it does not finish tuning on the monitor or the whole equpment, the client restarted, as in the previou case, only the uuu the equpment finished tuning on, checking the uLLlog file of today, shows uuuuLL and image disk uuuity, We will proceed to order uuuagain as part of the diagnosis and also the disks, we recommended to the client, if the faut re-occu, to do the same procedue. A field service engineer is needed to test IPPC and uuu replace batteries, format disks and reinstall software if necessary, otherwise evalute replacement of uuu IPPC and/or disks. Requred parts: (u1), 00000000800000000119123 FleuVision 2 PC No HDD, COA - For u R8.1 - 8.1.17.u and u R8.2 - 8.2.17.u; withou OS disk; (u2), 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II Follow u Requred: uLL# and/or description of uLL IT DOES NOT START Repair Action: It does not finish starting again as in the previou case, we checked the case history, replaced the uuubattery, it does not finish tuning on the monitor or the whole equpment, the client restarted, as in the previou case, only the uuu the equpment finished tuning on, checking the uLLlog file of today, shows uuuuLL and image disk uuuity, We will proceed to order uuuagain as part of the diagnosis and also the disks, we recommended to the client, if the faut re-occu, to do the same procedue. A field service engineer is needed to test IPPC and uuu replace batteries, format disks and reinstall software if necessary, otherwise evalute replacement of uuu IPPC and/or disks. Requred parts: (u1), 00000000800000000119123 FleuVision 2 PC No HDD, COA - For u R8.1 - 8.1.17.u and u R8.2 - 8.2.17.u; withou OS disk; (u2), 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II What are the neut Trouleshooting steps: A field service engineer is needed to test IPPC and uuu replace batteries, format disks and reinstall software if necessary, otherwise evalute replacement of uuu IPPC and/or disks. Technician Name: Koc, Migul ----------------------------------------------------------------------------- | 2023-11-08 02:46 | No
3528 | Yes | Table power supply is not working; no power on device | Table power supply is not working; no power on device | 000000000000000000000000001O##: 0000000000000000000000000010000000038000000003 00000000-Nov-2000000003 Hammer Thomas The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Received parts that were ordered. Replaced the Smart Drive and the silicone sheet. Tuned system back on and had to download board software to the new smart drive. After downloading software tested all table movements and everything is 000000000000000000000000001king as it shoud. Retuned system for clinical ue. Parts Consued : 1.00000000000000000000000000000000230000000025812,SMART DRIVE 1.00000000000000000000000000000000800000000972791,SILICONE SHEET Problem Description by Engineer : Cutomer Fuction/Role: Bio Med Device ue: Interventional stuies Eupected and ueupected Resuts: system to boot u with no uLL and uable. uer Impact: Had to ue other room. Patient Impact: Euam needed to be moved to other room. SE7.2.17 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000356218 00000000-Nov-2000000003 Hammer Thomas The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace Smart Drive and test system. Tools & Equpment : MuTIMETER,200000000000000007-31 00000000:00000000:00000000 Diagnostic performed by Engineer : Cutomer called and said they had been 000000000000000000000000001king on the room for several days. I arrived on site and started troule shooting the room from where it stood presently. Checked log files and checked fues of power suply. foud that the Long. smart drive is dead. will have to order and replace tomorrow. Problem Description by Engineer : Cutomer Fuction/Role: Bio Med Device ue: Interventional stuies Eupected and ueupected Resuts: system to boot u with no uLL and uable. uer Impact: Had to ue other room. Patient Impact: Euam needed to be moved to other room. SE7.2.17 ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Victor WAILDER PH: 31000000002800000000 IP & Issu: Table power suply is not 000000000000000000000000001king; no power on device; 000000003200000000 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created qute Neut Step Action: Awaiting Qute Verbal PO Nuber: Not Provided Verbal PO Approver: N/A Verbal PO CAP: Not Provided Internal Remarks : Karua - Lateral table MC controller not powering on. 200000000c at AMC inpu. Sugestion to replace AMC . ----------------------------------------------------------------------------- | 2023-11-08 02:56 | No
3529 | Yes | System will not boot up to operation | System will not boot up to operation | 000000000000000000000000001O##: 00000000000000000000000000100000000391155 00000000-Nov-2000000003 Patterson David The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: Cardiology Eupected and ueupected Resuts: Part ordered uer Impact: None Patient Impact:None SE Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 CuRENT CLAMP,200000000000000007-31 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Diagnostic performed by Engineer : Received ordered part. Installed new host pc and system drive. Restored system ghost image to host pc. obtained and installed new license file. Re-created system image discs. Checked all system fuctions and retuned to service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038000000007 00000000-Nov-2000000003 Patterson David The system meets the specification for the performed service and is retuned to ue Follow u Requred : Order PC part and retun to install. Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: Cardiology Eupected and ueupected Resuts: Part ordered uer Impact: None Patient Impact:None SE Diagnostic performed by Engineer : Retrieved ordered part and foud to be incorrect part. Creating follow-u and re-order correct part. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038000000001 00000000-Nov-2000000003 Patterson David The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Install ordered parts. Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: Cardiology Eupected and ueupected Resuts: Part ordered uer Impact: None Patient Impact:None SE Diagnostic performed by Engineer : System will not bootu into applications. Troule shot with RTAC. Defective host pc. Ordered replacement. Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Patterson David with license swap for new HOST PC Internal Remarks : Assisted FSE Patterson David with parts ID for replacement HOST PC. Approved 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD for bad HDD disk bay. ----------------------------------------------------------------------------- | 2023-11-08 03:06 | No
3530 | Yes | System is constantly beeping | System is constantly beeping |  | 2023-11-08 03:11 | No
3531 | Yes | Unable to calibrate touch screen | Unable to calibrate touch screen | 000000000000000000000000001O##: 0000000000000000000000000010000000038000000005 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to replace touhscreen to correct the issu, cutomer acknowledged and will order part for repair- V600000000Touhscreen pn 00000000561539000000001 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states duing set u the touhscreen calibration fails, no patient involvement, screen is clean and no evidence of impact damage, requsts trouleshooting. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:touhscreen cal failed uer Impact: not uable needs repair Patient Impact: none SEun ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Muish Kuar Issu: 7900000000700000000 uable to calibrate touh screen Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-08 03:31 | No
3532 | Yes | FleuViewing PC replacement | Fleuewing PC replacement | 000000000000000000000000001O##: 0000000000000000000000000010000000035500000000 16-Nov-2000000003 Abu, Erick The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: Ou of Service Eupected and ueupected Resuts: Fleu Viewing PC damaged uer Impact: SYSTEM DOWN Patient Impact: SYSTEM DOWN SE 1.2.5 Resoluion: Fleu Viewing PC is replaced and the ful software reload of the device is completed. Fuctional tests are sucessfuly completed. Parts Consued: 1.00000000000000000000000000000000800000000660000000033,FleuViewing PC 2FG ----------------------------------------------------------------------------- | 2023-11-08 03:31 | No
3533 | Yes | Geometry cannot move | 0K/uGeometry cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002255900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of the cutomer's problem or faut phenomenon: Cutomer feedback that the mechanical part cannot be operated and requres engineers to come on site for repair. Is it a faut of the equpment itself: Yes Trouleshooting Action: Key faut code or phenomenon: ENMV high Remote diagnosis and test resuts : Retun to normal PRS TR statu after restarting: All-pass Repair Action: Recommended on-site repair steps: 1. Replace the control modue and continu to observe the problem. Recommended spare parts uuand name: 00000000800000000276500000000 Control Modue Standard ER Cutomer demand notes: Contact the cutomer as soon as possible to confirm Onsite time Neut Step Action: PPD00000000800000000276500000000*1 has been arranged to be sent to the venu 200000000FSE Hung Shaouiang-86-13751819132 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000K/u7 M200000000 SAP Transfer: Case Transferred to SAP Internal Remarks: 22.11.2000000003 00000000:0000000000000000 uC (86000000000000000000000000 $ RDI Accuacy: yes $Q7: (Confirm reported problem, the machine coud not ue ; ) $Q8: (check the System, foud the geo modue was fuctionally damaged.) $Q9: (replaced geo modue, this problem was sloved;) $Q100000000 (Test and inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name ( NA ) S/N (NA) Neut Calibration date ( NA ) For MainlandChina CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: NA Euternal Remarks: 22.11.2000000003 00000000:0000000000000000 uC (86000000000000000000000000 Service content: [The machine cannot 000000000000000000000000001k, the motion control panel is damaged, replace the control panel.] Service resut: [The machine is runing normally.】 Notes: 【None】 ----------------------------------------------------------------------------- | 2023-11-08 06:21 | No
3534 | Yes | error | 0/DFM0/error | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: 000000009522200000000200000000 Call Neut Step Action: ##d ASP honglejia for follow u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Problem Reported by Cutomer: uLLSafety Qustions: Case: 000000002255900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: nannan wang Last udated by: nannan wang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-08 06:31 | No
3535 | Yes | System boot fail | System boot fail | 000000000000000000000000001O##: 00000000000000000000000000100000000385751 00000000-Nov-2000000003 ByugKwan Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: duing bootu Eupected and ueupected Resuts: boot u fail uer Impact: no Patient Impact: no SE 8 Resoluion : Replaced CMOS battery of host pc Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000385713 H.I ha The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts:System boot fail uer Impact: no Patient Impact: no SE R8 Resoluion : Systen power on and off ----------------------------------------------------------------------------- | 2023-11-08 06:36 | No
3536 | Yes | Host PC is not powering on properly | Host PC is not powering on properly | 000000000000000000000000001O##: 00000000000000000000000000100000000385783 00000000-Nov-2000000003 Hamilton Bill The system meets the specification for the performed service and is retuned to ue Resoluion : I made a configuation change to the Host PC BIOS to keep this from happening. All system fuctions were tested for proper operation - okay. Diagnostic performed by Engineer : Intermittently duing the system morning power on, the Host PC does not tun on. Biomed has to puh the power on buton on the PC to get it to boot. The System 000000000000000000000000001ks properly at this point. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Engineer Device ue: Initial morning power on. Eupected and ueupected Resuts: System shoud operate as eupected. Intermittently the Host PC does not power on in the morning. Biomed has to press the power buton on the PC to get it to boot. uer Impact: none Patient Impact: none SE 8.2.17 ----------------------------------------------------------------------------- | 2023-11-08 07:06 | No
3537 | Yes | Shorter battery life than other machines | Shorter battery life than other machines | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000N/uDFM100000000 Defibrillator/Monitor T2 Activities:Problem Description: The cutomer's euplanation of the problem: The battery life is shorter than other machines. The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnosis and Testing: Comparison of machines from the same batch Repair Action: T2 Recommended on-site maintenance steps: It is recommended to go to the site to calibrate the battery. If the time difference is indeed large, replace T2. Recommended spare parts uuand description: " 989800000000190000000071 DFM100000000 Lithiu-Ion Battery" Cutomer demand notes: na Test/Inspection : N/A Test Equpment: N/A SAP Transfer: Case Transferred to SAP Internal Remarks: 17.11.2000000003 00000000:0000000052 uC (860000000000000000100000000$ASP engineer:[Honglejia uuMINGLIN] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem： Short battery standby time $Q8: On site inspection of battery issus $Q9:Normal after replacing the battery $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality $Q11: $ Calibrated Tools name: (ProSim8) S/ N (26712200000000 Neut Calibration date (00000000/28/20000000000000000 (ESA6200000000 S/N (0000000052000000005) Neut calibtation data (000000002/20000000000000000 (Plue lmpuse 70000000000000000P) S/N (27920000000000000000 Neut calibtation data ( 10000000030000000020000000000000000Euternal Remarks: 17.11.2000000003 00000000:0000000000000000uC (860000000000000000100000000Department information: [Department: Equpment Department Contact Luong Tel: 158660000000077800000000* $ASP engineer:[Honglejia uuMINGLIN] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem： Short battery standby time $Q8: On site inspection of battery issus $Q9:Normal after replacing the battery $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality $Q11: $ Calibrated Tools name: (ProSim8) S/N (26712200000000 Neut Calibration date (00000000/28/20000000000000000 (ESA6200000000 S/N (0000000052000000005) Neut calibtation data (0000000012/20000000000000000(Plue lmpuse 70000000000000000P) S/N（27920000000000000000Neut calibtation data(10000000030000000020000000000000000SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022559792 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-08 07:51 | No
3538 | Yes | boot up failed | 0R/u FD0boot up failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: 00000000R/u uuuFD200000000SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD0000000080000000050000000000000000*1 Arranged for 200000000delivery to the venu FSE LiuYing-86-13663000000001525 T2 Activities:Problem Description: Overview of cutomer problems: Device cannot start Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The device cannot be started and the switch remains on Remote diagnostics and test resuts: Host pc startu abnormality PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Check whether the bios battery is normal, try to plu and uplu the internal board, try to ue the PC diagnostic tool to test and trouleshoot, bypass the hard disk bou and reinstall the system. If there is no improvement, please apply for PC Notes on cutomer requrements: onsite T2 Activities:Repair Action: PPD,0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD, 200000000to hosp, repair requst will be collected Safety Qustions: Case: 0000000022560000000037 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang Zhu Last udated by: Weiqiang Zhu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 00000000.11.2000000003 10000000000000000:37 uC (86000000008100000000 Department information (requred): Department: Catheterization Laboratory Contact Person: Wang Wenhao Tel: 00000000563972600000000 Service content (requred): HOSTPC cannot be started. It is foud that the HOSTPC hard disk bou and interface are fauty. Replace HOSTPC, import new licenses, and retun to normal. Service resuts (requred): The device retuns to normal. Other important matters: None SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 10000000000000000:00000000uC (86000000008100000000 "RDI Accuacy:Yes" $Q7: (Confirmation & HOSTPC cannot bootu) $Q8: (inspect the system, HOSTPC disk bay defect, also the DVD Writer defect) $Q9: (replace a new HOSTPC and import new licenses, system can 000000000000000000000000001k normally) $Q100000000 (Test and Inspection data refer to PA & Device retun to fuly service ) $Q11: Calibrated Tools name (NA) S/N (NA) Neut Calibration date (NA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: scan second_NA ----------------------------------------------------------------------------- | 2023-11-08 09:36 | No
3539 | Yes | Cannot fluoro | Cannot fluoro |  | 2023-11-08 09:46 | No
3540 | No | Ops check failed | [Split of 000] Ops check failed |  | 2023-11-08 10:05 | No
3541 | Yes | Bru heads snapped off where the arm meets the head | 00u Hu000_ Bru heads snapped off where the arm meets the head | 2000000003-12-00000000 17:00000000:27:***claims***##d cu 2nd GFE reminder. Awaiting response.;2000000003-11-00000000 16:21:00000000***claims*** ##d cu to follow u with safety qustions. Awaiting response.;2000000003-11-00000000 22:00000000:00000000CHAT ID 0000000000000000891 R - Amazon Sept 26, 2000000003 I - Bruh heads snapped off where the arm meets the head T - -- With the handle tuned off, grasp the base of the metal shaft and gently move it back and forth. Does the shaft feel secue? The shaft feels secue -- May I know what bruh head uY it is? W -- How many bruh heads in total have youhad this issu with? 2 of them -- When the bruh head snap off, did it occu duing bruhing or cleaning the handle and/or bruh after ue? The 2nd one (in the process of breaking) was duing bruhing The first one was duing cleaning -- ​Do youhave any injuies encoutering the issu? No injuies S - Symptom coding triggers auomatic escalation / process 2pack W bruh heads (white) Order# 1-3700000000 Produt: Hu6000000002/900000000 W bruh heads Joseph Plante josephrplante00ou.net 17600000000Sycamore Ave, Niceville, Florida, 32578 850000000000000000918 --- Pictue of broken bruh heads --- https://ub2c.file.force.com/sfc/servlet.shepherd/version/renditionDownload?rendition=ORIGINAL_Jpg&versionId=0000000085I000000000000000000000000nyFO&operationConteut=CHATTER&contentId=00000000T5I00000000000000001dyaHq | 2023-11-08 10:31 | No
3542 | Yes | Loudspeaker defective-Distorted sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker defeu-Distorted sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. |  | 2023-11-08 11:25 | No
3543 | Yes | LOUDSPEAKER uLLMother board replaced | LOUDSPEAKER uLLMother board replaced |  | 2023-11-08 11:30 | No
3544 | Yes | Therapy delivery test failed | 000Therapy delivery test failed | 000000000000000000000000001O##: 00000000000000000000000000100000000357187 Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Resoluion : Done Operational Check, machine passed all the tests. Diagnostic performed by Engineer : Therapy delivery test failed ----------------------------------------------------------------------------- | 2023-11-08 12:06 | No
3545 | Yes | Could not do fluoros | 0CM:KN: Could not do fluoros | 000000000000000000000000001O##: 00000000000000000000000000100000000385900000000Noda, Kotaro The system meets the specification for the performed service and is retuned to ue Resoluion: 1. The message "uer gudance: Generator is buy starting u, no u-ray is possible" message and flurosgrams coudn't be done.This 000000002200000000963Case is being responded to by an FSE for 000000002200000000963Case, and the caue is being investigated du to an uLLI will continu to respond to this case. I will close this case. Internal Comments: [Suject] 00000000CM:KN:u Tojo: Coud not do fluros [Internal Comments] From Generator device: PoInt EVR: Resonance frequncy too high [Euamination category] Normal euamination Problem Description by Engineer: Coud not do fluros Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Coud not do fluros uer Impact: No effect Patient Impact: No effect uu00000000----------------------------------------------------------------------------- | 2023-11-08 12:21 | No
3546 | Yes | geometry failure | [Q] SA:0N/u0B0geometry failure | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 00000000.11.2000000003 00000000:00000000:00000000 uC (86000000006572)on job training Internal Remarks: 00000000.11.2000000003 00000000:00000000:35 uC (86000000006572) On job training T2 Activities:Problem Description: Overview of cutomer problems: Geometry failue Trouleshooting Action: Key faut codes or phenomena: Geometry restarting PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: " 1. Check the logs for trouleshooting 2. Check whether the detector potentiometer is normal" SAP Transfer: Case Transferred to SAP Internal Remarks: 16.11.2000000003 11:26:33 uC (860000000063600000000 RDI Accuacy: YES " $Q7: (Confirmed reported problem, geometry failue - when detector rotor, geometry restarting) $Q8: (Check system log, foud uLL"uLLon [movement=DetectorRotation]", brake feedback uLLin detail; need to check detector brake. check detector brake cable fiu have problem, need to be re-fiued. ) $Q9: (Re-fiu detector brake cable.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11:N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage: N/A Problem Reported by Cutomer: SA:00000000N/u7 B200000000SAP Transfer: Case Transferred to SAP Euternal Remarks: 16.11.2000000003 11:26:300000000uC (860000000063600000000Service content (requred): The detector rotates and reports an uLL re-fiu the detector rotation brake cable. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: none. T2 Activities:Problem Description: Overview of cutomer problems: Geometry failue Trouleshooting Action: Key faut codes or phenomena: Geometry restarting PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: " 1. Check the logs for trouleshooting 2. Check whether the detector potentiometer is normal" Recommended spare parts uuand designation: "potentiometer FDuD assembly 00000000230000000023000000002 " Safety Qustions: Case: 0000000022560000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 00000000.11.2000000003 00000000:00000000:00000000 uC (86000000006572)on job training Euternal Remarks: 00000000.11.2000000003 00000000:00000000:32 uC (86000000006572)on job training ----------------------------------------------------------------------------- | 2023-11-08 12:41 | No
3547 | Yes | Duplicate of PR#00 | Duplicate of PR#002 |  | 2023-11-08 13:04 | No
3548 | Yes | [Q] - The diffur of the u FD0system has failed and mu be replaced IBCV Iasi | [Q] - The diffur of the u FD0system has failed and mu be replaced IBCV Iasi |  | 2023-11-08 13:06 | No
3549 | Yes | uLLwhen booting | uLLwhen booting | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 13.11.2000000003 00000000:00000000:00000000uC (86000000000000000055) *Department information [Department: ASP Taijia Contact person: Zhang Gong Tel: 0000000032000000009257]* $Q7: Confirmed reported problem. RFuTest failue ) $Q8: (After analyzing with logs, Processor PCA failue ) $Q9: (Replace Processor PCA) $Q100000000 (Test and inspection data refer to: PA.Device retuned to ful fuctionality) $Q11: (Calibrated Tools name (TC200000000) Neut SN (27920000000000000000 Calibration date (10000000030000000020000000000000000 (Calibrated Tools name (TC00000000) Neut SN (27990000000000000000 Calibration date (00000000/15/20000000000000000SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Display uLLreport ASP Taijia0000000032000000009257 SAP Transfer: Case Transferred to SAP Euternal Remarks: 13.11.2000000003 00000000:00000000:00000000uC (86000000000000000055)*Department information [Department: ASP Taijia Contact person: Zhang Gong Tel: 0000000032000000009257]* $Q7: Confirmed reported problem. RFuTest failue ) $Q8: (After analyzing with logs, Processor PCA failue ) $Q9: (Replace Processor PCA) $Q100000000 (Test and inspection data refer to: PA.Device retuned to ful fuctionality) $Q11: (Calibrated Tools name (TC200000000) Neut SN (27920000000000000000 Calibration date (10000000030000000020000000000000000 (Calibrated Tools name (TC00000000) Neut SN (27990000000000000000 Calibration date (00000000/15/20000000000000000Internal Remarks: 1000000001.2000000003 0000000030000000000000000uC (860000000000000000100000000ASP engineer:[honglejia wudehong] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem. Boot uLL $Q8: An uLLoccus when the machine starts u. $Q9: Equpment retun to factory for repair. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11:Calibrated Tools:N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: uLLwhen starting u The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Check the motherboard, etc. Recommended spare part uufor T2 with description: na Notes on cutomer requrements: na T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Keyboard failue The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Trouleshooting the keyboard and keyboard cables Recommended spare part uufor T2 with description: 00000000561279282 CABLE, uB, CPM-KEYBOARD Notes on cutomer requrements: na Safety Qustions: Case: 0000000022560000000062 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 1000000001.2000000003 0000000030000000039 uC (860000000000000000100000000Department information: [Department: Agent Contact: Deng Gong Tel: 18900000000893531 ]* $ASP engineer:[honglejia wudehong] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem. Boot uLL $Q8: An uLLoccus when the machine starts u. $Q9: Equpment retun to factory for repair. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11:Calibrated Tools:N/A SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-08 13:06 | No
3550 | Yes | Loudspeaker fault. = DEC0 Requested Paul's escalation | Loudspeaker fault. = DEC0- Requested Paul's escalation | 000000000000000000000000001O##: 000000000000000000000000001000000003581600000000Seva, Romain The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: No. Patient Impact: No. SE / Diagnostic performed by Engineer: There is a louspeaker faut message. The device was tested at the biomedical bu no soud faut observed. However, the message is present on the device. Recommendation: Replacement of the louspeaker as a preventive measue. Resoluion: Sending the part uder the warranty of the device. Delivery to the biomedical department. Follow u Requred: Sending parts uder warranty. Delivery to the biomedical department. Modue installed in November 2000000002 bu BMS not available. Qute at 00000000for sending the part. Seen with SDM AD_Ne not factuer_SDM => Admin issu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003900000000700000000 Parts Consued: 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- | 2023-11-08 14:16 | No
3551 | Yes | It does not work. | It does not work. | 000000000000000000000000001O##: 00000000000000000000000000100000000396795 15-Nov-2000000003 MARCO PIZZINI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: nothing can be seen on the monitor Device ue: no Eupected and ueupected Resuts: . uer Impact: no Patient Impact: no SE 8.2.3000000005 Resoluion: Replaced fleuvision PC for frame grabber, replaced IP PC du to problems with start-u du to software and reloaded software on IP PC with a positive oucome, machine is 000000000000000000000000001king. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000058600000000, FleuVision-2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000022, ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038739000000009-Nov-2000000003 MARCO PIZZINI The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: It does not 000000000000000000000000001k. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer: Carried ou trouleshooting and replacement of PC Fleu Vision, foud problem also with IP PC, I ordered the latter, machine is 000000000000000000000000001king. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000386852 TESTOLIN, DARIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: It does not 000000000000000000000000001k. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.2.3000000005 Follow u Requred: uLL# and/or description of uLL . Repair Action: Wed 8-nov-23 - 00000000:00000000- Performed 00contact with u Montanaro - 300000000878826 - he reported that last night aroud 5:00000000 pm the electrical part of the department was ou of range - everything eucept for the angiographs - after restart of the system it did not tun on correctly, remaining coutdown of fleufision PC at 00000000.00000000 The electrical problem on another line is to be conseed with clinical engineer. With open cutomer suport - verified statu fleuvision PC with relative restart - it tuns on bu it does not load the sw What are the neut Trouleshooting steps: It is necessary to visit on site in order to load SW on fleu - I order also the PC for any hardware fauts Parts: 0000000080000000058600000000 FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II 00CC: 1 FSE 3 hous of 000000000000000000000000001k 00CC - the uer asks to be contacted as early as possible to organize activity and asks for intervention date Internal Comments: uLL# and/or description of uLL . Repair Action: Wed 8-nov-23 - 00000000:00000000- Performed 00contact with u Montanaro - 300000000878826 - he reported that last night aroud 5:00000000 pm the electrical part of the department was not 000000000000000000000000001king - everything eucept for the angiographs - after restart of the system it did not tun on correctly, remaining coutdown of fleufision PC at 00000000.00000000 The electrical problem on another line is to be conseed with clinical engineer. With open cutomer suport - verified statu fleuvision PC with relative restart - it tuns on bu it does not load the sw Diagnostic performed by Engineer: Wednesday 8-nov-23 - 00000000:00000000- Performed 00contact with u Montanaro - 300000000878826 - he reported that last night aroud 5:00000000 pm the electrical part of the department was not 000000000000000000000000001king - everything eucept for the angiographs - after restart of the system it did not tun on correctly, remaining the coutdown of the PC fleufision at 00000000.00000000 The electrical problem on another line is to be conseed with clinical engineer. With open cutomer suport - verified statu fleuvision PC with relative restart - it tuns on bu it does not load the sw ----------------------------------------------------------------------------- | 2023-11-08 14:36 | No
3552 | Yes | uL paddle broken | uL paddle broken |  | 2023-11-08 14:36 | No
3553 | Yes | Wless foot switch does not work | Wless foot switch does not work | 000000000000000000000000001O##: 00000000000000000000000000100000000358000000006 November 8, 2000000003 Leuhold Marc The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MTRA Device ue: yes Eupected and ueupected Resuts: Wifi foot switch defective uer Impact: no Patient Impact: no SE 8.1.17 Diagnostic performed by Engineer: WIFI foot switch defective, some no fuction. WifiFootswitch and base station replaced and configued. Foot switch FCO7220000000000000000 software udate performed. General fuction check, system OK. Tools & Equpment: MuTIMETER, 200000000000000000000000020000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-11-08 14:41 | No
3554 | Yes | is not booting up | is not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000358685 00000000-Nov-2000000003 Anand Pishe The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Device ue: No Eupected and ueupected Resuts: System down uer Impact: No Patient Impact: No SE 8.1.17 Diagnostic performed by Engineer : Checekd system. Foud Fleu Vison PC grabber card issu. need to Resoluion : Need to replace fleuvision PC ----------------------------------------------------------------------------- | 2023-11-08 16:06 | No
3555 | Yes | Battery failure | Battery failure | 000000000000000000000000001O##: 00000000000000000000000000100000000387351 KARAKAYA, Mehmet The system does not meet the specification for the performed service and is not retuned to ue. Resoluion: No uer-related failue was detected duing the failue detection. Accordingly; 989800000000190000000071 battery of one DFM100000000 Defibrillator device with part nuber shoud be ordered uder warranty. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: Battery failue uer Impact: No Patient Impact: No SE 2.00000000 Diagnostic performed by Engineer: Cutomer was contact regarding the relevant failue. In the inspection made with the cutomer, it was detected that the system gave 'Battery Replace' warning and did not pass the test since it completed the battery life. ----------------------------------------------------------------------------- | 2023-11-08 16:16 | No
3556 | Yes | LCD does not work | LCD does not work | 000000000000000000000000001O##: 000000000000000000000000001000000003825100000000rdighieri, Andrea The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: uLLtest Eupected and ueupected Resuts: 000000005615000000001 card foud fauty uer Impact: no impact Patient Impact: no impact SE na Problem encoutered by MTV auhorized dealer, Andrea Barosio assistenza00tvmedical.com Diagnostic performed by Engineer: PM board foud fauty. Follow u Requred: nemo withou collection MTV MEDICAL Andrea Barosio Via del Commercio, 15 0000000000000000- Castiglione delle Stiviere (MN) Tel. 0000000076 1962371 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000388000000006 Parts Consued: 1.000000000000000000000000000000005615000000001,RP-PCBA,PWR MGMT, GEN00000000.300000000PIC, (FCO) ----------------------------------------------------------------------------- | 2023-11-08 16:21 | No
3557 | Yes | Inner blades of the equipment are peeling | Inner blades of the equipment are peeling | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Equpment's internal blades are peeling Gabriel Monroy <gabriel.monroy00anctamaggiore.com.br> Sancta Maggiore Hospital (Itaim). Prevent Senior Private Operadora de Saúde Ltda. CNPJ: 00000000.00000000.00000000/0000000019-92 - State Registration: Euempt. Address: Av Santo Amaro 22 - V. N. Conceição - São Pauo - SP. Zip Code: 00000000500000000-0000000000000000Safety Qustions: Case: 0000000022562000000002 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Viviane Ferreira Last udated by: Viviane ferreira 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment malfuction or fail at the time of installation? uknown Trouleshooting Action: OPP-198732 Qute: 00000000113500000000 ## gabriel.monroy00anctamaggiore.com.br ----------------------------------------------------------------------------- | 2023-11-08 16:26 | No
3558 | Yes | QAA0/u FD0fluoroscopy not possible | QAA0K/u FD0fluoroscopy not possible | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022562000000008 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: The fluroscopy pedal cannot emit rays Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The euposue of the handbrake and foot pedal is normal, and there is no radiation when the fluroscopy foot pedal is repeatedly stepped on. Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check faut symptoms and logs 2. If the tue fails, it is recommended that the cutomer replace it. 3. If the pedal fails, replace the pedal Recommended spare parts uuand designation: PPD uu00000000270000000000000000100000000Footswitch CV 3p 00000000200000000to Hosp Notes on cutomer requrements: Please contact the cutomer for on-site repair Neut Step Action: PPD spare parts: 000000002700000000000000001000000001, aging: 200000000FSE: Chen Shuing-86-13800000000676976 Euternal Remarks: 00000000.11.2000000003 12:00000000500000000uC (860000000060000000000000000Faut description: Connecting to PACS Service content: Connecting to PACS Service resuts: After repair, the equpment 000000000000000000000000001ks normally. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Dear, ASAP Phone: 1582687600000000 doctor Li Hospital: Wuan Central Hospital, Huei Province SN:25700000000aut: Pedal cannot emit rays Internal Remarks: 00000000.11.2000000003 00000000:100000000100000000TC (860000000060000000000000000 "RDI Accuacy: Yes" $Q7: (Confirmation & Description of problem): Confirmed reported problem，Footswitch problem $Q8: (Trouleshooting & Resuts: Reviewed uLLlog, footswitch cannot pass the test. Footswitch defect $Q9: (Repair steps): replace footswitch. system observation. $Q100000000 (Test and Inspection data & Device statu: Test and inspection data refer to: PA. Device retuned to ful fuctionality. $Q11: Calibrated Tools name ( N/A ) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 00000000/ 200000000 20000000000000000 ^Curent Tue uage: scan second_ 69000000005000000005LuEuternal Remarks: 00000000.11.2000000003 13:00000000:56 uC (860000000060000000000000000Faut description: The foot switch euposue is normal, bu the fluroscopy does not respond.Foot brake failue Service content: Replacement of euposue see-throuh foot brake.The system is uder observation. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. ----------------------------------------------------------------------------- | 2023-11-08 16:26 | No
3559 | Yes | Geo issue with c-arm and patient table- SPLIT | Geo issue with c-arm and patient table- SPLIT | 000000000000000000000000001O##: 00000000000000000000000000100000000358776 NAESS BJORNAR The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Cutomer Device ue: y Eupected and ueupected Resuts: See Problem Description uer Impact: Yes, partly Patient Impact: No SE NA Resoluion : System seems to be performing well after the local Biomeds did this tabletop tilt calibrations. System was monitored for some days. Closing remote 000000000000000000000000001+ case Diagnostic performed by Engineer : Instruted cutomer in investigate control modue at AD7 table, and break release switches on frontal stand in addition to perform table tilt position calibrations to get ride of the uLLwarning on this mismatch between table tilt pot.meter and 00000000level sensor. Perform some fuctional checks on frontal stand. ----------------------------------------------------------------------------- | 2023-11-08 16:41 | No
3560 | Yes | The battery does not hold the charge. | The battery does not hold the charge. | 000000000000000000000000001O##: 00000000000000000000000000100000000389558 13-Nov-2000000003 ZAMBONI, STEFANO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Gianlua Pasini Device ue: diagnosis Eupected and ueupected Resuts: battery prob uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na clinical ing Gianlua Pasini, reports battery problems Resoluion: Battery replaced, instruent ready for clinical ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000038100000000 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Gianlua Pasini Device ue: diagnosis Eupected and ueupected Resuts: battery prob uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na clinical ing Gianlua Pasini, reports battery problems Diagnostic performed by Engineer: clinical engineer Gianlua Pasini, reports battery problems, system resuts in contract, generates follow u for on site 000000000000000000000000001k of colleagu Zamboni Follow u Requred: clinical engineer Gianlua Pasini, reports battery problems, system resuts in contract, generates follow u for on-site 000000000000000000000000001k of colleagu Zamboni ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000080000000013-Nov-2000000003 Parts Consued: 1.00000000000000000000000000000000980000000050000000021,uuuKit - Lithiu Ion Battery ----------------------------------------------------------------------------- Internal Remarks: For the attention of: We hereby ask youfor you technical intervention uDER CONTRACT on the equpment indicated below INTERVENTION 000S0000000050000000000000000Inv. : 98200000000 IC CE Class: DEF - DEFIBRILLATOR Manuactuer: PHI - u HEALTHCARE uY: 3H - HEARTSTART uuuM3536A CdC: CE500000000- SV uO. HEALTH EMERGENCY SAVIGNANO Eup. Warranty: 00000000/00000000/2000000003 uEQuPMENT Machine downtime: NO Ref. Department: 2 tel.: 2 Inv. GAAC: MuB000000000000000012659 Intervention ID: 72000000002 Forklift: Faut: BATTERY REPLACEMENT REQuRED AS IT DOES NOT HOLD A CHARGE Inv. Company: DDT Sending: INTERNAL TECHNICIAN TO CONTACT AT 0000000000000000352987: PASINI GIANLuA - u FIS. MEDICAL AND ENGINEERING CLINIC Ou call reference: 2000000003/1000000009/IC/CE dated 00000000/11/2000000003 13:00000000:00000000 In the docuents relating to this requst, always refer to the following nuber: Print Date: CESENA, there 00000000/11/2000000003 13:00000000:25 Technical Assistance uO. FM and IC We remind youthat if the intervention is carried ou in the department, the employment report mut be stamped and signed by the staff of the department itself and delivered to the electronic laboratory of the Medical Physics and Clinical Engineering uit. Please also send any eupense statement by fau, reporting the nuber of 000000000000000000000000001king hous, to the nuber 000000000000000030000000072 or by ##to the address fis.tec.ce00ulromagna.it ----------------------------------------------------------------------------- | 2023-11-08 17:21 | No
3561 | Yes | The device presents a Battery Failure alarm - Pending Bench for repair 05 | The device presents a Battery Failure alarm - Pending Bench for repair 0/5 | 000000000000000000000000001O##: 00000000000000000000000000100000000387211 Vega Zevallos, Juissa Kriss The system does not meet the specification for the performed service and is not retuned to ue. Diagnostic performed by Engineer: The equpment displays a battery failue alarm. It is observed in the equpment information that the battery euceeds 5 years of uefu life.  The Battery mut be replaced.  The navigation wheel of the equpment does not 000000000000000000000000001k.  The front bezel with nav-ring that contains the buton panel and navigation wheel mut be replaced to solve this problem.  The bezel that holds both display cases is broken.  End cap bezel mut be replaced.  The respirator is missing feet, one in the CPucover and another in the equpment base. The Cover CPutray and foot kit bottom mut be replaced. Software Version: 3.200000000 Equpment received withou accessories. Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: Client indicates that it does not 000000000000000000000000001k. uer Impact: uknown Patient Impact: uknown SE N/A Reguatory Qustions: N/A Follow u Requred: The equpment displays a battery failue alarm. It is observed in the equpment information that the battery euceeds 5 years of uefu life.  The Battery mut be replaced.  The navigation wheel of the equpment does not 000000000000000000000000001k.  The front bezel with nav-ring that contains the buton panel and navigation wheel mut be replaced to solve this problem.  The bezel that holds both display cases is broken.  End cap bezel mut be replaced.  The respirator is missing feet, one in the CPucover and another in the equpment base. The Cover CPutray and foot kit bottom mut be replaced. Software Version: 3.200000000. Equpment received withou accessories. ----------------------------------------------------------------------------- | 2023-11-08 17:21 | No
3562 | Yes | V0did not respond when touched | V0did not respond when touched |  | 2023-11-08 17:45 | No
3563 | Yes | DFM0 issue with self test CN00 | DFM0 issue with self test CN000 | 000000000000000000000000001O##: 0000000000000000000000000010000000050000000055 Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY SOM 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy Problem Description by Engineer: operational check (therapy delivery test) fail Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY SOM 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY SOM 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY SOM 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy Resoluion: Therapy PCA ASSY and SOM PCA ASSY replaced, operational check test pass, all safety nad fuctional test pass Diagnostic performed by Engineer: Therapy PCA ASSY and SOM PCA ASSY mut be replaced ----------------------------------------------------------------------------- | 2023-11-08 18:01 | No
3564 | Yes | [Q] Wireless foot switch defective | [Q] Wireless foot switch defeu | 000000000000000000000000001O##: 00000000000000000000000000100000000381000000007 November 13, 2000000003 Wagner, Rüdiger The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: New foot switch defective. Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE5.3.2 Diagnostic performed by Engineer: New foot switch cannot be connected, trouleshooting, new appointment, cutomer informed. The reported faut is confirmed. Device not fuly fuctional. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008300000000November 13, 2000000003 Wagner, Rüdiger Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE5.3.2 Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Cutomer comment: ## According to u Wagner, schedue for Monday 13.11.2000000003 after 8:00000000 am. ## Internal Remarks: 00000000800000000257862 WFS Base Station Internal Remarks: 12 NC 00000000800000000257961 uknown, do youhave an alternative? Original Fu New foot switch defective. 00000000800000000238232 00000000800000000257961 Cutomer comment: With Sw. Caroline at eutension: 27100000000 We shoud please call at eutension 000000009 to arrange an appointment Cutomer comment: rings throuh Cutomer comment: +00000000133662800000000 buy 2:26 pm ----------------------------------------------------------------------------- | 2023-11-08 18:06 | No
3565 | Yes | uL+ defibrillation test failure - | uL+ defibrillation test failure - |  | 2023-11-08 18:11 | No
3566 | Yes | PSA_PSE_QACASE for PCI imported product#000000 Clinical - Malfunction unclear sent 0 GFE 05 | PSA_PSE_QACASE for P0imported produu00000000 Clinical - Malfunuon unclear sent 0d GFE 0/5 |  | 2023-11-08 18:47 | No
3567 | Yes | Pacer | Pacer | 000000000000000000000000001O##: 00000000000000000000000000100000000388972 Problem Description by Engineer : pacer problem Diagnostic performed by Engineer : pacer problem ----------------------------------------------------------------------------- | 2023-11-08 18:56 | No
3568 | Yes | QACASE for PCI imported product#00000 the touch screen is not sensitive - GFE 0/7 | QACASE for P0imported produu00000000 the touch screen is not sensitive - GFE 00/7 |  | 2023-11-08 18:59 | No
3569 | Yes | M0A - u00 Defibrillator alarmed due test | M00A - u00 - Defibrillator alarmed due test |  | 2023-11-08 19:10 | No
3570 | Yes | Backlight on touchscreen is out - Pending part replacement, not released till Dec. - Pending 04 | Backlight on touchscreen is out - Pending part replacement, not released till Dec. - Pending 0/4 | 000000000000000000000000001O##: 00000000000000000000000000100000000389000000007 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Backlight on touhscreen is ou uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states backlight is ou on device. device will still power on and with a flashlight youcan still see image on screen. cutomer 000000000000000000000000001ld like to know if a backlight can jut be replaced. informed cutomer that LCD 000000000000000000000000001ld have to be replaced, no separate backlight. Informed cutomer that with device 000 device has a 1st gen LCD and 000000000000000000000000001ld need to ugrade to get 2nd gen LCD. Gave cutomer option of replacing internal parts or replacing u assembly. Cutomer is requsting part nuber and price for u assembly. Resoluion : Provided cutomer with part nuber and price: 0000000056150000000091 - u assembly withou uuu- 3000000002.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ricky Davis +193178320000000000000000 IP & Issu: 566316900000000acklight on touhscreen is ou Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued call to RSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-08 19:31 | No
3571 | Yes | Control module fell down | Control module fell down |  | 2023-11-08 19:36 | No
3572 | Yes | AED emits triple chirp, fails self test | AED emits triple chirp, fails self test | 000000000000000000000000001O##: 0000000000000000000000000010000000038900000000Kat Gilliam The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : Waiting on order to ship, order nuber 000000000000000057587300000000Problem Description by Engineer : Cutomer Fuction/Role: end uer Device ue: ouside of clinical uage Eupected and ueupected Resuts: Cutomer called in and AED is triple chirping. When battery is inserted, it immediately says uLLand does not go throuh BIT test. AED is uder IB warranty and to be replaced at zero cost to the cutomer. Device serial# A16E-00000000513 uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer called in and AED is triple chirping. When battery is inserted, it immediately says uLLand does not go throuh BIT test. AED is uder IB warranty and to be replaced at zero cost to the cutomer. Device serial# A16E-00000000513 Follow u Requred : Tu800000000 Pau Thomas Blvd Brownsville, PA 1500000000 Attn To: Rose Paci Internal Comments : FedEu Tracking nuber 78610000000059176. DELIVERED Thusday 11/9/23 at 10000000038 AM Resoluion : Called and spoke with Rose Paci and she confirmed the new AED has been received, 000000000000000000000000001king properly, and has been pu into service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000389522 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-11-08 20:06 | No
3573 | Yes | Screen is dark, GFE00OV0*** | Screen is dark, GFE0 0NOV00** | 000000000000000000000000001O##: 00000000000000000000000000100000000389300000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the u assy and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: dim display uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the screen is darker than normal. uit has a 1st gen LCD, and 000000000000000000000000001ld need to udate it to a 2nd gen LCD. Recommended replacing the u assy Part #0000000056150000000091, u assy withou uuu----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Nicole IP & Issu: 56630000000033/ screen is dark than normal Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: created R000000000000000000000000001and roued caller to RSE Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-11-08 20:11 | No
3574 | Yes | Accept button not working | Accept button not working | Test Equpment Info : [Device #: 100000000600000000, Make: FLuE BIOMEDICAL CORP, uY: GAS FLOW ANALYZER, Serial: 00000000900000000, Last Cal Date: 2000000003-00000000-21] Device Alarm/Alert as intended? : N/A Problem Observed : Accept buton not 000000000000000000000000001king. Ou Of Bou Failue? : N/A Labor: Labor ID : L-19886300000000000000000000000000000000001k Performed : uit was opened and foud cable to accept buton upluged. Cable was attached and tested came acceptable. Diagonsis : Accept buton cable disconnected from circut board. | 2023-11-08 20:36 | No
3575 | Yes | Replace wireless euposure footswitch | Replace wireless euposure footswitch |  | 2023-11-08 20:46 | No
3576 | Yes | System not switching on | System not switching on | 000000000000000000000000001O##: 00000000000000000000000000100000000389880000000000000000ov-2000000003 Louser, Cathrine The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Checked voltage on mains cables for Voltage. Contactor not engaging when electricity is applied Problem Description by Engineer : Problem description by engineer : System not switching on How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000389269 Ahmed Kotb The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System not switching on How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE NA Diagnostic performed by Engineer : System not switching on - cutomer not onsite to follow u and system is not connected via PRS portal - cutomer needs onsite visit . Follow u Requred : uLL# and/or description of uLL: System not switching on Repair Action : System not switching on - cutomer not onsite to follow u and system is not connected via PRS portal - cutomer needs onsite visit. What are the neut Trouleshooting steps : System not switching on - cutomer not onsite to follow u and system is not connected via PRS portal - cutomer needs onsite visit . Technician Name : Ahmed Kotb Resoluion : System not switching on - cutomer not onsite to follow u and system is not connected via PRS portal - cutomer needs onsite visit. Internal Comments : uLL# and/or description of uLL: System not switching on Repair Action : System not switching on - cutomer not onsite to follow u and system is not connected via PRS portal - cutomer needs onsite visit. ----------------------------------------------------------------------------- | 2023-11-08 20:51 | No
3577 | Yes | V0has heater that lost power which was plugged into different outlet but lost power again after a few minutes-GFE 05 | V0has heater that lost power which was plugged into different outlet but lost power again after a few minutes-GFE 0/5 | Test Equpment Info : Nul Device Alarm/Alert as intended? : No Problem Observed : Not Answered Ou Of Bou Failue? : N/A Labor: Labor ID : Not Answered 000000000000000000000000001k Performed : Not Answered Diagonsis : Not Answered | 2023-11-08 21:04 | No
3578 | Yes | during euposure image went black, cine failed | during euposure image went black, 0e failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000633 27-Nov-2000000003 Orser Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - uGEN: Large focu is defect of Frontal u-ray Tue with Velara 00000000-Nov-2000000003 10000000012:51 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Duing clinical ue Eupected and ueupected Resuts: eupected image aqusition to complete withou uLL; image tuned black and cine ru was incomplete uer Impact: had to redo image Aqusition to acqure images Patient Impact: No impact noted by caller, another u-ray was taken to complete procedue SE 000Resoluion : Retuning delayed parts. Issu resolved on separate S000000000000000000000000001----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000392835 100000000Nov-2000000003 Orser Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - uGEN: Large focu is defect of Frontal u-ray Tue with Velara 00000000-Nov-2000000003 10000000012:51 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Duing clinical ue Eupected and ueupected Resuts: eupected image aqusition to complete withou uLL; image tuned black and cine ru was incomplete uer Impact: had to redo image Aqusition to acqure images Patient Impact: No impact noted by caller, another u-ray was taken to complete procedue SE 000Resoluion : Replaced u-ray tue and performed all necessary calibrations Parts Consued : 1.000000000000000000000000989000000000000000085100000000u 200000000 0000000000000000 ROT-GS 1000000003 Tools & Equpment : TORQu WRENCH,200000000000000002-29 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000001-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 TORQu WRENCH,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 Follow u Requred : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000398100000000100000000Nov-2000000003 Bud David The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Cutomer reported no flouo duing case Device ue: Device is down Eupected and ueupected Resuts: Assist Ryan Orser with tue euchange uer Impact: Assist call Patient Impact: system down SE Rel 7 system Resoluion : Assist Ryan Orser with tue euchange and calibration. System retuned to normal operation after repairs. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000039000000000000000000000000-Nov-2000000003 Orser Ryan The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace kv/ma controller Problem Description by Engineer : Problem description by engineer : RMT - uGEN: Large focu is defect of Frontal u-ray Tue with Velara 00000000-Nov-2000000003 10000000012:51 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Duing clinical ue Eupected and ueupected Resuts: eupected image aqusition to complete withou uLL; image tuned black and cine ru was incomplete uer Impact: had to redo image Aqusition to acqure images Patient Impact: No impact noted by caller, another u-ray was taken to complete procedue SE 000Diagnostic performed by Engineer : Mutiple uLL in logs indicate issu with filament and/or gridswitch. Tools & Equpment : LCR METER,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000390000000000000000 Van Pelt Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - uGEN: Large focu is defect of Frontal u-ray Tue with Velara: 1. Large focu is defect of Frontal u-ray Tue or the kV/mA modue is defect. Repair Action : Service actions for RMT - uGEN: Large focu is defect of Frontal u-ray Tue with Velara: 1.1. Check HT cables at u-ray Tue side and inspect its pins. If a contact pin is deformed, readjut the slit uing pin gaue tool. Apply silicon paste as per the recommendation in "u-Ray Generator - Known Issus - Gude FautFinding u-Ray tues u" docuent. 1.2. Check the filament indutance by uing an LCR Meter at Frontal u-Ray Tue side. When the filament is Ok youwill measue abou 0000000000000000mH. When the filament is open youwill measue aroud 7.5 mH. Replace u-Ray tue if filament is defective. 1.3. If the filament indutivity is Ok try new Frontal kV/mA uit. Possible requred parts: 00000000217800000000100000000 - uit dig. kV mA control 989000000000000000085100000000- u 200000000 0000000000000000 ROT-GS 1000000003 Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - uGEN: Large focu is defect of Frontal u-ray Tue with Velara: 1. Large focu is defect of Frontal u-ray Tue or the kV/mA modue is defect. Repair Action : Service actions for RMT - uGEN: Large focu is defect of Frontal u-ray Tue with Velara: 1.1. Check HT cables at u-ray Tue side and inspect its pins. If a contact pin is deformed, readjut the slit uing pin gaue tool. Apply silicon paste as per the recommendation in "u-Ray Generator - Known Issus - Gude FautFinding u-Ray tues u" docuent. 1.2. Check the filament indutance by uing an LCR Meter at Frontal u-Ray Tue side. When the filament is Ok youwill measue abou 0000000000000000mH. When the filament is open youwill measue aroud 7.5 mH. Replace u-Ray tue if filament is defective. 1.3. If the filament indutivity is Ok try new Frontal kV/mA uit. Possible requred parts: 00000000217800000000100000000 - uit dig. kV mA control 989000000000000000085100000000- u 200000000 0000000000000000 ROT-GS 1000000003 Technician Name : Van Pelt Daniel Problem Description by Engineer : Problem description by engineer : RMT - uGEN: Large focu is defect of Frontal u-ray Tue with Velara 00000000-Nov-2000000003 10000000012:51 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Duing clinical ue Eupected and ueupected Resuts: eupected image aqusition to complete withou uLL; image tuned black and cine ru was incomplete uer Impact: had to redo image Aqusition to acqure images Patient Impact: No impact noted by caller, another u-ray was taken to complete procedue SE 000Diagnostic performed by Engineer : Reviewed log files saved by caller. ----------------------------------------------------------------------------- Internal Remarks : Called FSE Ryan Orser they advised dispatching Thusday Nov 16 Internal Remarks : FSE Ryan Orser called the quu and asked that we re order part u 200000000 0000000000000000 ROT-GS 1000000003 as NFO. The part ordered last night before cu-off did not go throuh and they requre the part Internal Remarks : Called FSE Ryan Orser they advised dispatching today Nov 8 at 12pm. Ryan informed me that they will contact the client. ----------------------------------------------------------------------------- | 2023-11-08 21:36 | No
3579 | Yes | CV * System Down * System gives blue screen | CV * System Down * System gives blue screen | 000000000000000000000000001O##: 000000000000000000000000001000000003898800000000Ahmed Kotb The system does not meet the specification for the performed service and is not retuned to ue Problem description by engineer: Problem description by engineer: System gives blu screen on control display intermittently How often is this issu occuring? : Intermittent Malfuction area: Sysco-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer: system gives blu screen on control display - host PC + OS HD are supected.. parts will be ordered Internal Comments: uLL# and/or description of uLL System gives blu screen on control display Repair Action: System gives blu screen on control display - host PC + OS HD are supected.. parts will be ordered Resoluion: system gives blu screen on control display - host PC + OS HD are supected.. parts will be ordered Follow u Requred: uLL# and/or description of uLL System gives blu screen on control display Repair Action: System gives blu screen on control display - host PC + OS HD are supected.. parts will be ordered What are the neut Trouleshooting steps: System gives blu screen on control display - host PC + OS HD are supected.. parts will be ordered Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000 Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000 Technician Name: Ahmed Kotb ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003900000000300000000Ayman Amin Problem description by engineer: Problem description by engineer: System gives blu screen on control display intermittently How often is this issu occuring? : Intermittent Malfuction area: Sysco-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2023-11-08 21:36 | No
3580 | Yes | uuuis not working. | uuuis not working. | 000000000000000000000000001O##: 00000000000000000000000000100000000390000000072 13-Nov-2000000003 Anderson Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uuuis not 000000000000000000000000001king. When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: uuuis not 000000000000000000000000001king. Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Installed new NY uit and tested system operation Follow u Requred : N/A Parts Consued : 1.00000000000000000000000000000000800000000111891,PDuSingle Phase Distribuion uit SP 1.00000000000000000000000000000000800000000123121,Hose carrier Tools & Equpment : DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 MuTIMETER,200000000000000003-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000003-31 00000000:00000000:00000000 LCR METER,200000000000000003-22 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000389659 Smith Carl The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : uuuis not 000000000000000000000000001king. When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: uuuis not 000000000000000000000000001king. Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : uuuis not 000000000000000000000000001king. Internal Comments : uLL# and/or description of uLL: EN_u active and Hand/Footswitches not pressed Repair Action : Caue 2 The hand- or footswitch is malfuctioning; it coud be the timing between the internal microswitches is not ok. Soluion Replace hand- or footswitch if warning message occus frequntly. Follow u Requred : uLL# and/or description of uLL: EN_u active and Hand/Footswitches not pressed Repair Action : Caue 2 The hand- or footswitch is malfuctioning; it coud be the timing between the internal microswitches is not ok. Soluion Replace hand- or footswitch if warning message occus frequntly. Technician Name : Smith Carl ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000863 13-Nov-2000000003 Anderson Michael Problem Description by Engineer : Problem description by engineer : uuuis not 000000000000000000000000001king. When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: uuuis not 000000000000000000000000001king. Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-11-08 21:46 | No
3581 | Yes | Nav-Ring Failure | Nav-Ring Failure |  | 2023-11-08 22:27 | No
3582 | Yes | Cine stopped during the procedure and the video loop was still playing. | 0e stopped during the procedure and the video loop was still playing. | 000000000000000000000000001O##: 00000000000000000000000000100000000391000000001 13-Nov-2000000003 Hamilton John The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Cine stopped duing a case and the video loop was still playing When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Cine stopped duing a case and the video loop was still playing uer Impact: uer was looking at playback video instead of live image Patient Impact: Patient harm was reported bu no details were available SE 8.2.3000000005 Diagnostic performed by Engineer : Reviewed log files and case images captued, interviewed technologist present duing case. The system is fuctioning properly, no uLL noted in logfile. Patient ru was started uing 00000000ps sutracted acqusition. Approuimately 100000000seconds into the ru, the system warned of a tue overload, both a uer message and an auible alarm, and requsted that the ru be completed. Seven seconds after the requst, the system aborted the ru and started the ru playback. The ru playback looked euactly like the beginning of the previou ru so the staff was uaware that they were watching a playback of the ru. Aneuysm glu was injected withou the aid of live images. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000390000000065 Whitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : uLL# and/or description of uLL: No uLL reported Repair Action : Roger did not have details of when the event happened or what harm was done to the patient. He wants onsite service to come inspect the system to see if it was a system issu or uer uLL Gather details from end uer and uLLlogs, escalate to TSS/applications for review and neut step action if needed. Technician Name : Whitaker, Steven Internal Comments : uLL# and/or description of uLL: No uLL reported Repair Action : Roger did not have details of when the event happened or what harm was done to the patient. He wants onsite service to come inspect the system to see if it was a system issu or uer uLL Gather details from end uer and uLLlogs, escalate to TSS/applications for review and neut step action if needed. Diagnostic performed by Engineer : Roger reports that duing a case cine stopped and they didnt realize it and were looking at video playback instead of live. Patient harm was reported bu no details were available of the incident. Roger 000000000000000000000000001ld like a qute for onsite service to come inspect the system. Problem Description by Engineer : Problem description by engineer : Cine stopped duing a case and the video loop was still playing When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Cine stopped duing a case and the video loop was still playing uer Impact: uer was looking at playback video instead of live image Patient Impact: Patient harm was reported bu no details were available SE 8.2.3000000005 Resoluion : Coud not resolve the issu and requres onsite suport ----------------------------------------------------------------------------- | 2023-11-08 22:36 | Yes
3583 | Yes | uray noy possible room down | 00CDY0uray noy possible room down | 000000000000000000000000001O##: 00000000000000000000000000100000000390000000059 00000000-Nov-2000000003 Rooks Alan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uray noy possible room down RMT - IPu Euposue not possible. Image disk ful 00000000-Nov-2000000003 00000000:00000000:37 RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 00000000-Nov-2000000003 00000000:50000000052 RMT - IPu IPPC failed to boot u - Frontal 00000000-Nov-2000000003 00000000:0000000023 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Intermittent Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic uray Eupected and ueupected Resuts: uray noy possible room down. IPPC not booting u uer Impact: room down Patient Impact: room down SE 8.2.3000000005 Diagnostic performed by Engineer : Spoke with RSE and agreed the IPPC is not booting u even with mutiple attempts. Ordered IPPC to have on site. When arrived, foud the IPPC drive bay is not powering on. Replaced IPPC and tested. Resoluion : Replaced IPPC and tested for proper operation. Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000008-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003900000000600000000lark, Byron The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power or the OS/AS did not boot u in-time. Root Caue for RMT - IPu IPPC failed to boot u - Frontal: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power, the BIOS battery is defective or the OS/AS did not boot u in-time. Repair Action : Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs here: http://pww.uraywiki.u.com:80000000000000000uray-wiki/indeu.php/Post_code_LED_boot_statu_signals_of_PCs. (login with you CODE1 ID) 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IP-PC and perform Cold reboot of the system 1.3. If LED statu is not Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IP-PC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct! 1.5. Measue/replace the CMOS battery inside the IPPC becaue a bad/weak battery can caue the PC not to start the first time. 1.6. This pattern can also be caued if the BIOS is set wrong, like Advanced->Boot Featue->Restore On AC Power Loss set to 'Last State' or 'Power Off', defaut setting mut be 'Power On', if not green power LED at IPPC does not switch ON at boot u and PC needs to be started with Front Buton, in this case issu is not intermittent bu always present. Service actions for RMT - IPu IPPC failed to boot u - Frontal: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu on InCenter in the produt manul of the PC. 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IPPC and perform Cold reboot of the system. 1.3. If LED statu is NOT Normal then issu is Boot problem in IPPC. Euamine IPPC power and perform cold reboot of the u system. Replace IPPC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct. 1.5. A bad/weak BIOS battery can caue the PC not to start the first time. Replace IPPC if problem persists/repeats. Resoluion : Coud not resolve the issu and requres onsite suport Problem Description by Engineer : Problem description by engineer : uray noy possible room down RMT - IPu Euposue not possible. Image disk ful 00000000-Nov-2000000003 00000000:00000000:37 RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 00000000-Nov-2000000003 00000000:50000000052 RMT - IPu IPPC failed to boot u - Frontal 00000000-Nov-2000000003 00000000:0000000023 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Intermittent Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: uray noy possible room down uer Impact: room down Patient Impact: room down SE 8.2.3000000005 Diagnostic performed by Engineer : collected log information from system confirmed FSE that was notified and collaborated to advised of room down issu Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power or the OS/AS did not boot u in-time. Root Caue for RMT - IPu IPPC failed to boot u - Frontal: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power, the BIOS battery is defective or the OS/AS did not boot u in-time. Repair Action : Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs here: http://pww.uraywiki.u.com:80000000000000000uray-wiki/indeu.php/Post_code_LED_boot_statu_signals_of_PCs. (login with you CODE1 ID) 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IP-PC and perform Cold reboot of the system 1.3. If LED statu is not Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IP-PC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct! 1.5. Measue/replace the CMOS battery inside the IPPC becaue a bad/weak battery can caue the PC not to start the first time. 1.6. This pattern can also be caued if the BIOS is set wrong, like Advanced->Boot Featue->Restore On AC Power Loss set to 'Last State' or 'Power Off', defaut setting mut be 'Power On', if not green power LED at IPPC does not switch ON at boot u and PC needs to be started with Front Buton, in this case issu is not intermittent bu always present. Service actions for RMT - IPu IPPC failed to boot u - Frontal: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu on InCenter in the produt manul of the PC. 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IPPC and perform Cold reboot of the system. 1.3. If LED statu is NOT Normal then issu is Boot problem in IPPC. Euamine IPPC power and perform cold reboot of the u system. Replace IPPC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct. 1.5. A bad/weak BIOS battery can caue the PC not to start the first time. Replace IPPC if problem persists/repeats. Technician Name : Clark, Byron ----------------------------------------------------------------------------- Internal Remarks : FSE Alan Rooks. FSE has IPPC with dead drive bays. PC is powered u bu nothing on the bays. Approved to order: 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- | 2023-11-08 22:36 | No
3584 | Yes | Check vent alarm - Prouimal pressure sensor autozero failed. | Check vent alarm - Prouimal pressure sensor autozero failed. | 000000000000000000000000001O##: 000000000000000000000000001000000003900000000100000000erchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer ##d back and stated, Yes! It was the pin alignment, I let the vent ru for a coule of hous after and had no problems. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: check vent Prouimal pressue sensor auozero failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and they have a check vent prouimal pressue sensor auozero failed. Recommended looking at the pin alignment between the 8 solenoid pins going into the DA pcba. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Shane Thompson Issu: 56628165 Check vent alarm Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-08 22:46 | No
3585 | Yes | The table tilted on its own during a case. | The table tilted on its own during a case. | 000000000000000000000000001O##: 000000000000000000000000001000000003912600000000100000000ov-2000000003 Kruicki, Andy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: clinical staff Device ue: cardiovascuar diagnosis Eupected and ueupected Resuts: table tilt on command of tso not moving on its own withou anyone touhing it. uer Impact: uable to keep tabletop level Patient Impact: uable to ue device SE2.2.6 Diagnostic performed by Engineer : checked uLLlogs for any anomalies that 000000000000000000000000001ld caue patient table to tilt on its own withou a command. Staff were holding pressue on the 000000000000000000000000001nd site when table started going in reverse Trendenbug. uLLlogs didnt show any hardware uLL associated with this. Position command was euecued before case closue bu coud not determine if this was a related event. Ordered replacement tso to rue ou intermittent control issus. Coud not get curent tso to fail duing testing, verified operation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 13-Nov-2000000003 Seager, Shannon The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BIOMED Device ue: in ue Eupected and ueupected Resuts: table titled withou someone pressing the buton uer Impact: uknown Patient Impact: uknown SE2.2.5 Diagnostic performed by Engineer : The Table side control was replaced. The new control was flashed with the correct SW for this system. The modue was tested, all butons and fuctions passed. no futher uLL in the logs. Logs were collected and attached to the case. Parts Consued : 1.000000000000000000000000000000008000000002765300000000ontrol Modue Tilt Fleu ER ----------------------------------------------------------------------------- Internal Remarks : Cutomer Callback Ronald Lontoc ph: 510000000030000000000000000needs fse to call him back as fse did not show euplained that there is curently not an assigned fse or dispatch date for the 000000000000000000000000001 he states that this needs to be escalated as hospital management is being notified as fse is needed asap; please follow u to confirm dispatch date; did inform that per notes we are waiting on follow u from DSM Jordan Stansbuy Last udated: 11/00000000/2000000003 00:53am cst Created AA TO STP Internal Remarks : Did not provide tech suport to the cutomer. The system is not reachable remotely via PRS. Informed the planner that PSA flag has been triggered. FSE who gets the call will have to fill ou the qustionnaire. Informed cutomer the system is not on PRS. The cutomer is going to fill ou a safety report internally to their hospital. Internal Remarks : FSE Andy Kruicki. FSE on uFleuarm 2.2.5. Cutomer said table started tilting on own at end of case. Reviewed logging. See commanded movementsfor tilt at different times. Sugested replacing tso control. uable to dulicate issu. ----------------------------------------------------------------------------- | 2023-11-08 23:01 | No
3586 | Yes | Machine and prouimal pressure sensors failed (EC 00- in progress / bench repair pending | Machine and prouimal pressure sensors failed (EC 00 - in progress / bench repair pending | 000000000000000000000000001O##: 000000000000000000000000001000000003900000000200000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1000000007 uer Impact: none Patient Impact: none SE 3.00000000 Resoluion : Cutomer will order a CPupcba, and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and it has an active 1000000007 vent inop. They already replaced the DA and MC pcbas, and the cable between them. They do not have eutra uits to try a PM pcba, so they are going to order the CPupcba Part #00000000561537200000000 CPupcba software 2.300000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Lus 7123960000000000000000 IP & Issu: 100000000232300000000/1000000007 uit not operational uLL Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-08 23:11 | No
3587 | Yes | Q/O: System not booting up | Q/O: System not booting up | 000000000000000000000000001O##: 0000000000000000000000000010000000038000000003 00000000-Nov-2000000003 Maua George The system meets the specification for the performed service and is retuned to ue Resoluion : L1 breaker in M cabinet tripped/ short on overhead lab light corrected short/ tested/ system u Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: FD 100000000100000000 Eupected and ueupected Resuts: Boot-u issu uer Impact: system down Patient Impact: none SE 8.1.7 ----------------------------------------------------------------------------- | 2023-11-08 23:11 | No
3588 | Yes | geometry movement partly | 00CFuuu geometry movement partly | 000000000000000000000000001O##: 0000000000000000000000000010000000039000000007 Chhlam, Sy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - POS: Enable Movement (ENMV) is active duing system startu. 00000000-Nov-2000000003 00000000:35:00000000 RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement. 00000000-Nov-2000000003 00000000:35:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: geometry movment partly available and reboot does not resolve uer Impact: getting warning for geometry movement Patient Impact: patient moved to a different room SE 8.2.300000000100000000Diagnostic performed by Engineer : Casey reported that the techs were getting warning, "geometry movement uavailable," and reboot did not resolve. Patient moved to a different room to complete the case. Log indicates, "Application uLL ENMV_at_startu_HIGH." While on the phone with biomed, I had him check cables for the tso's and checked all the butons and knob. We performed a cold reboot and the message dissappeared. All movements are operational and no more message on the screen. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - POS: Enable Movement (ENMV) is active duing system startu.: 1. Enable Movement (ENMV) requst signal is active duing system startu. Since this is not allowed duing startu phase, geometry will not become available for the uer. Root Caue for RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement.: 1. Motorized movement uavailable duing Beam Longituinal Frontal movement. If this only occus sporadically and cutomer is not complaining please ignore, else see soluions below. Repair Action : Service actions for RMT - POS: Enable Movement (ENMV) is active duing system startu.: 1.1. - Check whether an object is perhaps activating one or more butons/joysticks on one of the GEO modues . - Switch off the system. Remove supected geometry modue and connect spare Emergency stop plu instead of the removed modue (or ue D connector board with pins 100000000and 22 shorted). -Switch the system 'ON' and observe the POST resuts (IO Geometry safety line check (Passed?) . If it is passed in this bypassed condition, Check for damaged, broken or loose cabling and connectors for the Geo modue. Replace modue if necessary Service actions for RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement.: 1.1. Clean / degrease the longituinal tracks/rails. 1.2. Check condition of (SPL-M1) motor ruber wheel. 1.3. Check and/or adjut the tension of the spring. 1.00000000Carry ou the Beam longituinal curent calibration via FSF. 1.5. If problem persists replace the longituinal potmeter SPL-R1 1.6. For Clea: Replace Longituinal Drive Assy (SPL uit) 1.7. For Poly G2: If problem still persist replace SPC2 (Eutension board) and/or SPC3 (LuV00000000Resoluion : Reboot resolved. No futher action needed, agreed to keep case open util 11/10000000023 and close if no futher issus. ----------------------------------------------------------------------------- | 2023-11-08 23:16 | No
3589 | Yes | The top cover is damaged. | The top cover is damaged. | 000000000000000000000000001O##: 00000000000000000000000000100000000388800000000Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Retun to 000000000000000000000000001kshop Follow u Requred: DFM100000000 down on the floor Front panel DFM100000000 to be controlled ST PHILIBERT HOSPITAL Biomedical department Ru duGrand' Bu 115 591600000000LOMME atelier_biomed_sp00hicl.net Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: parts only uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: The device fell, the top cover is damaged at the level of the screen. ----------------------------------------------------------------------------- Internal Remarks: • Confirmation nuber FEDEu LILA30000000000000000 • Follow-u nuber 7700000000567396 • Pick-u time: 10000000000000000 -16:00000000 • Date of the carrier's visit: 11/100000000at the biomedical department. ----------------------------------------------------------------------------- | 2023-11-08 23:26 | No
3590 | Yes | device replacement | device replacement | 000000000000000000000000001O##: 00000000000000000000000000100000000391000000008 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to acct address Attn Kelly Fogarty Resoluion : shipped serial # B22H-0000000000000000, tracking 786126800000000100000000 DELIVERED Thusday 11/9/23 at 9:500000000M Signed for by: K.GREEN Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: not making contact with batteries uer Impact: none Patient Impact: none SE E.00000000.000000007 See ref Case # 00000000225865800000000 Cut called stating that the new battery they puchased has difficuty in making contact with the AED and 000000000000000000000000001ld not power on. Diagnostic performed by Engineer : They tried several 000000000000000000000000001king batteries and encoutered the same issu. If they remove the battery halfway and it makes contact, it then powers on. Advised cut to remove device and will replace device uder warranty. 000000008900000000 Install before date: 2000000003-00000000-00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000391000000005 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-11-08 23:36 | No
3591 | Yes | Monitoring Interrupt' & then 'Performing Self Test' (patient 0 | Monitoring Interrupt' & then 'Performing Self Test' (patient 0 | Internal Remarks : All remote suport provided by Neil Paice - Produt Suport Engineer & Roongruchai Chongolnee - Produt Suport Specialist ----------------------------------------------------------------------------- | 2023-11-08 23:41 | No
3592 | Yes | Screen is too dark to view - Per GFE, pending part 06 | Screen is too dark to view - Per GFE, pending part 0/6 | 000000000000000000000000001O##: 000000000000000000000000001000000003910000000000000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer has a V600000000and the screen is too dark to view. uit has a 1st gen LCD, and 000000000000000000000000001ld need to udate it to a 2nd gen LCD. Recommended replacing the u assy Part #0000000056150000000091, u assy withou uuuProblem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: too dark to view LCD uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order the u assy and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: 817-329-2836 Preston IP & Issu: Screen is too dark to view Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-08 23:51 | No
3593 | Yes | Device has display issues - this is an incoming loaner that also requires repairs in addition to FFU service. (FFU... | Device has display issues - this is an incoming loaner that also requires repairs in addition to FFU service. (FFU... | 000000000000000000000000001O##: 000000000000000000000000001000000003993200000000nyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v0000000028 Diagnostic performed by Engineer : Confirmed complaint the uits Display has a distorted image intermittently. To correct the issu we will replace the Display Assembly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Resoluion : Replaced Display Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000399279 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v0000000028 Diagnostic performed by Engineer : Confirmed complaint the uits Display has a distorted image intermittently. To correct the issu we will replace the Display Assembly. Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair ----------------------------------------------------------------------------- | 2023-11-08 23:51 | No
3594 | Yes | Screen is broken. | Screen is broken. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 Resoluion : Replaced Display Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits Screen is cracked. To correct the issu the Display Assembly will need to be replaced. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000391188 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits Screen is cracked. To correct the issu the Display Assembly will need to be replaced. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000391113 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Follow u Requred : Problem Description: biomed reports the display on the device is broken, sending to bench Form ued: Information Requst Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE 00000000.300000000roblem Description/ueupected Resuts: biomed reports the display on the device is broken, sending to bench Diagnostic performed by Engineer : Problem Description: biomed reports the display on the device is broken, sending to bench Form ued: Information Requst Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Zach Bowers IP & Issu: 98586938 Screen is broken. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : 11/13/2000000003 Requsted Loaner has been shipped to Cut with FedEu Tracking# 716991710000000067. QuST SO# 20000000065 CASE#: 0000000022593122 PN: 989700000000000000000000000082 SN: 616136 Railings: Both Cutomer SN: 6167300000000 Turownsbug Fire Territory 00000000 E Northfield Dr Brownsbug, IN 0000000012-2113 Attn: Zach Brown 317-00000000-00000000200000000---------------------------------------------------------------------------- | 2023-11-08 23:56 | No
3595 | Yes | Auuiliary alarm supply failed | Auuiliary alarm supply failed | 000000000000000000000000001O##: 000000000000000000000000001000000003911300000000ccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to replace MC pcb to correct, cutomer acknowledged and will order part, issu resolved- Material Nuber 00000000561537531 - RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:Auillary alarm suply failed message uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing uit set u the uit alarms Auiliary alarm suply failed on screen, no patient involvement, in biomed verified uLL1115 in event log and no other codes to note, verified PM pcb has recently been replaced per FCO86600000000000000006 and is pn 1153636, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Brad Olson IP & Issu: " 69387886 getting check vent auiliary failed." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-09 00:12 | No
3596 | Yes | Battery defective | Battery defeu | 000000000000000000000000001O##: 00000000000000000000000000100000000383315 ----------------------------------------------------------------------------- | 2023-11-09 00:16 | No
3597 | Yes | Battery defect mu be replaced - GFE 0/8 | Battery defeumu be replaced - GFE 00/8 |  | 2023-11-09 00:16 | No
3598 | Yes | Battery defect mu be replaced - Pending quote for repair 04 | Battery defeumu be replaced - Pending quote for repair 0/4 |  | 2023-11-09 00:21 | No
3599 | Yes | Battery defect mu be replaced - GFE 0/8 | Battery defeumu be replaced - GFE 00/8 |  | 2023-11-09 00:21 | No
3600 | Yes | Battery defect mu be replaced - pending quote for repair | Battery defeumu be replaced - pending quote for repair |  | 2023-11-09 00:21 | No
3601 | Yes | Battery defective | Battery defeu |  | 2023-11-09 00:21 | No
3602 | Yes | Intermittent picture | Intermittent piure |  | 2023-11-09 00:30 | No
3603 | Yes | SPLIT - CRACKED SCREEN, NO POWER INDICATOR LIGHT, PHILLIPS WILL NOT SELL PARTS FOR THESE UNITS, NEEDS DISPLAY | SPLIT - CRACKED SCREEN, NO POWER INDICATOR LIGHT, PHILLIPS WILL NOT SELL PARTS FOR THESE UNITS, NEEDS DISPLAY | 000000000000000000000000001O##: 00000000000000000000000000100000000385191 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: RDT Sales Rep Jake Nyhart Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Cracked screen. No power indicator light. Needs display board and power board.No patient was involved in this case. Diagnostic performed by Engineer : Problem Description: Cracked screen. No power indicator light. Needs display board and power board. No patient was involved in this case. Cracked screen. No power indicator light. Needs display board and power board.No patient was involved in this case. Form ued: Bench Trouleshooting Steps Taken: NA Follow u Requred : Cracked screen. No power indicator light. Needs display board and power board. Resoluion : Sent to Bench ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Eric Bertelson IP & Issu: 900000000000000001 CRACKED SCREEN, NO POWER INDICATOR LIGHT, PHILLIPS WILL NOT SELL PARTS FOR THESE uITS, NEEDS DISPLAY Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-09 01:21 | No
3604 | Yes | Getting battery failure uLLduring PBCA replacement (FCO) | Getting battery failure uLLduring PBCA replacement (FCO) | Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Joshu Hendren - phone 715 6000000000000000000000000 IP & Issu: 57871000000005 - Getting battery failue uLLduing PVCA replacement Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Other: per FSE issu has been resolved, no trouleshooting needed. ----------------------------------------------------------------------------- | 2023-11-09 02:16 | No
3605 | Yes | Internal battery failure - GFE 0/0 | Internal battery failure - GFE 00/0 | 000000000000000000000000001O##: 00000000000000000000000000100000000392211 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer installed a good battery, and verified it operated fine. He ran it on battery, and it was able to charge properly too. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and they are getting an internal battery failue. The previou tech removed the battery so we were not able to trouleshoot the issu. Recommended installing a good battery and test the internal battery and charging of the battery. He will reply with the resuts. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: internal battery failue uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Tim Blazek - phone 00000000 70000000000000000100000000 IP & Issu: 6800000000200000000- Internal battery failue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-09 02:26 | No
3606 | Yes | uLL0 - battery failure - Sent 0 GFE 05 | uLL00 - battery failure - Sent 0d GFE 0/5 | 000000000000000000000000001O##: 0000000000000000000000000010000000039200000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer stated he placed the order for internal battery and received them. He plans to installed them today or when available. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: battery faut uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the uit has a battery faut uLLcode 11200000000 While researching the uit, in Auut 2000000003, the FCO66 was implemented, and the battery manuactued date was Oct 2000000005. Euplained to him it will need a battery and it coud solve the issu. Also euplain, he 000000000000000000000000001ld need to bu the batteries from u since Parts Souce’s battery do not operate with ou new PM pcba’s part #989800000000626900000000 internal battery ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Jorge Mosquda Issu: 68273252 uLL11200000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-09 03:16 | No
3607 | Yes | Battery failed, Send Letter, IS complete | Battery failed, Send Letter, IS complete | 000000000000000000000000001O##: 00000000000000000000000000100000000392700000000 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: battery failue uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states device is giving battery failue uLLcode 11200000000Confirmed device is new PM board with firmmware 1.300000000 Cutomer also states the battery is 3rd party. informed cutomer that 3rd party batteries do not commuicate with 1.300000000SW. Cutomer states he will order replacement battery and call back if futher assistance is needed. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jerry Solis 8189522200000000 IP & Issu: 2000000000000000000000000900000000/ Battery failed check uLL main battery replaced, power cycling appears to be neut remedy bu needs help. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created 000000000000000000000000001and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-09 03:31 | No
3608 | Yes | System does not turn on | System does not turn on |  | 2023-11-09 04:11 | No
3609 | Yes | System freezes | System freezes |  | 2023-11-09 04:16 | No
3610 | Yes | No image is shown | No image is shown |  | 2023-11-09 04:26 | No
3611 | Yes | Battery failed uLLmessage. | Battery failed uLLmessage. | 000000000000000000000000001O##: 00000000000000000000000000100000000389800000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer that battery may be defective and cauing the alarm, cutomer will order a new battery to replace, no futher follow u needed, issu resolved. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts:Battery failed alarm after FCO66 uer Impact:not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states FCO86600000000000000006 PM pcb was recently performed and at power on the uit alarms Battery failed on screen, no patient involvement, verified battery is u brand with 2000000001 date and is charged, requsts trouleshooting for repair. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Andrea Choate IP & Issu: 72100000000273 Battery failed uLLmessage. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-09 04:36 | No
3612 | Yes | CV u uLLFLUOSTR-IMAGEDS Image storage unavailable | CV u uLLFLUOSTR-IMAGEDS Image storage unavailable | 000000000000000000000000001O##: 0000000000000000000000000010000000038510000000000000000-Nov-2000000003 Orser Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: technologist Device ue: in ue Eupected and ueupected Resuts: image storage uavailable uer Impact: uable to ue u-ray Patient Impact: na SE8.2 Resoluion : Replaced IPPC and hard drives Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-11-09 04:36 | No
3613 | Yes | SPLIT_-Death The cuomer reported that patient died-needs logs. Apnea alarm failed. . Pending RMR 0ec0 | SPLIT_-Death The cuomer reported that patient died-needs logs. Apnea alarm failed. . Pending Ru0ec0/ |  | 2023-11-09 04:39 | Yes
3614 | Yes | AED is emitting the triple chirp, fails self-test | AED is emitting the triple chirp, fails self-test | 000000000000000000000000001O##: 000000000000000000000000001000000003931300000000impson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED was not powering on cutomer replace the battery and pads. AED kept triple chirping and cutomer removed battery . Instrut cutomer to insert the battery, AED went straight into test-mode AED failed self-test with uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 11 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A12I-00000000992 Battery Install Before# 2000000009/00000000/300000000 Lot#00000000000000005P Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. ----------------------------------------------------------------------------- | 2023-11-09 04:41 | No
3615 | Yes | UNIT WILL NOT POWER ON, PHILLIPS WILL NOT SELL PARTS FOR THESE UNITS, | UNIT WILL NOT POWER ON, PHILLIPS WILL NOT SELL PARTS FOR THESE UNITS, | Internal Remarks : Interaction Type: New Case Caller: Eric Bertelson IP & Issu: "900000000000000005 uIT WILL NOT POWER ON, PHILLIPS WILL NOT SELL PARTS FOR THESE uITS," Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-09 05:01 | No
3616 | Yes | Audio&Sounds/Audio Confirmed: Speaker Malfunction: Mainboard/HW: Replaced parts. | Audio&Sounds/Audio Confirmed: Speaker Malfunuon: Mainboard/HW: Replaced parts. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000917 22-Nov-2000000003 Bugannavar, Praveen The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000056000000007371,SS PCA VS3-VS00000000ain Board D5 FS Tools & Equpment : PATIENT SIMuATOR,200000000000000005-22 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-19 00000000:00000000:00000000 Diagnostic performed by Engineer : Confirmed faut. Has the Speaker Mal Fuction uLLafter replacing the speaker and clearing the uLLlog. needs a main board replacement. Resoluion : Replaced the Main Board and the NIBP modue. udated the SW and Config. Ful PM passed. Retuned to service. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Patient Monitoring Eupected and ueupected Resuts: Speaker malfuction uLL and no auio uer Impact: Alternate device ued Patient Impact: no impact on patient SE 7.33 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000039300000000 17-Nov-2000000003 Andrew Vincent The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace mainboard Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Patient Monitoring Eupected and ueupected Resuts: Speaker malfuction uLL and no auio uer Impact: Alternate device ued Patient Impact: no impact on patient SE 7.33 Diagnostic performed by Engineer : Able to recreate "Auio failed" and "speaker malf" uLL. uable to recreate actul loss of auio/alarms. Replaced speaker and cleared uLLlog as per service manul recommendations however the failed to rectify the issu. Mainboard to be ordered. Parts Consued : 1.0000000000000000000000000000000056000000008000000001,SS CBL Speaker Assy FS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000386600000000Alhad Apte The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace speaker assembly Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Ou of ue Eupected and ueupected Resuts: Speaker malfuction uer Impact: No impact on uer Patient Impact: no impact on patient SE Not provided Diagnostic performed by Engineer : Speaker is not 000000000000000000000000001king. Replacement requred by field enginer. ----------------------------------------------------------------------------- | 2023-11-09 05:51 | No
3617 | Yes | Space low | 0R/u FD0Space low | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Ou of memory SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 12.11.2000000003 11:57:53 uC (86000000008285) RDI Accuacy Yes $Q7: (confirmed reported problem, Space low $Q8: (Trouleshooting & Resutst: patient image store space offten low, ippc hardisk have problem) $Q9: (Repair steps:replace ippc hardisk, format new hardisk) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools NA SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 12.11.2000000003 11:57:500000000uC (86000000008285) Service content: IPPC storage space is often insuficient, replace the ippc hard drive Service resuts: The equpment retuned to normal ue after repair Other important matters: None T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that image storage space is low and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Free disk space is too low. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Euecue Database Consistency Test and Repair, delete Repaired patients, and see if the space is restored. 2. If there is no effect, euecue Recreate Image Store. 3. If there is no effect, try Bypass hard disk, and then reinstall the IPPC software. 00000000If there is no effect, ue PC. Diagnostic Tool performs diagnostic tests and replaces corresponding spare parts 5. Check machine uLLreports and perform faut repairs Recommended spare part uuand name: " 0000000080000000050000000092 IP PC Revised_II no HDD &0000000080000000050000000000000000Hard disk spare part " Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022587900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD 00000000800000000500000000000000002 200000000Yang Fan-86-1338000000001985 Normal shipment ----------------------------------------------------------------------------- | 2023-11-09 07:41 | No
3618 | Yes | OCAP:0/u FD0System cannot make euposure and geometry movement problem and collimator problem | OCAP:0R/u FD0System cannot make euposure and geometry movement problem and collimator problem | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Locked & no rays emitted Internal Remarks: 2000000001.2000000003 16:00000000:18 uC (8600000000000000009) $Q1:(Y)$Q2:(N)$Q3:()$Q00000000)$Q5:() $Q6:( Can't euposue ) $ Q7: (Confirmation & Description of problem: Confirmed, Can't euposue sometimes. ) $Q8: (Trouleshooting & Resuts: Check the logfile, CAN commuication missed. BOu TBCB and SIBBOu uit are supiciou. After replace BOu TBCB board, the problem happened after 3 days, so, the SIBBOu uit is most supiciou. ) $Q9: (Repair steps: Replace BOu TBCB Board AD7, and SIBBOu uit, download SIBBOu uit software. ) $Q100000000 (Test and Inspection data&Device statu: Device retrued to ful fuctionality. Test and Inspection data refer to PA. ) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Neut Calibration date (N/A) For Mainland China CV/uR, CT engineers only. Please fill in: Curent Tue: 56300000000uInstalled data: 00000000/00000000/2000000003 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 2000000001.2000000003 15:58:00000000uC (8600000000000000009) Service content: The perspective cannot be euposed intermittently, and the mechanical movement is not available.CAN commuication interruted.Replace SIBBOu uit, replace BOu TBCB Board AD7. Service resuts: After several weeks of observation, the faut has not recured and the machine can be ued normally. Safety Qustions: Case: 0000000022588000000009 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Overview of cutomer problem: The equpment cannot be euposed, the mechanical movement fails, and the condenser reports an uLLEngineers are requred to be present. Is the faut caued by the device itself? Yes Trouleshooting Action: Remote diagnosis and test resuts: Withou PRS and logs, the caue of the faut cannot be determined.Engineers are requred to be on site PRS TR statu: None PRS Repair Action: Recommended onsite repair steps: Engineers are on site to perform repairs based on log uLLreports and indicator light statu, especially the common parts of these three fauts, suh as SIB problems Cutomer demand notes: Onsite Neut Step Action: 00000000918869000000009 Teacher Hus landline called ----------------------------------------------------------------------------- | 2023-11-09 07:51 | No
3619 | Yes | check the device | uheck the device | BuQ&R Feedback: No Reason: Internal Suport Problem Reported by Cutomer: crash SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: 1823130000000000000000called and said that 000000000000000000000000001kers still need to be dispatched, and RSE has been notified SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problem: The equpment crashed last night and was uable to emit rays. Youneed to go on site to check whether the equpment itself is fauty: Yes Trouleshooting Action: Key faut code or phenomenon: The equpment crashed last night and was uable to emit rays. u-ray conditions, restart is curently normal, youneed to go on site to check the equpment remote diagnosis and test resuts: n/a PRS TR statu: All-pass Repair Action: Recommended on-site repair steps: Go to the site to check whether the euternal power suply is normal, whether there is a phase loss, and whether the MIuis damaged Cutomer Remarks: Contact the repair person as soon as possible to confirm the arrival time Safety Qustions: Case: 0000000022588118 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang ZhuLast udated by: Weiqiang Zhu1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: 1823130000000000000000called again and asked when the engineer coud come. Engineer Deng called back and said he 000000000000000000000000001ld call the cutomer back. Internal Remarks: 1000000001.2000000003 00000000:28:00000000 uC (860000000082700000000 $Q1: ( N )$Q2: (N) $Q3: (na) $Q00000000(na) $Q5: (na) $Q6: ( check the device. ) "RDI Accuacy: Yes". $Q7: (Confirmation & Description of problem: check the device. ) $Q8: ( Check the device and measue the power suply. ) $Q9: (Check the device and measue the power suply. ) $Q100000000 ( Check the system fuction and it 000000000000000000000000001ks all well. Test and inspection data refer to: [ PA ] & Device retuned to service with ful fuctionality. ) $Q11:( calibrated tools refer to PA ) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: MM/ DD/ YYYY ^Curent Tue uage: scan second_ SAP Transfer: Dispatched for Onsite Suport Neut Step Action: 1823130000000000000000Incoming calls requre engineers to check the backgroud Euternal Remarks: 11.11.2000000003 00000000:18:00000000uC (860000000082700000000 Service content: Detection of high-voltage modues and power suply of measuing equpment. Service resuts: The repair is completed and the equpment is runing normally. Other important matters: Pay attention to the stability of the power suply. ----------------------------------------------------------------------------- | 2023-11-09 08:16 | No
3620 | Yes | Defibrillator/Monitor/Treatment port failure | 0N/Effi0 DFM0 Defibrillator/Monitor/Treatment port failure | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 16.11.2000000003 00000000:53:39 uC (860000000000000000100000000 $ASP engineer:[honglejia wudehong] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.The cable connector latch cannot be reset $Q8: The bukle cannot be reset after the cable connector is connected $Q9: Replacement therapy socket $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: $ Calibrated Tools name: (ProSim8) S/N (26712200000000 Neut Calibration date (00000000/28/20000000000000000 (ESA6200000000 S/N (0000000052000000005) Neut calibtation data (000000002/20000000000000000 (Plue lmpuse 70000000000000000P) S/N (27920000000000000000 Neut calibtation data (10000000030000000020000000000000000 Euternal Remarks: 16.11.2000000003 00000000:53:36 uC (860000000000000000100000000Department information: [Department: Agent Contact: Deng Gong Tel: 18900000000893531]* $ASP engineer:[honglejia wudehong] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.The cable connector latch cannot be reset $Q8: The bukle cannot be reset after the cable connector is connected $Q9: Replacement therapy socket $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: $ Calibrated Tools name: (ProSim8) S/N (26712200000000 Neut Calibration date (00000000/28/20000000000000000 (ESA6200000000 S/N (0000000052000000005) Neut calibtation data (000000002/20000000000000000 (Plue lmpuse 70000000000000000P) S/N（27920000000000000000Neut calibtation data(10000000030000000020000000000000000Safety Qustions: Case: 00000000225881200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Treatment port failue The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: upacking and testing mutiple machines, t000000000000000000000000001of them foud that the clips of the treatment terminals did not retun when the handles were installed, and the others were normal. Replacing the handles had no effect. Repair Action: T2 Recommended on-site repair steps: It is recommended to go to the site to check the treatment port and fill in the CFM and Escalation Recommended spare part uufor T2 with description: DFM100000000 Therapy Socket Assembly (Treatment Port)* 0000000056000000008971 Notes on cutomer requrements: na Test/Inspection: N/A Test Equpment: N/A Problem Reported by Cutomer: Card does not retun SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-09 08:16 | No
3621 | Yes | Hot case B:/0/uDFM0 Defibrillator/Monitor/Treatment port failure | Hot case B:/0N/Effi0 DFM0 Defibrillator/Monitor/Treatment port failure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Card does not retun SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 16.11.2000000003 00000000:59:200000000TC (860000000000000000100000000 $ASP engineer:[honglejia wudehong] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.The cable connector latch cannot be reset $Q8: The bukle cannot be reset after the cable connector is connected $Q9: Replacement therapy socket $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: $ Calibrated Tools name: (ProSim8) S/N (26712200000000 Neut Calibration date (00000000/28/20000000000000000 (ESA6200000000 S/N (0000000052000000005) Neut calibtation data (000000002/20000000000000000 (Plue lmpuse 70000000000000000P) S/N (27920000000000000000 Neut calibtation data (10000000030000000020000000000000000 Euternal Remarks: 16.11.2000000003 00000000:59:22 uC (860000000000000000100000000Department information: [Department: Agent Contact: Deng Gong Tel: 18900000000893531]* $ASP engineer:[honglejia wudehong] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.The cable connector latch cannot be reset $Q8: The bukle cannot be reset after the cable connector is connected $Q9: Replacement therapy socket $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: $ Calibrated Tools name: (ProSim8) S/N (26712200000000 Neut Calibration date (00000000/28/20000000000000000 (ESA6200000000 S/N (0000000052000000005) Neut calibtation data (000000002/20000000000000000 (Plue lmpuse 70000000000000000P) S/N（27920000000000000000Neut calibtation data(10000000030000000020000000000000000Safety Qustions: Case: 0000000022588122 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Treatment port failue The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: upacking and testing mutiple machines, t000000000000000000000000001of them foud that the clips of the treatment terminals did not retun when the handles were installed, and the others were normal. Replacing the handles had no effect. Repair Action: T2 Recommended on-site repair steps: It is recommended to go to the site to check the treatment port and fill in the CFM and Escalation Recommended spare part uufor T2 with description: DFM100000000 Therapy Socket Assembly (Treatment Port)* 0000000056000000008971 Notes on cutomer requrements: na Test/Inspection: N/A Test Equpment: N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-09 08:16 | No
3622 | Yes | SA:0uDFM0 Defibrillator/Monitor/The electrode plate is unrecognizable | SA:0/Effi0 DFM0 Defibrillator/Monitor/The eleuode plate is unrecognizable | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022588163 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: SA:00000000/uDFM100000000 Defibrillator/Monitor/electrode panel cannot be recognized ----------------------------------------------------------------------------- | 2023-11-09 08:31 | No
3623 | Yes | M0A - u00 Defibrillation test failed | u Defibrillation test failed | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022588179 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: ECR:： 00000000/HeartStart uuumonitor/defib/ Neut Step Action: ##d ASP uiedong for follow u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-11-09 08:36 | No
3624 | Yes | Auto restart. | Auto restart. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: ECR:00000000/HS uL+/ Safety Qustions: Case: 0000000022588181 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Confirmed with u Chen that ASP Jingmao Zhonglian Buiness will follow u Transfer to sales handling, This case is considered invalid and withou any futher steps if qutation is not accepted, Will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-11-09 08:36 | No
3625 | Yes | A geometry uLLoccurred during a technique. | A geometry uLLoccurred during a technique. | 000000000000000000000000001O##: 000000000000000000000000001000000003938900000000amata, Takayui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Geometry uLLoccured duing a techniqu. u Ikinose 0000000036719335 ・Aroud 11:100000000 He 000000000000000000000000001ld like u to check the log. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Geometry uLLoccured duing the operation. u Ikinose 0000000036719335 ・Aroud 11:100000000 He 000000000000000000000000001ld like u to check the log. uer Impact: No effect Patient Impact: No effect SE 8.2.3000000005 Resoluion: I checked the log. At 11:0000000028, "EStopActive = 1 changed" was sudenly entered, and the Geometry power was off. - Becaue the timing was suh that table operations were not being performed, I don't think there is a modue or cable condition. - I think the problem was caued by uintended mistaken operation. ・I reported the above to u Ichonase and received his/her consent. ----------------------------------------------------------------------------- | 2023-11-09 09:16 | No
3626 | Yes | IIA_0-0-0- information on the label was wiped out after cleaning and disinfection test | IIA_00-0000- information on the label was wiped out after cleaning and disinfeuon test |  | 2023-11-09 10:24 | No
3627 | Yes | IIA_0-0-0 - 'incompatible battery' alarm with an audible tone | IIA_00-0000- 'incompatible battery' alarm with an audible tone |  | 2023-11-09 10:34 | No
3628 | Yes | :The equipment does not move(SPLIT) | :The equipment does not move(SPLIT) |  | 2023-11-09 11:06 | No
3629 | Yes | 0 - u00 Operational check failed | 00- u000- Operational check failed | 000000000000000000000000001O##: 00000000000000000000000000100000000386800000000 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Diagnostic performed by Engineer : EOL produt part not available. Informed the same to cutomer and requsted them to go for Equpment level ugradation. Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A ----------------------------------------------------------------------------- | 2023-11-09 12:36 | No
3630 | Yes | SA:0uDFM0 Defibrillator/Monitor/Alarm:Batery need to be replaced | SA:0/Effi0 DFM0 Defibrillator/Monitor/Alarm:Batery need to be replaced | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Consutation Internal Remarks: Call not answered Neut Step Action: ##d to ASP Honglejia Buiness for follow-u Neut Step Action: Mingping District Hospital of Integrated Traditional Chinese and Western Medicine Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 00000000225892600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: LiuWeiming, 18916700000000559 ----------------------------------------------------------------------------- | 2023-11-09 13:36 | No
3631 | Yes | Audio test Failed on operational check | Audio test Failed on operational check |  | 2023-11-09 14:06 | No
3632 | Yes | WFS problem | WFS problem | 000000000000000000000000001O##: 000000000000000000000000001000000003900000000000000009-Nov-2000000003 Sang000000000000000000000000001 lee The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: radiologist Device ue: in ue Eupected and ueupected Resuts: WFS problem uer Impact: inconvenience problem Patient Impact: none SE 8 Tools & Equpment : MuTIMETER, 200000000000000001-27 00000000:00000000:00000000 Resoluion: Reconnected wireless footswitch: passed. Verified all the fuctions 000000000000000000000000001k properly. The system is fuly operational. ----------------------------------------------------------------------------- | 2023-11-09 14:51 | No
3633 | Yes | SA:0uDFM0 Defibrillator/Monitor/Electrode plate issue | SA:0/Effi0 DFM0 Defibrillator/Monitor/Eleuode plate issue | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 00000000225896600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Hongqin Tang Jingyi 0000000030000000061800000000 called consut ----------------------------------------------------------------------------- | 2023-11-09 14:56 | No
3634 | Yes | Battery failure | Battery failure | 000000000000000000000000001O##: 00000000000000000000000000100000000390000000000000000KARAKAYA, Mehmet The system does not meet the specification for the performed service and is not retuned to ue. Resoluion: No uer-related failue was detected duing the failue detection. Accordingly; 989800000000190000000071 battery of one DFM100000000 Defibrillator device with part nuber shoud be ordered uder warranty. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: Battery failue uer Impact: No Patient Impact: No SE 2.00000000 Diagnostic performed by Engineer: Cutomer was contact regarding the relevant failue. Duing the inspection with the cutomer, it was detected that the system coud not pass the test as it completed the battery life. ----------------------------------------------------------------------------- | 2023-11-09 15:06 | No
3635 | Yes | Security Issue_Legal Issue_Importing patients from Worklist, complaints about miued images between patients | [Q] PSA_PSE_Security Issue_Legal Issue_Importing patients from Worklist, complaints about miued images between patients |  | 2023-11-09 15:28 | No
3636 | Yes | The LS gives a DPM error | The LS gives a DPM error |  | 2023-11-09 15:31 | No
3637 | Yes | Latency of the alarms report =Case in progress as of DEC00 | Latency of the alarms report =Case in progress as of DEC0/03 | 000000000000000000000000001O##: 00000000000000000000000000100000000387729 Thierry LE GONIDEC Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Reguatory Qustions: No/No Diagnostic performed by Engineer: No/No ----------------------------------------------------------------------------- | 2023-11-09 15:46 | No
3638 | Yes | DFM0 showed low battery. | DFM0 showed low battery. | 000000000000000000000000001O##: 00000000000000000000000000100000000500000000000000005 11-Dec-2000000003 Alejandro Rodrigo Mauier Córdoba The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Parts Consued: 2.0000000000000000000000000000000056000000009971,DFM100000000 ASSY-Label Set-Spanish Resoluion: Equpment located in a warehoue ou of ue. The cutomer tries to perform the fuctional tests and the battery is not charged. Equpment is left on charge. FCO860000000000000000600000000is carried ou according to protocol. Equpment left in service Problem Description by Engineer: Cutomer Fuction/Role: Talk to suervisor Patricia Briñon 638900000000857 Device ue: Not in clinical ue Eupected and ueupected Resuts: Battery failue uer Impact: None Patient Impact: None SE 2.00000000.21 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000767 Rau Cordero Calderon The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Talk to suervisor Patricia Briñon 638900000000857 and she says that the DFM100000000 indicates low battery when tuned on. Photo of the battery in the case is attached, it is an old battery 18216-000000002-S-REV H Assign FSE. Read SB861000000002800000000 ugrade to 2.00000000.32 Order part: 989800000000190000000071 - DFM100000000 Lithiu Ion Battery u2nd Floor, left zone, control C uI Coronary Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Talk to suervisor Patricia Briñon 638900000000857 Device ue: Not in clinical ue Eupected and ueupected Resuts: Battery failue uer Impact: None Patient Impact: None SE 2.00000000.21 Diagnostic performed by Engineer: Talk to suervisor Patricia Briñon 638900000000857 and she says that the DFM100000000 indicates low battery when tuned on. Photo of the battery in the case is attached, it is an old battery 18216-000000002-S-REV H Assign FSE. Read SB861000000002800000000 ugrade to 2.00000000.32 Order part: 989800000000190000000071 - DFM100000000 Lithiu Ion Battery u2nd Floor, left zone, control C uI Coronary ----------------------------------------------------------------------------- | 2023-11-09 16:06 | No
3639 | Yes | Battery failure 00 | Battery failure 0-00 | 000000000000000000000000001O##: 0000000000000000000000000010000000050000000099000000001-Dec-2000000003 Alejandro Rodrigo Mauier Córdoba The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: Battery failue uer Impact: Inability to ue the equpment in battery mode and Portuuse languge Patient Impact: The equpment is ued with limitations SE Resoluion: Equpment located in a warehoue ou of ue. The cutomer tries to perform the fuctional tests and the battery is not charged. Equpment is left on charge. FCO860000000000000000600000000is carried ou according to protocol. Equpment left in service Parts Consued: 2.0000000000000000000000000000000056000000009971,DFM100000000 ASSY-Label Set-Spanish ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000771 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: Battery failue uer Impact: Inability to ue the equpment in battery mode and Portuuse languge Patient Impact: The equpment is ued with limitations SE Diagnostic performed by Engineer: The client (Jun Lus Blanco 60000000000000000286) tells u that the equpment's battery is damaged and the equpment has the Portuuse languge. The material order is managed and the case is transferred onsite so that the equpment can be reviewed by FSE. Resoluion: Case is transferred onsite. Follow u Requred: The client (Jun Lus Blanco 60000000000000000286) tells u that the device's battery is damaged and the device has the Portuuse languge. - Equpment located on the 2nd floor, left zone, control C Coronary uI - Assign FSE - Requst cited material ----------------------------------------------------------------------------- | 2023-11-09 16:06 | No
3640 | Yes | uCVu w/ u0 u- System often gives an uLLmessage 05 | uCVu w/ u00 u- System often gives an uLLmessage 05 | 000000000000000000000000001O##: 0000000000000000000000000010000000039610000000017-Nov-2000000003 Tawast, Erik The system does not meet ful Specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Cardiologist Device ue: Imaging Eupected and ueupected Resuts: Systems freezes when uing TEE transduers uer Impact: Needs to reboot the system Patient Impact: Longer procedue SE 9.000000005 Resoluion: uable to Dulicate the reported problem on the EPIQ CVu [SN: u823B1761] that was originally reported by the cutomer becaue it was in ue. u8-2t transduers [SN: F00000000PDQ and F00000000G00000000 tested on another device. u8-2t transduer with 000F00000000PDQ had a light vertical stripe in the center of the image and therefore it needs to be replaced. Follow u Requred: Replacement u8-2t transduer needed. Internal Comments: Connecting u8-2t transduer with 000F00000000G00000000caued an uLL[DiagCode: 000000005]. After cleaning the pins of the connector the transduer 000000000000000000000000001ked withou any issus. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000039000000006 BENCH 000000000000000000000000001D The system does not meet ful Specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: CCC: Please schedue 2h onsite with cutomer. Pasi can go here on Monday while he is onsite already. Please tell the cutomer that transduers F00000000PDQ and F00000000G00000000need to be with the system on the visit! FSE: Please clean u8-2ts: F00000000PDQ and F00000000G00000000and system transduer ports. Take logs. Check also Wifi connection! Problem Description by Engineer: Cutomer Fuction/Role: Cardiologist Device ue: Imaging Eupected and ueupected Resuts: Systems freezes when uing TEE transduers uer Impact: Needs to reboot the system Patient Impact: Longer procedue SE 9.000000005 Diagnostic performed by Engineer: I sent e-mails with cutomer. New EPIQ CVu has a lot of issus with TEE transduers. When a transduer is connected, the system often gives uLLmessage 000000005 or uLL0000000000000000Foreign debris or Moistue has been detected in the transduer connector". These uLL were seen with transduers F00000000PDQ and F00000000G00000000 This issu needs an onsite visit and both transduer connectors and system connectors shoud be cleaned (Acqusition modue was replaced Oct 17th). ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000020000000017-Nov-2000000003 Parts Consued: 1.0000000000000000000000000000000056180000000077, Transduer, u8-2t ----------------------------------------------------------------------------- | 2023-11-09 16:56 | No
3641 | Yes | DESAT alarm was not called for room 0 | DESAT alarm was not called for room 0. | 000000000000000000000000001O##: 00000000000000000000000000100000000396000000005 Martin, John The system meets the specification for the performed service and is retuned to ue Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Tele Tech has some sectors that are not showing the alarm limits for HR and SpO2. Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Foud the sector action butons were redued. Instruted to hold down Ctrl + u arrow to restore butons. Alarm limits are now showing for all sectors and all rows. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Lauen moore IP & Issu: 7700000000532 not alarming and vitals are not appearing Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & 000000000000000000000000001rtac Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-09 17:01 | No
3642 | Yes | Defibrillator discharge failure | Defibrillator discharge failure |  | 2023-11-09 17:26 | Yes
3643 | Yes | The generator does not start. | The generator does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000396339 November 13, 2000000003 Alejandro Lopardo The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The 200000000power suply of the generator coud be failing, it is recommended to replace it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No impact. uer Impact: No impact Patient Impact: No impact SE 000Resoluion: The client says that the equpment did not show the uLL after repairing the 0000000000000000000000000010000000039600000000. Equpment is operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003963100000000Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The 200000000power suply of the generator coud be failing, it is recommended to replace it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No impact. uer Impact: No impact Patient Impact: No impact SE 000Diagnostic performed by Engineer: The 200000000power suply of the generator coud be failing, it is recommended to replace it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The 200000000power suply of the generator coud be failing, it is recommended to replace it. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: The 200000000power suply of the generator coud be failing, it is recommended to replace it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-11-09 17:31 | No
3644 | Yes | Battery Failure - pending repair (pending repair quote) | Battery Failure - pending repair (pending repair quote) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000622 18-Nov-2000000003 Lus Esteves The system does not meet the specification for the performed service and is not retuned to ue Resoluion: Qute sent to replace the battery that is damaged and the potentiometer that controls the navigation buton Diagnostic performed by Engineer: . Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Follow u Requred: MO carried ou 000000005h ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000029 18-Nov-2000000003 Lus Esteves Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2023-11-09 17:46 | No
3645 | Yes | /startup failture | 0R/u FD0startup failture | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Neut Step Action: Send order to FSE Gu Yingjie Euternal Remarks: 00000000.11.2000000003 12:20000000000000000TC (86000000008398) Service content: The host cannot start intermittently, the measued BIOS battery voltage is low, replace the battery Service resut: The device retuns to normal Problem Reported by Cutomer: Cannot start when booting, need to restart manully Contacted FSE Zhang Chen Neut Step Action: Gu YuIE Internal Remarks: 00000000.11.2000000003 12:26:00000000uC (86000000008398) RDI Accuacy:Yes $Q7: (Confirmed reported problem& bootu uLL) $Q8: (checked indicate: the voltage of BIOS battery was low) $Q9: (replace the battery) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated tools name(NA) S/N (NA) Neut calibration date (NA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: second: N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000225900000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: Startu failue Detailed description of cutomer problem or faut phenomenon: Cannot start when booting, need to restart manully Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Host PC manul boot Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Replace host pc bios battery Notes on cutomer requrements: ----------------------------------------------------------------------------- | 2023-11-09 17:51 | No
3646 | Yes | QACASE for PCI imported product#000000Power off | QACASE for P0imported produu00000000 Power off |  | 2023-11-09 17:52 | No
3647 | Yes | AED triple chirp - | AED triple chirp - | 000000000000000000000000001O##: 00000000000000000000000000100000000391600000000 Kat Gilliam The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : waiting on order to ship, order nuber 000000000000000057588917 Problem Description by Engineer : Cutomer Fuction/Role: end uer Device ue: ouside of clinical uage Eupected and ueupected Resuts: Cutomer called and stated the AED was giving an uLL After trouleshooting the AED emitted a triple chirp and said to remove and reinsert battery. AED is uder IB Warranty and to be replaced at zero cost to the cutomer. Device serial# B17G-000000001800000000ser Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer called and stated the AED was giving an uLL After trouleshooting the AED emitted a triple chirp and said to remove and reinsert battery. AED is uder IB Warranty and to be replaced at zero cost to the cutomer. Device serial# B17G-000000001800000000ollow u Requred : Tu00000000100000000uningdale Dr Rancho Mirage, u2700000000 Attn To: Henry Goldman Internal Comments : FedEu tracking 7861270000000000000000. DELIVERED Thusday 11/9/23 at 100000000100000000AM Resoluion : Cutomer has not responded back regarding AED being back in service. Closing 000000000000000000000000001----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000391621 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-11-09 18:11 | No
3648 | Yes | After device boot up, black screen | After device boot up, black screen |  | 2023-11-09 18:12 | No
3649 | Yes | The screen has damage and the battery does not charge anymore. Bou u2 | The screen has damage and the battery does not charge anymore. Bou Tempu0 |  | 2023-11-09 18:16 | No
3650 | Yes | It does not start. | It does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000396765 15-Nov-2000000003 Jun Fombella Vela The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : SutePC is failing, it needs to be replaced. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Does not start uer Impact: Cannot ue the device Patient Impact: No impact SE 1.2.5.00000000Resoluion: Sute PC is replaced and system software is reloaded. Parts Consued: 1.0000000000000000000000000000000080000000070000000000000000,uR1.2.5 SW Pack 1.00000000000000000000000000000000800000000660000000083,PC Sute Tools & Equpment : DOSE-KVP-LIGHT METER, 2000000008-00000000-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000396672 Adrian Sancho Lopez The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : SutePC is failing, it needs to be replaced. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Does not start uer Impact: Cannot ue the device Patient Impact: No impact SE 1.2.5.00000000Diagnostic performed by Engineer: SutePC is failing, it needs to be replaced. Internal Comments: uLL# and/or description of uLL: N/A Repair Action : SutePC is failing, it needs to be replaced. Follow u Requred: uLL# and/or description of uLL: N/A Repair Action : SutePC is failing, it needs to be replaced. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-11-09 18:36 | No
3651 | Yes | LS DPM uLLpacer issue | LS DPM uLLpacer issue |  | 2023-11-09 18:51 | No
3652 | Yes | AED defect | AED defect | Internal Remarks: Referred cutomer to specialty dealer Internal Remarks: udeliverable: u retail partner - System: uknown Ventilation produt - processing nuber: 00000000225912300000000 ##corrected for mike.messerschmidt00fas.bud.de and sent again ----------------------------------------------------------------------------- | 2023-11-09 19:06 | No
3653 | Yes | uL broken cable | uL broken cable |  | 2023-11-09 19:09 | No
3654 | Yes | Screen calibration | Screen calibration | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000238 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king correctly even after calibration to correct the issu we will replace the Display assembly. Resoluion : Replaced Display Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000397100000000 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king correctly even after calibration to correct the issu we will replace the Display assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000396900000000 Al Hammad Lamiae The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Problem Description: Biomed states touh screen not responding Biomed states touh screen not responding Form ued: Bench Trouleshooting Steps Taken: cutomer states:We have had mutiple problems with this monitor losing the screen calibration. We have to enter the service mode to reset the calibration. We have followed the directions, removing the screen protector, calibrating the screen, and applying a new screen protector. It will hold the calibration for 2-3 restarts, then it begins to not 000000000000000000000000001k again. This is especially a problem on the right side of the screen when entering patient's age, etc. Cannot select the 6 key. Also an issu with mauimizing the image for video laryngoscopy or the captue buton when it is in ue. Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: Biomed states touh screen not responding Diagnostic performed by Engineer : Problem Description: Biomed states touh screen not responding Biomed states touh screen not responding Form ued: Bench Trouleshooting Steps Taken: cutomer states:We have had mutiple problems with this monitor losing the screen calibration. We have to enter the service mode to reset the calibration. We have followed the directions, removing the screen protector, calibrating the screen, and applying a new screen protector. It will hold the calibration for 2-3 restarts, then it begins to not 000000000000000000000000001k again. This is especially a problem on the right side of the screen when entering patient's age, etc. Cannot select the 6 key. Also an issu with mauimizing the image for video laryngoscopy or the captue buton when it is in ue. ----------------------------------------------------------------------------- Cutomer comment : 000000003000000000000000000000000Gb5QMQAZ ----------------------------------------------------------------------------- | 2023-11-09 19:21 | No
3655 | Yes | M0A - u00 Functional check failed | M00A - u00 - Funuonal check failed | 000000000000000000000000001O##: 00000000000000000000000000100000000397330000000013-Nov-2000000003 Matteo Carraro The system meets the specification for the performed service and is retuned to ue. Resoluion: Check faut; discharge tests were carried ou with a positive oucome. Fuctional check was carried ou with a positive oucome. The system is 000000000000000000000000001king perfectly again. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Laua Ghirardello Device ue: Diagnosis Eupected and ueupected Resuts: problem on verification uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from Clinical E.Ghirardello Laua who reported problems on defibrillator. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000396000000009 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Follow u Requred: I heard from clinical engineering Laua Ghirardello who reported problems on defibrillator, they ask for the presence of the technician for any calibrations and futher fuctional tests, from system it is uder contract, assign on-site, equpped with tools Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Laua Ghirardello Device ue: Diagnosis Eupected and ueupected Resuts: problem on verification uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from Clinical E.Ghirardello Laua who reported problems on defibrillator. Diagnostic performed by Engineer: I heard from clinical engineering Laua Ghirardello who reported problems on defibrillator, they ask for the presence of the technician for any calibrations and futher fuctional tests. ----------------------------------------------------------------------------- | 2023-11-09 19:56 | No
3656 | Yes | Battery is down now and then | Battery is down now and then |  | 2023-11-09 19:56 | No
3657 | Yes | It is not possible to acquire the trace through ucable | It is not possible to acquire the trace through ucable | 000000000000000000000000001O##: 00000000000000000000000000100000000397591 22-Nov-2000000003 Raele, antonino The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineering Device ue: patient monitoring Eupected and ueupected Resuts: Trace cannot be acqured via ucable uer Impact: no Patient Impact: no SEna Parts Consued: 1.0000000000000000000000000000000056000000009171,DFM100000000 Measuement Modue uSpO2+CO2 Resoluion: Replaced DFM100000000 Measuement Modue u+ Sp00000000 + CO2. Performed fuctional tests with personal test cable with positive resuts. Cable suplied to the department not 000000000000000000000000001king detected; codes are released to the coordinator who signs at the bottom to order the new one. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003960000000000000000ALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: u Massimo Rotella euplains to me that they are uable to acqure the utrace via cable (also tried to replace). He tells me it's ugent. Order parts and plan technical. Diagnostic performed by Engineer: u Massimo Rotella euplains to me that they are uable to acqure the utrace via cable (also tried to replace). He tells me it's ugent. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineering Device ue: patient monitoring Eupected and ueupected Resuts: Trace cannot be acqured via ucable uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- | 2023-11-09 20:21 | No
3658 | Yes | Patient circuit occluon will not populate, GFE00OV0 Response received/IS complete | Patient 0cuit occluon will not populate, GFE0 0NOV00Response received/IS complete | 000000000000000000000000001O##: 00000000000000000000000000100000000390000000031 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Patient circut occluion will not popuate. uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 989800000000617100000000- flow sensor assembly - 1217.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states: Patient circut occluion will not popuate. troule shot with cutomer. device responses properly when circut is disconnected bu does not give circut occlued uLL Informed cutomer manuactuers recommendation is to replace flow sensor assembly. cutomer is requsting part nuber for replacement flow sensor ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Carlo Marasigan IP & Issu: 62222000000002 Patient circut incluion will not popuate. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-09 21:36 | No
3659 | Yes | AED fails self-test, "Remove and reinsert battery" message, single chirps - | AED fails self-test, "Remove and reinsert battery" message, single chirps - | 000000000000000000000000001O##: 00000000000000000000000000100000000397965 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to acct address, Attn Kelly Fogarty Resoluion : shipped serial # B23A-19253, FedEu tracking 78617100000000878 DELIVERED Friday 11/10000000023 at 9:00000000AM Signed for by: K.GREEM Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: not passing self test uer Impact: none Patient Impact: none SE E.00000000.000000007 See ref Case # 00000000225866200000000 Cut called stating that device not passing self tests. Diagnostic performed by Engineer : Cut said that they were getting an uLLmsg “remove and reinsert battery” followed by single chirps. Advised cut to remove device from service and will replace device uder warranty. 000000008900000000 Install before date: 2000000003-00000000-00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000398000000001 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-11-09 21:36 | No
3660 | Yes | M0A / u00 Functional test failed. - | M00A / u000/ Funuonal test failed. - | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000199 13-Nov-2000000003 FALASCA, IVO The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Technician Matteo Carrara who carried ou the check foud that condenser needs to be replaced. Order the spare part and schedue a technician. Since the spare part is blocked du to obsolescence, check if it can be ordered throuh alternative channels. Diagnostic performed by Engineer: Technician Matteo Carrara who carried ou the investigation foud a problem on condenser. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Laua Ghiraldello Device ue: Diagnosis Eupected and ueupected Resuts: problem on fuctional tests uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from Clinical E.G. Laua Ghiraldello, who reports problems on fuctional test. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000039800000000000000003-Nov-2000000003 Matteo Carraro The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: The system is not able to suply the requsted energy. Awaiting repair at the office of the head nuse. Diagnostic performed by Engineer: Performed discharge test with a negative oucome. Follow-u is schedued to order parts. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Laua Ghiraldello Device ue: Diagnosis Eupected and ueupected Resuts: problem on fuctional tests uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from Clinical E.G. Laua Ghiraldello, who reports problems on fuctional test. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000003959800000000ELONI, FABRIZIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: I heard from clinical engineering Laua Ghiraldello, who reported problems on fuctional test, from system it is uder contract, they requst the presence of the technician for any calibration and the neut fuctional tests, equpped with tools, to assign on site. Diagnostic performed by Engineer: I heard from clinical engineering Laua Ghiraldello, who reported problems on fuctional test, the system is uder contract, they requst the presence of a technician for any calibration and the following fuctional tests. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Laua Ghiraldello Device ue: Diagnosis Eupected and ueupected Resuts: problem on fuctional tests uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from Clinical E.G. Laua Ghiraldello, who reports problems on fuctional test. ----------------------------------------------------------------------------- | 2023-11-09 21:41 | No
3661 | Yes | bench repair, charger port broken | bench repair, charger port broken | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000957 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE uknown Diagnostic performed by Engineer : Confirmed complaint the uits Charging port is physically damaged. To correct the issu we will replace the Front Case Assembly. Resoluion : Device was repaired by u technician Chongolnee, Roongruchai with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE uknown Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair Diagnostic performed by Engineer : Confirmed complaint the uits Charging port is physically damaged. To correct the issu we will replace the Front Case Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000390000000055 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Tech states the charging port is broken and the uit will not charge; needs to send in for bench repair. Verified shipping address: Appomattou Couty Rescu Squd ATTN: Tom Walton 763 Confederate Blvd APPOMATTOu, VA 2000000002-000000001 u Resoluion : Verified cutomers contact and shipping information is correct and ##d the u bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001agnostic performed by Engineer : Tech states the charging port is broken and the uit will not charge; needs to send in for bench repair. Verified shipping address: Appomattou Couty Rescu Squd ATTN: Tom Walton 763 Confederate Blvd APPOMATTOu, VA 2000000002-000000001 u Problem Description by Engineer : Cutomer Fuction/Role: Tech states the charging port is broken and the uit will not charge; needs to send in for bench repair. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: tom walton IP & Issu: 90000000039800000000bench repair, charger port broken Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-09 22:16 | No
3662 | Yes | The navigation ring did not respond. - Pending repair 08 | The navigation ring did not respond. - Pending repair 0/8 |  | 2023-11-09 22:37 | No
3663 | Yes | system is stuck in reboot cycle and will not completely turn on | system is stuck in reboot cycle and will not completely turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007900000000100000000ov-2000000003 O'Connor Thomas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System is not starting u When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system down reboot does not resolve uer Impact: system down Patient Impact: no impact SE Not provided Resoluion : Boot issus ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 13-Nov-2000000003 Mackey, Daryl The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Install licenses, drivers, and uuudatabase. Problem Description by Engineer : Problem description by engineer : System is not starting u When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system down reboot does not resolve uer Impact: system down Patient Impact: no impact SE Not provided Diagnostic performed by Engineer : Replaced HDD and Host PC. FSE did not have access to service mode to install drivers, licenses, and uuudatabase. Other FSE will complete service activity tomorrow morning. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.00000000000000000000000000000000800000000200000000612,Nehalem Host PC Monoplane Rev_III NO_HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000039800000000 00000000-Nov-2000000003 Mackey, Daryl The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace Hard Drive Disk for Host PC and install the Operating System. If HDD replacement does not resolve issu, replace Host PC. Problem Description by Engineer : Problem description by engineer : System is not starting u When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system down reboot does not resolve uer Impact: system down Patient Impact: no impact SE Not provided Diagnostic performed by Engineer : Performed Host PC and HDD trouleshooting with assistance from an NSS. HDD foud to be fauty even after bypassing HDD slot connection. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000398189 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System is not starting u When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system down reboot does not resolve uer Impact: system down Patient Impact: no impact SE Not provided Diagnostic performed by Engineer : James reported that the system is not completing bootu and reboot does not resolve. Per biomed, the system is not able to access the operating system and it's trying to look for dhcp bu system is set for static ip. No PRS and system is stuk looking for client mac address. Internal Comments : uLL# and/or description of uLL: System not starting u. Repair Action : Check if the pc in the cabinet is starting u or not and check the LEDS. Replace if necessary. 00000000800000000200000000611 - Nehalem Host PC Monoplane Rev_III NO_HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: System not starting u. Repair Action : Check if the pc in the cabinet is starting u or not and check the LEDS. Replace if necessary. 00000000800000000200000000611 - Nehalem Host PC Monoplane Rev_III NO_HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: JAMES JERIAN IP & Issu: 6500000000998/system is stuk in reboot cycle and will not completely tun on Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Performed fingerprint swap for host pc replacement. Internal Remarks : Interaction Type: Cutomer Call Back Caller: James IP & Issu: CCB checnking the ETA Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Advised FSE will be onsite today 9:300000000 Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : FSE Daryl Mackey. FSE onsite and system not seeing boot device. Tried reseating drive, then guded FSE to bypass drive bay. Still not being seen. FSE got into BIOS to see boot priority and HD not an option. Appears HDD is bad. Also sugested to get Host PC to have jut in case, bu ue the HDD first. Approved p/n: 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD 0000000080000000050000000000000000- Hard disk spare part. Internal Remarks : uable to see host boot-u on display after hard drive replacement. Recommended FSE either locate spare monitor and connect directly to host pc to verify PC fuctionality or replace host pc with new host that curently is located onsite. ----------------------------------------------------------------------------- | 2023-11-09 22:51 | No
3664 | Yes | touchscreen misalignment | touchscreen misalignment | 000000000000000000000000001O##: 00000000000000000000000000100000000398398 Miller Jason The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Cutomer states touhscreen isn't calibrating and giving calibration failed when attempted, ,cutomer sees marks on touhscreen itself and states u y auel isn't lining u in diagnostic mode. cutomer is requsting part nuber for replacement touhscreen. Resoluion : Provided cutomer with part nuber: 00000000561539000000001 - touhscreen Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Parts ID uer Impact: system restricted Patient Impact: none reported SE uknown ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Scott Buzard IP & Issu: 67169325/Cutomer Requsts Suport - Parts ID Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001and roued cu for suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-09 22:56 | No
3665 | Yes | Low uuuitor value - Op check fail (split from 00 | Low uuuor value - Op check fail (split from 000) |  | 2023-11-09 23:00 | No
3666 | Yes | Touch screen is not working or calibrating - Sent 0 GFE 05 | Touch screen is not working or calibrating - Sent 0d GFE 0/5 | 000000000000000000000000001O##: 0000000000000000000000000010000000039800000000 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Parts ID touhscreen uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber: 00000000561539000000001 - touhscreen Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states touhscreen isn't 000000000000000000000000001king properly and giving bad touh uLLmessage. cutomer is requsting replacement part nuber for touhscreen. ----------------------------------------------------------------------------- Internal Remarks : Caller: Brue 26000000003000000001131// Brue.Boissoneaut00gilitihealth.com IP & Issu: 6000000007000000009 - Touh screen is not 000000000000000000000000001king or calibrating Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-11-09 23:21 | No
3667 | Yes | The touchscreen is intermittently unresponsive. - GFE 0/6 | The touchscreen is intermittently unresponsive. - GFE 00/6 | 000000000000000000000000001O##: 00000000000000000000000000100000000398171 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Parts ID touhscreen uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states touhscreen isn't 000000000000000000000000001king properly and giving bad touh uLLmessage. cutomer is requsting replacement part nuber for touhscreen. Resoluion : Provided cutomer with part nuber: 00000000561539000000001 - touhscreen Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: David Shoemaker - phone 937 638 71300000000 IP & Issu: 800000000278100000000- The touhscreen is intermittently uresponsive. Cut said the front bezel was replaced by FSE in 2000000002 becaue of the recall. Cut wants to know if touhscreen was inclued in the bezel replacement. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-09 23:21 | No
3668 | Yes | charging faulty | charging faulty | 000000000000000000000000001O##: 00000000000000000000000000100000000398700000000 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000056000000005000000001,RDT Front Enclosue kit (for 00000000-100000000u-R & Tools & Equpment : PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 Resoluion : Faut confirmed, fails to charge throuh DC inpu. Front enclosue replaced, tested OK. Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: The device is not getting charged and no indications at all while connecting with the charger. End uer replaced the battery, cleaned the charging port and checked bu issu remained. Foud duing PVT. uer Impact: n/a Patient Impact: n/a SE n/a Reguatory Qustions : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000395900000000 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u Front Enclosue 0000000056000000005000000001 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: The device is not getting charged and no indications at all while connecting with the charger. End uer replaced the battery, cleaned the charging port and checked bu issu remained. Foud duing PVT. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u Front Enclosue 0000000056000000005000000001 ----------------------------------------------------------------------------- Internal Remarks : Hi Mohammed , Please can youconfirm a sutable collection date and address for you device to be collected by uS CASE 0000000022600000000257 SN 61000000009 TEMPu PRO Collection & Retuns Address : London Ambuance Service uit 7 Consu Avenu Rainham RM13 8HY Internal Remarks : Tracking Nuber:1Z3R600000000669500000000529 COLLECT TODAY ----------------------------------------------------------------------------- | 2023-11-09 23:26 | No
3669 | Yes | Backup alarm failed uLL GFE00OV0*** | Backup alarm failed uLL GFE0 0NOV00** | 000000000000000000000000001O##: 00000000000000000000000000100000000398800000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to replace CPuto address complaint, cutomer acknowledged and will order part then reprogram CPuwhen installed, issu resolved- Material Nuber 00000000561535700000000- RP-PCBA,CPuGEN2, (SW 2.100000000, V600000000Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:Backu alarm failed alarm message uer Impact: not uable needs repair Patient Impact: none SE2.100000000Diagnostic performed by Engineer : Cutomer states duing uit set u the uit alarms Backu alarm failed message on screen, no patient involvement, verified uLL1100000000in event log, no other codes noted, uit has 2.100000000software, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: pandy perry ph: 757330000000000000000 IP & Issu: Device is making a constant alarm or noise; backu alarm failed uLLreceived; 10000000000000000000000007 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created qute Neut Step Action: Awaiting Qute Verbal PO Nuber: not provided Verbal PO Approver: not provided Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-10 00:11 | No
3670 | Yes | code 0 OVP Circuit Failure | code 00 OVP 0cuit Failure | 000000000000000000000000001O##: 0000000000000000000000000010000000039900000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order a MC pcba and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and it has a OVP circut failue code 1127. Recommended replacing the MC pcba Part #00000000561537531, MC pcba, list price $1198.00000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: OVP alarm, code 1127 uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- | 2023-11-10 00:21 | No
3671 | Yes | M0A - u0 - Blades Damaged | M00A - u00 - Blades Damaged | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Equpment problem: Damaged equpment blades and plates. Priority: Restricted search for: José Felipe Carvalho de Deu Sector: Clinical Engineering Phone: 551135000000002100000000 Phone: 551135000000002100000000 E-mail: Engclinicabp00p.org.br uit: Real and Benememer Portuuse Benefit Association Address: R MAREST CARDIM 769 SAO PAuO-SP #PII_AJHFHIJJJJJJJJ Safety Qustions: Case: 0000000022593366 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Nicole Carmo Ferreira Last udated by: Nicole Carmo Ferreira 1. The equpment was in clinical ue at the time the problem was detected ? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT T2 Activities:Problem Description: PROBLEM DESCRIPTION: Damaged equpment blades and plates WHEN WAS THE FIRST OCCuRENCE? THERE IS... Week HOW FREQuNT IS THE PROBLEM? Continuuly IN WHICH PART OF THE PROCESS DOES THE PROBLEM OCCu? Boot u WHAT RECENT EVENTS HAVE OCCuRED? Duing Normal ue DID THE CuTOMER CREATE A SYSTEM-SPECIFIC REPORT FOR THE PROBLEM? No ALTERNATIVE 00IF DIFFERENT FROM ONEEMS: No Trouleshooting Action: NOTE # uLLAND/OR PROBLEM DESCRIPTION: Equpment Blades and Plates Damaged WHAT STEPS FOLLOWED DuING DIAGNOSIS? In contact with uJosé, I was informed that the equpment's blades and plates were damaged. WHAT ARE THE NEuT TROuLESHOOTING STEPS FOR THE PROBLEM? It is necessary to change the blades and plates. Once the cutomer sends the qute for the parts, the cutomer will carry ou the replacement. Estimated downtime: N/A Repair Action: IF NECESSARY PART(S) - 12 NC RECOMMENDED(S): 000000005600000000771 - uuuKit Paddle Pocket Plates 0000000056300000000991 - Paddle Repl. Water Resistant Kit RECOMMENDED REPAIR ACTION(S) In line with the cutomer shipping as NEMO. uon receiving the parts, the cutomer will contact the remote to monitor the installation. RECOMMENDED ACTIONS FOR THE FSE BEFORE THE VISIT: N/A CSR: Raqul Freitas Test/Inspection: N/A Test Equpment: N/A SAP Transfer: Case Transferred to SAP Trouleshooting Action: Qute sent SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to sending a buget SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-10 00:26 | No
3672 | Yes | M0A / u0 / Worn out paddles (ref. split PR 00). | M00A / u00 / Worn out paddles (ref. split PR 000). |  | 2023-11-10 00:36 | No
3673 | Yes | Display failed, GFE00EC0/IS complete*** | Display failed, GFE0 0DEC00IS complete*** | 000000000000000000000000001O##: 0000000000000000000000000010000000039300000000 29-Nov-2000000003 Melgar, Alvaro The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Waiting on part: uu 00000000561528891 - RP-LCD, Assy,2nd GEN, V600000000to become available for neut cutomer's visit. Problem Description by Engineer : Cutomer Fuction/Role: Ed Magee Device ue: V600000000Eupected and ueupected Resuts: Backlight failed on display. uer Impact: None Patient Impact: None SE N/A Dark screen Tools & Equpment : TORQu SCREWDRIVER,2000000003-12-31 00000000:00000000:00000000 MuTIMETER,2000000003-12-31 00000000:00000000:00000000 STATIC CONTROL DEVICE,200000000000000005-31 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000056153600000000,RP-Cable, LCD, NEC, 2nd Gen, V600000000V68000000001.00000000000000000000000000000000561528911,RP-Cable, u PCBA to LCD, V600000000Diagnostic performed by Engineer : Cutomer complained that the backlight had failed on display. uit operates and touhpad screen 000000000000000000000000001ks bu it remains dark. FSE confirmed cutomer's concern while onsite. FSE replaced RP-cable, LCD, NEC, end Gen, V600000000V6800000000 and RP-cable, u PCBA to LCD, V600000000 Replacement of both cables did not fiu the issu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000538971 29-Nov-2000000003 Melgar, Alvaro Diagnostic performed by Engineer : N/A Problem Description by Engineer : Cutomer Fuction/Role: Ed Magee Device ue: V600000000Eupected and ueupected Resuts: Backlight failed on display. uer Impact: None Patient Impact: None SE N/A Dark screen ----------------------------------------------------------------------------- | 2023-11-10 00:36 | No
3674 | Yes | Intermittent uLLmessage 0. | Intermittent uLLmessage 00. | 000000000000000000000000001O##: 00000000000000000000000000100000000399000000001 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to replace blower to correct, cutomer acknowledged and will order part, issu resolved- Material Nuber 00000000561512761 - RP-Blower Assy,V600000000Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:Blower temp high alarm uer Impact: removed from service Patient Impact: therapy delayed SEun Diagnostic performed by Engineer : Cutomer states while uit was in ue it alarmed Blower temperatue high alarm on screen, uit continud to ventilate bu was removed from patient ue, no patient harm reported, in biomed complaint was not dulicated, cooling fan is fuctional and Blower temp in Pneuatics is within specs, noted uLLcode 1122 in event log, requsts trouleshooting to correct issu. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: " Robert Ward" IP & Issu: "80000000069155 Intermittent uLLmessage 1122." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-10 00:41 | No
3675 | Yes | Blower temperature high (EC 0) | Blower temperature high (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000399313 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: high blower temp uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber: 00000000561512761 - blower assembly Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states device is giving high blower temp uLLcode, 1122. Cutomer is requsting part nuber for replacement blower assembly. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Kim Bell Issu: 58000000002200000000uLLMessage : Blower to high Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-10 01:16 | No
3676 | Yes | low disk space after patients deleted | 00CFNuu low disk space after patients deleted | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000383 Sulivan Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : spoke to Zack ,inhose engineer , he has called back for additional suport on clearing low disk space message on system When was the first occurence? Within... : 100000000Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Cutomer is is performing replacement of disks and compuer and additional qustions uer Impact: uable to store images Patient Impact: no patient involed SE 8.2.3000000005 Diagnostic performed by Engineer : speaking to Zack he asked if there is additional steps after replacing ip pc and disks , he is following the recommeded du. Board software load took longer than cutomer eupected .software download completed with green ready statu . cutomer did cold restart . still message on system Rse sugested a image disk recreate be performed . this cleared message and cine rus were being stored to system .System is operational accord to cutomer Internal Comments : uLL# and/or description of uLL: disk space message on system Repair Action : gued cutomer thrusoftware load and restore process Resoluion : Will hold and check with cutomer then complete case system is u and futional at this timecall back spoke to zack no futher issus ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000039900000000 Van Pelt Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : RMT - IPu Free disk space is too low. 00000000-Nov-2000000003 00000000:00000000:200000000MT - IPu IPPC failed to boot u - Frontal 00000000-Nov-2000000003 00000000:51:100000000 How often is issu occuring? : Intermittent Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: Eupected to be able to take more images however system disk space is low, even after deleting patients from the system uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Reviewed log files and discused issu with biomed. Reoccuring issu from case 000000002200000000900000000Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Free disk space is too low.: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Root Caue for RMT - IPu IPPC failed to boot u - Frontal: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power, the BIOS battery is defective or the OS/AS did not boot u in-time. Repair Action : Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu IPPC failed to boot u - Frontal: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu on InCenter in the produt manul of the PC. 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IPPC and perform Cold reboot of the system. 1.3. If LED statu is NOT Normal then issu is Boot problem in IPPC. Euamine IPPC power and perform cold reboot of the u system. Replace IPPC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct. 1.5. A bad/weak BIOS battery can caue the PC not to start the first time. Replace IPPC if problem persists/repeats. Resoluion : Provided part nubers for replacement PC and hard disks. 0000000080000000050000000000000000- Hard disk spare part 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD. Biomed had M-cabinet repair manul open and has replacement procedue available to gude him on replacement of PC and/or disks. Recommend replacing PC and disks du to intermittent failue of IPPC to boot u as well as the disk uLL/free space coud be caued du to bad or defective image disk. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000399678 Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- Internal Remarks : Assisted cutomer with getting transferred to parts and placing parts order. Parts ordered for neut day early AM. ----------------------------------------------------------------------------- | 2023-11-10 01:31 | No
3677 | Yes | When the power button was pressed at the time of connecting the AC power, the device was not activated. | When the power button was pressed at the time of conneung the AC power, the device was not auted. |  | 2023-11-10 02:27 | No
3678 | Yes | FCO 0failure upon power on - Pending onsite repair 00 | FCO 0 failure upon power on - Pending onsite repair 0/0 | 000000000000000000000000001O##: 00000000000000000000000000100000000525386 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: the air and o2 flow sensors are not reading in the vent info screen uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : ASP has a V600000000and they are not seeing the air and o2 information on the vent info screen. She did not test the uit, bu it seems to be runing fine. The air and o2 information is not displayed on the vent info screen. She stated the cable is good and properly connected to the DA pcba. Recommended to replace the GDS. Part #00000000561512191 GDS. ASP was on site on 12-5-23 and replaced the GDS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000053000000008 Harper, Stephanie Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000312 Harper, Stephanie Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-11-10 03:01 | No
3679 | Yes | Device will not power on | Device will not power on | 000000000000000000000000001O##: 00000000000000000000000000100000000398895 16-Nov-2000000003 Walker, Craig The system meets the specification for the performed service and is retuned to ue Tools & Equpment : TORQu SCREWDRIVER,200000000000000000000000031 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000000000000031 00000000:00000000:00000000 MuTIMETER,200000000000000000000000031 00000000:00000000:00000000 Diagnostic performed by Engineer : * Resoluion: What was done to complete the repair?See FCO 0000000000000000000000000010000000017800000000. Power Mgmt Brd resolved the problem. What tests were performed to ensue the repair is complete? See T&V attached to this 000000000000000000000000001from FCO. What were the resuts of the tests performed?See T&V attached to this 000000000000000000000000001from FCO. And did they pass? YES. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: cannot power on uit uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000398387 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: cannot power on uit uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and he is having issus with tuning the uit on. We were able to power on the uit, bu it take a few minues. We verified there was 116 AC volts going into the power suply, and 200000000C volts going from the power suply to the PM pcba. Cutomer stated it was an intermittent issu for several months. Recommended the PM pcba FCO66 and uuuFCO57 to be remediated. Created an AA for both FCOs to be implemented. This uit was repaired on the FCO66 case #0000000020000000022127, 0000000000000000000000000010000000017800000000 by ASP on 11-15-23 Follow u Requred : Cutomer has a V600000000and he is having issus with tuning the uit on. We were able to power on the uit, bu it take a few minues. We verified there was 116 AC volts going into the power suply, and 200000000C volts going from the power suply to the PM pcba. Cutomer stated it was an intermittent issu for several months. Recommended the PM pcba FCO66 and uuuFCO57 to be remediated. Created an AA for both FCOs to be implemented. ----------------------------------------------------------------------------- Internal Remarks : Contact Name: emsar Method: ##nsa What was discused: cutomer called in abou this case. they were requsting to have this fco implemented on the device. I went ahead and udated the contact information on both case and 000000000000000000000000001O#level. i udated the nsa as well. i will reach ou to emsar to schedue the down uit/fco. incident case- 000000002260000000000000000/fco case 0000000020000000022127 Neut Step Action: n/a Internal Remarks : Interaction Type: New Case Caller: Michael McAnnally IP & Issu: 70000000090000000000000000Device will not power on Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-10 03:11 | No
3680 | Yes | uLLlog and fluoro failure warning | 0hh uLLlog and fluoro failure warning | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000822 22-Nov-2000000003 Stone, Matthew The system meets the specification for the performed service and is retuned to ue Resoluion : Installed new converter calibrated did t000000000000000000000000001ademption and EDL uit passed withou any uLL ready for clinical ue Problem Description by Engineer : Problem description by engineer : 00000000hw & 00000000hw uLLlog and fluro failue warning When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times per week Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Reboot of system does not resolve issu uer Impact: uable to continu with clinical applications Patient Impact: No futher information is available as per the cutomer SE 7.2.9.00000000Diagnostic performed by Engineer : 00000000HH uLLand generator logs pointed to Q2 converter, puling down the other converter Parts Consued : 1.00000000000000000000000000000000210000000016291,Converter 8E Velara Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000003-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000003-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000100000000ov-2000000003 Brown, Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: uable to perform service du to the cutomer uing the room all day then uable to access the service area du to it being badge reader SE Diagnostic performed by Engineer : did not gain access to the room for the reported repair cutomer ued the room all day ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 100000000ov-2000000003 Stone, Matthew The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : 00000000hw & 00000000hw uLLlog and fluro failue warning When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times per week Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Reboot of system does not resolve issu uer Impact: uable to continu with clinical applications Patient Impact: No futher information is available as per the cutomer SE 7.2.9.00000000Follow u Requred : Trouleshoot air logs, install parts on hand for the in-houe reschedue if needed Diagnostic performed by Engineer : Arrived on scene on site at zero 8:300000000AM was told uit was in clinical operations we 000000000000000000000000001ld not have access to the room util 1500000000 at 100000000 inqured abou the 1500000000 timeframe was told it 000000000000000000000000001ld be after 17300000000before room 000000000000000000000000001ld be releasedadvised in houe management to reschedue for repairs ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000339 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : 00000000hw & 00000000hw uLLlog and fluro failue warning When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Mutiple times per week Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Reboot of system does not resolve issu uer Impact: uable to continu with clinical applications Patient Impact: No futher information is available as per the cutomer SE 7.2.9.00000000Diagnostic performed by Engineer : Kevin says generator uLL have been occuing on 11/7/23 preventing fluroscopy. RSW and LOD are showing uLL in reported problem. Biomed is requsting onstite service tomorrow or weekend to resolve. Duing debrief and diagnostics, Kevin Anderson called stating that he has already spoken to the cutomer and they have no agreeded to have onsite dispatched on Monday, 11/13/23 Internal Comments : uLL# and/or description of uLL: 00000000hu, 00000000hw Repair Action : Guded by DSM for Monday dispatch Follow u Requred : uLL# and/or description of uLL: 00000000hu, 00000000hw Repair Action : Guded by DSM for Monday dispatch What are the neut Trouleshooting steps : Pattern Name: RMT - uGEN: uLL(00000000Hu) Asymmetrical kV detected - Frontal Root Caue 1: Fluroscopy failed du to 00000000Hu uLL(kV asymmetrical) Service Action 1: Check the detailed uLLlog and check which side Anode or Cathode is too low or too high. Service Action 2: Swap the HV cable Anode/Cathode, if the problem moves try new HV cable. Service Action 3: Perform tue adaptation and conditioning. Service Action 00000000If the deviation is only a little bit above the allowed limit, try new HV tank. Service Action 5: Swap Anode with Cathode converter, if the problem moves to other side try new converter. Service Action 6: If the problem does NOT move after swapping the converter, try new kVmA uit. Service Action 7: If all this doesn't solve the problem try new u-ray tue. Technician Name : Kuis Peter ----------------------------------------------------------------------------- | 2023-11-10 04:21 | No
3681 | Yes | Transfer from QN – HS0EFOA "The voice is distorted when the self testing." | Transfer from QN –u"The voice is distorted when the self testing." |  | 2023-11-10 04:36 | No
3682 | Yes | wireless footswitch not working | wireless footswitch not working | 000000000000000000000000001O##: 00000000000000000000000000100000000393116 Stout Chad Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Internal Remarks : FSE Chad Stout. FSE doing wireless footswitch fco and charged uit . Will not even tun on. Approved for replacement. ----------------------------------------------------------------------------- | 2023-11-10 04:36 | No
3683 | Yes | The AED beeped once - Log analysis requested | The AED beeped once - Log analysis requested | 000000000000000000000000001O##: 00000000000000000000000000100000000393587 Suui, Kosue The system meets the specification for the performed service and is retuned to ue Resoluion: When the main body log was analyzed at the manuactuing factory in the uited States, an alarm was detected du to the pad impedance euceeding the allowable valu. No record of uit failue was confirmed. Based on the above, the pad is defective, so we will replace the pad and retun the produt withou any problems. [Simutaneou FCO 000000000000000000000000001k] There are no FCOs that requre 000000000000000000000000001k on this equpment. December 11, 2000000003 Fuua Denshi KitatohokuHanbai We will send the 000000000000000000000000001k report along with the equpment to the person in charge. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit the 000000000000000000000000001k report on paper? <YES>[Final confirmation] Perform operation check after 000000000000000000000000001k based on prescribed instrutions YES Tools & Equpment : MuTIMETER, 200000000000000003-15 00000000:00000000:00000000 Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Fuua Electronic Sales Representative Device ue: Not ued Eupected and ueupected Resuts: 1 beep uer Impact: None Patient Impact: None SE None Diagnostic performed by Engineer: When received in standby state, the reproduibility of abnormal alarms coud not be confirmed. We have confirmed that youwill pass the euam at uT. We have confirmed that the adhesive strength of the heat seal on the pad has decreased and it is beginning to peel off. ----------------------------------------------------------------------------- Internal Remarks: Failue Details 2 The alarm has not stopped since 11/2. I jut changed the battery last year and haven't ued the pads yet. I also fiued the hemispherical pad point, bu it didn't change. I 000000000000000000000000001ld like to requst an inspection. LOT Adut Pad 220000000000000000-0000000000000000attery M00000000000000008-2222 ----------------------------------------------------------------------------- | 2023-11-10 06:16 | No
3684 | Yes | QAA0/u FD0cannot euposure | QAA0R/u FD0cannot euposure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.11.2000000003 00000000:00000000:00000000uC (86000000008287) "SuT. Ready for Clinical ue local time" : " 2000000003-11-00000000 13:300000000 SAP Transfer: Case Transferred to SAP Euternal Remarks: 1000000001.2000000003 00000000:0000000036 uC (86000000008287) Service content: The image cannot be played back, uplu the image compuer transmission board. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. Safety Qustions: Case: 0000000022595000000001 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: QAA00000000R/u uuuFD100000000SAP Transfer: Case Transferred to SAP Internal Remarks: 1000000001.2000000003 00000000:0000000038 uC (86000000008287) "RDI Accuacy: Y" $Q7: (Confirmed reported problem, the images cannot be displayed on review monitor.) $Q8: (Check the system onsite, finding the image transmission card LEDs of PCs are not bright. It caued the image transmission problem.) $Q9: (Reconnect the image transmission card on Host PC.) $Q100000000 (Test and Inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools (N/A) SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine coud not be euposed and images coud not be viewed, and engineers needed to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: The machine cannot be euposed and cannot browse images. Remote diagnostics and test resuts: NA PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC starts normally, check whether the statu of the POST LED behind the IPPC is normal, and ue the PC Diagnostic Tool to perform a diagnostic test. 2. Euecue Recreate Image Store. If it fails to recreate, try to bypass the hard disk bou or format the hard disk. 3. Check the IPPC and Is the commuication net000000000000000000000000001k cable between the Host PC normal? 00000000Check the machine uLLand perform faut repairs. Recommended spare part uuand name: " 0000000080000000050000000000000000Hard disk spare part &0000000080000000050000000092 IP PC Revised_II no HDD Cutomer demand notes: Contact cutomers promptly to confirm Onsite time ----------------------------------------------------------------------------- | 2023-11-10 06:31 | No
3685 | Yes | Automatic power on | Automatic power on | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Auomatic power on Neut Step Action: The provincial governor uuJianhong has been contacted to give ASP Honglejia buiness follow-u Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 000000002259500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: uiao Chunbao 13976000000006738 ----------------------------------------------------------------------------- | 2023-11-10 06:46 | No
3686 | Yes | The device powered on automatically, GFE00OV0 Response received/IS complete | The device powered on automatically, GFE0 0NOV00Response received/IS complete | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Auomatic power on Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: uiao Chunbao 13976000000006738 Neut Step Action: The provincial governor uuJianhong has been contacted to give ASP Honglejia buiness follow-u Safety Qustions: Case: 000000002259500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-10 06:46 | No
3687 | Yes | Defib is beeping and has uLL: Therapy not available | Defib is beeping and has uLL: Therapy not available | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000992 Bronte Westaway The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : front case replacement and T&V Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Device ue: Ou of ue Eupected and ueupected Resuts: Therapy not available uLLuer Impact: No impact on uer Patient Impact: no impact on patient SE Not provided Diagnostic performed by Engineer : Confirmed therapy disable on arrival. Ru op-check which remedied issu. Checked logs for additional fauts - no critical requring intervention. Front case dislodged from rear houing - damage foud on front casing internally - front case to be replaced. all other fuctions 000000000000000000000000001king as eupected ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000039000000001 Alhad Apte The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Assessment and repair Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Device ue: Ou of ue Eupected and ueupected Resuts: Therapy not available uLLuer Impact: No impact on uer Patient Impact: no impact on patient SE Not provided Diagnostic performed by Engineer : Therapy not available uLL Bench assessment and repair is requred ----------------------------------------------------------------------------- | 2023-11-10 06:46 | No
3688 | Yes | Defibrillation is not possible | Defibrillation is not possible | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor/ Safety Qustions: Case: 0000000022595200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-10 07:41 | No
3689 | Yes | SPLIT - Unit does not get past the self-test (Ref 00 | SPLIT - Unit does not get past the self-test (Ref 00 |  | 2023-11-10 08:00 | No
3690 | Yes | NUH ¦ u FD0¦ ISSUE ¦ Main bldg 0not powering up | NUH ¦ u FD0¦ ISSUE ¦ Main bldg 0 not powering up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000977 100000000Nov-2000000003 KOK WAI CHAN The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Radiograher Device ue: no Eupected and ueupected Resuts: system cannot power u uer Impact: no Patient Impact: no SE v8.2.17 Resoluion : Foud host pc was not power u . Reseat all connector and perform a proper shudown procedue. Reposed u system and system power according. Test system ok ----------------------------------------------------------------------------- | 2023-11-10 08:16 | No
3691 | Yes | 0 / u0 / AC Power indication not showing (Informed cuomer device is EoL). | 00/ u000/ AC Power indication not showing (Informed cuomer device is EoL). | 000000000000000000000000001O##: 00000000000000000000000000100000000390000000000000000nanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Resoluion : Checked and foud Bezel LED has problem, Informed cutomer abou EOL abou this produt. Diagnostic performed by Engineer : AC Power indication not showing ----------------------------------------------------------------------------- | 2023-11-10 08:31 | No
3692 | Yes | Getting Initialization error | Getting Initialization error | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000783 100000000Nov-2000000003 VEMBARASAN K The system meets the specification for the performed service and is retuned to ue Resoluion : System software corruted, replaced new hard disk and reloaded sw system made fuctional Problem Description by Engineer : Problem description by engineer : As per uer the blu dup is coming and system software not booting Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000397971 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uLL# and/or description of uLL: As per uer the blu dup is coming and system software not booting Repair Action : FSE to visit and reset the HDD and reload the ghost and check with PC diagnostic tool What are the neut Trouleshooting steps : FSE to visit and reset the HDD and reload the ghost and check with PC diagnostic tool Technician Name : Kuar, Atu Internal Comments : uLL# and/or description of uLL: As per uer the blu dup is coming and system software not booting Repair Action : FSE to visit and reset the HDD and reload the ghost and check with PC diagnostic tool Diagnostic performed by Engineer : As per uer the blu dup is coming and system software not booting Problem Description by Engineer : Problem description by engineer : As per uer the blu dup is coming and system software not booting Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000----------------------------------------------------------------------------- | 2023-11-10 09:01 | No
3693 | Yes | Audio&Sounds/No-Sound: Loudspeaker defective: Speaker/HW: EOL/Provided Information | Audio&Sounds/No-Sound: Loudspeaker defeu: Speaker/HW: EOL/Provided Information |  | 2023-11-10 09:10 | No
3694 | Yes | SPLIT- broken CO0eceptacle - REFERENCE PR # 00 | SPLIT- broken CO0receptacle - REFERENCE PR # 000 |  | 2023-11-10 09:31 | No
3695 | Yes | QA00000 - Device cannot enter the operation interface | QA00000000- Device cannot enter the operation interface |  | 2023-11-10 10:50 | No
3696 | Yes | QA00000 - uwaveform incorrect | QA0000000 - uwaveform incorrect |  | 2023-11-10 11:01 | No
3697 | Yes | SA:0/u FD0Euposure not possible | SA:0K/u FD0Euposure not possible | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 10000000011.2000000003 00000000:00000000:22 uC (8600000000800000000) "SuT. Ready for Clinical ue local time" : " 2000000003-11-00000000 20000000000000000 Internal Remarks: 10000000011.2000000003 00000000:3000000006 uC (86000000009567) Only on job training Euternal Remarks: 10000000011.2000000003 00000000:30000000000000000TC (86000000009567) only on job training T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine was euposed and reported uLL, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Euposue not possible Remote diagnosis and test resuts: Restart has no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC is normal. 2. Check whether the MPD control board is normal. 3. Check whether the high-voltage commuication is normal. 00000000Check whether the FDC commuication is normal. 5. Check the machine uLLand perform faut repairs. Recommended spare part uuand name: " 0000000080000000050000000000000000Hard disk spare part &0000000080000000050000000092 IP PC Revised_II no HDD Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Internal Remarks: 13.11.2000000003 11:31:55 uC (8600000000800000000)$Q1: (Y) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem. images disk uLL can't write disk.) $Q8: (Performed system test, images disk uLL can't write disk) $Q9: (replace IPPC and install systerm) $Q100000000 (Sytem TEST OK. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (N/A) ^Curent Tue installed date: N/A ^Curent Tue uage: N/A Euternal Remarks: 13.11.2000000003 11:31:52 uC (8600000000800000000) Service content: The device cannot be euposed, go to the site to replace the IPPC and reinstall the system. Service resut: The equpment is ued normally. SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: FSE uu Yu 1330000000010000000011 Call for repair Retun visit failue ##mm SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002259600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-10 11:11 | No
3698 | Yes | Device may be prematurely draining batteries - | Device may be prematurely draining batteries - | 000000000000000000000000001O##: 00000000000000000000000000100000000396193 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: IREEDEM - DEALER AuHORIZED Device ue: Test on defibrillator Eupected and ueupected Resuts: the device discharges the batteries after a few days uer Impact: no impact On Patient Impact: no impact SE no problem verified after normal maintenance check, no report of patient impact. Diagnostic performed by Engineer: The device quckly discharges the batteries. Follow u Requred: nemo withou collection IREDEEM Spa Piazza dei Martiri 19000000001900000000n.1, 0000000021 Bologna - T +39 000000001 0000000035879 tel. 300000000855700000000c/a. Domenico Giordano ----------------------------------------------------------------------------- | 2023-11-10 13:26 | No
3699 | Yes | System does not go up | System does not go up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000732 November 13, 2000000003 Malhas, Ali The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Problem Description by Engineer: Problem description by engineer: System does not go u. Hangs at 00000000:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Resoluion: The reported faut is confirmed. Trouleshooting: No. Soluion: FleuVision PC replaced, OS and APP installed and the konfig. replayed. System restart and fuction check sucessfuly completed. FleuVision layous checked with the uer - > i. O. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Baugarten, Rüdiger The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: System does not go u. Hangs at 00000000:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: Cutomer is to be System does not go u. Hangs at 00000000:00000000 Internal Comments: uLL# and/or description of uLL System does not go u. Hangs at 00000000:00000000 FleuVision compuer does not start Repair Action: Te with part Follow-u requred: uLL# and/or description of uLL System does not go u. Hangs at 00000000:00000000 FleuVision compuer does not start Repair Action: Te with part Technician name: Baugarten, Rüdiger ----------------------------------------------------------------------------- Cutomer comment: Ms. Döring +0000000000000000272126800000000aren.doering00ankt-gertrauen.de asks abou a repair appointment and asks for info Internal Remarks: uu 0000000080000000012200000000 - u Haswell Fleuvision 2 PC no HDD Cutomer ETA: November 13, 2000000003 Cutomer comment: ### uer has info abou service on 13.11. aroud 11:00000000/12:00000000 hrs. ----------------------------------------------------------------------------- | 2023-11-10 13:42 | No
3700 | Yes | QAASA：0/u FD0OR Table/System cannot startup | QAASA：0R/u FD0OR Table/System cannot startup | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problem: The device couts down to 00000000:00000000 and cannot enter the system normally. I restarted it three times, bu it still doesn't 000000000000000000000000001k.Engineers are requred to be present. Is the faut caued by the equpment itself? Yes Trouleshooting Action: Remote diagnosis and test resuts: The device couts down to 00000000:00000000 and cannot enter the system normally. I restarted it three times, bu it still doesn't 000000000000000000000000001k.Requres engineer presence PRS TR Statu: No PRS Repair Action: Recommended onsite repair steps: Check whether IPPC or IPC is not tuned on, check whether the PC's power suply or PC's BIOS battery is discharged. Cutomer demand notes: Onsite SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 10000000011.2000000003 00000000:10000000033 uC (86000000009500000000) $Q7: (Confirmation & Description of problem/the system cannot boot u) $Q8: (cutomer report that the system cannot boot u, check and foud the TS problem ) $Q9: (reconnect the cable, reconnect the mainboard, then restart system OK) $Q100000000 (Test system OK and inspection data refer to: PA. Device retuned to ful fuctionality. ) $Q11: Calibrated Tools:N/A Euternal Remarks: 10000000011.2000000003 00000000:10000000027 uC (86000000009500000000) "SuT. Ready for Clinical ue local time" : " 2000000003-11-00000000 23:00000000" Service content (requred): The host cannot start. Check the TS net000000000000000000000000001k cable, plu and uplu the host board, and the system will retun to normal. Service resuts (requred): [After repair, the equpment retuned to normal operation.] Other important matters: [none] Safety Qustions: Case: 0000000022598000000005 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Cannot boot normally SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-10 13:42 | No
3701 | Yes | GEH ¦ DFM0 ¦ uleads on patient it shows "learning u unable to u at the critical time | GEH ¦ DFM0 ¦ uleads on patient it shows "learning u unable to u at the critical time | 000000000000000000000000001O##: 00000000000000000000000000100000000395700000000100000000ov-2000000003 K000000000000000000000000001 SING MAK The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Nuse Device ue: DFM100000000 Eupected and ueupected Resuts: uable to detect ureadings uer Impact: No Patient Impact: No SE Rev 2.00000000Tools & Equpment : PATIENT SIMuATOR,200000000000000005-25 00000000:00000000:00000000 Resoluion : Cutomer Fuction/Role: Nuse Device ue: DFM100000000 Eupected and ueupected Resuts: usim test - Passed, Eutract log - No uLLfoud. uer Impact: No Patient Impact: No SE Rev 2.00000000----------------------------------------------------------------------------- | 2023-11-10 13:42 | No
3702 | Yes | QAA0u FD0movement failure | QAA0/u FD0movement failure | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: The cutomer called and needed to replace the motor uu 00000000800000000162371 Motor linear drive. Please check whether the contract can provide spare parts and dispatch 000000000000000000000000001k. Problem Reported by Cutomer: QAA00000000L/u uuuFD200000000 Safety Qustions: Case: 0000000022598671 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-10 14:21 | No
3703 | Yes | System not booting up. | System not booting up. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000857 22-Nov-2000000003 Oshiro Kenneth The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed FSE Device ue: Cathlab Eupected and ueupected Resuts: GeoPC doesn't tun on uer Impact: System Down Patient Impact: System Down SE R000Resoluion : Replacement of GEO PC and reload of software Parts Consued : 1.00000000000000000000000000000000222000000007883,Bootable SOFTWARE PKT. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 100000000ov-2000000003 Oshiro Kenneth The system meets the specification for the performed service and is retuned to ue Follow u Requred : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomed FSE Device ue: Cathlab Eupected and ueupected Resuts: No geometry movements available uer Impact: System was down Patient Impact: System was down SE R8.1.17 Resoluion : Replacement of Geometry PC, reloaded software and performed fuctional testing. Parts Consued : 1.00000000000000000000000000000000800000000000000007813,GEO Q35 IPC COA Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000398985 100000000Nov-2000000003 Oshiro Kenneth The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace Geometry PC and SCC Power uit if requred Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 Follow u Requred : IWS R1.5.1 Monitor is Blu Resoluion : Foud GEO PC is not powered u. Possible short circut internal to the PC. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Engineer Device ue: Cardiac Cathlab Eupected and ueupected Resuts: System will not bootu uer Impact: System is down Patient Impact: System is down SE R8.1.17 ----------------------------------------------------------------------------- File uloaded by CS Portal uer : File Name: IMG_00000000300000000jpg. Comments: File(s) uloaded duing case creation Internal Remarks : CCB requsting statu change to P2. Advised RP in chat. File uloaded by CS Portal uer : File Name: IMG_000000003000000001.jpg. Comments: File(s) uloaded duing case creation Internal Remarks : Ken Oshiro - qustion abou creating uB stick for Geometry IPC. Internal Remarks : Planner Initials: GM Action Taken: FSE Ken Oshiro accepted call for 11/900AM Oucome::dispatched Internal Remarks : FSE Ken Oshiro. FSE fouf blown Fue on u200000000, F12. Now has power to geo pc bu will not power on. Auhorized geo PC. Futher checking, SCC does not power on at all and verified power into it. Auhorized 00000000230000000032500000000- POWER(&CONTROL) uIT WITHOu BOARDS as well in case its needed. Internal Remarks : Planner Initials: GM Action Taken: FSE Ken Oshiro accepted call for 11/900AM Oucome::dispatched ----------------------------------------------------------------------------- | 2023-11-10 14:31 | No
3704 | Yes | uCV 0troubleshooting. | uCV 0troubleshooting. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000923 Ahmed Kotb The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Dear Team, Good evening. We have an issu on an u CV 200000000sn 177. The cutomer euuuenced geometry partly available and no motorized movement uLL Once I was at the clinic, I coud see that the stand and table were not switched on after system boot How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE NA Resoluion : fiued Internal Comments : uLL# and/or description of uLL: Dear Team, Good evening. We have an issu on an u CV 200000000sn 177. The cutomer euuuenced geometry partly available and no motorized movement uLL Once I was at the clinic, I coud see that the stand and table were not switched on after system boot Repair Action : Dear Team, Good evening. We have an issu on an u CV 200000000sn 177. The cutomer euuuenced geometry partly available and no motorized movement uLL Once I was at the clinic, I coud see that the stand and table were not switched on after system boot Diagnostic performed by Engineer : Dear Team, Good evening. We have an issu on an u CV 200000000sn 177. The cutomer euuuenced geometry partly available and no motorized movement uLL Once I was at the clinic, I coud see that the stand and table were not switched on after system boot ----------------------------------------------------------------------------- | 2023-11-10 15:16 | No
3705 | Yes | Battery eupired | Battery eupired | 000000000000000000000000001O##: 00000000000000000000000000100000000397872 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Catania Hospital At ATI GE, u; conmed. C/O Garibaldi di Nesima. Via Palermo, 636 95121, Catania, floor - 3 beside radiotherapy ALFONSO BuGARELLO 300000000650000000000000000Problem Description by Engineer: Cutomer Fuction/Role: biome dept Device ue: no Eupected and ueupected Resuts: Battery of 2000000006 uer Impact: no Patient Impact: no SE na The battery is from 2000000006. Diagnostic performed by Engineer: Heard from technician Bugarello: battery to be replaced, year of prodution 2000000006. Shipment follows. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Parts Consued: 1.0000000000000000000000000000000098000000000000000000000000DFM100000000 Lithiu Ion Battery ----------------------------------------------------------------------------- Internal Remarks: part code: 989800000000190000000071 Internal Remarks: uuic Po Garibaldi Center. IBA has been aware of the requst. Information was asked from the cutomer. ----------------------------------------------------------------------------- | 2023-11-10 15:41 | No
3706 | Yes | System will not boot up | System will not boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000960000000013-Nov-2000000003 Gary Kane The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System will not boot u uer Impact: NA Patient Impact: NA SE8.2 Resoluion : Replaced fleuvision PC. Issu resolved and fuctional check completed. Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000079 100000000Nov-2000000003 Gary Kane The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace fleuvision pc Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System will not boot u uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : System not powering u and screen stuk at 00000000time cout down ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000029 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to trouleshoot futher . Supect FV PC however uable to speak to the cutomer to verify where the system is failing in the boot stage Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System will not boot u uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : No RSW findings. RSE has made mutiple attempts to contact the cutomer withou sucess ----------------------------------------------------------------------------- Internal Remarks : Called nuber 3 times , no answer. 2nd nuber not in service ----------------------------------------------------------------------------- | 2023-11-10 16:06 | No
3707 | Yes | BENCH: Audio& Sounds: u0larm for a speaker uLLinop. No sound: Speaker/HW/ Replaced 000 | BENCH: Audio& Sounds: u0alarm for a speaker uLLinop. No sound: Speaker/HW/ Replaced 00000 | 000000000000000000000000001O##: 000000000000000000000000001000000000000000096 Fahager, Andreas The system meets the specification for the performed service and is retuned to ue Resoluion: Replaced speaker assembly and did PM sucessfuly. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: Broken speaker uer Impact: low Patient Impact: mediu SE na Diagnostic performed by Engineer: Problem acknowledged. Parts Consued: 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000003-12-12 00000000:00000000:00000000 PATIENT SIMuATOR, 2000000003-12-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000022 BENGTSSON NICKLAS The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Biomed reports that they get inop for broken speaker. Biomed confirms that there is no soud from the device. Recommended part: 0000000056000000003831 Mu_100000000/u3 Speaker assembly or 0000000056000000000000000081 Mu_100000000/u3 Main board Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: Broken speaker uer Impact: low Patient Impact: mediu SE na Diagnostic performed by Engineer: Biomed reports that they get inop for broken speaker. Biomed confirms that there is no soud from the device. Resoluion: Biomed 000000000000000000000000001ld like to send device for faut finding/repair at u bench. ----------------------------------------------------------------------------- | 2023-11-10 16:11 | No
3708 | Yes | Charging button reacted with a delay, visible defect on the connection of the pad cable. | Charging button reaud with a delay, visible defeuon the conneuon of the pad cable. |  | 2023-11-10 16:16 | No
3709 | Yes | Charging port intermittently not charging Split from PR00 | Charging port intermittently not charging Split from PR000) |  | 2023-11-10 16:30 | No
3710 | Yes | Defib is alarming Asystole despite patient having and displaying an urythm | Defib is alarming Asystole despite patient having and displaying an urythm |  | 2023-11-10 16:39 | No
3711 | Yes | Request for Therapy PCA High Voltage Cable | Request for Therapy PCA High Voltage Cable | 000000000000000000000000001O##: 000000000000000000000000001000000000000000032 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Domenico Giordano Device ue: Diagnosis Eupected and ueupected Resuts: problem with cable uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I have heard from the Clinical E.P. Domenico Giordano. He reported problems with the voltage cable. Resoluion: heard from clinical en Domenico Giordano, he reports problems on voltage cable, tests were done, foud part nuber uefu to him. I close the call remotely. ----------------------------------------------------------------------------- Internal Remarks: Good morning, For Intrepid s/n CN73900000000300000000I 000000000000000000000000001ld need the spare part in photo. It is a Therapy PCA High Voltage Cable, part nuber 0000000056000000001271. I cannot find it for sale in the u shop site. Cordially Domenico Giordano Person in charge of Technical Assistance Domenico.Giordano00redeem.it IREDEEM S.p.A. Piazza dei Martiri 19000000001900000000n.1, 0000000021 Bologna - T +39 000000001 0000000035879 F +39 000000001 0000000035882 ----------------------------------------------------------------------------- | 2023-11-10 17:06 | No
3712 | Yes | L arm shell falls off_SPLIT | L arm shell falls off_SPLIT | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 12.11.2000000003 100000000:00000000 uC (86000000008978) "RDI Accuacy: YES" $Q7: (Confirmed reported problem was that the system cover comes off ) $Q8: (Checked and foud the tape of the cover had problem. ) $Q9: (Re-installed the cover and fiued the other 3 machines. Did some tests, system was OK. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality ) $Q11: Calibrated Tools name ( N/A ) S/N ( N/A ) Neut Calibration date ( N/A ) SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: Cover faut Detailed description of cutomer problem or faut phenomenon: L arm shell falls off Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: n/a Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Install the fiued shell, check and eliminate equpment safety hazards Recommended spare parts uuand designation: uu00000000230000000000000000811 cover, left side L-arm Notes on cutomer requrements: Please contact the cutomer for on-site repair Safety Qustions: Case: 000000002260000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 12.11.2000000003 100000000:00000000uC (86000000008978) Faut phenomenon: The machine shell falls off Service content: Check the casings of 00000000achines and reinstall and fiu them Service resuts: After the repair is completed, the machine can be ued normally. Problem Reported by Cutomer: QAASA:OCAP:00000000R/u uuuFD200000000Neut Step Action: Gu FEI SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-10 17:26 | No
3713 | Yes | Duplicate of PR# 07 | Duplicate of PR# 007 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000267 Assad MAJEED The system does not meet the specification for the performed service and is not retuned to ue Resoluion: Requst for 000000000000000000000000001kshop retun for complete inspection of the equpment. Follow u Requred: The cutomer reports that when the ward nuse tried to disconnect the u3 modue from the monitor, the nuse was electrified at the handle, on the arm holding the modue handle, and susequntly ended u in the emergency department. Not knowing whether the problem originated from the monitor or the modue, the cutomer asked for the equpment to be retuned to the 000000000000000000000000001kshop for a complete check-u. CSR: MAJEED Assad Name: u Florent MOuTARDIER Phone: +330000000068000000006 ## secteu2.biomed300000000chunimes.fr PICK u: Location Name: CHuNIMES Street: PL OF PROFESSOR ROBERT DEBRE City: NIMES State: 300000000Zip: 30000000000000000 Coutry: France Biomedical Service Hall 5, level -2 Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: In ue Eupected and ueupected Resuts: uer electrification uer Impact: Injuy Patient Impact: None SE uknown Diagnosis performed by Engineer: The cutomer reports that when the nuse in the department wanted to uplu the later foud in the emergency department. Not knowing whether it comes from the monitor or the modue, the cutomer wants the equpment retuned to the 000000000000000000000000001kshop for a complete check. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000592000000008 ----------------------------------------------------------------------------- Internal Remarks: • FEDEu confirmation nuber: I Pu300000000 • Tracking nuber 7700000000366000000001 • Pick-u time: 10000000000000000 - 16:00000000 • Carrier delivery date: 11/100000000t BIOMEDICAL SERVICE ----------------------------------------------------------------------------- | 2023-11-10 17:51 | Yes
3714 | Yes | Problems with the automatic movement of the C-arc | Problems with the automatic movement of the C-arc | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000199 200000000Nov-2000000003 FEDERICO, RAFFAELE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Problems are reported in the auomatic movement of the c-arm Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Resoluion: The shift motor has been replaced and diagnostic and fuctional tests have been carried ou with positive resuts. Parts Consued: 1.00000000000000000000000000000000800000000217000000002,FD SHIFT DRIVE ASSY PACKAGED ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006100000000200000000Nov-2000000003 GIuSANI, MARCO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Technician Device ue: ou of ue Eupected and ueupected Resuts: movement uer Impact: no impact Patient Impact: no impact SE2.3.2 Resoluion: Shift movement motor replaced, fuctional tests carried ou with pass resut ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000091 TESTOLIN, DARIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: Problems are reported in the auomatic movement of the c-arm Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer: 00contact made with u Salvatore - the arc from the parking position to the 000000000000000000000000001king position at the time of the control requst does not start bu they notice that the detector tries to get u bu fails - in some situtions it was necessary to replace the complete engine - Internal Comments: uLL# and/or description of uLL: . Repair Action: 00contact made with u Salvatore - the arc from the parking position to the 000000000000000000000000001king position at the moment of the control requst does not start bu they notice that the detector tries to get u bu fails - in some situtions it was necessary to replace the motor buiness sut - Follow u Requred: uLL# and/or description of uLL: . Repair Action: 00contact made with u Salvatore - the arc from the parking position to the 000000000000000000000000001king position at the moment of the control requst does not start bu they notice that the detector tries to get u bu fails - in some situtions it was necessary to replace the motor complete - What are the neut Trouleshooting steps : Requres onsite visit for part replacement Parts: 00000000800000000217000000002 FD shift drive assy packaged 00CC: 1 FSE 2h of 000000000000000000000000001k Technician Name : TESTOLIN, DARIO ----------------------------------------------------------------------------- | 2023-11-10 18:06 | No
3715 | Yes | M0 - CN00- Defibrillator paddles did not work - | M00 - CN0000- Defibrillator paddles did not work - |  | 2023-11-10 18:14 | Yes
3716 | Yes | Dianosing help for PC issues | Dianosing help for PC issues | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000162 Clark, Byron The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - ACQ: Image Store: Inconsistent Image Store - Frontal IPPC 00000000-Nov-2000000003 100000000:15 When was the first occurence? Within... : 00000000-Nov-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: disc uLLmessages for IPPC bu 000000000000000000000000001king uer Impact: system 000000000000000000000000001king Patient Impact: No patient impact reported by cutomer. SE 000Diagnostic performed by Engineer : 000000000000000000000000001ked with biomed who was creating log and seeing POST fails listed bu not showing any neut steps requred. on 11/9 uLLlisted shows DB and image issus.. Confirmed part nuber for IPPC and hard drives if necessary Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - ACQ: Image Store: Inconsistent Image Store - Frontal IPPC: 1. Software issu: there is an inconsistency in the patient Database wrt the stored images on the Frontal IPPC. Repair Action : per conversation Brent is going to re-create image store and them perform DB consistency repair and test uit Service actions for RMT - ACQ: Image Store: Inconsistent Image Store - Frontal IPPC: 1.1. Perform âVerify Image Store Data Consistency testâ in FSF. 1.2. If problem persists, Recreate Image Store in FSF. Don't forget to archive patient data if wanted before euecuing this step. 1.3. If the database consistency repair fails, replace the image disk. ClarityIQ_II IP PC no HDD 0000000080000000050000000023 Hard disk spare part 0000000080000000050000000000000000Resoluion : per conversation no CCB by Monday morning OK to close 000000000000000000000000001. NO CCB closing ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Brent IP & Issu: 56932000000008- Dianosing help for PC issus Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to quu. Adv cutomer of callback. Contacted On call Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-10 18:21 | No
3717 | Yes | Battery uLL GFE00OV0/IS complete*** | Battery uLL GFE0 0NOV00IS complete*** | BuQ&A Feedback: No Reason: Corrective Maintenance Euternal Remarks: 21.11.2000000003 1000000006:00000000POLARIS (POLARIS) ================= Gutavo de Assis Soua - 00000000000000005759635300000000100000000000000000000000038 =================== ==== Replaced V600000000series lug uLLbattery: 1000000001185600000000Battery installed lot ID: M0000000036P2 Equpment has been tested, released for ue, ok. Note: It is recommended to let the battery get a ful charge. NF: 00000000500000000Problem Reported by Cutomer: Battery uLLSafety Qustions: Case: 0000000022600000000272 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Carlos Ferreira Last udated by: Carlos Ferreira 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? AT Repair Action: schedued for 11/100000000 FSE Gutavo Neut Step Action: Analysis carried ou and directed to sending a buget SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-10 18:26 | No
3718 | Yes | PSA_ Saturation probes reading low and incorrect RMA Sent 00for return of Sensors | PSA_ Saturation probes reading low and incorreuRMA Sent 0/0 for return of Sensors |  | 2023-11-10 18:30 | No
3719 | Yes | No live image to be seen in the room | No live image to be seen in the room | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 13-Nov-2000000003 BASTIAAN VAN BEVEREN The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: FSE, Device ue: clinical Eupected and ueupected Resuts: after uer Impact: no live image Patient Impact: after SE System does not give live image anymore, Resoluion: Cables connected properly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000RONALD BOK The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE System does not give live image anymore, Diagnostic performed by Engineer: Ralph Lenior will check if something is connected wrong, Bastiaan replaced image PC yesterday Follow u Requred: Bastiaan will also check what is going on with image PC ----------------------------------------------------------------------------- | 2023-11-10 18:31 | No
3720 | Yes | QACASE for PCI imported product#00000 the device suddenly crashed | QACASE for P0imported produu0000000 the device suddenly crashed |  | 2023-11-10 18:42 | No
3721 | Yes | SPLIT - M0A - u0 - Plates Damaged - See reference PR 02 | SPLIT - M00A - u00 - Plates Damaged - See reference PR 002 |  | 2023-11-10 19:22 | No
3722 | Yes | uL+ failure in the defibrillation test (Replaced uL+ Kit-Therapy uuuitor) | uL+ failure in the defibrillation test (Replaced uL+ Kit-Therapy uuuor) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000799 November 13, 2000000003 Ruén Álvarez Villa The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000002-00000000 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER, 200000000000000001-00000000 00000000:00000000:00000000 Diagnostic performed by Engineer: The problem is in the condenser, we have a new one in stock and it is replaced. After corrective maintenance, the equpment is operational and meets the manuactuer's specifications. The necessary verifications have been performed to certify that the equpment is back to the operating conditions prior to the failue. Problem Description by Engineer: Cutomer Fuction/Role: We talk to José Lus García 666370000000065 from Electromedicine. Device ue: Not in clinical ue. Eupected and ueupected Resuts: The equpment does not pass the uer test, it is displaying a failue in the defibrillation test, and the statu indicator is red. uer Impact: None Patient Impact: None SE Not apply. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000867 Rau Cordero Calderón The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: We talk to José Lus García 666370000000065 from Electromedicine and he says that uL+ does not pass the uer test, it is displaying a failue in the defibrillation test and the statu indicator is red. 8612900000000HeartStart uL+ Defibrillator/Monitor obsolete since December 31, 2000000002 Assign a Field Service Service To see if it can be repaired. Order part: 0000000056000000006351 - uL+ Kit-Therapy uuuitor uN+00000000ugent Attention C, cart stops Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We talk to José Lus García 666370000000065 from Electromedicine. Device ue: Not in clinical ue. Eupected and ueupected Resuts: The equpment does not pass the uer test, it is displaying a failue in the defibrillation test, and the statu indicator is red. uer Impact: None Patient Impact: None SE Not apply. Diagnostic performed by Engineer: We talk to José Lus García 666370000000065 from Electromedicine and he says that uL+ does not pass the uer test, it is displaying a failue in the defibrillation test and the statu indicator is red. 8612900000000HeartStart uL+ Defibrillator/Monitor obsolete since December 31, 2000000002 Assign a Field Service Service To see if it can be repaired. Order part: 0000000056000000006351 - uL+ Kit-Therapy uuuitor uN+00000000ugent Attention C, cart stops ----------------------------------------------------------------------------- | 2023-11-10 19:26 | No
3723 | Yes | The device is not going through the test. | The device is not going through the test. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000895 29-Nov-2000000003 Buak Tezel The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device uage: Monitoring. Eupected and ueupected Resuts: Processor Pca is defective. uer Impact: No. Patient Impact: There was no effect on the patient. SE 2.00000000.32 Resoluion: The fauty Processor PCA part of the device was replaced with a new one. Duing the testing and control phase, it was determined that the Therapy Knob part was also fauty. According to the faut detection, 0000000056000000008981 - Mut be ordered. Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000003-16 00000000:00000000:00000000 MuTIMETER, 200000000000000002-16 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER, 200000000000000000000000000000000 00000000:00000000:00000000 Follow u Requred: The fauty Processor PCA part of the device was replaced with a new one. Duing the testing and control phase, it was determined that the Therapy Knob part was also fauty. According to the faut detection, 0000000056000000008981 - Shoud be Ordered. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000Buak Tezel The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Duing the checks, it was determined that the Processor Pca part of the device was fauty. According to the faut detection, 0000000056000000009000000001 - Mut be ordered. Resoluion: Duing the checks, it was determined that the Processor Pca part of the device was fauty. According to the faut detection, 0000000056000000009000000001 - Mut be ordered. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device uage: Monitoring. Eupected and ueupected Resuts: Processor Pca is defective. uer Impact: No. Patient Impact: There was no effect on the patient. SE 2.00000000.32 ----------------------------------------------------------------------------- | 2023-11-10 19:36 | No
3724 | Yes | The device suddenly alarmed, Pressure sensor failure - Final 05 | The device suddenly alarmed, Pressure sensor failure - Final 0/5 |  | 2023-11-10 19:38 | No
3725 | Yes | ETCO0filter line disconnected' | ETCO0'filter line disconneud' | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000831 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : End uer provided with Fiued Price Repair details Device to be collected and sent to u Bench for Investigation & Repair once PO received 1u Single Board Compuer 000000005600000000100000000 1u Capnomete Modue Kit 000000005600000000000000001 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Patient Services Administrator Device ue: In ue Eupected and ueupected Resuts: Initially the ETCO2 displayed as it shoud do with waveform and figues, after approuimately 600000000seconds the display said ETCO2 'filter line disconnected', when had not been touhed and definitely not detached. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : End uer provided with Fiued Price Repair details Device to be collected and sent to u Bench for Investigation & Repair once PO received 1u Single Board Compuer 000000005600000000100000000 1u Capnomete Modue Kit 000000005600000000000000001 ----------------------------------------------------------------------------- Internal Remarks : Hi Emily , Please let u know you collection address and a sutable date for you Tempu Pro, BT DOCK SN 600000000100000000 Faut : ETCO2 'filter line disconnected' Internal Remarks : Tracking Nuber:1Z3R60000000066970000000000000000COLLECT FRIDAY 17.11.23 Devon Air Ambuance SAM SLATER EDWARD COR000000000000000000000000001D 5 Sandpiper Cout Harrington Lane EuETER Eu00000000S 000000000039200000000666 ----------------------------------------------------------------------------- | 2023-11-10 20:01 | No
3726 | Yes | High flow without flow. - MIR Goal: 00 | High flow without flow. - MIR Goal: 0004 |  | 2023-11-10 20:30 | No
3727 | Yes | The pins of the connector of the electrode adapter cable code M00are broken | The pins of the conneur of the eleuode adapter cable code M00A are broken | 000000000000000000000000001O##: 00000000000000000000000000100000000397715 GIuFRIDA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: We deliver in NEMO withou collection 3u 989800000000100000000981 Ship to P.O. GARIBALDI - NESIMA Via Palermo 636 95122 - Catania Clinical Engineering Service Att.ne u Bugarello 300000000650000000000000000Problem Description by Engineer: Cutomer Fuction/Role: clinical ing mvs Device ue: mutiparametric monitor Eupected and ueupected Resuts: broken cables uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer: We send material to NEMO withou collection ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000997 Parts Consued: 1.000000000000000000000000989800000000100000000981,HeartStart Hands-free Cable 1.000000000000000000000000989800000000100000000981,HeartStart Hands-free Cable 1.000000000000000000000000989800000000100000000981,HeartStart Hands-free Cable ----------------------------------------------------------------------------- Internal Remarks: uuic PO Garibaldi Centro. Iba aware, requst information from the Cutomer. ----------------------------------------------------------------------------- | 2023-11-10 20:36 | No
3728 | Yes | Touchscreen front bezel is not functional. | Touchscreen front bezel is not funuonal. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 Thierry LE GONIDEC The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: qute 0000000056153600000000 Bezel, front, Non-English (white/non-coated) for SN >= 100000000111553 & 200000000000000007800000000 Follow u Requred: qute Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: patient monitoring Eupected and ueupected Resuts: parts only uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: Hello, 000000000000000000000000001ld it be possible to send u a qute for a v600000000front panel, the rotary touh screen of the selection buton no longer 000000000000000000000000001ks. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003700000000 Parts Consued: 1.0000000000000000000000000000000056153600000000,RP-BezelFront, w/NAV, Wht, NCTD, V600000000----------------------------------------------------------------------------- | 2023-11-10 20:46 | No
3729 | Yes | CMEPP -the big screen is blank- cell number 000 | CMEPP -the big screen is blank- cell number 0 000 | 000000000000000000000000001O##: 00000000000000000000000000100000000398219 100000000Nov-2000000003 Dun Peter The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: u Device ue: Vascuar Eupected and ueupected Resuts: Live was not displaying on the Fleuvision monitor uer Impact: coud not see Live image on monitor in euam room Patient Impact: N/A SE 8.1 Resoluion : Changed inpu of Live signal to Fleuvision. Tested system, displaying correctly. ----------------------------------------------------------------------------- Internal Remarks : Cutomer called back to speak with FSE becaue she hadn't received a callback. I sent the FSE an ##asking him to contact the cutomer. I notice the call back nuber on the case was illegible and changed it to the one she gave me. ----------------------------------------------------------------------------- | 2023-11-10 20:51 | No
3730 | Yes | Pressure Accuracy | 00 - Prouimal Pressure Accuracy |  | 2023-11-10 21:11 | No
3731 | Yes | Tube arm is not centered | Tube arm is not centered | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 November 15, 2000000003 Mauimo, Migul The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: The arc movements do not 000000000000000000000000001k correctly. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Z1 motor is replaced, calibrations and tests are sucessfuly performed. The equpment is 000000000000000000000000001king properly. Parts Consued: 1.00000000000000000000000000000000800000000000000006900000000FLEuARM RZ1 MOTOR DRIVE ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000797 November 15, 2000000003 Morelos, David The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: in ue Eupected and ueupected Resuts: The equpment is auhorized. uer Impact: uknown Patient Impact: uknown SE ukn bios Resoluion: Suport in installation of Z movement motor. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000373 November 1000000002000000003 Mauimo, Migul The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Waiting for Motor z1, which had already been ordered. Problem Description by Engineer: Cutomer Fuction/Role: Engineer Device ue: uEupected and ueupected Resuts: Movements blocked uer Impact: no Patient Impact: no SE 2.2 Diagnostic performed by Engineer: We performed what they recommended from the factory and conclued that the motor has problems which had already been ordered and waiting for it. Equpment waiting for spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000359 November 13, 2000000003 Hernandez, Manul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: The arc movements are disabled. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Tests recommended by the factory are performed in order to determine the specific failue. According to the tests, it is determined that the Z1 movement motor is cauing the failue, so it will be requsted for replacement. Equpment waiting for spare part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003800000000ovember 13, 2000000003 Mauimo, Migul The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: It is requred to continu with the check-u and spare parts are requsted by MH. Problem Description by Engineer: Cutomer Fuction/Role: Engineer Device ue: uEupected and ueupected Resuts: Movements blocked uer Impact: no Patient Impact: no SE 2.2 Diagnostic performed by Engineer: Suport is provided in MC to Manul Henez. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000525 November 11, 2000000003 Hernandez, Manul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Check the uLLlog, and identify problems with the motor, it does not perform the calibration of the Z1 movement, the problem is escalation and waiting for the response from the factory. We ordered the motor, waiting for a response from the factory. Equpment waiting for spare part. Problem Description by Engineer: The arc movements do not 000000000000000000000000001k correctly. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Diagnosis is performed and problems with Z1 movement are foud. We were able to calibrate bu the problem persists. It is an intermittent failue, the problem will be escalate to the factory. We will continu with the diagnosis on Monday. Equpment is supended to continu with the diagnosis. ----------------------------------------------------------------------------- | 2023-11-10 21:16 | No
3732 | Yes | Response received/u0 Patient images and data miu-up/INV00 | Response received/u0- Patient images and data miu-up/INV006 | 000000000000000000000000001O##: 00000000000000000000000000100000000398000000001 Sebastien LE GALL The system meets the specifications for the performed service and is retuned to ue. Resoluion: Case closed. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: duing euams Eupected and ueupected Resuts: data miuing (images) between 2 euams uer Impact: None Patient Impact: None SE 1.7.1 Diagnostic performed by Engineer: Description of the problem: the cutomer told me that images were miued between t000000000000000000000000001euaminations. When did the first occurence take place? Tusday, November 1000000002000000003 At what frequncy does this problem occu? Once uiqu At what stage of the 000000000000000000000000001k process does the problem occu? Acqusition I informed the cutomer that the problem is identified and that he was informed by the letter FSN79500000000532. He told me that the uer is new to the system and that he will take the system uage flow with him. I am sending the letter that was already sent previouly to the cutomer. ----------------------------------------------------------------------------- | 2023-11-10 21:16 | No
3733 | Yes | Bench Repair Tracking | Bench Repair Tracking | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000Kaan Ozkan The system does not meet the specification for the performed service and is not retuned to ue. Follow-u Requred: Repair on the atol Tools & Equpment : MuTIMETER, 200000000000000003-31 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER, 200000000000000002-00000000 00000000:00000000:00000000 Diagnostic performed by Engineer: The inspection performed detected that the device gave low uuuitor valu uLL uuuitor is fauty. DFM100000000 ASSY uuuitance 0000000056000000009000000001 shoud be quted. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: Therapy uLL uer Impact: no Patient Impact: no SE D.00000000.21 ----------------------------------------------------------------------------- | 2023-11-10 21:41 | No
3734 | Yes | LIFE Quotes / Wireless Base Station and Wireless Footswitch | LIFE Quotes / Wireless Base Station and Wireless Footswitch |  | 2023-11-10 22:01 | No
3735 | Yes | Mid scan the system fully froze and they were unable to eupose | Mid scan the system fully froze and they were unable to eupose | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000000000006-Nov-2000000003 Mark Bradley The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Mid scan the system fuly froze and they were uable to eupose and were uable to tun the machine again bu after 00000000ard restarts they got it to 000000000000000000000000001k bu then it froze again uer Impact: NA Patient Impact: NA SE8.2 Resoluion : Lateral IPPC replaced. System tested. No problems encoutered. Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell IP PC no HDD Tools & Equpment : MuTIMETER,20000000000000000000000009 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000008 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 15-Nov-2000000003 Mark Bradley The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Mid scan the system fuly froze and they were uable to eupose and were uable to tun the machine again bu after 00000000ard restarts they got it to 000000000000000000000000001k bu then it froze again uer Impact: NA Patient Impact: NA SE8.2 Reguatory Qustions : N/A Follow u Requred : Fit a replacement Lateral IPPC Diagnostic performed by Engineer : Replacement Lateral IPPC fitted. Bu ufortuately this had a net000000000000000000000000001k faut. New lateral IPPC ordered. Tools & Equpment : MuTIMETER,20000000000000000000000009 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000008 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000916 100000000ov-2000000003 Mark Tuker The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace lateral IPPC AND reload software Reguatory Qustions : N/A Diagnostic performed by Engineer : No symptoms on visit. uLLlog indicates lateral imaging failue. Recent lateral Imaging PC replacement to be replaced again Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Mid scan the system fuly froze and they were uable to eupose and were uable to tun the machine again bu after 00000000ard restarts they got it to 000000000000000000000000001k bu then it froze again uer Impact: NA Patient Impact: NA SE8.2 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000819 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to faut find .Please encouage staff to provide a mobile nuber when logging any cases going forward Reguatory Qustions : N/A Diagnostic performed by Engineer : No PRS/RSW , du to incorrect tel nuber RSE uable to speak to end uers to verify what happened. System had new Haswell Host/Ip < 3 weeks Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Mid scan the system fuly froze and they were uable to eupose and were uable to tun the machine again bu after 00000000ard restarts they got it to 000000000000000000000000001k bu then it froze again uer Impact: NA Patient Impact: NA SE8.2 ----------------------------------------------------------------------------- Internal Remarks : Confirmed with Monica from GE for 1000000001 Internal Remarks : Booked FSE for tomorrow 15/11 tried to call site no answer ----------------------------------------------------------------------------- | 2023-11-10 22:26 | No
3736 | Yes | The system show message “Generator is bu starting up, No u-ray is possible.” | The system show message “Generator is bu starting up, No u-ray is possible.” | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000851 100000000Nov-2000000003 CHATCHAI SANTANA The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technician. Device ue: In clinical ue. Eupected and ueupected Resuts: The system show message “Generator is buy starting u, No u-ray is possible.” uer Impact: No impact. Patient Impact: No impact. SE 1.2.3 Diagnostic performed by Engineer : uLL: From Generator device: uSC: Tue Curent ou of range. Main breaker of Air condition is trip. Temperatue in technical room is too high, Generator 000000000000000000000000001king not properly. Resoluion : Tun-ON main breaker of Air condition and waiting for temperatue in technical room is decrease. The system test and retun to ful fuction. ----------------------------------------------------------------------------- | 2023-11-10 22:46 | No
3737 | Yes | SPLIT-battery does not charge anymore-See reference PR 00 | SPLIT-battery does not charge anymore-See reference PR 000 |  | 2023-11-10 22:50 | No
3738 | Yes | SPLIT-The buttons of the front cover are also damaged-See reference PR 00 | SPLIT-The buttons of the front cover are also damaged-See reference PR 000 |  | 2023-11-10 22:51 | No
3739 | Yes | The wireless foot pedal intermittently works | The wireless foot pedal intermittently works | 000000000000000000000000001O##: 00000000000000000000000000100000000399555 100000000ov-2000000003 Campbell Jim The system meets the specification for the performed service and is retuned to ue Resoluion : the wireless footswitch has a known issu. There is a connection issu with the hardware in the system. An FCO is coming ou in early Februry 20000000000000000o address the problem. I have instruted the cutomer that I will escalate their case so it gets addressed first when the FCO to fiu the problem is released. In the meantime the wired footswitch is the best option. Problem Description by Engineer : Cutomer Fuction/Role: technologist Device ue: stepping on the foot pedal to take u-ray Eupected and ueupected Resuts: the foot pedal 000000000000000000000000001ks intermittently uer Impact: need to switch to wired foot switch Patient Impact: none SE00000000.5 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: David Seeram IP & Issu: 80000000090000000035;The wireless foot pedal intermittently 000000000000000000000000001ks Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create case and onsite 000000000000000000000000001Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-11 01:06 | No
3740 | Yes | when rotating the device the unit free floats and moves itself. | when rotating the device the unit free floats and moves itself. | Internal Remarks : Interaction Type: New Case Caller: Michael Fewkes IP & Issu: "76890000000056 when rotating the device the uit free floats and moves itself." Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and created onsite suport service 000000000000000000000000001for cutomer (Cutomer stated he is not the rad tech and has very limited knowledge on the system) Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-11 01:11 | No
3741 | Yes | Service level 0ccess is needed for Fleu PC | Service level 0access is needed for Fleu PC | 000000000000000000000000001O##: 00000000000000000000000000100000000519813 00000000-Dec-2000000003 Peterson Daniel The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,200000000000000002-29 00000000:00000000:00000000 Resoluion : tested system and foud longituinal drive motor slipping. Attempted calibration mutiple times to get the curent calibration to pass. Recommend replacing drive motor. Foud cradle on the table no fuctioning correctly. The zero offset is replaced with an incorrect valu when the system is powered off while the table is cradled. Recommend replacing the smart drive controller. Provided 3rd party part nuber. Foud IWS pc not booting. Recommend reloading software and possibly replacing hard drive. 3rd party notified Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: clinical Eupected and ueupected Resuts: lateral drive issu uer Impact: none Patient Impact: none SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000519811 00000000-Dec-2000000003 Diaz Fernando The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : Tested system, able to determine longitue motor for C-arm drive slipping on eutended rails cauing calibration to fail. Completed calibrations on C-arm by rotating C-Arm in opposite direction as indicated in calibration setu. Also determined that the table center potentiometer and potentiometer readou modue to be fauty in table, along with table cradle potentiometer readou modue. The table is operational bu fail calibrations. Also foud 3D PC not booting application and will need futher evalution to determine the problem. 3D option on system is curently not being ued as per tech. System is operational. Informed cutomer 626 Imaging manager John DiMarzo that the mentioned parts will need to be replaced and calibrations perfomed for system optimal performance. System repairs will be completed as per 626 Imaging manger John DiMarzo 626 being udated of issus. Antonio Lagnada Biomed Lead at VA Loma Linda was also udated by these finding. Foot switch has been replaced on system, part ordered by 626 Imaging and FCO for foot switch will be completed by Dan Peterson. System settings have been backed u. Problem Description by Engineer : Cutomer Fuction/Role: Lead Device ue: yes Eupected and ueupected Resuts: System will not pass C-Arm longitue calibration uer Impact: limited Patient Impact: no SE n/a ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000560000000021-Nov-2000000003 Diaz Fernando The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Lead Device ue: no Eupected and ueupected Resuts: System requres beam longitue calibration and backus. uer Impact: yes Patient Impact: no SE n/a Diagnostic performed by Engineer : Attempted to complete stand beam curent longitue calibration and it failed at the eutended rail section. Calibrated stand within the rails withou the eutension. Stand can be manully puhed back. System is fuctional, with the euception of the new park position. Case has been escalated to have NSS help evalute system. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000000000003-Nov-2000000003 Diaz Fernando The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Tested system observed that the system 000000000000000000000000001ld boot u and stay at 00000000:00000000 coutdown and 000000000000000000000000001ld not proceed. Diagnosed Fleu PC to be bad uing PC analyzer. Informed cutomer Antonio Lagnada and 3rd party servicing company of findings. Service will be completed by 626 imaging. System is down. Problem Description by Engineer : Cutomer Fuction/Role: 3rd Party 626 Device ue: no Eupected and ueupected Resuts: System boots u and stops cout down 00000000:00000000 uer Impact: yes Patient Impact: no SE n/a Diagnostic performed by Engineer : Tested system observed that the system 000000000000000000000000001ld boot u and stay at 00000000:00000000 coutdown and 000000000000000000000000001ld not proceed. Diagnosed Fleu PC to be bad uing PC analyzer. Informed cutomer Antonio Lagnada and 3rd party servicing company of findings. Service will be completed by 626 imaging. System is down. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Fernanado Diaz (FSE) IP & Issu: FSE called to get IGT NSS Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Consuted NSS and roued FSE for suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: John PH# 631800000000000000002 IP & Issu: Missing a lot of components / Needs to get FSE ou again to help finish replacing parts needed. Need adjuted qute if needed also. Related to case 0000000022600000000795 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Created 000000000000000000000000001and Paged QA Team Neut Step Action: Awaiting Entitlement Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Fernado, can't pass long ceiling curent calibration after restoring 3rd party service system, incluing ne licensee, new PC/ghost etc. Learned it has head end eutended ceiling rail and configued as suh. Log shows movement not available. Went thrulong pot, position area and position adjutment good, bu curent cal fails at close to the high end. Advised to replace the long motor first, also maybe long pot, Their part nubers. provided. Internal Remarks : Planner Initials: GM Action Taken: FSE Fernando Diaz requsted call to be moved to 12/00000000o time change having issus with go app Oucome: dispatched Internal Remarks : Interaction Type: New Case Caller: John Dimarzo IP & Issu: " 56300000000879 Service level 2 access is needed for Fleu PC" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: sumitted qute requst for onsite; informed cutomer to eupect contact Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : TPR for several parts Internal Remarks : Interaction Type: Cutomer Call Back Caller: " John Dimarzo 631-800000000-000000002" IP & Issu: 56300000000879 Service level 2 access is needed for Fleu PC Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: Qute Action Taken: Created Qute Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : FSE Fernando Diaz. Verified correct Fleuvision 2 PC p/n. Provided 00000000800000000119123. ----------------------------------------------------------------------------- | 2023-11-11 03:21 | No
3742 | Yes | Alarms-Failure to Alarm / Did not Alarm VTAC- No trouble found-working to spec - Provided information. | Alarms-Failure to Alarm / Did not Alarm VTAC- No trouble found-working to spec - Provided information. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009700000000Youg, Maria The system meets the specification for the performed service and is retuned to ue Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Clinical Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE Rev C Problem Description/ueupected Resuts: The Vtach was not alarming. Resoluion : Problem Description: The Vtach was not alarming. Form ued: Trouleshooting/Investigation Object: Jut this one patient Not Object: Software Version: Rev C Date and Time of Issu: uer Actions: Where is the Deviation: Trouleshooting Steps Taken: This patient was not alarming Vtach when the tech seen the ru of PVCs Confirmed that the Vtach settings were HR 1300000000and PVCs >5 The patients HR did not go over 122 So it did not violate the HR portion of the Vtach alarm They did get yellow vent rhythm alarm for the PVCs Euplained to monitor tech both 000000000000000000000000001ld need to be violated at the same time to produe the Vtach alarm. Was PRS ued: No PRS Permission: Is Caue Known: Yes Caue: The HR did not violate the Vtach setting ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: virginia griffith IP & Issu: ccb states RSE advised issu if more of a clinical issu Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created clinical 000000000000000000000000001and uferred - udated case statu to Critical P1 Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: New Case Caller: virgina griffith 600000000000000000000000000000000eut81763 on behalf of Terri Reed IP & Issu: "88553278 The system is not alarming VTAC. and the patient connected to it is VTAC" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-11 04:01 | No
3743 | Yes | Monitor Froze during care of patient | Monitor Froze during care of patient | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000519 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Follow u Requred : Problem Description: tech reports the Tempu Pro froze on 11/9/2000000003 aroud 7:00000000 a.m. when attempting to print, device was removed from service, sending to bench for repair Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repair Confirmed delivery address ##d bench repair form Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE Problem Description/ueupected Resuts: tech reports the Tempu Pro froze on 11/9/2000000003 aroud 7:00000000 a.m. when attempting to print, device was removed from service, sending to bench for repair Diagnostic performed by Engineer : Problem Description: tech reports the Tempu Pro froze on 11/9/2000000003 aroud 7:00000000 a.m. when attempting to print, device was removed from service, sending to bench for repair Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form ----------------------------------------------------------------------------- Internal Remarks : Received Teams message on 11/9/23 from RSE - Devon Adams asking that I accept 000000000000000000000000001O#from quu and call cuback to assist Replied with agreement and accepted 000000000000000000000000001O# --- Called cuback at (00000000) 272-0000000099, left a voic## sent an ## 1st attempt - 11/00000000 - 6:17 PM - Left voic##- Sent ##Internal Remarks : Checked ##on 11/21/23, foud reply from cuon 11/9/23: We will be ou of the office util Monday du to the holiday. I will call and follow u on Monday. If youneed futher before then, I am available by ## Thank you --- Foud reply from cuon 11/1000000003: Please call or respond to the case online so we can keep this moving forward. I tried calling the nuber in you response below bu was uable to remain on hold for longer than 300000000minues. We need to get this monitor in for evalution. --- Foud additional remote service and Bench 000000000000000000000000001O# where my co-000000000000000000000000001ker - Kim Jordan assisted cuon 11/1000000003 with sending device to the Bench for repair --- Sent reply ##to cuapologizing for my delay in response: Hi Stephen. I apologize for my delay in responding to you Thank youfor you patience. I see that my co-000000000000000000000000001ker - Kim 000000000000000000000000001ked with youon 11/100000000o get you device sent to ou Bench for repair. Feel free to contact u shoud futher assistance be requred for this repair. Thank youagain Internal Remarks : Interaction Type: New Case Caller: Stephen West IP & Issu: 1000000000000000071 & Monitor Froze duing care of patient When printing the 12-lead the screen froze and 000000000000000000000000001ld not let me select anything for abou 300000000seconds. I tried to send to Fairview suport center bu it 000000000000000000000000001ld not let me click anything. The screen then tuned off and then the monitor restarted and ran a self check. All of this was happening while the paper was still printing. Then it went back to monitoring the patient like nothing happened. It then popped u an “monitoring interruted” message. The monitor is fuly charged and was not on the monitor mout. I have a pictue of the uLLmessage if needed. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Cutomer comment : Other I called Alvin Harris at 800000000-722-9377 in response to an ##received regarding this case. Ou office was closed at the time he tried to contact u. I am following u to keep this case moving forward. File uloaded by CS Portal uer : The information/requst pertains to: Technical information File Name: m7 monitor froze message.jpg. Comments: Here is the uLLmessage that was displayed Internal Remarks : Checked ##on 11/28/23, foud reply ##on 11/22/23 from cu Alvin, No need to apologize, nothing needed at this point. Please take a good look at this device, it’s never a good thing to have a cardiac monitor freeze u while youare 000000000000000000000000001king on a patient. Thanks! ----------------------------------------------------------------------------- | 2023-11-11 04:16 | No
3744 | Yes | uPro ¦ AC inlet where power cord is inserted to provide AC power and charge battery is broken | TempuPro ¦ AC inlet where power cord is inserted to provide AC power and charge battery is broken | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000Harris, Alvin The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Other: Chief/Officer Device ue:Not in Clinical ue Eupected and ueupected Resuts:NA uer Impact:Can't power device uing AC or charge installed battery Patient Impact:No impact SEv00000000.300000000roblem Description/ueupected Resuts: AC inlet where power cord is inserted to provide AC power and charge battery is broken CT:G Diagnostic performed by Engineer : Problem Description: AC inlet where power cord is inserted to provide AC power and charge battery is broken AC inlet where power cord is inserted to provide AC power and charge battery is broken Form ued: Bench Trouleshooting Steps Taken: Received call on 11/9/23 from Distribuor - Chad Callison - chad00edessentials.com via direct connect uing 00000000000000000000000039 Cueuplained that the end uer at the facility was Jon Simons - 815 222 2819 - jsimon00pfire.com Confirmed there was no patient involvement Cuwanted end uer to receive a loaner device with SW rev v00000000.300000000ith all featues eucept for IBP Custated he had the device that requred repair and wanted to receive a shipping label Verified cus ship to address MELROSE PARK FIRE DEPT 3600000000 W Lake St Melrose Park IL 6000000006000000002700000000 ATT: Jon Simon Resoluion : Sent to Bench --- ##d Irvine CA Bench form on 11/9/23 to cu Sumitted requsts for cuto receive shipping label and end uer to receive loaner from Bench Internal Comments : Checked case on 11/29/23, foud several ## were sent with a qute to end uer/cutomer - Jon Simon informing him that the servicing of the device at the Bench 000000000000000000000000001ld requre puchase order approval Replied to each ## directing my responses to Distribuor - Chad Callison with the qute attached: Hi Chad. I apologize for the delayed response. While closing ou my 000000000000000000000000001O# today, I noticed that ou Bench has sent ## to the end uer/cutomer - Jon Simon seeking a puchase order for the servicing of the Tempu Pro monitor w/SN 615500000000. There appears to be no warranty/contract coverage for this device. The most recent contract was CON000000000000000096 and was active from 1/11/2000000002 to 1/1000000002000000003. Please take a look and respond to the Bench. Thank you --- Copied on reply from Chad to cu Good morning Jon, I recently shipped the Tempu Pro 615500000000 into the repair facility for the Charger Port being broken. They mentioned they have ##d youa repair qute and are awaiting you response. ufortuately, you uits do not have an eutended warranty so they need you approval before fiuing the uit. Can youplease let me know if youreceived the qute? ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE uknown Diagnostic performed by Engineer : Confirmed complaint the uits Charging port is physically damaged and to correct the charging port The Front Case will have to be replaced. Additionally the Rear Cover is physically damaged and the Kickstand is missing. To correct the issu we will need to replace the Rear Cover Assembly Follow u Requred : 00000000356000000005000000001-Front Case Assembly 0000000056000000005000000001-Rear Cover 00000000ous for repair ----------------------------------------------------------------------------- | 2023-11-11 04:36 | No
3745 | Yes | direct line,0 room , lab 0cath lab: upon powering up, all the monitors are shutting down | direuline,0h room , lab 0 cath lab: upon powering up, all the monitors are shutting down | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000762 100000000Nov-2000000003 Gregory Christopher The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,200000000000000008-31 00000000:00000000:00000000 Resoluion : Find that system image processing system is uavailable. Diagnostics point to image processing board. Check system board software; IPB lost its firmware. Re-flash ipb and reboot system. System tested on after mutiple reboots. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not able to ue Eupected and ueupected Resuts: system fluro monitors all black uer Impact: room down Patient Impact: none SE 7.u ----------------------------------------------------------------------------- | 2023-11-11 05:06 | No
3746 | Yes | faulty footswitch, not able to program footswitch | faulty footswitch, not able to program footswitch | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000 Marc Skelton Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-11-11 05:41 | No
3747 | Yes | Stand cannot move | Stand cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: QAA00000000/u uuuFD200000000Safety Qustions: Case: 0000000022600000000500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-11 06:01 | No
3748 | Yes | Host PC cannot boot up | Host PC cannot boot up | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 10000000011.2000000003 00000000:0000000019 uC (86000000008900000000"SuT. Ready for Clinical ue local time" : " 2000000003-11-10000000011:00000000" SAP Transfer: Case Transferred to SAP Euternal Remarks: 10000000011.2000000003 00000000:36:21 uC (86000000008900000000 Service content: uable to boot, re-insert and uplu the host pc memory modue and board, replace the motherboard battery, and the device retuns to normal. Service resuts: After repair, the equpment is normal. Others: none SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: SA：00000000R/u uuuFD200000000 SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Overview of cutomer problem: Hos tPC device cannot be tuned on. Detailed description of cutomer problem or faut phenomenon: Hos tPC device cannot be tuned on. Is it a faut of the device itself? Yes Trouleshooting Action: Key faut codes or phenomena: The Hos tPC device cannot be powered on. Remote diagnosis and test resuts: The Hos tPC device cannot be powered on. Check that the power suply is normal. It is supected that the hostpc is fauty. PRS TR statu: All pass Repair Action: Recommended onsite repair steps: 1. Check the HostPC power suply, 2. Reformat the HostPC OS disk and reinstall the system, 3. Replace the host PC. Notes on cutomer requrements: Onsite Safety Qustions: Case: 000000002260000000092 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 10000000011.2000000003 00000000:36:200000000TC (86000000008900000000 "RDI Accuacy: YES" $Q7: (Confirmed reported problem, Host PC cannot bootu) $Q8: (Check the Host pc power, it is ok, maybe the host pc need to be reseat.) $Q9: (Open the host pc, reinstall the memory and graphics card, replace the BIOS battery.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools (NA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: NA ----------------------------------------------------------------------------- | 2023-11-11 11:37 | No
3749 | Yes | OCAP:0/u FD0Post processing is not operational | OCAP:0R/u FD0Post processing is not operational | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 11.11.2000000003 13:00000000:00000000 uC (8600000000600000000) Service content (requred): The equpment cannot be euposed. Check to determine if the power suply of the image disk is fauty. After replacing the power suply, test and ue the normal machine to resue normal ue. Service resuts (requred): [After repair, the equpment retuned to normal operation.] Other important matters: [none] Problem Reported by Cutomer: OCAP:00000000R/u uuuFD200000000 SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 11.11.2000000003 13:00000000:00000000uC (8600000000600000000) $Q7: (Confirmation & Description of problem/Post processing is not operational) $Q8: (cutomer report device report uLLmessage Post processing is not operational, check and foud power suply image disk defect) $Q9: (replace power suply image disk then test system fuction ok, system OK) $Q100000000 (Test system ru ok and inspection data refer to: PA. Device retuned to ful fuctionality. ) $Q11: N/A SAP Transfer: Case Transferred to SAP T2 Activities:Repair Action: PPD/ 00000000230000000000000000900000000 POWER SuPLY IMAGE DISK SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002260000000027 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problem: The device can only see throuh, cannot eupose, cannot save images, cannot retrieve previou images, and engineers need to be present. Is it the faut of the device itself? Yes Trouleshooting Action: Remote diagnostics and test resuts: Image access failue PRS TR statu: None PRS Repair Action: Recommended onsite repair steps: It is recommended to plu and uplu the ISB board, plu and uplu the image hard disk, and replace broken accessories. Cutomer demand notes: Onsite ----------------------------------------------------------------------------- | 2023-11-11 11:37 | No
3750 | Yes | Defibrillator/Monitor/ucannot be measured | 0N/Effi0 DFM0 Defibrillator/Monitor/ucannot be measured | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Repair Action: Service Content: uable to detect uReplacement of ulead set via normal spare parts service. Service resuts: Problem Reported by Cutomer: 00000000N/uDFM100000000 Defibrillator/Monitor Safety Qustions: Case: 000000002260000000073 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-11 12:26 | No
3751 | Yes | uuushock test failed | u shock test failed | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: ECR:00000000/HeartStart uuumonitor/defib/cannot discharge Safety Qustions: Case: 0000000022600000000000000002 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-11 12:36 | No
3752 | Yes | Generator is bu starting up, no u-ray is possible | Generator is bu starting up, no u-ray is possible | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 1000000001.2000000003 00000000:300000000 uC (860000000089900000000 RDI Accuacy: yes Q7: (Confirmed reported problem. tue filament was opened, Cutomer cannot ue system to diagnostic.) Q8: (On site review uLLlog, foud system report HV uLLi ue mutimeter to measue filament resistance of tue, filament was opened. So need replace a new tue to fiu it.) Q9: (replace a new tue and perform the calibration, device retun normal ) Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) Q11:Calibrated Tools name (Mutimeter Flue 287) S/N (2700000000000000001) Neut Calibration date (2000000003/8/15) Neut Step Action: PPD 989000000000000000085100000000 *1, the tue needs to be approved before shipment, please send 000000000000000000000000001k first FSE: Gu Yuie -86-151581800000000500000000SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022600000000200000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problem: After the device is tuned on, it reports "Generator is buy starting u, no u-ray is possible". The system cannot be euposed. After a while, the device retuns to normal. Cutomer problem description or faut phenomenon: Equpment After powering on, "Generator is buy starting u, no u-ray is possible" is reported, and the system cannot be euposed. After a while, the equpment retuns to normal. Engineers need to check whether the caue of the failue is the equpment itself: Yes Trouleshooting Action: Key faut code Or phenomenon: "Generator is buy starting u, no u-ray is possible" is reported after the device is tuned on. The system cannot be euposed. After a while, the device retuns to normal. Engineers need to check the caue of the faut. Remote diagnosis and test resuts are reported: "Generator is buy starting u, no u-ray is possible" is reported after the device is tuned on. "Generator is buy starting u, no u-ray is possible", the system cannot be euposed, and after a while, the equpment retuns to normal. Engineers need to check the caue of the faut. It is supected that the high voltage has an abnormality duing the self-test. PRS TR statu: uable Repair Action : Recommended on-site maintenance steps: 1. Check the high-voltage log to confirm whether there is any self-test abnormality. 2. Check whether the SPI commuication line between uSC and each susystem is in good contact. Notes on cutomer requrements: Onsite Problem Reported by Cutomer: uable to boot SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 1000000001.2000000003 00000000:300000000 uC (860000000089900000000 Service content: Equpment euposue uLL it is juged that the small focu filament of the tue is open and the large focu filament is aging, and a new tue needs to be replaced. Service resuts: After service, the equpment 000000000000000000000000001ks normally. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-11 12:51 | No
3753 | Yes | Software get stuck | Software get stuck | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 12.11.2000000003 12:51:59 uC (860000000013100000000 Service content: Intermittent crash, check foud that there is a problem with the host host system, restore the host host system. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Overview of cutomer problems: The software is stuk and needs to be reinstalled.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: The software is stuk and needs to be reinstalled.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: The software is stuk and needs to be reinstalled.Engineers are requred to be present Notes on cutomer requrements: Onsite Safety Qustions: Case: 0000000022600000000312 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000R/u uuuFD200000000 SAP Transfer: Case Transferred to SAP Internal Remarks: 12.11.2000000003 12:52:00000000 uC (860000000013100000000 "RDI Accuacy:Yes" $Q7:Confirmed reported problem.The application software ru slowly $Q8: Check the machine and find the system of host PC have problem. $Q9: Restore the host pc system Image. $Q100000000 Test and inspection data refer to: PA. Device retuned to ful fuctionality. $Q11: NA. Curent tue install dated:(N/A) curent tue uage(:N/A) Curent tue install dated:(N/A) curent tue uage(:N/A) ----------------------------------------------------------------------------- | 2023-11-11 12:51 | No
3754 | Yes | The charger button of the paddle run out of service and fail to the operational check | 000-The charger button of the paddle run out of service and fail to the operational check | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 000000002260000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : 16925725-The charger buton of the paddle ru ou of service and fail to the operational check SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-11 12:56 | No
3755 | Yes | SELF TEST FAILED .. | SELF TEST FAILED .. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : CM - SELF TEST FAILED Safety Qustions : Case: 0000000022600000000600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable ----------------------------------------------------------------------------- | 2023-11-11 13:16 | No
3756 | Yes | The AED beeped - Log analysis requested | The AED beeped - Log analysis requested | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 Suui, Kosue The system does not meet the specification for the performed service and is not retuned to ue Resoluion: After analyzing the main uit log at the manuactuing factory in the uited States, we foud a record of failue of the high voltage (HV) calibration test. It was necessary to replace the main uit du to a diagnostic failue du to an electrical failue in a high voltage element within the main uit, bu the repair has been cancelled, so we will retun the complete set. On December 1, 2000000003, we received you requst to cancel the repair via ## [Simutaneou FCO 000000000000000000000000001k] There are no FCOs that requre 000000000000000000000000001k on this equpment. December 00000000, 2000000003 Fuua Denshi Niigata Hanbai We will send the 000000000000000000000000001k report along with the equpment to the person in charge. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit the 000000000000000000000000001k report on paper? <YES>[Final confirmation] Perform operation check after 000000000000000000000000001k based on prescribed instrutions NO Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Fuua Electronic Sales Representative Device ue: Not ued Eupected and ueupected Resuts: A beep soud was heard, bu the caue is uknown. uer Impact: None Patient Impact: None SE None Diagnostic performed by Engineer: After receiving the produt, we confirmed that it passed the battery installation test. Even in standby monitoring, the reproduibility of abnormal alarms coud not be confirmed. Tools & Equpment : MuTIMETER, 200000000000000003-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks: Faut details 2 A beep soud was heard, bu the caue is uknown. Please analyze the caue. LOT Adut Pad 220000000000000000-0000000057 Adut Pad 221000000003-0000000056 Battery 89159P ----------------------------------------------------------------------------- | 2023-11-11 13:36 | No
3757 | Yes | not turning on at all | not turning on at all | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000751 21-Nov-2000000003 Martinic, Chris The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP Problem Description by Engineer : Cutomer Fuction/Role: radiographer Device ue: angiography Eupected and ueupected Resuts: no uer Impact: none Patient Impact: none SE Resoluion : Replaced us controller and battery. Both were requred for the repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 13-Nov-2000000003 Martinic, Chris The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: radiographer Device ue: Angie Eupected and ueupected Resuts: None uer Impact: None Patient Impact: None SEuknown Resoluion : Load imbalance faut on single phase uS in M cabinet. Bypassed uS temporarily while parts are raised. Cutomer informed, uit tested and retuned to ue. ----------------------------------------------------------------------------- | 2023-11-11 13:46 | No
3758 | Yes | Machine is not booting up | Machine is not booting up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000261 Kuar, Atu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : as per cutomer Application is not booting u and struked at 0000000000000000 and uable able to ue Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : FSE vipin suported cutomer remotely and foud that IPPC was not booting u. . Internal Comments : uLL# and/or description of uLL: No uLL Repair Action : FSE vipin suported cutomer to switch ON the IPPC manully now IPPC is 000000000000000000000000001king fine. Observed the system for few days now system is 000000000000000000000000001king fine Resoluion : FSE vipin suported cutomer to switch ON the IPPC manully now IPPC is 000000000000000000000000001king fine. Observed the system for few days now system is 000000000000000000000000001king fine ----------------------------------------------------------------------------- | 2023-11-11 14:11 | No
3759 | Yes | Pads Adapter Cable not working. | Pads Adapter Cable not working. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000---------------------------------------------------------------------------- | 2023-11-11 14:21 | No
3760 | Yes | System powered out. Clarity computers will ont power up. L0reaker was tripped | System powered out. Clarity computers will ont power up. L0breaker was tripped | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000776 23-Nov-2000000003 Botts, Andrew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In ue Eupected and ueupected Resuts: Clarity PC 000000000000000000000000001’t power on after power hit uer Impact: Down Patient Impact: Delays SENA Parts Consued : 1.0000000000000000000000000000000080000000050000000023,ClarityIQ_II IP PC no HDD Resoluion : DEFOA PC retuned. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000100000000ov-2000000003 Botts, Andrew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In ue Eupected and ueupected Resuts: Clarity PC 000000000000000000000000001’t power on after power hit uer Impact: Down Patient Impact: Delays SENA Resoluion : Cutomer stated IPPC 000000000000000000000000001ld not boot u after power hit that tripped the L2 breaker bu power was present at the cable. Ordered new IPPC for neut day Satuday delivery. Installed new IPPC and it 000000000000000000000000001ld power u the recreate image store 000000000000000000000000001ld fail uing the new and old drives. Tested with PC diagnostics tool and the PC failed the hard disk test. Ordered another IPPC NFO DEFOA. The second PC fuctioned properly and resolved the issu. Tested and verified system as well as performed a ghost backu. Retuned system to service. Tools & Equpment : MuTIMETER,200000000000000002-29 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Follow u Requred : Part picku ----------------------------------------------------------------------------- Internal Remarks : Case: 000000002260000000000000000 FSE: Botts, Andrew FSE looking for correct part nuber to IPPC. Instruted FSE to place part order for 0000000080000000050000000022 - ClarityIQ_II IP PC no HDD. Internal Remarks : Interaction Type: New Case Caller: Chris IP & Issu: "38356- System powered ou. Clarity compuers will ont power u. L2 breaker was tripped" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created qute, cutomer wants AH Neut Step Action: awaiting qute Verbal PO Nuber: 10000000027500000000erbal PO Approver: Chris Valera Verbal PO CAP: N/A Internal Remarks : FSE Andrew called in. Help[ed FSE via phone. FSE replaced IPPC and drives. Able to load the IPPC in download board sw. uable to recreate image store on drives. FSE tried original drives too. tested with PC diagnostics tool and shows a drive mismatch, not seeing one of the drives. Sugested FSE to get another IPPC. ----------------------------------------------------------------------------- | 2023-11-11 14:26 | No
3761 | Yes | He was not able to u the wireless foot switch to take a picture. | 0-00T&M - Order parts FD0 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000866 200000000Nov-2000000003 Anh Hu Ho The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000975000000001,u uuu(OR Table) R 8.2.17 sw pack 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V Problem Description by Engineer : Cutomer Fuction/Role: Technical Device ue: u uuuFD200000000Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE8.2.1.17 Resoluion : Reinstall OS and Application SW Replace PD SCOMP DCPS 2000000002-5V ----------------------------------------------------------------------------- | 2023-11-11 14:31 | No
3762 | Yes | uL+ therapy disabled alarm | uL+ therapy disabled alarm | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: No Faut Foud: uconfirmed Problem / Failue T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . SAME DAY HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on of day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLLAlarm:Therapy disabled, Ru OP check WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Instrut cutomer to ru the operation check and all items are passed. uL+ retun to ful fuctionality Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: NA Test/Inspection: Power on test and visul test are passed, device retun to ful fuctionality Test Equpment: NA Safety Qustions: Case: 0000000022600000000269 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Reserved disabled SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-11 14:56 | No
3763 | Yes | 0 / u0/ uL + Power supply issue (Operation check corrected the issue) | 00/ u00 / uL + Power supply issue (Operation check correud the issue) | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes : This PO is a Reference Nuber only SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : 00000000R_000000002_CV_uL plu_Power suply issu u5130000000000000000 Safety Qustions : Case: 0000000022600000000632 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Vickie Hung Last udated by: Vickie Hung 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-11 15:36 | No
3764 | Yes | No u-ray with footswitch intermittently | No u-ray with footswitch intermittently | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/u uuuFD200000000T2 Activities:Repair Action: Recommended on-site repair steps: "nemo,00000000270000000000000000100000000Footswitch CV 3p 00000000200000000to hosp, the repair person will collect it" Service Content: The foot pedal is not sensitive to euposue. Send the foot pedal throuh spare parts eupress service. Safety Qustions: Case: 0000000022600000000812 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang Zhu Last udated by: Weiqiang Zhu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-11 15:56 | No
3765 | Yes | starts by itself displays message check fan program crc test failure. | starts by itself displays message check fan program crc test failure. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000Assad MAJEED The system does not meet the specifications for the performed service and is not retuned to ue. Resoluion: Procedue to reload the software in the V600000000sent to the cutomer. If the failue persists, the cutomer will contact u to perform on-site service to replace the PCBA CPu Diagnostic performed by Engineer: The cutomer reports that duing start-u of the V600000000 after a few seconds, the V600000000restarts sudenly and displays an uLLmessage "Check fan: CRC test failue of the program". It seems that the failue is related to the software of the V600000000or to the PCuPCBA board. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Reguatory Qustions: No/No ----------------------------------------------------------------------------- | 2023-11-11 15:56 | No
3766 | Yes | Electrocardiographic monitoring failure | Eleuocardiographic monitoring failure | 000000000000000000000000001O##: 00000000000000000000000000100000000521358 00000000-Dec-2000000003 Nuo, Vilas Boas The system does not meet ful specification for the performed service and is retuned to ue with as limitation accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: awaiting factory response Follow u Requred: awaiting factory response Problem Description by Engineer: Cutomer Fuction/Role: SIE Device ue: ou side of ue Eupected and ueupected Resuts: ufailue uer Impact: no impact Patient Impact: no impact SEn\a Diagnostic performed by Engineer: waits for factory response ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000018 200000000ov-2000000003 Nuo, Vilas Boas The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Waiting for factory report Problem Description by Engineer: Cutomer Fuction/Role: SIE Device ue: ou side of ue Eupected and ueupected Resuts: ufailue uer Impact: no impact Patient Impact: no impact SEn\a Diagnostic performed by Engineer: Waiting for factory report Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000008-23 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-11-11 16:01 | No
3767 | Yes | There was a message MPM application.eue | There was a message MPM application.eue |  | 2023-11-11 16:01 | No
3768 | Yes | The screen is dark at the top in comparison to the lower part | The screen is dark at the top in comparison to the lower part |  | 2023-11-11 16:06 | No
3769 | Yes | ouygen leak | ouygen leak |  | 2023-11-11 17:16 | No
3770 | Yes | Presenting Error | Presenting Error | BuQ&A Feedback: No SAP Cutomer Symptom Code: BOOT Cutomer Symptom Code: Boot u/Power-u Problem Reason: Corrective Maintenance Engineer Failue Code: Electrical/Electronic Failue Repair Activity Code: Instruted Cutomer SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022600000000500000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sthefany Pereira Last udated by: Sthefany Pereira 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT T2 Activities:Problem Description: Problem description by engineer : Commuication uLLwith the generator Trouleshooting Action: What steps are taken duing the system diagnosis : Checking the logs uLL# and/or description of uLL: Generator commuication failed Repair Action: ► After restart Twice the machine started 000000000000000000000000001king again; ► The client Fernanda saved a log and we analyzed it. The logs show that there was a power flutution and therefore the high voltage generator restarted, cauing commuication with the machine to fail. ► The analysis was passed on to the client and the procedue was euplained if it occured again. The machine mut be tuned off, wait for the power to stabilize and tun it on again; ► Left uder observation for 3 days and everything was normal. Cutomer Fernanda is aware that the call has been closed as resolved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK Test Equpment: N/A OneEMS internal remarks: Problem description by engineer : Generator commuication uLLWhat steps are taken duing the system diagnosis : Checking the logs uLL# and/or description of uLL: Generator commuication failed Repair Action "Euternal" : ► After restarting 2 times the machine started 000000000000000000000000001king again; ► The client Fernanda saved a log and we analyzed it. The logs show that there was a power flutution and therefore the high voltage generator restarted, cauing commuication with the machine to fail. ► The analysis was passed on to the client and the procedue was euplained if it occured again. The machine mut be tuned off, wait for the power to stabilize and tun it on again; ► Left uder observation for 3 days and everything was normal. Cutomer Fernanda is aware that the call has been closed as resolved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK Test Equpment: N/A Eutra Internal comment field: RSE MARcel Ogassawara Problem Reported by Cutomer: showing restricted uLLName: Fernanda Dassie Sector: clinical engineering 00¹: 27 998113129 00²: E-mail: Suervisorhmc00ngemedica.com.br HEMODINAMICA MERIDIONAL address: Ru Sao Joao Batista 200000000 Cariacica-ES 29156-200000000BR ----------------------------------------------------------------------------- | 2023-11-11 17:21 | No
3771 | Yes | Intrepid Monitor/Defibrillator CN00D-SV000 | Intrepid Monitor/Defibrillator CN000D-SV000 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 Problem Description by Engineer : Intrepid for repair / service Resoluion : Co2 pup calibrated , Op check passed. EST attached , uit tested and serviced to manuactuer's specifications. ----------------------------------------------------------------------------- | 2023-11-11 17:31 | No
3772 | Yes | Does not trigger and display is disabled defribilation | Does not trigger and display is disabled defribilation | BuQ&A Feedback: No Reason: Corrective Maintenance Trouleshooting Action: Qute sent Problem Reported by Cutomer: Equpment Problem: does not fire and display is disabled defribilation Search By: Schalon Seraf Sector: Eng 00¹: +551967598215 00²: ## schalon00otmail.com Priority: uit: HOSPITAL BELA VISTA Address: R. Antônio Carlos, 122 - Consolação, São Pauo - SP, Euternal Remarks: 21.11.2000000003 19:0000000029 POLARIS (POLARIS) ================= Gutavo de Assis Soua - 00000000000000005760000000077900000000100000000000000000000000038 =================== ==== DFM100000000 series: CN32600000000218 does not pass the operational test, fails the Ther apy Delivery Test. Part nuber change requred. Equpment not ok. Trouleshooting Action: Schedued for FSE Gutavo on 21-11 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 21.11.2000000003 19:0000000029 POLARIS (POLARIS) ================= Gutavo de Assis Soua - 00000000000000005760000000077900000000100000000000000000000000038 =================== ==== GG6: Operational tests approved. GG13: No calibratable tools were ued. Neut Step Action: Analysis carried ou and directed to sending a buget Safety Qustions: Case: 0000000022600000000200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jose Carlos Alves Sena Last udated by: Jose Carlos Alves Sena 1. The equpment was in clinical ue at the time the problem occured was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-11 19:06 | No
3773 | Yes | No sound. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. | No sound. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000297 Belen Juien The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: yes Eupected and ueupected Resuts: message no soud on the u3 uer Impact: no Patient Impact: no SE n Resoluion: replacement of the motherboard, tests OK, fuctional retun to service Parts Consued: 1.0000000000000000000000000000000056000000000000000081,Mu_100000000/u3 Main board ----------------------------------------------------------------------------- Internal Remarks: 0000000056000000000000000081 Mu_100000000/u3 Main board arrived ----------------------------------------------------------------------------- | 2023-11-11 19:11 | No
3774 | Yes | Low Leak-CO0ebreathing Risk (EC 00 | Low Leak-CO0Rebreathing Risk (EC 00) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000639 200000000ov-2000000003 Vishnu CR The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Checked the machine. Need to replace the Flow sensor assembly. Problem Description by Engineer : Cutomer Fuction/Role: Na Device ue: no Eupected and ueupected Resuts:na uer Impact:no Patient Impact: no SEna Resoluion : Replaced new RP-flow sensor assy, air & o2, V600000000 Now the machine 000000000000000000000000001king properly . Handed over the equpment to cutomer in good 000000000000000000000000001king condition. Parts Consued : 1.000000000000000000000000989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000Tools & Equpment : FLOW METER,2000000003-12-29 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000001000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000110000000000000000ov-2000000003 Vishnu CR The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : V600000000flow sensor assy air &o2- 989800000000617100000000Problem Description by Engineer : showing CO2 rebreathing uLL Problem Description by Engineer : Cutomer Fuction/Role: Na Device ue: no Eupected and ueupected Resuts:na uer Impact:no Patient Impact: no SEna Diagnostic performed by Engineer : Cross checked the GDS. Need to replace the air &O2 flow sensor assy. handed over the equpment to cutomer in initial condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 Nandha, Rahu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Main board to be ordered. Problem Description by Engineer : Cutomer Fuction/Role: Na Device ue: no Eupected and ueupected Resuts:na uer Impact:no Patient Impact: no SEna Diagnostic performed by Engineer : Main board to be ordered. ----------------------------------------------------------------------------- | 2023-11-11 19:51 | No
3775 | Yes | M0A / u00/ uLLtest failed euternal plates. | M00A / u000 / uLLtest failed euternal plates. | 000000000000000000000000001O##: 00000000000000000000000000100000000600000000915 28-Nov-2000000003 Ardighieri, Andrea The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: eng. clinical Device ue: uuutest Eupected and ueupected Resuts: failed euternal paddles uer Impact: no impact Patient Impact: no impact SE na Resoluion: The necessary replacement is no longer available as the defibrillator is ou of suport. System not repairable, ou of suport letter sent. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 28-Nov-2000000003 Peconio, Alberto The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: department Device ue: defibrillator test Eupected and ueupected Resuts: euternal paddles test failed uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Plate cable replaced bu the problem persists intermittently. uuuitor replacement is recommended Parts Consued: 1.0000000000000000000000000000000056300000000991,Paddle Repl. Water Resistant Kit Follow u Requred: Plate cable replaced bu the problem persists intermittently. uuuitor replacement is recommended ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000112 Ardighieri, Andrea The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: department Device ue: defibrillator test Eupected and ueupected Resuts: euternal paddles test failed uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Euternal plates failed test, device ou of suport, we proceed with checking the availability of spare parts. Follow u Requred: After consuting the coordinator, the device fails the euternal plate test. The test has been performed both with the plates inserted into the slots and with the test load, bu it often fails. We proceed with an attempted order of euternal plates and condenser, in case of uavailability it will be necessary to issu an EOS letter. ----------------------------------------------------------------------------- | 2023-11-11 20:01 | No
3776 | Yes | Artery alarms are automatically reset to off-alarm statu No Trouble Found-Evaluation Confirmed Device Working. | Artery alarms are automatically reset to off-alarm statu No Trouble Found-Evaluation Confirmed Device Working. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 27-Nov-2000000003 BASHIR, MouamadouSharif The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: no Patient Impact: no SE / Resoluion: Having not been able to observe a problem on site, we requsted a specific euample of the problem (time, patient, etc.) from the medical staff in order to better analyze it. The cutomer will come back to u when the problem arises again on D-7 in order to recover the logs on the control uit. Concerning training, the cutomer has already requsted training for himself as well as for the service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002900000000eva, Roman The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: no Patient Impact: no SE / Diagnostic performed by Engineer: Blood Pressue alarms auomatically stop when the patient's condition is ou of alarm. Requst for assistance to check the monitor's alarm settings and modify them if necessary. Wants the nusing staff to acknowledge the alarm and not for the alarm to acknowledge auomatically. Resoluion: ##sent to the cutomer with the setting and behavior of blocked alarms (auio blocked and/or visul blocked). The client is uder contract. He 000000000000000000000000001ld like the intervention of a technician so that he can check the alarm settings of the monitors with the service and also provide a reminder on the operation of the alarms (visuls blocked/Auio blocked, etc.). Follow u Requred: Requst for RR for on-site intervention. ATTENTION, see joint intervention with file n°0000000022665317 (same site/client). File n°000000002261200000000= Intervention in the service to check the alarm settings of the monitors as well as provide a reminder on the operation of the alarms (visuls blocked/Auio blocked, etc.) Check with the cutomer/department if they want auio blocked for red alarm or red & yellow alarm. Bu the organic wants the alarm not to be auomatically acknowledged when the patient is no longer in alarm condition. File n°0000000022665317 = Intervention for the replacement of the MSL connector of the rack. The cutomer already has the part. ----------------------------------------------------------------------------- | 2023-11-11 20:21 | No
3777 | Yes | AED does not turn on | AED does not turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000317 MESNIL, Mauime The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Reguatory Qustions: No/No Resoluion: The battery is less than 3 years old (installed in 2000000001 with the defibrillator) and the defibrillator is ued normally and was never ued on a patient. The battery seems defective. Ordered a new battery placed with the Cutomer Care Department. Diagnostic performed by Engineer: Futher to the replacement of the electrodes, the device does not start anymore, nothing happens futher to the insertion of the battery. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Device does not start. uer Impact: No. Patient Impact: No. SE N/A ----------------------------------------------------------------------------- | 2023-11-11 20:21 | No
3778 | Yes | Multifunction button faulty | 0 E0Multifunuon button faulty | 000000000000000000000000001O##: 000000000000000000000000001000000000000000021 100000000ov-2000000003 Nuo, Vilas Boas The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Front Panel & Selector Replacement Diagnostic performed by Engineer: Diagnostic, sumit qute Problem Description by Engineer: Cutomer Fuction/Role: Hélio Rodrigus +3512390000000000000000 Device ue: uknown Eupected and ueupected Resuts: Mutifuction buton fauty. uer Impact: uknown Patient Impact: uknown SE Not apply ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000952 Rau Cordero Calderón The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Faut: Mutifuction buton fauty Youwill need to plan an on-site visit to carry ou the repair. If youneed to requst 0000000056000000009261 - DFM100000000 ASSY Front Case 0000000056000000000000000051 - DFM100000000 ASSY-Label Set-Euopean Portuuse Diagnostic performed by Engineer: Faut: Mutifuction buton fauty Youwill need to plan an on-site visit to carry ou the repair. If youneed to requst 0000000056000000009261 - DFM100000000 ASSY Front Case 0000000056000000000000000051 - DFM100000000 ASSY-Label Set-Euopean Portuuse Problem Description by Engineer: Cutomer Fuction/Role: Hélio Rodrigus +3512390000000000000000 Device ue: uknown Eupected and ueupected Resuts: Mutifuction buton fauty. uer Impact: uknown Patient Impact: uknown SE Not apply ----------------------------------------------------------------------------- | 2023-11-11 20:41 | No
3779 | Yes | The FleuVision computer does not start automatically | The Fleusion computer does not start automatically | 000000000000000000000000001O##: 00000000000000000000000000100000000521887 28-Nov-2000000003 Garcia, Carlos Alberto The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedical Technician Device ue: Runing-mut restart Eupected and ueupected Resuts: Runing-mut restart uer Impact: Runing-mut restart Patient Impact: Runing-mut restart SE 8.1.300000000Diagnostic performed by Engineer: Fleuvision-PC replacement. Software loading and configuation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000858 28-Nov-2000000003 Magdy Maarou Khalil The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Medical Device ue: NO Eupected and ueupected Resuts: key screen does not respond correctly uer Impact: no Patient Impact: no SE 8.1.3000000005 Resoluion: I have changed the PC and installed the SW again. I have left the system runing and low observation. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000060000000013-Nov-2000000003 Magdy Maarou Khalil The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: install PC and SW Problem Description by Engineer: Cutomer Fuction/Role: Medical Device ue: NO Eupected and ueupected Resuts: key screen does not respond correctly uer Impact: no Patient Impact: no SE 8.1.3000000005 Reguatory Qustions: N/A Diagnostic performed by Engineer: I have investigated the PC, installed the Hard Drive and everything is 000000000000000000000000001king fine for the moment. ----------------------------------------------------------------------------- | 2023-11-11 21:11 | No
3780 | Yes | C-arch does not want to make roll scan | C-arch does not want to make roll scan | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008100000000os lubers The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: n.a Device ue: in ue Eupected and ueupected Resuts: uer Impact: n.a Patient Impact: n.a SE n.a martin versloot checked this duing maintenance, it 000000000000000000000000001ks as it shoud.27-11-2000000003 Diagnostic performed by Engineer: Martin Verloot checked this duing maintenance and it 000000000000000000000000001ks as it shoud.27-11-2000000003 ----------------------------------------------------------------------------- | 2023-11-11 21:51 | No
3781 | Yes | Email 000 Telemetry did not alarm in the euchange - GFE0sent 000 | Email 00 00Telemetry did not alarm in the euchange - GFE0 sent 0/0/0 | 000000000000000000000000001O##: 00000000000000000000000000100000000512900000000 00000000-Dec-2000000003 Carlos Manul Pinheiro The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: monitoring Center Eupected and ueupected Resuts: view all equpment uer Impact: name Patient Impact: None SEB.00000000.13 Resoluion: We were uable to reprodue the issu and detect the telemetry of the reported event. The central logs were removed for futher evalution. Some information is necessary for u to be able to follow u internally. The date and time and the telemetry 000are essential data for u to be able to assess whether there is a problem. For a more complete assessment, whenever a faut is detected, we recommend checking the telemetry and the control uit. If the telemetry is alarming? If this alarm is auible in telemetry and in the control panel. It is also advisable to see if this alarm is registered in the patient's alarm log. Tests were carried ou on the control uit and the alarms were checked. Operational equpment and meets factory specifications. ----------------------------------------------------------------------------- | 2023-11-11 22:06 | Unknown
3782 | Yes | Response documented/uCVu/Device freezes during studies caung corrupt images // CA query added 07 | Response documented/uCVu/Device freezes during studies caung corrupt images // CA query added 0/7 | Internal Remarks : Interaction Type: New Case Caller: Hawraa Aldoui +1857278100000000 IP & Issu: 960000000000000000Device freezes duing stuies cauing corrut images Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Adv of Holiday; Qute req created and adv of ##Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-12 01:21 | No
3783 | Yes | Displays message that motorized movement is not available | Displays message that motorized movement is not available | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000579 200000000ov-2000000003 Mauimu, Migul The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: TR Device ue: u Eupected and ueupected Resuts: no movement uer Impact: no Patient Impact: no SE 8.1 Parts Consued: 1.00000000000000000000000000000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS Resoluion: Mirror Case 0000000019000000009500000000- for consuption correction, with instrutions sent by suply. Equpment operating correctly ----------------------------------------------------------------------------- | 2023-11-12 02:06 | No
3784 | Yes | Foot pedal stopped working and have no uuuand no record_SPLIT_ | Foot pedal stopped working and have no uuuand no record_SPLIT_ | 000000000000000000000000001O##: 000000000000000000000000001000000000000000097000000005-Nov-2000000003 Orser Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: technologist Device ue: in ue Eupected and ueupected Resuts: intermittent loss of Geo movements uer Impact: na Patient Impact: na SE 8.1 Resoluion : Replaced damaged operator controls Parts Consued : 2.000000000000000000000000989600000000182673,TABLE CLAMP+BELLY BAR 1.00000000000000000000000000000000800000000151715,GEO MODuE WP Kit 1.000000000000000000000000000000008000000001517300000000MAGING MODuE KIT WP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000860000000000000000ov-2000000003 Orser Ryan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace Geo and imaging modues Problem Description by Engineer : Cutomer Fuction/Role: technologist Device ue: in ue Eupected and ueupected Resuts: intermittent loss of Geo movements uer Impact: na Patient Impact: na SE 8.1 Diagnostic performed by Engineer : Logs indicate emergency stop was pressed. Requsted movements not logged. Geometry and imaging modues are damaged and need to be replaced ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tara Grey 5000000006000000003800000000 IP & Issu: 60000000057662/Foot pedal stopped 000000000000000000000000001king and have no uuuand no record. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-12 03:36 | No
3785 | Yes | Left side foot pedal not working | Left side foot pedal not working | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 00000000-Dec-2000000003 MuH, JOE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad Techs / Doctor Device ue: Fluro pedal not 000000000000000000000000001king on the Wired Footswitch Eupected and ueupected Resuts: Ful Footswitch fuctionality uer Impact: Delay Patient Impact: None SE 8.2.300000000100000000Resoluion : No onsite 000000000000000000000000001k was performed on this 000000000000000000000000001 This 000000000000000000000000001contained part management only. Parts Consued : 1.00000000000000000000000000000000270000000000000000100000000Footswitch CV 3p 00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000010000000012-Nov-2000000003 MuH, JOE The system meets the specification for the performed service and is retuned to ue Follow u Requred : Part management Resoluion : Arrived onsite and confirmed the reported issu of the left side Fluro pedal not 000000000000000000000000001king on the Wired Footswitch. Ran diagnostics within FSF Service Mode; Fluro pedal Failed / Acqusition pedal Passed / One-Shot pedal Passed. Reviewed diagnostic logs to confirm nothing else was cauing the reported issu; Logs looked good. Replaced the Wired Footswitch with a brand new Wired Footswitch. Tested fuctionality of new part; All tests Passed. System fuly fuctional and retuned to service. Problem Description by Engineer : Cutomer Fuction/Role: Rad Techs / Doctor Device ue: Fluro pedal not 000000000000000000000000001king on the Wired Footswitch Eupected and ueupected Resuts: Ful Footswitch fuctionality uer Impact: Delay Patient Impact: None SE 8.2.300000000100000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: John IP & Issu: 85700000000957 left floral foot pedal not 000000000000000000000000001king Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason For Entitlement Claim N/A Action Taken: created case, sq's, onsite 000000000000000000000000001O#Neut Step Action: N/A Verbal PO Nuber: LIFECYCLE Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-12 07:01 | No
3786 | Yes | Operation was fine following yesterday's planned power cut, but today the power does not come on. | 0DS:AW: Operation was fine following yesterday's planned power cut, but today the power does not come on. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 13-Nov-2000000003 Ishikawa, SatoruThe system meets the specification for the performed service and is retuned to ue Resoluion: The cutomer is not there, so he cannot obtain a signatue.November 13, 2000000003 / I told him abou the 000000000000000000000000001k and he agreed. Reguatory Qustions: N/A Problem Description by Engineer: Operation was fine following yesterday's planned power cu, bu today the power does not come on. He said that the peripheral equpment, incluing the IW, is 000000000000000000000000001king well, bu he said that even when he responded following a planned power cu and tuned the power distribuion board off/on, it did not look like power was being suplied to the transformer. * He said that the IW is set u to acqure power from the wall. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: Operation was good following yesterday's planned power cu, bu today the power does not come on. He said that the peripheral equpment, incluing the IW, is 000000000000000000000000001king well, bu he said that even when he responded following a planned power cu and tuned the power distribuion board off/on, it did not look like power was being suplied to the transformer. * He said that the IW is set u to acqure power from the wall. uer Impact: Euamination canceled Patient Impact: No effect SE 2 Resoluion: The equpment power does not come on.When the M cabinet is opened all the LEDs tun off. A problem occured in the uit that controls the power suply (fan and power tray) for the entire system, which caued the problem. I replaced the uit and confirmed that the equpment power tuned on and that it was operating normally. Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS Tools & Equpment : MuTIMETER, 200000000000000003-17 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000723 Watanabe, Atsuhi The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: [Suject] 00000000DS:AW:u Kawamua: Operation was fine after yesterday's planned power ouage bu today the power can't be tuned on. [Internal Comments] Parts (no nearest domestic warehoue) Cutomer requst/complaint * I asked him to contact the person in charge at the facility again, bu he said that it was only the u so he asked that u look at it. Other: It is possible that there is a problem with the power suply in the hospital, and the part has not been comes ou of the equpment. The following items are not in stock in Tokyo warehoue. 0000000080000000000000000900000000 PDuFanTray and DCPS 0000000080000000000000000900000000 PDuFues Oupus [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [TetsuIshikawa] Ishikawa, Satoru - FSE1 Visit date and time 2000000003/11/1311:00000000 - FSE1 000000000000000000000000001k time 2 [Euamination category] Normal euamination Follow u Requred: * Problem Description by Engineer: Operation was fine following yesterday's planned power cu, bu today the power does not come on. He said that the peripheral equpment, incluing the IW, is 000000000000000000000000001king well, bu he said that even when he responded following a planned power cu and tuned the power distribuion board off/on, it did not look like power was being suplied to the transformer. * He said that the IW is set u to acqure power from the wall. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: Operation was good following yesterday's planned power cu, bu today the power does not come on. He said that the peripheral equpment, incluing the IW, is 000000000000000000000000001king well, bu he said that even when he responded following a planned power cu and tuned the power distribuion board off/on, it did not look like power was being suplied to the transformer. * He said that the IW is set u to acqure power from the wall. uer Impact: Euamination canceled Patient Impact: No effect SE 2 Diagnostic performed by Engineer: RC Response: Cold restart: OK Restarted from The Brainer: OK I euplained how to do it: OK Euplanation regarding failue locations: OK An FSE will respond on site. ----------------------------------------------------------------------------- | 2023-11-13 05:51 | No
3787 | Yes | System Not Booting up. | System Not Booting up. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 13-Nov-2000000003 Radan IVKOVIC The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : A cutomer called u to report the System boot-u issu. When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system Shoud Boot u to clinical Interface. uer Impact: cant ue the system Patient Impact: Not provided SE 8.2.17.00000000 System start u problem Resoluion : Check the bios battery on host pc and foud uder 2V. Replaced the battery with spare one, reset OS disk and ram memory slots. Tuned system off/on several times sucessfuly. Will follow u neut morning with cutomer. System ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000Lokeh Vinay The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : A cutomer called u to report the System boot-u issu. When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system Shoud Boot u to clinical Interface. uer Impact: cant ue the system Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer : Discused and confirmed the issu with the cutomer. A cutomer called u to advise that the system was not booting u, and the Data Monitor was not showing u any stages of the bootu sequnce. bu had some backlight and power to it. TSM was trying to initialize, bu was stopped at IP ping, Thorouhly asked the cutomer qustions abou boot-u. based on their euplanation, The caue was supected to be HOST PC. Got help from the cutomer to open u the M-cabinet and tried to Manully power u the Host Compuer, pressing the buton on the Host PC to manully tun that on 000000000000000000000000001ked, cutomer started the BIOS Screen with Americal megatrends. followed by some DOS commands which are part of the reguar boot-u sequnce and then booted u to Clinical Interface. based on this behavior of the host PC not starting u on its own i supect a dead or a low voltage Mother board battery as the supect. FSE will be requred to site to replace the 3V battery on the Host compuer. Internal Comments : uLL# and/or description of uLL: No Log files were available duing the time of the issu as the host Pc was not powered on. Repair Action : FSE requred to site: 1. Please try and replicate the issu and measue the voltage of the motherboard battery uing the voltmeter. 2. Please replace the 3V battery and then check the time and date mut be pass000000000000000000000000001d protected, pass000000000000000000000000001d: tpi00000000, he last digit I think is capital I for India bu coud be lower case l for lion please try both and find ou the correct one. Follow u Requred : uLL# and/or description of uLL: No Log files were available duing the time of the issu as the host Pc was not powered on. Repair Action : FSE requred to site: 1. Please try and replicate the issu and measue the voltage of the motherboard battery uing the voltmeter. 2. Please replace the 3V battery and then check the time and date mut be pass000000000000000000000000001d protected, pass000000000000000000000000001d: tpi00000000, he last digit I think is capital I for India bu coud be lower case l for lion please try both and find ou the correct one. Technician Name : Lokeh Vinay ----------------------------------------------------------------------------- | 2023-11-13 06:11 | No
3788 | Yes | He was not able to u the wireless foot switch to take a picture. | 0CS:SS: He was not able to u the wireless foot switch to take a piure. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000192 Saka, Shuji The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: He was not able to ue the wireless foot switch to take a pictue. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: He was not able to ue the wireless foot switch to take a pictue. uer Impact: No effect Patient Impact: No effect SE uknow Resoluion: RC Response: The foot switch with the cable is normal. It was possible to tun the wireless foot switch off/on. As it was the first problem, I promised to respond if it recus and finished the response for today Internal Comments: [Suject] 00000000S:SS:u Okua: He was not able to ue the wireless foot switch to take a pictue. [Euamination category] Other: ----------------------------------------------------------------------------- | 2023-11-13 07:21 | No
3789 | Yes | RR- System can't startup application | RR- System can't startup application | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000SANGBOONSONG MuDGROM The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Follow u Requred : System can't startu application system starting 00000000:00000000 Diagnostic performed by Engineer : sugest cutomer restart system Cutomer need FSE onsite Problem Description by Engineer : Cutomer Fuction/Role: Device ue: no Eupected and ueupected Resuts: System can't startu application system starting 00000000:00000000 uer Impact: no Patient Impact: no SE ----------------------------------------------------------------------------- | 2023-11-13 10:06 | No
3790 | Yes | Machine is not switching on | Machine is not switching on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000200000000Nov-2000000003 Viswanath Buti, Hemanth The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technicina Device ue: No Eupected and ueupected Resuts:NA uer Impact:No Patient Impact: No SE Resoluion : Replaced new HV cable with , carried ou conditioning , calibration passed , Level 1 passed , verification done, tested 000000000000000000000000001king, fuctional checked and handed over system in ful 000000000000000000000000001king condition Parts Consued : 1.0000000000000000000000000000000080000000038700000000,PD.SCOMP.PCM_DC200000000- PCM powersuply ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000237 Kuar, Vipin The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technicina Device ue: No Eupected and ueupected Resuts:NA uer Impact:No Patient Impact: No SE Diagnostic performed by Engineer : Checked and foud no power LEd on review modue uer is uable to open cabinet so need FSE visit system is down Follow u Requred : FSE to visit and check inpu suply . Check M cabinet fues and replace if its blown. check earthing as well. ----------------------------------------------------------------------------- | 2023-11-13 11:11 | No
3791 | Yes | The system startup stops with "0:0" displayed. | 0DS:TK: The system startup stops with "0:0" displayed. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000020000000013-Nov-2000000003 Sato, Taketo The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Reguatory Qustions: N/A Problem Description by Engineer: System startu stops at "00000000:00000000" display. ・u Fuuhima Eutension: 800000000 ・I restarted the system twice, bu there was no improvement. - Only a frame is displayed on the Fleuvision screen. - uuuodue displays "Call service". Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: System startu stops at "00000000:00000000" display. ・u Fuuhima Eutension: 800000000 ・I restarted the system twice, bu there was no improvement. - Only a frame is displayed on the Fleuvision screen. - uuuodue displays "Call service". uer Impact: No impact Patient Impact: No impact SE uknow Diagnostic performed by Engineer: [Requst details] System startu stops at "00000000:00000000" display. [000000000000000000000000001k details] We have confirmed the phenomenon. We have confirmed an uLLin the grabber card of the FleuvisionPC that controls the large monitor in the euamination room. To improve the problem, it will be necessary to replace the FleuvisionPC, bu since we confirmed that the problem was restored when the device was restarted, we decided not to replace it today. Curently, the operation of the PC to Fleuvision is ustable, so there is a possibility of commuication failue duing device startu. If youtun off the power to the device at that time, there is a possibility that it will not start u when yourestart it, so please do not tun off the power to the device while youare preparing for an inspection. If youeuuuence problems suh as no image being displayed on Fleuvision or freezing duing the inspection, please ue the backu monitor (Live, Ref) on the Fleuvision monitor to perform the inspection. Please be aware that in the event of a recurence, youwill need to replace you PC. [Other failues/measues] None [Equpment statu] Equpment uable [000000000000000000000000001k statu] 000000000000000000000000001k completed ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000192 Kamata, Takayui The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: [Suject] 00000000DS:TK: u Fuuhima: System startu stops at "00000000:00000000" display. [Internal Comments] COMP [Parts ou of stock] 0000000080000000012200000000:u Haswell Fleuvision 2 PC no HDD QTY:1 ZREP/Traceable Nagoya(u1000000000000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Nagoya(u100000000[Dispatch] 1 RC ----- ------------FSE registration information----------------- -Visit FSE1 Name [Shuto Sato] Sato, Taketo -FSE1 Visit date and time 2000000003/ 11/1316:00000000 -FSE1 000000000000000000000000001king time 3 [Shipping instrutions] -----------------Shipping instrutions #1------------------- -- -Delivery conditions CPu(no warehoue pick-u/transport arrangements) -Delivery date Mut arrive at 2000000003/11/13 15:15 -Delivery address: Nagoya warehoue/FSE picks u -Recipient name: Shuto Sato -0000000000000000007797-96200000000-Others -Client name: Takayui Kamata [Inspection category] Others Follow u Requred: Taketo Sato 16:00000000 3h Problem Description by Engineer: System startu stops at "00000000:00000000" display. ・u Fuuhima Eutension: 800000000 ・I restarted the system twice, bu there was no improvement. - Only a frame is displayed on the Fleuvision screen. - uuuodue displays "Call service". Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: System startu stops at "00000000:00000000" display. ・u Fuuhima Eutension: 800000000 ・I restarted the system twice, bu there was no improvement. - Only a frame is displayed on the Fleuvision screen. - uuuodue displays "Call service". uer Impact: No impact Patient Impact: No impact SE uknow Diagnostic performed by Engineer: - FleuvisionPC statu Pwr:on HD:off Lan1:flash Lan2:off - Tuning the PC off and on does not improve the problem. - Requsted response from FSE u Sato. ----------------------------------------------------------------------------- | 2023-11-13 11:31 | No
3792 | Yes | System Hang up | System Hang up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000331 13-Nov-2000000003 K.H lee The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Operator Device ue: not in ue Eupected and ueupected Resuts: system hangu uer Impact: no Patient Impact: no SE R8 Resoluion : Checked system logs. Replaced point. Implemented tue calibration. Fuctional check : pass. Parts Consued : 1.00000000000000000000000000000000210000000013887,Power Inverter (PoInt) Certeray Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 H.I ha The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : CLARITYIQ_II IP PC NO HDD & POWER INVERTER (POINT) CERTERAY change test Problem Description by Engineer : Cutomer Fuction/Role: Operator Device ue: not in ue Eupected and ueupected Resuts: system hangu uer Impact: no Patient Impact: no SE R8 Resoluion : CLARITYIQ_II IP PC NO HDD change ----------------------------------------------------------------------------- | 2023-11-13 12:01 | No
3793 | Yes | The monitor installed in the wall fell | 0DS:AW: The monitor installed in the wall fell | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000278 13-Nov-2000000003 Tsuinaka, Yamato The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiology Device ue: Yes Eupected and ueupected Resuts: Wall-mouted monitor fell uer Impact: None Patient Impact: None SE8.1 Resoluion: We have removed the wall-mouted monitor. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000275 Watanabe, Atsuhi The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: [Suject] 00000000DS:AW:u Nakatani: The monitor installed on the wall has come off. [Internal Comments] Others It is uclear whether the installation is managed by u.Details will be confirmed on-site. [Dispatch] 1 RC -----------------FSE Registration Information----------------- -Visit FSE1 Name [Yamato Tsuinaka ] Tsuinaka, Yamato -FSE1 Visit date and time 2000000003/11/13 17:00000000 -FSE1 000000000000000000000000001king time 2 [Inspection category] Other Diagnostic performed by Engineer: RC Response details: Euplanation of operation method: OK Euplanation of faut location identification: OK FSE will be on-site to confirm the sitution. Problem Description by Engineer: The monitor installed on the wall has come off. It is hanging from the cable, so there is no malfuction. It was installed on the wall near the entrance of the euamination room, and a separate u monitor was installed. Cutomer Fuction/Role: Radiology Technician Device ue: uder euamination Eupected and ueupected Resuts: The monitor installed on the wall has come off. It is hanging from the cable, so there is no malfuction. It was installed on the wall near the entrance of the euamination room, and a separate u monitor was installed. uer Impact: No impact Patient Impact: No impact SE 8.1.3000000005 ----------------------------------------------------------------------------- Internal Remarks: CFMPR0000000071500000000Pulished. ----------------------------------------------------------------------------- | 2023-11-13 13:11 | No
3794 | Yes | System cannot be started | System cannot be started | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000636 November 1000000002000000003 Maris, Ezra The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called u Meyer +0000000050000000060000000052 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: Determined that the buton of the hard drive in the host PC is defective. Informed the uer that the PC mut be replaced. Bu since the uer wanted to continu 000000000000000000000000001king, I euplained to him a temporary soluion with which he coud continu to 000000000000000000000000001k temporarily util the new PC is there. He implemented this soluion. Informed the client that the installation of the new PC will take 1.5 days since this PC is not the same as the old one and the entire software mut be reinstalled accordingly. Follow-u requred: Replacement of the host PC and hard drive incluing loading of new software ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called u Meyer +0000000050000000060000000052 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: System does not start correctly Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: TE service to replace OS hard drive, if necessary schedue entire host PC 3 hous of 000000000000000000000000001k If the compuer is not available, service only makes sense with hard drive. Technician name: Robert Kern ----------------------------------------------------------------------------- Internal Remarks: Dear Market, We have checked you critical order for part # 00000000800000000200000000612. Please be noted, there's no stock for this part in local RDC. Please refer below to MST and order the alternate part and cancel this one. In case of any qustions, please reach ou to IGT helpdesk. (see attachment) Cutomer comment: ## Appt. arranged with u Meyer for 1000000001. from approu. 9:00000000 am. Internal Remarks: Via Robert Kern: Hello Nils, The specified nuber is the curently installed 12 NC. I wrote that a service withou compuer makes sense. If you000000000000000000000000001ld like to order the sucessor, that is open to you According to SAP, youcan order the Haswell PC. With kind regards, Robert Kern Internal Remarks: Via u Gerdes: Please do not schedue the compuer replacement yet (follow-u) - the cutomers cannot make the system available for 1.5 days, he will clarify again with u Maris. Internal Remarks: According to u Gerdes, the compuer will not be replaced. Internal Remarks: uuEscalation 00000000800000000200000000612 Stakeholder Stefan Riu ----------------------------------------------------------------------------- | 2023-11-13 13:16 | No
3795 | Yes | Sometime can’t flu and flu delay. | Sometime can’t flu and flu delay. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000932 13-Nov-2000000003 ATIP PuKHAO The system meets the specification for the performed service and is retuned to ue Resoluion : Recreate image ippc Problem Description by Engineer : Cutomer Fuction/Role: cathlab Device ue: uray Eupected and ueupected Resuts: sometimes can’t fluand uray delay uer Impact: no Patient Impact: no SE8.1 ----------------------------------------------------------------------------- | 2023-11-13 13:41 | No
3796 | Yes | CONTACTEL 000 00 | CONTAuL 0/0/00000 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000817 00000000-Dec-2000000003 Jeremy CHATENET The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: »=Cutomer Fuction/Role: nuse or bioDevice ue: iuupected and ueupected Resuts: none uer Impact: nonePatient Impact: noneSE8.2 Resoluion: Replacement pedal and antenna Parts Consued: 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.0000000000000000000000000000000080000000000000000811,Pictogram Sheet Footswitch Family 1.00000000000000000000000030000000000000000660000000000000000,Pictogram Sheet Footswitch u1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Nicolas WAGNER The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Follow u Requred: CDP+HT Jeremy CHATENET 00000000/12/2000000003 10000000000000000 a.m. Please have it delivered to the biomedical department, for the u technician Jérémy Chatenet: wireless footswitch 3p kit 00000000800000000733000000001 qty: 1 pictogram sheet footswitch family 0000000080000000000000000811 qty: 1 Wireless base station (WBS) 00000000800000000257862 qty: 1 Problem Description by Engineer: »=Cutomer Fuction/Role: nuse or bioDevice ue: iuupected and ueupected Resuts: none uer Impact: nonePatient Impact: noneSE8.2 Diagnosis performed by Engineer: The wifi pedal is defective, restarting the room allowed it to pair with the base bu when the uer tried to do fluroscopy the wifi light came on. I order the necessary parts and schedue the local RT. ----------------------------------------------------------------------------- | 2023-11-13 13:51 | No
3797 | Yes | System does not start anymore | System does not start anymore | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 21-Nov-2000000003 Klein, Steffen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: uknown Eupected and ueupected Resuts: Eupected uer Impact: no Patient Impact: no SE n.a. Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion: Parts reported ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000996 21-Nov-2000000003 Klein, Steffen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : The uLLmessage is confirmed. no RR Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Resoluion: Settings made on FleuVision. HD display test checked ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000872 100000000ov-2000000003 Klein, Steffen The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: HD images are not displayed Follow u Requred: Reporting the PC Problem Description by Engineer: Problem description by engineer : The uLLmessage is confirmed. no RR Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: uuureplaced, layous restored. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000727 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer : The uLLmessage is confirmed. no RR Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: Pre-clarification by TSS M Kammerer System stops at 0000000000000000. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow u Requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: FSE S Klein is integrated throuh TSS and provides ET. Please plan for him with 3 hous of free time. Technician Name: Robert Kern ----------------------------------------------------------------------------- Cutomer comment: ## Appointment with Ms. Laur for November 21st. 11 a.m. agreed ## Internal Remarks: FSE Klein will contact youas soon as he has foud the compuer in Reulingen. Cutomer comment: Contact person: u Schmid Tel.: +000000001212000000001163 Internal Remarks: uu 0000000080000000012200000000 - u Haswell Fleuvision 2 PC no HDD Cutomer ETA: 16-NOV-2000000003 ----------------------------------------------------------------------------- | 2023-11-13 14:06 | No
3798 | Yes | Total failure system. | Total failure system. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000337 November 16, 2000000003 Simon, Aleuander The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MTRA Device ue: System down Eupected and ueupected Resuts: System Down uer Impact: - Patient Impact: - SE 8.2.3000000005 Tools & Equpment : DOSE-KVP-LIGHT METER, 2000000003-12-21 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000003-12-21 00000000:00000000:00000000 Resoluion: Host PC installation; ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000962 November 16, 2000000003 Simon, Aleuander The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MTRA Device ue: System down Eupected and ueupected Resuts: System down uer Impact: - Patient Impact: - SE 8.2.300000000Parts Consued: 1.00000000000000000000000000000000800000000122391,u Haswell Monoplane Host PC no HDD Diagnostic performed by Engineer: Host PC installation; ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000515 November 1000000002000000003 Brand, Dirk The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: System does not start uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Host PC does not start. PC replacement failed since the delivered hardware was not compatible. Renewed service with new PC requred, is already being planned. The state of the old PC is too ucertain to allow patient operation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000118 November 1000000002000000003 Kuchel, Oliver The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Ms. Olbrich +00000000221 0000000082000000001 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: The delivered compuer coud not be installed. A follow-u service was schedued Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000688 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Ms. Olbrich +00000000221 0000000082000000001 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: FSE K. Rödder has already pre-clarified with bulding services. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: Schedue TE service with host PC and hard drive, 2 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Internal Remarks: Also schedued u Brand Internal Remarks: Feedback from FSE, the service will probably be schedued all day since it will amout to a complete reinstallation. Please schedue in suh a way. Cutomer comment: ## Appointment info to Ms. Bayer, 15.11. from approu. 9:300000000# Internal Remarks: Via u Simon System is still at a stand still, he has to go to the cutomer again today. Internal Remarks: Please schedue new service with 1 u00000000800000000122391 llua Haswell Monoplane Host PC no HDD and 3 hous. Cutomer comment: ## Appointment info to uer. 1000000001 from approu. 9:00000000 am## ----------------------------------------------------------------------------- | 2023-11-13 14:46 | No
3799 | Yes | After normal FCO euecution, the equipment works normally. Two days after normal FCO euecution, At present, the device is | After normal FCO euecution, the equipment works normally. Two days after normal FCO euecution, At present, the device is |  | 2023-11-13 15:27 | No
3800 | Yes | machine is completely down | machine is completely down | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000129 100000000Dec-2000000003 Mahmou, Mari The system meets the specification for the performed service and is retuned to ue Resoluion : - Part retun 000000000000000000000000001oblem Description by Engineer : Problem description by engineer : Point voltage not available Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Point voltage not available uer Impact: None Patient Impact: None SE 8.2.17.00000000Diagnostic performed by Engineer : - Part retun 000000000000000000000000001--------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 22-Nov-2000000003 Mahmou, Mari The system meets the specification for the performed service and is retuned to ue Resoluion : - System Host-PC has been replaced. - System is handed in a good 000000000000000000000000001king condition. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Diagnostic performed by Engineer : - When the system is switched on, there is boot sequnce on the monitor. - Host-PC foud fauty, it's powered bu Hard disk power is down, hence no boot. - Host-PC needs replacement. Problem Description by Engineer : Problem description by engineer : complete system doesn't bootu How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000978 13-Nov-2000000003 Roble Ramanan The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Point voltage not available Problem Description by Engineer : Problem description by engineer : Point voltage not available Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Point voltage not available uer Impact: None Patient Impact: None SE 8.2.17.00000000Follow u Requred : uLL# and/or description of uLL: Point voltage not available Repair Action : mains voltage OK MIuvoltage OK Point voltage not available. supecting Point issu Technician Name : Roble Ramanan Internal Comments : uLL# and/or description of uLL: Point voltage not available Repair Action : mains voltage OK MIuvoltage OK Point voltage not available. supecting Point issu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000030000000013-Nov-2000000003 Jamil Hossain The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Host PC not booting , not power u Motherboard supected Follow u Requred : Host pc replacement and Sw download Problem Description by Engineer : Problem description by engineer : complete system doesn't bootu How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002700000000hmed Kotb The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uLL# and/or description of uLL: complete system doesn't bootu Repair Action : complete system doesn't bootu -- it seems that the host PC doesn't boot u .. an onsite visit is needed What are the neut Trouleshooting steps : complete system doesn't bootu -- it seems that the host PC doesn't boot u .. an onsite visit is needed Technician Name : Ahmed Kotb Resoluion : complete system doesn't bootu -- it seems that the host PC doesn't boot u .. an onsite visit is needed Internal Comments : uLL# and/or description of uLL: complete system doesn't bootu Repair Action : complete system doesn't bootu -- it seems that the host PC doesn't boot u .. an onsite visit is needed Problem Description by Engineer : Problem description by engineer : complete system doesn't bootu How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : complete system doesn't bootu -- it seems that the host PC doesn't boot u .. an onsite visit is needed ----------------------------------------------------------------------------- | 2023-11-13 15:31 | No
3801 | Yes | DOA HS0 device does not pass first Battery INsertion Test at cuomer, fault teut is: pads not uble | DOA HS0- device does not pass first Battery INsertion Test at cuomer, fault teut is: pads not uble |  | 2023-11-13 15:55 | No
3802 | Yes | It is damaged after a fall. | It is damaged after a fall. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006100000000November 15, 2000000003 José Lus Acero Rivas The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: The large fue holder fell off and it is damaged. An engineer needs to go and check it, see if it 000000000000000000000000001ks and see if it has repair or if it needs to be changed. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostics Eupected and ueupected Resuts: The large chassis buky is damaged after a fall. uer Impact: They cannot do breast euams. Patient Impact: They have to do large breasts in another mammography device. SE Not provided. Resoluion: 20000000000000000buky chassis holder is repaired. Verification of correct operation. Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000355 David Diaz Paniagu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The large fue holder fell off and it is damaged. An engineer needs to go and check it, see if it 000000000000000000000000001ks and see if it has repair or if it needs to be changed. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostics Eupected and ueupected Resuts: The large chassis buky is damaged after a fall. uer Impact: They cannot do breast euams. Patient Impact: They have to do large breasts in another mammography device. SE Not provided. Diagnostic performed by Engineer: The large fue holder fell off and it is damaged. An engineer needs to go and check it, see if it 000000000000000000000000001ks and see if it has repair or if it needs to be changed. Internal Comments: uLL# and/or description of uLL The large fue holder fell off and it is damaged. An engineer needs to go and check it, see if it 000000000000000000000000001ks and see if it has repair or if it needs to be changed. Repair Action: The large fue holder fell off and it is damaged. An engineer needs to go and check it, see if it 000000000000000000000000001ks and see if it has repair or if it needs to be changed. Follow u Requred: uLL# and/or description of uLL The large fue holder fell off and it is damaged. An engineer needs to go and check it, see if it 000000000000000000000000001ks and see if it has repair or if it needs to be changed. Repair Action: The large fue holder fell off and it is damaged. An engineer needs to go and check it, see if it 000000000000000000000000001ks and see if it has repair or if it needs to be changed. Technician Name: David Diaz Paniagu ----------------------------------------------------------------------------- | 2023-11-13 16:11 | No
3803 | Yes | PSA_PSE_the central scope alarm did not sound. Potential duplicate of PR# 00GFE0sent 0ec0 | PSA_PSE_the central scope alarm did not sound. Potential duplicate of PR# 000 GFE0 sent 0Dec03 |  | 2023-11-13 16:54 | Yes
3804 | Yes | PSA_PSE_ISCV Contains a bug which delaying archiving of images to PACS - Cuomer has reported a patient has passed away | PSA_PSE_ISCV Contains a bug which delaying archiving of images to PACS - Cuomer has reported a patient has passed away |  | 2023-11-13 17:16 | Yes
3805 | Yes | uPro ¦Damage to button overlay/front enclosure SPLIT | TempuPro ¦Damage to button overlay/front enclosure SPLIT |  | 2023-11-13 17:25 | No
3806 | Yes | Misalignment in the reacting position on the touch panel | Misalignment in the reaung position on the touch panel |  | 2023-11-13 17:32 | No
3807 | Yes | The device did not power on | The device did not power on |  | 2023-11-13 18:20 | No
3808 | Yes | The device booted on its own and they failed to power it off. - pending repair | The device booted on its own and they failed to power it off. - pending repair |  | 2023-11-13 18:35 | No
3809 | Yes | Battery failure | Battery failure | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003500000000ARAKAYA, Mehmet The system does not meet the ful specification for the performed service and is retuned to ue with limitation as accepted by the cutomer, no follow-u requsted. Resoluion: No uer-related failue was detected duing the failue detection. Accordingly; 989800000000190000000071 battery of one DFM100000000 Defibrillator device with part nuber is uder warranty Shoud be ordered within scope. Diagnostic performed by Engineer: Cutomer was contact regarding the relevant failue. Duing the inspection with the cutomer, the battery of the system Since it completed its service life, it was detected that it did not pass the test. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: Battery failue uer Impact: No Patient Impact: No SE 2.00000000 ----------------------------------------------------------------------------- | 2023-11-13 18:36 | No
3810 | Yes | "Sensor Failed"was displayed | "Sensor Failed" (Check Vent: Machine Pressure Sensor Auto-Zero Failed, Diagnosis Code 00) was displayed |  | 2023-11-13 18:56 | No
3811 | Yes | cannot hard reset , stuck on reset screen | 00CDT0cannot hard reset , stuck on reset screen | 000000000000000000000000001O##: 000000000000000000000000001000000000000000020000000013-Nov-2000000003 Autin Kenneth The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Stuk at 0000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE Not provided Resoluion : Replaced bad cmos battery in fleuvision pc. Now boots normally ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000039 Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Stuk at 0000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer : System stuk at 0000000000000000. Had Biomed shu down to breaker. Reset system and watch power u. System stuk at 0000000000000000 again. No layou bou on fleu monitor, so thinking is Fleuvision PC issu. Dispatching case Internal Comments : uLL# and/or description of uLL: Stuk at 0000000000000000 Repair Action : Verify issu. Possibly Flkeu PC issu. Repair as needed Possible parts: 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: Stuk at 0000000000000000 Repair Action : Verify issu. Possibly Flkeu PC issu. Repair as needed Possible parts: 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Mcnamara Pau ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: autin erickson IP & Issu: 6700000000726/cannot hard reset , stuk on reset screen Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech/ informed of cb Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-13 19:16 | No
3812 | Yes | They cannot deliver rays | They cannot deliver rays | 000000000000000000000000001O##: 000000000000000000000000001000000000000000088 17-Nov-2000000003 BRAN, ROBERTO The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : Cannot deliver beams with "image disk problem" uLLmessage Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion: Replacement of PC hard disk for image storage - IPPC (parts present in local stock). Fuctional tests performed with favorable resuts. The system is released for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006900000000TESTOLIN, DARIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : Cannot deliver beams with "image disk problem" uLLmessage Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer: 00contact made with u Enrico Bedini - some images at the time of switching on. Remote connection made - recreate image store - After discusion with Bran we plan onsite with IPPC - Internal Comments: uLL# and/or description of uLL: . Repair Action: 00contact made with u Enrico Bedini - some images at the time of switching on. Remote connection made - recreate image store - After discusion with Bran we plan onsite with IPPC - Follow u Requred: uLL# and/or description of uLL: . Repair Action: 00contact made with u Enrico Bedini - some images at the time of switching on. Remote connection made - recreate image store - After discusion with Bran we plan onsite with IPPC - What are the neut Trouleshooting steps: Onsite visit requred for part replacement Parts: 0000000080000000050000000092 IP PC Revised_II no HDD 00CC: 1 FSE 3h of 000000000000000000000000001k Technician Name: TESTOLIN, DARIO ----------------------------------------------------------------------------- Internal Remarks: I attach a cutomer requst for completing the application ----------------------------------------------------------------------------- | 2023-11-13 19:21 | No
3813 | Yes | QACASE_for PCI imported product#00000The non-invasive uLLin u has an alarm indicating | QACASE_for P0imported produu00000000The non-invasive uLLin u has an alarm indicating |  | 2023-11-13 19:31 | No
3814 | Yes | uLLMessage: Motor stalled - GFE 0/6 | uLLMessage: Motor stalled - GFE 00/6 | 000000000000000000000000001O##: 000000000000000000000000001000000000000000037 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer has a V600000000and the blower stalled. He verified with code 11300000000any times in the significant log. Recommended the blower, then the MC pcba Part #00000000561512761, blower assy Part #00000000561537531, MC pcba Cutomer will order a blower assy and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: blower stalled, code 11300000000 uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Dale Giddens Issu: 5660000000000000000 uLLMessage: Motor stalled Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-13 19:36 | No
3815 | Yes | u0- System changes the images from 0atient to another or it does not send them or they disappear | u0- System changes the images from 0patient to another or it does not send them or they disappear | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000November 1000000002000000003 Santiago Cases Lopez The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine. Device ue: Euam to patient. Eupected and ueupected Resuts: The acqured images were being saved in the previou patient. uer Impact: He had to start the euam again. Patient Impact: Nothing. SE 5.000000002 Diagnostic performed by Engineer: Software is reinstalled (v. 5.000000002.). After corrective maintenance, the equpment is operational and meets the manuactuer's specifications. The necessary verifications have been performed to certify that the equpment is back to the operating conditions prior to the failue. Parts Consued: 1.00000000000000000000000000000000562000000006221,5.000000002,SW,Affiniti700000000Application ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000085 November 13, 2000000003 Santiago Cases Lopez The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: udate software version. to 5.000000002. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine. Device ue: Euam to patient. Eupected and ueupected Resuts: The acqured images were being saved in the previou patient. uer Impact: He had to start the euam again. Patient Impact: Nothing. SE 1.5.2 Diagnostic performed by Engineer: The condition of the equpment is checked. Hardware tests are passed correctly. Checked the history of health alerts that affect the equpment and observed that the FCO79500000000529 corrected that uLL among others. The software version of the equpment mut be udated to version 5.000000002. ----------------------------------------------------------------------------- | 2023-11-13 19:46 | No
3816 | Yes | QACASE for PCI imported product#000000the equipment in the cath room suddenly failed | QACASE for P0imported produu00000000 the equipment in the cath room suddenly failed |  | 2023-11-13 19:47 | No
3817 | Yes | Equipment aborting shock | Equipment aborting shock | BuQ&A Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 00000000226189300000000afety Qustion Statu: Sumitted PSA: false PSE: tru PSF Statu: In Process Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? Yes 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Was the equpment defective or failed at the time of installation? uknown SAP Transfer: Dispatched for Onsite Suport Trouleshooting Action: schedued for FSE Cifone - 13/11 Problem Reported by Cutomer: Equpment problem: Equpment aborting shock Priority: STOPPED uit: HOSPITAL DAS CLINICAS DA FACuTADE DE MEDICINA DA uS P Address: Av Dr Eneas Carvalho de Aguar 255 Sao Pauo-SP 0000000000000000-900000000 Search for: Daiane Castro Eng clinic 11 26618500000000 engenhariaclinica.icr00c.fm.up.br SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-13 20:11 | Yes
3818 | Yes | Maintenance uLL. | Maintenance uLL. | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 29.11.2000000003 13:21:56 POLARIS (POLARIS) ================= Gutavo de Assis Soua - 00000000000000005765667500000000100000000000000000000000038 =================== ==== GG6: Operational tests approved. GG13: No calibratable tools were ued. SAP Transfer: Case Transferred to SAP Trouleshooting Action: OPP-198883 Qute: 00000000113676 ## engenharia.santana00ospitalsaocamilosp.org.br SAP Transfer: Case Transferred to SAP Euternal Remarks: 29.11.2000000003 13:21:55 POLARIS (POLARIS) ================= Gutavo de Assis Soua - 00000000000000005765667500000000100000000000000000000000038 =================== ==== uuuDefibrillator: u00000000579000000001 has "device uLL Maintenance" applied a shock to Qed.6 at 1500000000 and only reached 125J, which indicates that the uuuitor has valus ​​below eupectations. It will need to be replaced. Equpment not ok. Trouleshooting Action: Schedued for FSE Gutavo on 11/28 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: EQuPMENT SHOWING INTERMITTENT DEFIBRILLATION uLLDuING OPERATIONAL TEST EQuPMENT WITH MSG OF (DEVICE uLL MAINTENANCE) Neut Step Action: Analysis carried ou and directed to sending a buget Safety Qustions: Case: 0000000022615000000006 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Engenharia Clínca Santana Engenharia Clínca Santana Last udated by: Engenharia Clínca Santana Engenharia Clínca Santana 1. The equpment was in clinical ue in the when the problem was detected? #2. Was the patient or uer injued? No. 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? Yes 00000000Was the equpment defective or failed at the time of installation? At the Trouleshooting Action: Qute sent SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-13 20:16 | No
3819 | Yes | Lead shield hanger arm has elbow joint piece fell off | Lead shield hanger arm has elbow joint piece fell off | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 28-Nov-2000000003 KuTLY, MOHAMMAD The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: Plastic piece fell off the arm holding the lead shield. Small cover that did not hit anyone or anything. uer Impact: None Patient Impact: None uuResoluion : Replaced Mavig u-ray shield missing covers with new ones Parts Consued : 1.0000000000000000000000000000000080000000086100000000,Rear Cover 1.00000000000000000000000000000000800000000828731,Cover Eutension Arm ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000373 George Owens The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Repair or replace cover as needed. Diagnostic performed by Engineer : Debriefed caller. She 000000000000000000000000001ld like to let fse know a small plastic cover from one of the elbows on the arm holding the lead shield came off. Everything is still 000000000000000000000000001king normally, so when time allows they 000000000000000000000000001ld like it replaced or reinstalled. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: Plastic piece fell off the arm holding the lead shield. Small cover that did not hit anyone or anything. uer Impact: None Patient Impact: None uu----------------------------------------------------------------------------- | 2023-11-13 20:41 | No
3820 | Yes | PSA_SPLIT 0Medical device occasionally fails acquiring usignal. | PSA_SPLIT 00Medical device occasionally fails acquiring usignal. |  | 2023-11-13 20:45 | No
3821 | Yes | Medical device occasionally fails acquiring usignal. | Medical device occasionally fails acquiring usignal. |  | 2023-11-13 20:46 | No
3822 | Yes | QACASE for PCI imported product#00000_the table could not move | QACASE for P0imported produu00000000_the table could not move |  | 2023-11-13 20:47 | No
3823 | Yes | PSA_SPLIT 0Medical device occasionally fails acquiring usignal. | PSA_SPLIT 00Medical device occasionally fails acquiring usignal. |  | 2023-11-13 20:48 | No
3824 | Yes | SPLIT 0Medical device occasionally fails acquiring usignal. | SPLIT 00Medical device occasionally fails acquiring usignal. |  | 2023-11-13 20:49 | No
3825 | Yes | QACASE for PCI imported product#00000_The non-invasive uLLin u has an alarm indicating | QACASE for P0imported produu00000000_The non-invasive uLLin u has an alarm indicating |  | 2023-11-13 21:07 | No
3826 | Yes | QACASE for PCI imported product#000000he non-invasive uLLin u alarms and indicates "main ala | QACASE for P0imported produu00000000 The non-invasive uLLin u alarms and indicates "main ala |  | 2023-11-13 21:14 | No
3827 | Yes | M0A - u00- it does not work on battery despite changing a new battery | u- it does not work on battery despite changing a new battery |  | 2023-11-13 21:15 | No
3828 | Yes | AED is constantly in self-test loop | AED is constantly in self-test loop | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000857 Hendrik Gritzka The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Nemo: Recipient/## BLIV: No Other: Please provide a gurantee, as the case was reported on 2000000008.2000000003 in Case 0000000022179163. Gurantee went u to 26.00000000.2000000003 Stck/uuAT: 1u 0000000056000000005311 FRu (C B) Defib, German 00000000, Euch in AT Problem Description by Engineer: CCutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: Device in self-test loop uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: Device is in self-test loop. Specialist dealer reported on 2000000008.2000000003 in case 0000000022179163. Gurantee went u to 26.00000000.2000000003 ----------------------------------------------------------------------------- Internal Remarks: 00uteAdmin: Please grant a warranty or 100000000% discout becaue the report was made within the gurantee, bu the case coud not be processed du to a blocked accout. Cutomer comment: Continu with Case 0000000022752700000000. This case is closed becaue it was opened incorrectly. ----------------------------------------------------------------------------- | 2023-11-13 21:31 | No
3829 | Yes | uCVu- - u0 Probe doesn't articulate with hand controls while in u / NTF / Initial Complete | uCVu- - u00 Probe doesn't articulate with hand controls while in u / NTF / Initial Complete | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000000000005-Nov-2000000003 Williams Pau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : u8-2t manul controls are not 000000000000000000000000001king, 000000000000000000000000001't articuate the tip When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: for tip to articuate normally uer Impact: uable to move the tip of the probe with the controls Patient Impact: Not provided SE Not provided Diagnostic performed by Engineer : uable to verify probe articuated, change degree from 00000000to 1800000000vise versa cutomer will monitor ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Wiggins, Aaron The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: none Repair Action : RSE trouleshooting: Cut advised that the probe is imaging fine bu 000000000000000000000000001't move the tip with the controls on the handle. Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : uLL# and/or description of uLL: none Repair Action : RSE trouleshooting: Cut advised that the probe is imaging fine bu 000000000000000000000000001't move the tip with the controls on the handle. What are the neut Trouleshooting steps : Evalute the probe Replace if needed Cut will provide SN# of the probe to the FSE. Probe was in ue at the time of the call. 0000000056180000000077 - Transduer, u8-2t Technician Name : Wiggins, Aaron Problem Description by Engineer : Problem description by engineer : u8-2t manul controls are not 000000000000000000000000001king, 000000000000000000000000001't articuate the tip When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: for tip to articuate normally uer Impact: uable to move the tip of the probe with the controls Patient Impact: Not provided SE Not provided Diagnostic performed by Engineer : Cut advised that the probe is imaging fine bu 000000000000000000000000001't move the tip with the controls on the handle. ----------------------------------------------------------------------------- | 2023-11-13 21:51 | No
3830 | Yes | Connector problem | Conneur problem | 000000000000000000000000001O##: 000000000000000000000000001000000000000000025 Gullaue PANET The system does not meet the specification for the performed service and is not retuned to ue Resoluion: When tuning on the device, the cutomer follows the device's instrutions. “apply the electrodes” “plu in the connector” The cutomer plus in the connector bu the device gives the same auible indications all the time. We did a battery insertion test. Regarding what the device indicates: Self-test passed. Pads ued: M3713A Eupiration date: 2000000002-10000000000000000The cutomer will connect a new valid set of electrodes and test again. Resoluion: Hello, I did a test with another valid set of electrodes. That didn't solve the problem. Kind regards, Thomas Lemoine Resoluion: Hello, Thank youfor you feedback. 000000000000000000000000001ld it be possible for youto do a test with a simuator? The idea 000000000000000000000000001ld be to simuate a patient with this tester and see if the device passes the stage of connecting the electrodes. In this case, this involves “sacrificing” a set of electrodes or taking a ued one and puting banana plus on the end to create a test cable. Kind regards, Gullaue PANET Resoluion: Hello, I did the test that youasked me, the device still does not detect the connector. Kind regards, Thomas Lemoine Resoluion: u LEMOINE, In this case, the device has a problem identifying the connection of the electrodes. ufortuately, the device is uuable and cannot be repaired or replaced. I then invite youto contact ou AVF distribuor who will be able to direct youtowards another AED uY. 000000000000000000000000001ld youlike to receive an obsolescence ## Kind regards, Gullaue PANET Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient ue Eupected and ueupected Resuts: uknown uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the defibrillator. ----------------------------------------------------------------------------- Internal Remarks: The client wishes to speak to Gullaue PANET ----------------------------------------------------------------------------- | 2023-11-13 22:36 | No
3831 | Yes | shut down on its own and came back on by itself/ uLL 00. | shut down on its own and came back on by itself/ uLL 00.. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001800000000Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient Eupected and ueupected Resuts: RT reports uit shu down then powered back u on its own. uer Impact: swapped vent Patient Impact: vent was swapped with no patient issus SE uknown Resoluion : Cutomer has ran the uit for several hous with no issus. code 1100000000 and 3000000007 logged. Reviewed manul and advised that no futher action is requred. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: TIM ELLIS IP & Issu: shu down on its own and came back on by itself Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-13 22:36 | No
3832 | Yes | Failing prouimal pressure sensor auto zero, GFE00EC0*** | Failing prouimal pressure sensor auto zero, GFE0 0DEC00** | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to replace Sol 3 and 00000000long with seals to address uLLcode, provided part nubers and cutomer will order, no futher follow u needed, issu resolved- Solenoid valve (puge, auozero, crossover) (GDS) 00000000561511911 (Sol 3 and 00000000quntity 2) Seal, solenoid valve, package of 00000000manifold) 00000000561511700000000Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:Prouimal pressue sensor auozero failed alarm uer Impact: removed from service Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states while in ue the uit alarms Prouimal pressue sensor auozero failed alarm on screen, uit continud to ventilate patient bu was removed from ue withou incident, no patient harm reported, in biomed noted uLL1100000000 in log, no recent components replaced, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Chad Ekern IP & Issu: Failing prouimal pressue sensor auo zero Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-13 22:41 | No
3833 | Yes | Recording System is not booting up | Recording System is not booting up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 Vilkaukas Viktorija The system meets the specification for the performed service and is retuned to ue Resoluion : Tyler called back and spoke to front line agent Brandon Floyd. Issu is resolved. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: Not provided by cutomer Eupected and ueupected Resuts: System to boot u/System does not boot u uer Impact: System does not boot u Patient Impact: Not provided by cutomer SE 1.2.5 Diagnostic performed by Engineer : Tyler is reporting that the recording system is not booting u. Du to contract miu u, cutomer was called on 11/13 at 100000000m and left a voic## Also called and left a voic##message for Ray Kates, who is normally and onsite biomed who takes care of the equpment. udated management of the activity for this case. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: cut callback Caller: Tyler Harris IP & Issu: Recording System is not booting u, 781500000000200000000Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: sent AA and messaged RSE on teams to let them know issu is resolved and case shoud be closed per cut requst. uable to close R000000000000000000000000001in progress. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-13 22:41 | No
3834 | Yes | Device is switching between ac power to internal power simultaneouy | Device is switching between ac power to internal power simultaneouy | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000383 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer had an eutra power cord and installed it and it solved the issus. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and it switches between AC and battery. Cutomer was shaking the power cord and the uit was swtiching from AC to battery and back and forth. It coud be a fauty power cord. Part #98980000000061300000000, power cord. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit is switching from AC to battery uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Walter Long IP & Issu: 86512928 Device is switching between ac power to internal power simutaneouly. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-13 22:51 | No
3835 | Yes | Split case : visible defect on the connection of the pad cable( See Reference PR: 00)-Need GFE confirm split | Split case : visible defeuon the conneuon of the pad cable( See Reference PR: 00 )-Need GFE confirm split |  | 2023-11-13 23:13 | No
3836 | Yes | SOMETIMES IT DOES NOT PASS THE CHECK-UP (split from 00 | SOMETIMES IT DOES NOT PASS THE CHECK-UP (split from 00) |  | 2023-11-13 23:27 | No
3837 | Yes | System will not power on | 00CDN0System will not power on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000367 15-Nov-2000000003 Mandeville Nathan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : system will not power on, biomed confirmed uS u and power present; no blinking lights on control modue after holding power on buton How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: eupected system to power on; system will not power on uer Impact: System is down Patient Impact: No futher information is available as per the cutomer SE 8.1.17 Resoluion : System not powering on. Foud defective mpd control board. After board replacement, system powers on again. Parts Consued : 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000778 15-Nov-2000000003 Mandeville Nathan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace MPDucontroller Problem Description by Engineer : Problem description by engineer : system will not power on, biomed confirmed uS u and power present; no blinking lights on control modue after holding power on buton How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: eupected system to power on; system will not power on uer Impact: System is down Patient Impact: No futher information is available as per the cutomer SE 8.1.17 Diagnostic performed by Engineer : System is not powering on. Neither the review modue or the MPDubutons 000000000000000000000000001k. Diagnosed the MPDucontroller as defective. Ordering replacement. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000685 Van Pelt Daniel The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : system will not power on, biomed confirmed uS u and power present; no blinking lights on control modue after holding power on buton How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: eupected system to power on; system will not power on uer Impact: System is down Patient Impact: No futher information is available as per the cutomer SE 8.1.17 Diagnostic performed by Engineer : Discused problem with biomed, confirmed no lights blinking on control modue; system us is on and power present to system Internal Comments : uLL# and/or description of uLL: system will not power on Repair Action : Check if shunt trip has been activated. Check mains breaker in the MPDu If power LEDs at MPDuflash, check cable between MDY-u100000000nd NPu300000000. If no LED, check fue NP-F100000000 Possible requred parts: 00000000230000000016000000007 - MPD CONTROL uIT 00000000800000000000000001500000000- PFS272 POWERTRAY Follow u Requred : uLL# and/or description of uLL: system will not power on Repair Action : Check if shunt trip has been activated. Check mains breaker in the MPDu If power LEDs at MPDuflash, check cable between MDY-u100000000nd NPu300000000. If no LED, check fue NP-F100000000 Possible requred parts: 00000000230000000016000000007 - MPD CONTROL uIT 00000000800000000000000001500000000- PFS272 POWERTRAY Technician Name : Van Pelt Daniel ----------------------------------------------------------------------------- Internal Remarks : 0000000022620000000061 Mandeville Nathan FSE 000000000000000000000000001king on u 8.1 where system 000000000000000000000000001t stay on. FSE has trouleshot to MPDucontrol not keeping system on. Gave FSE ok to order the MPDucontrol board. ----------------------------------------------------------------------------- | 2023-11-13 23:31 | No
3838 | Yes | Image disk full | Image disk full | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000668 Chhlam, Sy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - IPu Free disk space is too low. 13-Nov-2000000003 00000000:00000000:15 RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 13-Nov-2000000003 00000000:18:00000000 When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : Mutiple times per week Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: image disk ful reboot did not resolve uer Impact: cannot make new images Patient Impact: patient moved to a different room to complete euam SE 000Diagnostic performed by Engineer : John reported that the image disk is ful and they are not able to make new images or do a fluro save even if the all the euams were deleted. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Free disk space is too low.: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Root Caue for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1. The Host PC (TH) is uable to commuicate with the Frontal IPPC (TP). Either the IPPC did not boot u du to lack of (Mains) power, a defective BIOS battery or the OS/AS did not boot u in-time. Repair Action : Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1.1. Check the back of the IPPC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs in the Trouleshooting manul of the IPPC on InCenter. 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related. Then euamine the ethernet wiring (THu12 - TPu100000000 between Host PC and IPPC and perform Cold reboot of the system. 1.3. If LED statu is NOT Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IPPC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct. 1.5. A defective BIOS battery coud be a reason why the IPPC does not boot. Then replace the Frontal IPPC (TP). Dialed into the system throuh PRS and performed a database inconsistency check and repair and then a recreate image store. After a reboot John was able to make new images. Resoluion : System operational after reboot and we agreed to keep case open util 11/15/23 and close if no call backs. ----------------------------------------------------------------------------- | 2023-11-13 23:36 | No
3839 | Yes | SPLIT-charge battery is broken-See Reference PR 00 | SPLIT-charge battery is broken-See Reference PR 001 |  | 2023-11-13 23:54 | No
3840 | Yes | Wireless foot switch is intermittently not working | 00CDT0Wireless foot switch is intermittently not working | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000961 18-Nov-2000000003 Powers Tim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - DTC: Command response: No EN_u. 13-Nov-2000000003 00000000:58:00000000 When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Wireless foot switch intermittently not releasing euposue uer Impact: Have to press wireless foot switch pedals mutiple times Patient Impact: No impact uu00000000Resoluion : Replaced intermittent wireless footswitch. Paired footswitch with base uit. Verified operation of footswitch. Performed Flouo and Cine. System retued to clinical ue. Parts Consued : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000030000000000000000660000000000000000,Pictogram Sheet Footswitch uTools & Equpment : MuTIMETER,2000000005-00000000-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000755 Sean Kiernan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - DTC: Command response: No EN_u.: 1. Possible issu with footswitch, possible too muh delay in activation of the 2 internal switches in footpedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the footswitch are not switching in correct sequnce/in time wrt each other. Root Caue for RMT - DTC: Command response: No EN_u.: 1. Enable u-ray is disabled bu footswitch is pressed. This coud be du to pedal tapping. Repair Action : Service actions for RMT - DTC: Command response: No EN_u.: 1.1. Youcan try to adjut the mechanical timing of the microswitches ('Command' and 'EN_u signal' ) in the footpedal with the adjutable screws at the backside. 1.2. Replace footswitch Service actions for RMT - DTC: Command response: No EN_u.: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. 00000000800000000238232 - Wireless Footswitch 3P Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - DTC: Command response: No EN_u.: 1. Possible issu with footswitch, possible too muh delay in activation of the 2 internal switches in footpedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the footswitch are not switching in correct sequnce/in time wrt each other. Root Caue for RMT - DTC: Command response: No EN_u.: 1. Enable u-ray is disabled bu footswitch is pressed. This coud be du to pedal tapping. Repair Action : Service actions for RMT - DTC: Command response: No EN_u.: 1.1. Youcan try to adjut the mechanical timing of the microswitches ('Command' and 'EN_u signal' ) in the footpedal with the adjutable screws at the backside. 1.2. Replace footswitch Service actions for RMT - DTC: Command response: No EN_u.: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. 00000000800000000238232 - Wireless Footswitch 3P Technician Name : Sean Kiernan Diagnostic performed by Engineer : uer called and reported that the wireless foot switch is having intermittent issus where they press the pedal and it does not release radiation. Puled sytsem log files and foud several instances of the EN_u signal not being active even thouh the pedal is pressed. Will recommend replacement of the wireless foot switch du to curent known issus with the foot switches and according to uLL from the log files. Problem Description by Engineer : Problem description by engineer : RMT - DTC: Command response: No EN_u. 13-Nov-2000000003 00000000:58:00000000 When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Wireless foot switch intermittently not releasing euposue uer Impact: Have to press wireless foot switch pedals mutiple times Patient Impact: No impact uu00000000----------------------------------------------------------------------------- | 2023-11-14 00:01 | No
3841 | Yes | prouimal pressure auto sensor failed uLLmessage | prouimal pressure auto sensor failed uLLmessage | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: prouimal pressue sensor auozero failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it has a prouimal pressue sensor auozero failed. He stated there is a code 1100000000 in the significant log. 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA. Part #00000000561511911, solenoid. Sol 3 and sol 00000000 You000000000000000000000000001ld need to bu 2. Part # 00000000561512561, DA pcba Resoluion : Cutomer will order solenoids first, then a DA pcba to repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Bryce Kegebein IP & Issu: " 820000000070000000000000000& prouimanol pressue auo sensor failed uLLmessage" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created r000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-14 01:01 | No
3842 | Yes | PSA & PSE_ur electrification. Bench repair pending | PSA & PSE_ur eleuification. Bench repair pending |  | 2023-11-14 01:51 | Yes
3843 | Yes | Capno Not Working | Capno Not Working | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000732 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confiimed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits CO2 connector is physically damaged. Additionally, the uits charging port is physically damaged. To correct both issus the Capnometer Assembly will need to be replaced and the Front Case Assembly will be ordered. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 000000005600000000000000001-Capnometer 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000077 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uuProblem Description/ueupected Resuts: biomed reports the CO2 port is broken, sending to bench Follow u Requred : Problem Description: biomed reports the CO2 port is broken, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form Diagnostic performed by Engineer : Problem Description: biomed reports the CO2 port is broken, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form ----------------------------------------------------------------------------- | 2023-11-14 02:01 | No
3844 | Yes | Genarator is bu. Can not make u-ray. | Genarator is bu. Can not make u-ray. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000619 13-Nov-2000000003 Ramirez, Carlos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Ryan Chapman Device ue: in ue Eupected and ueupected Resuts: establish u-ray uer Impact: none Patient Impact: no patient involved SE2.2 Resoluion : Actions taken: measued oupu voltages and phase 2 and 3 not present. Verified all fues were good. Foud F6 in 500000000m fue F6 fue was arching. Resoluion: tightened all sockets and assued no short circut visible. System repaired sucessfuly. No futher service requred. No patient involved ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000Sharvari Karua The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : The generator is not booting u. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Per the logs, the MIuis in a faut state. Internal Comments : uLL# and/or description of uLL: 200000000SuGC0000000030000000000000000From Generator device: MIu Phase faut Repair Action : 00000000210000000019971 - Mains Interface uit (MIu Certeray Resoluion : Coud not resolve the issu and requres onsite suport. Check the mains voltage suply on frontal MIuu30000000000000000 Check to see if one or more phases are missing. If the power is present. Replace the MIu ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action Taken: FSE Carlos Ramirez accepted call for 11/1300:00000000called cu spoke to Jonny Snachez and confirmed Oucome::dispatched ----------------------------------------------------------------------------- | 2023-11-14 02:01 | No
3845 | Yes | PSA & PSE_ur electrification. Bench repair pending | PSA & PSE_ur eleuification. Bench repair pending |  | 2023-11-14 02:25 | Yes
3846 | Yes | Touchscreen issue - Pending part replacement per GFE 05 | Touchscreen issue - Pending part replacement per GFE 0/5 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000169 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: duing PM service Eupected and ueupected Resuts: touhscreen issus uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order a touhscreen and repair the uit “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and it has an issu with the touhscreen. Duing PM service, they noticed the uit 000000000000000000000000001ld not respond to the touh. Calibration did not solve the issu either. part #00000000561539000000001, touhscreen ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Dewayne Logan IP & Issu: 60000000000000000200000000utomer Requsting Suport - Parts ID. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-14 03:31 | No
3847 | Yes | unable to uray, uLLtube problem | 00CDT0nable to uray, uLLtube problem | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 200000000ov-2000000003 Shula Chintan The system meets the specification for the performed service and is retuned to ue Resoluion : Assisted PFSE with repair. Problem Description by Engineer : Cutomer Fuction/Role: in-houe imaging tech Device ue: uknown Eupected and ueupected Resuts: not able to produe u-ray. uer Impact: uknown Patient Impact: uknown SE2.2.3 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000257 16-Nov-2000000003 Weir, Matthew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal 13-Nov-2000000003 16:30000000028 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: uable to uray, uLLtue problem uer Impact: uable to make uray Patient Impact: case aborted/patient moved to another room uu00000000Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000210000000020000000000000000,IPS Certeray R5 Resoluion : Replaced generator internal power suply. Tested system. The system is now fuctioning properly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 19-Nov-2000000003 Weir, Matthew The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: CANT u RAY Device ue: NO Eupected and ueupected Resuts: CANT u RAY uer Impact: CANT uE SYSTEM Patient Impact: NONE uuTools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 Diagnostic performed by Engineer : Trouleshoot system and order parts. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001600000000etcherside Jeff The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal 13-Nov-2000000003 16:30000000028 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: uable to uray, uLLtue problem uer Impact: uable to make uray Patient Impact: case aborted/patient moved to another room uu00000000Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal: 1. If the system is switched On too quckly after a switch Off, the Certeray does not start u correctly sometimes. Root Caue for RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal: 1. The Sute PC coud not open or lost connection with the Frontal Certeray generator. Repair Action : Service actions for RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal: 1.1. Instrut cutomer to reboot system and neut time wait at least 2 minues to power-u after a switch Off. Service actions for RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal: 1.1. Check 2300000000ac Fues F1, F2, F3 on PDu(NY) in M-cabinet. 1.2. Check u00000000 (192.68.32.3) of u-segment Controller is connected to u3, port of Ethernet Switch (TS) . 1.3. Coud be caued du to a defective u-Ray tue Cooling uit (CLu. Check statu of the CLu In some cases it was caued du to a bad K3 relay. This relay is not latching when power is applied to the cooler. (case ref: 000000001985000000008000000001.00000000Check power to generator and whether it is switched-On at startu of the system. 1.5. uR<2.1: Euamine if 200000000is present in the generator. Check if LED 'LVPS ou' is lit on MIu If not then replace MIu uRâ¥ 2.1: Euamine if 200000000is present in the generator. Check if LED H3 of the Cooling uit (CLu is lit, indicating 200000000is present. If not then replace IPS. IPS Certeray R5 00000000210000000020000000000000000 uSC Certeray 00000000210000000020000000091 Technician Name : Ketcherside Jeff Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal: 1. If the system is switched On too quckly after a switch Off, the Certeray does not start u correctly sometimes. Root Caue for RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal: 1. The Sute PC coud not open or lost connection with the Frontal Certeray generator. Repair Action : Service actions for RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal: 1.1. Instrut cutomer to reboot system and neut time wait at least 2 minues to power-u after a switch Off. Service actions for RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal: 1.1. Check 2300000000ac Fues F1, F2, F3 on PDu(NY) in M-cabinet. 1.2. Check u00000000 (192.68.32.3) of u-segment Controller is connected to u3, port of Ethernet Switch (TS) . 1.3. Coud be caued du to a defective u-Ray tue Cooling uit (CLu. Check statu of the CLu In some cases it was caued du to a bad K3 relay. This relay is not latching when power is applied to the cooler. (case ref: 000000001985000000008000000001.00000000Check power to generator and whether it is switched-On at startu of the system. 1.5. uR<2.1: Euamine if 200000000is present in the generator. Check if LED 'LVPS ou' is lit on MIu If not then replace MIu uRâ¥ 2.1: Euamine if 200000000is present in the generator. Check if LED H3 of the Cooling uit (CLu is lit, indicating 200000000is present. If not then replace IPS. IPS Certeray R5 00000000210000000020000000000000000 uSC Certeray 00000000210000000020000000091 Diagnostic performed by Engineer : cutomer states uable to uray, uLLtue problem. tried to reboot the system bu still have the issu, moving the patient to another room. collected logfile from the system for analysis ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Callback Caller: Solomon Tekabe IP & Issu: wanted to confirm if case went to technical suport before being dispacthed onsite Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: held for 11-100000000ntitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Matt Weir - no commuication between uSC and TS Managed Ethernet Switch. Replacing net000000000000000000000000001k cable did not resolve the issu. Sugested replacing uSC. Internal Remarks : Interaction Type: Cutomer Call Back Caller: Betsy Devlin IP & Issu: " 93277231" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Reached ou to FSE Matt, spoke to DSM informed he is 000000000000000000000000001king on it. called client back with udate Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : checked log with fse and see uLLwith generator commuication 200000000 missing replace IPS having fse order IPS ----------------------------------------------------------------------------- | 2023-11-14 03:41 | No
3848 | Yes | Generator Bu, u-ray not possible | 00CDN0Generator Bu, u-ray not possible | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 22-Nov-2000000003 Marshall Taru The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Arcing at Hivo Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion : Replaced Hivo, Hivo cables and point Parts Consued : nul,98000000000000000063000000002,H.V. cable O3-O3, 300000000m nul,98000000000000000063000000002,H.V. cable O3-O3, 300000000m 1.00000000000000000000000000000000210000000013887,Power Inverter (PoInt) Certeray 1.00000000000000000000000000000000210000000011828,HiVo high voltage transformer 1.000000000000000000000000132253787800000000POWERSIL PASTE AP, 5 G ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000323 16-Nov-2000000003 Todd Charles The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System is giving an uLLin the middle of euposues. Generator is buy, ,no u-ray is possible. RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1 13-Nov-2000000003 1000000008:00000000 RMT - uGEN: Frontal - u-ray Tue Arcing detected. u13-Nov-2000000003 1000000008:16 When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : Intermittent Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System is giving an uLLin the middle of euposues. Generator is buy, ,no u-ray is possible. uer Impact: wait a bit to retry to get u-ray Patient Impact: Not provided SE 1.2.5.00000000Resoluion : Assisted FSE. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000729 17-Nov-2000000003 Ha Hoang The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End uer Device ue: Generator Eupected and ueupected Resuts: No uray uer Impact: na Patient Impact: na SE 1.2 Resoluion : Assisted FSEs ru new hi tension cables. Installed Hivo and Point. Calibrated system. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000030000000017-Nov-2000000003 Compas Mike The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: ou side Eupected and ueupected Resuts: N/A uer Impact: none Patient Impact: none SE Resoluion : Assist with puling HT cables ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 100000000ov-2000000003 Todd Charles The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace parts. Problem Description by Engineer : Problem description by engineer : System is giving an uLLin the middle of euposues. Generator is buy, ,no u-ray is possible. RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1 13-Nov-2000000003 1000000008:00000000 RMT - uGEN: Frontal - u-ray Tue Arcing detected. u13-Nov-2000000003 1000000008:16 When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : Intermittent Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System is giving an uLLin the middle of euposues. Generator is buy, ,no u-ray is possible. uer Impact: wait a bit to retry to get u-ray Patient Impact: Not provided SE 1.2.5.00000000Diagnostic performed by Engineer : Foud arcing on Anode HT stick on HT tank side, HT cable and HV tank need replacement. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 Clark, Byron The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1. It coud be this message occus as collateral uLLafter u-ray Tue arcing uLL/warnings. Root Caue for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1. MIu Short circut on rail voltage oupu to PoInt Root Caue for RMT - uGEN: Frontal - u-ray Tue Arcing detected. u 1. The Certeray generator detected u-ray tue arcings on the frontal channel. Repair Action : Service actions for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1.1. If it occus only occasionally youcan ignore it. Else see soluion 2. 1.2. Inspect HT candle at Cathode (-) and Anode (+) sides of u-ray tue: If black bun marks are noticed, then: - Clean HT candles uing isopropyl alcohol, Check its pins uing pin gage and apply silicon paste as per Service Manul Corrective Maintenance 000000002-98000000006700000000REV AC - "Gude Faut Finding on u-Ray Tue Assemblies u CV and SRM" - Perform tue conditioning & adaptation. - Perform tue yield calibration and EDL calibration - Observe whether the problem is resolved in the uLLog. 1.3. If no bun marks on HT candles and problem still persists replace PoInt of frontal generator. Service actions for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1.1. Replace the PoInt of the frontal u-ray generator Service actions for RMT - uGEN: Frontal - u-ray Tue Arcing detected. u 1.1. Inspect HT candle at Cathode (-) and Anode (+) sides of u-ray tue: If black bun marks are noticed, then: - Clean HT candles uing isopropyl alcohol, Check its pins uing pin gage and apply silicon paste as per recommendation in "u-Ray Generator - Known Issus - Gude FautFinding u-Ray tues u" docuent. 1.2. If tue arcing occus often, please carry ou u-ray Tue Conditioning and then carry ou u-ray Tue Adaptation. Problem Description by Engineer : Problem description by engineer : System is giving an uLLin the middle of euposues. Generator is buy, ,no u-ray is possible. RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1 13-Nov-2000000003 1000000008:00000000 RMT - uGEN: Frontal - u-ray Tue Arcing detected. u13-Nov-2000000003 1000000008:16 When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : Intermittent Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System is giving an uLLin the middle of euposues. Generator is buy, ,no u-ray is possible. uer Impact: wait a bit to retry to get u-ray Patient Impact: Not provided SE 1.2.5.00000000Diagnostic performed by Engineer : collected logs and discused description of issu Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1. It coud be this message occus as collateral uLLafter u-ray Tue arcing uLL/warnings. Root Caue for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1. MIu Short circut on rail voltage oupu to PoInt Root Caue for RMT - uGEN: Frontal - u-ray Tue Arcing detected. u 1. The Certeray generator detected u-ray tue arcings on the frontal channel. Repair Action : Service actions for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1.1. If it occus only occasionally youcan ignore it. Else see soluion 2. 1.2. Inspect HT candle at Cathode (-) and Anode (+) sides of u-ray tue: If black bun marks are noticed, then: - Clean HT candles uing isopropyl alcohol, Check its pins uing pin gage and apply silicon paste as per Service Manul Corrective Maintenance 000000002-98000000006700000000REV AC - "Gude Faut Finding on u-Ray Tue Assemblies u CV and SRM" - Perform tue conditioning & adaptation. - Perform tue yield calibration and EDL calibration - Observe whether the problem is resolved in the uLLog. 1.3. If no bun marks on HT candles and problem still persists replace PoInt of frontal generator. Service actions for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1.1. Replace the PoInt of the frontal u-ray generator Service actions for RMT - uGEN: Frontal - u-ray Tue Arcing detected. u 1.1. Inspect HT candle at Cathode (-) and Anode (+) sides of u-ray tue: If black bun marks are noticed, then: - Clean HT candles uing isopropyl alcohol, Check its pins uing pin gage and apply silicon paste as per recommendation in "u-Ray Generator - Known Issus - Gude FautFinding u-Ray tues u" docuent. 1.2. If tue arcing occus often, please carry ou u-ray Tue Conditioning and then carry ou u-ray Tue Adaptation. Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1. It coud be this message occus as collateral uLLafter u-ray Tue arcing uLL/warnings. Root Caue for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1. MIu Short circut on rail voltage oupu to PoInt Root Caue for RMT - uGEN: Frontal - u-ray Tue Arcing detected. u 1. The Certeray generator detected u-ray tue arcings on the frontal channel. Repair Action : Service actions for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1.1. If it occus only occasionally youcan ignore it. Else see soluion 2. 1.2. Inspect HT candle at Cathode (-) and Anode (+) sides of u-ray tue: If black bun marks are noticed, then: - Clean HT candles uing isopropyl alcohol, Check its pins uing pin gage and apply silicon paste as per Service Manul Corrective Maintenance 000000002-98000000006700000000REV AC - "Gude Faut Finding on u-Ray Tue Assemblies u CV and SRM" - Perform tue conditioning & adaptation. - Perform tue yield calibration and EDL calibration - Observe whether the problem is resolved in the uLLog. 1.3. If no bun marks on HT candles and problem still persists replace PoInt of frontal generator. Service actions for RMT - MIu Short circut on rail voltage oupu to PoInt - Frontal_1: 1.1. Replace the PoInt of the frontal u-ray generator Service actions for RMT - uGEN: Frontal - u-ray Tue Arcing detected. u 1.1. Inspect HT candle at Cathode (-) and Anode (+) sides of u-ray tue: If black bun marks are noticed, then: - Clean HT candles uing isopropyl alcohol, Check its pins uing pin gage and apply silicon paste as per recommendation in "u-Ray Generator - Known Issus - Gude FautFinding u-Ray tues u" docuent. 1.2. If tue arcing occus often, please carry ou u-ray Tue Conditioning and then carry ou u-ray Tue Adaptation. Technician Name : Clark, Byron ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE with parts ID and TPR approval for damaged HIVO and HV cables from arcing. ----------------------------------------------------------------------------- | 2023-11-14 04:21 | No
3849 | Yes | speaker malfunction uLL = Training GGB_GFE sent 0ec_FU_0ec | speaker malfunuon uLL = Training GGB_GFE sent 0Dec_FU_0Dec | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000532 Sanchez, Karbin The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed called to report that the uit presented a "speaker malfuction" message. The message was gone when biomed power cycled the u3. Biomed asked whether or not there are known issus with the u3 speaker, and I let him know that I will need to research that and get back to him ASAP. Device ue: u3 Eupected and ueupected Resuts: Biomed called to report that the uit presented a "speaker malfuction" message. The message was gone when biomed power cycled the u3. Biomed asked whether or not there are known issus with the u3 speaker, and I let him know that I will need to research that and get back to him ASAP. uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer : Biomed called to report that the uit presented a "speaker malfuction" message. The message was gone when biomed power cycled the u3. Biomed asked whether or not there are known issus with the u3 speaker, and I let him know that I will need to research that and get back to him ASAP. Diagnostic performed by Engineer : The following ##to the cutomer sumarizes the findings Hi Tyru, I looked within ou system and there are no known issus with the speaker assembly. My sugestion is to test and monitor the uit in the shop, bu it is recommended to replace the speaker assembly before uing it on the floor. 0000000056000000003831 Mu_100000000/u3 Speaker Assembly Regards ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tyru Cabrera IP & Issu: speaker malfuction uLL 963000000006 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: free remote Reason for Entitlement Claim: N/A Action Taken: created case, R000000000000000000000000001Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-14 05:11 | No
3850 | Yes | uuudefective, GFE00OV0, Pending quote as of 0OV0*** | uuudefeu, GFE0 0NOV00 Pending quote as of 0NOV00** |  | 2023-11-14 05:51 | No
3851 | Yes | Unable to restart | Unable to restart | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002261500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: can't start system Trouleshooting Action: Key faut codes or phenomena: Client cannot provide any information PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Arrive at the scene to check the condition of the equpment 2. Check whether the power suply is normal, host system, etc. SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: uable to restart Internal Remarks: 13.11.2000000003 00000000:56:13 uC (860000000095300000000 $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem. The device fails to start.) $Q8: (Check that the circut insuance, commuication lines, and power suply of the device are normal. Download the log data and find no obviou euceptions that affect the device startu. Need to keep watching.) $Q9: (No maintenance action is performed, continu to observe ue.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: (N/A) ^Curent Tue installed date: (N/A) ^Curent Tue uage: (N/A) Euternal Remarks: 13.11.2000000003 00000000:56:100000000uC (860000000095300000000 Service content: The equpment reported by the cutomer for repair cannot be started. After arriving at the site, the equpment has been sucessfuly started and pu into ue. Preventive maintenance of the equpment is carried ou according to the cutomer's requrements. On-site inspections inclue normal circut insuance, normal commuication lines, and normal equpment power suply. Downloading log data has not foud any impact on the equpment. Obviou abnormalities in startu are juged to be accidental. Service resuts: The equpment is normal and fuly fuctional. ----------------------------------------------------------------------------- | 2023-11-14 05:56 | No
3852 | Yes | System startup failed | System startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 13.11.2000000003 00000000:0000000052 uC (86000000008321)"SuT. Ready for Clinical ue local time" : " 2000000003-11-13 16:00000000" SAP Transfer: Case Transferred to SAP Euternal Remarks: 1000000001.2000000003 00000000:20000000000000000uC (86000000008321) 000000000000000000000000001k content: The machine cannot be tuned on, there is a host problem, reconnect. Service resuts: After repair, the equpment retuned to normal operation. Other matters: none SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: 00000000N/u3 M15 Safety Qustions: Case: 000000002261500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Qingfeng Gu Last udated by: Qingfeng Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 1000000001.2000000003 00000000:200000000500000000uC (86000000008321) $Q7: (Confirmed reported problem. Cannot bootu) $Q8:（System cannot bootu, check sute PC, foud cannot power on) $Q9: (Re-connected sute PC and then check system fuction) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: NA SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: System startu fails when booting Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The screen has a backgroud light bu no display Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended that an engineer come on site to check the 000000000000000000000000001king statu of the Fantray DCPS.0000000080000000000000000900000000 ----------------------------------------------------------------------------- | 2023-11-14 06:01 | No
3853 | Yes | uuudefective | uuudefeu |  | 2023-11-14 06:06 | No
3854 | Yes | Teal UPS is off | Teal UPS is off | 000000000000000000000000001O##: 000000000000000000000000001000000000000000023 22-Nov-2000000003 Ramirez, Carlos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Nelson Figaroa Device ue: in ue Eupected and ueupected Resuts: replacement of Fleu PC uer Impact: uknown Patient Impact: no patient impact SE8.1.300000000Resoluion : Actions taken: verified logs and determined generator was not commuicating. Measued mains inpu to ENF1 and measued zero voltage, measued uS breakers and fue. Foud breaker to generator to have been tripped and temps in the cabinets high. Reestablished power to generator and requsted ac to cabinet he restored, Tested system with captuing u-ray. No futher service requsted. No patient involved ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 17-Nov-2000000003 Ramirez, Carlos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Nelson Figaroa Device ue: in ue Eupected and ueupected Resuts: replacement of Fleu PC uer Impact: uknown Patient Impact: no patient impact SE8.1.300000000Resoluion : Actions taken: replacement of Fleuvision PC was sucessfu. No futher service requred. A follow u to install Ru Monitoring Follow u Requred : Install Ru Monitoring Parts Consued : 1.0000000000000000000000000000000020000000000000000000000001,FRuFleu-PC EP/HD (with snapshot) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000856 200000000Nov-2000000003 Ramirez, Carlos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Teal uS is off How often is issu occuring? : Continuuly Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided Resoluion : Performed parts replacement and tested sucessfuly. No futher service requsted Parts Consued : 1.0000000000000000000000000000000020000000000000000000000001,FRuFleu-PC EP/HD (with snapshot) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 16-Nov-2000000003 Ramirez, Carlos The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Nelson Figaroa Device ue: in ue Eupected and ueupected Resuts: replacement of Fleu PC uer Impact: uknown Patient Impact: no patient impact SE8.1.300000000Resoluion : No resoluion: Actions taken/ PDufoud tripped and futher diagnosis is requred. A follow u activity has been created. Follow u Requred : Continu with PDuDiagnostic to power . ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 100000000ov-2000000003 Ramos, Anthony The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : continu trouleshooting Fleu PC Problem Description by Engineer : Problem description by engineer : Teal uS is off How often is issu occuring? : Continuuly Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer : uon arrival confirmed no power to system or uS. foud wall breaker was tripped. Cutomer stated there was a power ouage the day before. reset wall breaker and power was restored bu system was uable to fuly boot u . trouleshoot fleu PC . Engineer will follow u neut day to continu trouleshooting . ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 100000000ov-2000000003 Ward Randall The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Teal uS is off How often is issu occuring? : Continuuly Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer : 000000000000000000000000001ked with Anthony Ramos and facilities to trouleshoot no 00000000 inpu to the system. Foud that the mains 00000000V breaker was tripped and not providing power to the system. After resetting the breaker, the system will not completely boot u. There is still possibly a problem with the FleuVision PC that will requre additional trouleshooting. Follow u Requred : Trouleshooting no power to system ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000861 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Teal uS is off How often is issu occuring? : Continuuly Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer : Nelson says the PDuin the equpment room can't be bypassed with the rocker switch. The PDudoes not show u in the asset listing bu is likely the PDu0000000000000000 Rocker switch is no longer available. Internal Comments : uLL# and/or description of uLL: no uLL Repair Action : PDu0000000000000000Follow u Requred : uLL# and/or description of uLL: no uLL Repair Action : PDu0000000000000000 What are the neut Trouleshooting steps : Determine what modle of PDuand if it can be configued bypassed. May need to be coverted to LifeSoluion 000000000000000000000000001O# Technician Name : Kuis Peter ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action Taken: Follow u activity for FSE Carlos Ramirez for 11/18000000000 Oucome: dispatched Internal Remarks : 0000000022621817 Carlos Ramirez FSE calling in on u FD2000000008.1.30000000015 where system is stuk at 00000000:00000000. FSE trouleshot down to bad Fleu PC. FSE replaced and FSE says it 000000000000000000000000001ld intermittently come u or get stuk at 00000000:00000000, when it did come u review live 000000000000000000000000001ld flicker. FSE sent pictue of front of PC and one of the triangle warning indicators was amber. Also FSE sent pictue of back of PC and a red led was on between MSB/LSB. Looks like a DOA Fleu pc. FSE lukily had another on order and swapped ou and all is good now. FSE is testing and it looks like it u an runing. FSE switched off and on nuerou times. Internal Remarks : Approved to order replacement 0000000020000000000000000000000001 - FRuFleu-PC EP/HD (with snapshot) du to Fleu-PC not booting u after power failue at site. Internal Remarks : 0000000022621817 Carlos Ramirez FSE calling in on u FD2000000008.1.3000000005. FSE asking abou correct Fleu PC ordered. FSE also asking abou PDu000000000000000where FSE says the main breaker was tripped. FSE want to be able to say PDuis good. Advised FSE to tun off power to PDu check connections on top, open u bottom and inspect. Also to pu on a power monitor. Verified Fleu PC single link for FSE. FSE is also ordering a hard drive. FSE has a plan of action. Internal Remarks : Interaction Type: Internal Caller: Carlos IP & Issu: case 0000000022621817 NSS Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? N/A Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: transferred to IuR NSS +3118615 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Planner Initials: GM Action Taken: Follow u activity for FSE Carlos Ramirez for 11/17002pm Oucome: dispatched Internal Remarks : Planner Initials: GM Action Taken: FSE Carlos Ramirez requsted an assist for 11/10000000000000000M Oucome:: dispatched Internal Remarks : Interaction Type: Cutomer Call Back Caller: Nelson Figuroa IP & Issu: 51000000000000000000000000 called to check when service will be dispatched Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: NA Action Taken: sent alert in RP chat Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Caller: Lauel Kempt 3100000000773 17800000000P & Issu: CCB needs an udate on what was done to the uit and if the uit is ready for ue Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Posted in RP for some one to contact cutomer with an udate Action Taken: N/A Neut Step Action: Internal Remarks : Planner Initials: GM Action Taken: DSM George Becerra requsted call open and dispatched for FSE Carlos Ramirez for 11/1600AM Oucome: dispatched ----------------------------------------------------------------------------- | 2023-11-14 06:31 | No
3855 | Yes | The stand/table does not move. | [Q] 0CS:TK: The stand/table does not move. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000936 Kamata, Takayui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: The stand/table does not move. u Souou He restarted it, bu that does not improve the sitution. The message "Buton active..." is displayed. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The stand/table does not move. u Souou He restarted it, bu that does not improve the sitution. The message "Buton active..." is displayed. uer Impact: No effect Patient Impact: No effect uu00000000Resoluion: ・As the device was on the Controller in the Control Room, I had him remove it and that improved the sitution. ----------------------------------------------------------------------------- | 2023-11-14 06:41 | No
3856 | Yes | Tube current out of range | Tube current out of range | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: No euposue, high voltage startu is displayed SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Repair Action: ppd/989000000000000000086700000000u 200000000 + 000000007 ROT-GS 10000000000000000980000000060000000081 IGT Service Suport Tool/200000000to hospital/site1 Safety Qustions: Case: 0000000022615600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problem: tue curent ou of range Trouleshooting Action: Key faut code or phenomenon: [CHANNEL_INFO=u-Ray Channel Frontal][Info=[EuceptionId: 12116][MeasuedCurent[u]: 60000000025][ LoweruLLimit[u]: 6100000000][LowerWarningLimit[u]: 900000000200000000[uperWarningLimit[u]: 10000000028][uperuLLimit[u]: 176633][Focu: Smallest][Grid: WithGrid][TueVoltageSet: 70000000000000000][TueCurentSet: 0000000011900000000000000001000000001916][EuposueTimeSet: 0000000000000000873200000000000000006585331] PRS TR Statu: All-pass Repair Action: Recommended On-site maintenance steps: 1. Go to the site for filament test, and recommend spare parts uuand name based on the resuts: "ppd/ 989000000000000000086700000000u 200000000 + 000000007 ROT-GS 10000000000000000CONTRACT 0000000080000000060000000081 IGT Service Suport Tool/200000000to hospital/site1" Internal Remarks: 22.11.2000000003 00000000:50000000052 uC (86000000009000000005) $Q7:(Confirmed reported problem. The failue occured in Generation incluing tue. no u-ray. ) $Q8:(. check log, find uLLtue curent ou of range. Doing tue conditioning and adaptation does not 000000000000000000000000001k. Need replace the tue) $Q9:(Replaced u 200000000 + 000000007 ROT-GS. Performed perform replace the tue T2 sugestion effective.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11:$Calibrated Tools name (u2 RF Dectector, u2 Base uit, earth bonding) S/N (218000000007,257966,000000003000000002) Neut Calibration date (00000000/31/2000000005,00000000/31/2000000008,10000000010000000020000000000000000Neut Step Action: PPD9890000000000000000867000000001 Arrange shipment after approval FSE Liuuusen-86-1785310000000025 SAP Transfer: Case Transferred to SAP Euternal Remarks: 22.11.2000000003 00000000:500000000500000000uC (86000000009000000005) Service content: uable to go ou, tue curent euceeds limit, tue conditioning and tue adaptation are invalid.On-site replacement of the tue. Service resut: The equpment operates normally after repair. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-14 07:06 | No
3857 | Yes | Cannot startup_SPLIT | Cannot startup_SPLIT | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Neut Step Action: uable to boot, and there are many fauts at the same time, requring engineers to come to repair Safety Qustions: Case: 0000000022615616 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: Can't boot Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: uable to boot, and there are many fauts at the same time, requring engineers to come to repair Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Arrive to inspect PC Notes on cutomer requrements: Please contact the cutomer for on-site repair Problem Reported by Cutomer: Limited dispatch requres confirmation SS Ye Jinlong 00000000850000000015622 called to report repairs Neut Step Action: Need to confirm the dispatch of 000000000000000000000000001kers Internal Remarks: 1000000001.2000000003 13:0000000026 uC (86000000008266) RDI Accuacy: Yes $Q7: (Confirmation & GEO can't move and no message reported on the monitor) $Q8: (Reviewed the uLLlog, foud uLLlog abou AKD ;Performed visul inspection and it display A00000000and the normal is A2 & maybe AKD is defect ) $Q9: (reconnect the cable between AKD and GEO IPC, and it's ok then reconnect the cables of AKD then it retun normal and restart several times and system is ok, T2 sugestion effective.) $Q100000000 (Test and inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: (No Calibrated Tools ued) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 1000000001.2000000003 13:0000000000000000 uC (86000000008266) Service content: [The bed cannot move mechanically. After checking, the motor controller uder the bed does not commuicate normally. It retuned to normal after re-pluging and upluging the connection cable] Service resuts: [After service, the overall equpment operates normally] Other important matters: [None] ----------------------------------------------------------------------------- | 2023-11-14 07:06 | No
3858 | Yes | The pacing function test does not pass. | The pa0g funuon test does not pass. | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: Instruted Cutomer Neut Step Action: Langfang Fouth People’s Hospital Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Problem Reported by Cutomer: 00000000/HeartStart uL Defibrillator/Monitor/ SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: DESCRIPTION OF PROBLEM: The pacing fuction test not pass HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? na Trouleshooting Action: NOTE uLL# AND / OR DESCRIPTION OF uLL na WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Gude cutomer to ue euternal cable and test load re-test Repair Action: WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Gude cutomer to ue euternal cable and test load re-test RECOMMENDED REPAIR ACTION: The pacing test failed + it is recommended that the cutomer ue euternal electrode cables and connect the test load to reru the test After remote service, the equpment resues normal operation Test/Inspection: na Test Equpment: na Safety Qustions: Case: 0000000022615621 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-14 07:06 | No
3859 | Yes | Unable to take images due to low free space | Unable to take images due to low free space |  | 2023-11-14 08:11 | No
3860 | Yes | AED fails self-tests, "Not ready to u" uLLmessage with new battery and pads | AED fails self-tests, "Not ready to u" uLLmessage with new battery and pads | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003700000000dward Carson The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer swapped pads and battery however still not 000000000000000000000000001king. Faut to be with defib itself. uit not uder warranty anymore. Cutomer will need to puchase Defib Problem Description by Engineer : Cutomer Fuction/Role: Private practice Device ue: Ouside of ue Eupected and ueupected Resuts: As reported by cutomer uer Impact: uable to ue defib Patient Impact: uable to ue defib SE Not requred ----------------------------------------------------------------------------- | 2023-11-14 08:26 | No
3861 | Yes | Startup failure | Startup failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problem: uable to boot Cutomer problem description or faut phenomenon: System crash is caued by the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: n/a Remote diagnosis and test resuts: N/ A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Reinstall the system Cutomer demand notes: Please contact the cutomer for on-site repair Problem Reported by Cutomer: uable to boot Safety Qustions: Case: 000000002261600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 15.11.2000000003 1000000002:39 uC (8600000000900000000) Service content: Reinstall Host PC operating system and software, reconfigue equpment and calibrate. Service resuts: The device is powered on normally and the fluroscopic euposue is normal. Euternal Remarks: 13.11.2000000003 1000000003:52 uC (8600000000900000000) "SuT. Ready for Clinical ue local time" : " 2000000003-11-13 22:00000000" SAP Transfer: Case Transferred to SAP Internal Remarks: 15.11.2000000003 1000000002:00000000uC (8600000000900000000) $Q7: (Confirm the report problem that the device startu failtue) $Q8: (the OS and application is broken, need to reinstall) $Q9: ( 1. reinstall the OS and APP, configue the system. Device retuns to ful service. ) $Q100000000 Test and inspection data refer to PA. Device retuns to ful service. $Q11: Calibrated Tools:N/A. SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-14 08:36 | No
3862 | Yes | u-Ray not ready | u-Ray not ready | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000353 21-Nov-2000000003 WARANON KOTIRAM The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad tech Device ue: CV Eupected and ueupected Resuts: PDuMains Interface Modue Defective. uer Impact: u-Ray not ready. Patient Impact: No Impact. SE2.1 Tools & Equpment : MuTIMETER,200000000000000003-15 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuMains Interface Modue Diagnostic performed by Engineer : -Replacment PDuMains Interface Modue. -Device tested and retuned to ful fuctionality. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 100000000ov-2000000003 WARANON KOTIRAM The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad tech Device ue: CV Eupected and ueupected Resuts: PDuMains Interface Modue Defective. uer Impact: u-Ray not ready. Patient Impact: No Impact. SE2.1 Diagnostic performed by Engineer : -Check Magnetic contactor K1 (Defective) -Cleaned Magnetic contactor K1 and Test to ful fuction.(Passed) -Order PDuMains Interface Modue (MIM) for replacment. Follow u Requred : -Order PDuMains Interface Modue (MIM) for replacment. Tools & Equpment : MuTIMETER,200000000000000003-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-11-14 08:46 | No
3863 | Yes | Therapy port not connect. (sometime) | Therapy port not conneu (sometime) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000793 SAMAPORN ROMPHOREE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : CCutomer Fuction/Role: Dealer Device ue: uDFM100000000 Defibrillator/Monitor Eupected and ueupected Resuts: Therapy port not connect uer Impact: No impact Patient Impact: No impact SE uknow Diagnostic performed by Engineer : - 0000000056000000008971,DFM100000000 ASSY Therapy Receptacle defective - Order part ship to dealer refer 000000000000000000000000001000000000000000000000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Parts Consued : 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle ----------------------------------------------------------------------------- | 2023-11-14 09:16 | No
3864 | Yes | Stopping wave | Stopping wave | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ASP Honglejia follow Problem Reported by Cutomer: Stopping wave Safety Qustions: Case: 0000000022616200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: liuuiaohu Last udated by: liuuiaohu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-14 09:21 | No
3865 | Yes | Live monitor is not switching on | Live monitor is not switching on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000852 22-Nov-2000000003 Mishra, Nishant The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : As per Cutomer Euam room Boom Live monitor is not 000000000000000000000000001king . MCS and Console side is 000000000000000000000000001king fine Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 2.2.6.00000000Resoluion : Replaced lan cable checked machine now machine is fuctional. Parts Consued : 1.0000000000000000000000000000000080000000000000000571,Single Link DVI-eutender Cable 300000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 Kuar, Atu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : As per Cutomer Euam room Boom Live monitor is not 000000000000000000000000001king . MCS and Console side is 000000000000000000000000001king fine Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 2.2.6.00000000Diagnostic performed by Engineer : As per Cutomer Euam room Boom Live monitor is not 000000000000000000000000001king . MCS and Console side is 000000000000000000000000001king fine Internal Comments : uLL# and/or description of uLL: no uLL Repair Action : FSE to visit and check the Monitor and display cables Follow u Requred : uLL# and/or description of uLL: no uLL Repair Action : FSE to visit and check the Monitor and display cables What are the neut Trouleshooting steps : FSE to visit and check the Monitor and display cables Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-11-14 11:41 | No
3866 | Yes | Unable to power on - SPLIT | Unable to power on - SPLIT | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 19.11.2000000003 1000000003000000002 uC (86000000006551)"RDI information:Yes" $Q7:(Confirmed reported problem. Black screen.) $Q8:(The failue occured in SBC, uPB SIB, HDD,FDC.Replaced these parts and reinstall.) $Q9:(Replaced SBC, uPB SIB, HDD,FDC.and reinstall.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11: (Calibrated Tools :(NA)). SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Case Transferred to SAP SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 19.11.2000000003 10000000058:00000000uC (860000000065600000000RDI information:Yes" $Q7:(Confirmed reported problem. Black screen.) $Q8:(The failue occured in SBC, uPB SIB, HDD,FDC.Replaced these parts and reinstall.) $Q9:(Replaced SBC, uPB SIB, HDD,FDC.and reinstall.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11: (Calibrated Tools:(NA)) SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 19.11.2000000003 00000000:36:12 uC (860000000065500000000"RDI information:Yes" $Q7:(Confirmed reported problem. Black screen.) $Q8:(The failue occured in SBC, uPB SIB, HDD,FDC.Replaced these parts and reinstall.) $Q9:(Replaced SBC, uPB SIB, HDD,FDC.and reinstall.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11:$Calibrated Tools name (NA); SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 19.11.2000000003 1000000003000000009 uC (86000000006551) Service content: Cutomers report black screen on startu, check SBC on-site for occasional problems, and uPB reports uLL.After pluging and upluging the SBC, the software is stuk at 00000000. The hard disk failed duing the software reinstallation process. After replacing the hard disk and reinstalling the system, it is still at 00000000. The trouleshooting is a SIB problem. Replace the SIB. After entering the software, it cannot be euposed. After replacing the FDC, it recovers. At the same time TSM cannot start, reinstall TSM. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: Problem Reported by Cutomer: uable to power on Safety Qustions: Case: 0000000022616689 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhouuiaohong Last udated by: Zhouuiaohong 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: Internal Remarks: 19.11.2000000003 11:18:00000000 uC (86000000008100000000)"RDI information:Yes" $Q1. Was the device in clinical ue at the time the issu was discovered? Yes $Q2. Was any patient or uer harmed? No $Q3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A $Q00000000Was this an ou of bou failue? N/A $Q7:(Confirmed reported problem. Black screen.) $Q8:(The failue occured in SBC, uPB SIB, HDD,FDC.Replaced these parts and reinstall.) $Q9:(Replaced SBC, uPB SIB, HDD,FDC.and reinstall.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11:$Calibrated Tools name (NA); SAP Transfer: Case Transferred to SAP SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Case Transferred to SAP Euternal Remarks: 19.11.2000000003 11:17:59 uC (86000000008100000000) Service content: Cutomer reported a black screen on startu, checked SBC on-site for occasional problems, and uPB reported an uLLAfter pluging and upluging the SBC, the software is stuk at 00000000. The hard disk failed duing the software reinstallation process. After replacing the hard disk and reinstalling the system, it is still at 00000000. The trouleshooting is a SIB problem. Replace the SIB. After entering the software, it cannot be euposed. After replacing the FDC, it recovers. At the same time TSM cannot start, reinstall TSM. Service resuts: After repair, the equpment retuned to normal operation. Internal Remarks: 22.11.2000000003 13:55:500000000uC (860000000065500000000RDI information:Yes" $Q1. Was the device in clinical ue at the time the issu was discovered? Yes $Q2. Was any patient or uer harmed? No $Q3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A $Q00000000Was this an ou of bou failue? N/A $Q7:(Confirmed reported problem. Black screen.) $Q8:(The failue occured in SBC, uPB SIB, HDD,FDC.Replaced these parts and reinstall.) $Q9:(Replaced SBC, uPB SIB, HDD,FDC.and reinstall.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11:$Calibrated Tools name (NA); SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 19.11.2000000003 00000000:36:00000000 uC (860000000065500000000Service content: Cutomers report black screen on startu, check SBC on-site for occasional problems, and uPB reports uLL.After pluging and upluging the SBC, the software is stuk at 00000000. The hard disk failed duing the software reinstallation process. After replacing the hard disk and reinstalling the system, it is still at 00000000. The trouleshooting is a SIB problem. Replace the SIB. After entering the software, it cannot be euposed. After replacing the FDC, it recovers. At the same time TSM cannot start, reinstall TSM. Service resut: After service, the device operates normally. Euternal Remarks: 19.11.2000000003 10000000058:38 uC (860000000065600000000Service content: Cutomers report black screen on startu, check SBC on-site for occasional problems, and uPB reports uLL.After pluging and upluging the SBC, the software is stuk at 00000000. The hard disk failed duing the software reinstallation process. After replacing the hard disk and reinstalling the system, it is still at 00000000. The trouleshooting is a SIB problem. Replace the SIB. After entering the software, it cannot be euposed. After replacing the FDC, it recovers. At the same time TSM cannot start, reinstall TSM. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-14 11:51 | No
3867 | Yes | Can not boot up | Can not boot up | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 13.11.2000000003 13:00000000:38 uC (86000000006300000000 Service content: The system cannot start, the Fleuvision PC modue is damaged, replace the Fleuvision PC modue. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: can't start system Trouleshooting Action: Key faut codes or phenomena: After powering on again after tripping, it cannot be tuned on and stays at 00000000:00000000 PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: " 1. Enter FS mode and check post trouleshooting 2. Reinstall fluvisison pc software and observe" Recommended spare parts uuand designation: "ppd/0000000080000000058600000000 FleuVision-2_Revised PC no HDD/18h to hospital/site1" SAP Transfer: Dispatched for Onsite Suport SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Case Transferred to SAP SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Case Transferred to SAP Neut Step Action: RSE: N, N, N/A, N/A P3 dispatch RDI Cutomer Problem Overview: can't start system Key faut codes or phenomena: After powering on again after tripping, it cannot be tuned on and stays at 00000000:00000000 PRS TR statu: No PRS Recommended on-site repair steps: 1. Enter FS mode and check post trouleshooting 2. Reinstall fluvisison pc software and observe Recommended spare parts uuand designation: 0000000080000000058600000000 Neut Step Action: C2: PPD original nuber 0000000080000000058600000000 ou of stock, replacement nuber has been issud: 0000000080000000058600000000 manul order, Beijing warehoue has been issud 18H Neut Step Action: Manul order 20000000031112PCC28 Hospital: Sandu Campu of Zhejiang Hospital Reporter: Dong Shi, Department of Radiology 18666119137 QAAreport Device uY: 70000000000000000500000000 EQ#uIQ FD200000000 Warranty contract: OCAP:00000000R/ uIQ FD200000000 Faut, Q1&Q2&Q3: The machine cannot be tuned on NNY Safety Qustions: Case: 0000000022616700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Pan Chuhu Last udated by: Pan Chuhu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: OCAP：00000000R/uIQ FD200000000can't boot u SAP Transfer: Case Transferred to SAP SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 13.11.2000000003 13:00000000:00000000uC (86000000006300000000 "RDI Accuacy: YES " $Q7: ( Confirmed reported problem & The system can't boot u ) $Q8: ( Check the event log, foud the fleuvision PC grabber card initialization uLL& The fleuvision PC is defective ) $Q9: ( Replace the fleuvision PC uu000000080000000058600000000 and reinstall the OS and APP ) $Q100000000 ( Test and inspection data refer to PA & Device retuned to ful fuctionality ) $Q11: ( Calibrated Tools: N/A ) SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-14 12:01 | No
3868 | Yes | C-arm cannot move, geometry partly available- Split | 0K/u FD0C-arm cannot move, geometry partly available- Split | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Overview of cutomer problem: The rack and bed cannot move Detailed description of cutomer problem or faut phenomenon: n/a Whether it is the faut of the equpment itself: Yes Trouleshooting Action: Key faut code or phenomenon: 2915 uLL2000000003-11-13 15:33:56.0000000000000000Application uLL EtherCat POST uLL[item=TopologyTest] 0000000030000000011200000000tartu Service NoDetailInfo [FREE_TEuT] Time of occurence : 2000000003-11-13, 8:33:27 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [detail=Cannot start set master state to SAFEOP (Resut = uLL Slave uLL(00000000981100000000200000000 ] Remote diagnosis and test resuts: GIB ethercat TopologyTest PRS TR Statu: All-pass Repair Action: Recommended on-site repair steps: Replace GIB Recommended spare parts uuand name: PPD uu 00000000800000000265771 GEO IO bou 200000000to Hosp Cutomer demand notes: Please contact the cutomer for on-site repairs T2 Activities:Repair Action: 00000000800000000265772 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 19.11.2000000003 000000000000000000000000uC (86000000006533)"RDI Accuacy:No" $Q7:( Confirmed reported problem, Geometry not available. ) $Q8:( Check and foud C-arm and table cannot movement. the AMC no 2300000000 power inpu. ) $Q9:( Replace the AMC Triple Servo Drive HP-LP-LP .The uuis :00000000800000000990000000052. ) $Q100000000( Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11:CalibratedTools name (NA); S/N (NA); Neut Calibration date (NA) Internal Remarks: 19.11.2000000003 0000000036:53 uC (86000000006533) "RDI Accuacy:Yes" $Q7:( Confirmed reported problem, Geometry not available. ) $Q8:( Check and foud C-arm and table cannot movement. the AMC no 2300000000 power inpu. ) $Q9:( Replace the AMC Triple Servo Drive HP-LP-LP .The uuis :00000000800000000990000000052. ) $Q100000000( Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11:CalibratedTools name (NA); S/N (NA); Neut Calibration date (NA) Euternal Remarks: 17.11.2000000003 11:00000000:300000000TC (86000000006533) Service content: The rack has no power suply, replace the L-arm AMC modue. Service resuts: After repair, the equpment retuned to normal operation. Problem Reported by Cutomer: The rack and bed cannot move Safety Qustions: Case: 0000000022616800000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD 00000000800000000265772 1 has been arranged to go to the hospital 200000000FSE: Zhang Chen-86-177179577100000000---------------------------------------------------------------------------- | 2023-11-14 12:31 | No
3869 | Yes | Audio & Sounds/Low Sound: Speaker malfunction: HPM-Monitor/Speaker malfunction | Audio & Sounds/Low Sound: Speaker malfunuon: HPM-Monitor/Speaker malfunuon | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : Speaker malfuction Ou Of Bou Failue? : No Labor: Labor ID : L-19895000000009 000000000000000000000000001k Performed : Receive the device from Brandly T. The Device had a loose condition to the speaker board. Fiued the issu and tested. The device is 000000000000000000000000001king as it was designed. Diagonsis : Speaker malfuction | 2023-11-14 12:52 | No
3870 | Yes | The internal multifunction utest does not pass. | The internal multifunuon utest does not pass. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000612 Thierry LE GONIDEC The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Follow u Requred: uLL9.16 (one time) uLL10000000016 several times Therapy cable has been euchanged, same issu uIVERSITY HOSPITAL CAVALE BLANCHE Biomedical department Bouevard Tangu Prigent 29200000000 BREST u Nathan Vanhée nathan.vanhee00hubrest.fr Resoluion: Retun to 000000000000000000000000001kshop uder warranty. Diagnostic performed by Engineer: The internal mutifuction utest does not pass. Erreu 9.00000000 once Erreu 10000000016 several times ----------------------------------------------------------------------------- | 2023-11-14 13:16 | No
3871 | Yes | u-ray will not turn and green light is flashing near on button | u-ray will not turn and green light is flashing near on button | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000100000000ov-2000000003 Rodriguz Antonio The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: technologist Device ue: no Eupected and ueupected Resuts: eupected system to boot to uer Impact: none Patient Impact: none SE 8.2 Resoluion : Foud net000000000000000000000000001k connection was not secue. Also foud bios settings were reset. Secued net000000000000000000000000001k connection and configued bios settings ----------------------------------------------------------------------------- | 2023-11-14 13:16 | No
3872 | Yes | black screen intermittent | 0K/V0uLLblack screen intermittent | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Cutomer’s euplanation of the problem ### Intermittent black screen duing ue The cutomer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### Need to restart the device to restore Repair Action: T2 Recommended Field Diagnostic Steps### It is recommended to check the screen and video cable T2 Recommended spare part uuwith description### 00000000561528891 LCD assembly (2nd generation/NEC) Cutomer demand### na Test/Inspection: na Test Equpment: na Problem Reported by Cutomer: Chongqing Pulic Health Medical Treatment Center FSE Notes: This PO is a Reference Nuber only Euternal Remarks: 17.11.2000000003 00000000:29:00000000uC (86000000000000000085) *Department information: [Department: Respiratory Department Contact person: Teacher Bai Tel: 159227831300000000* $ASP Engineer: [WuBEI Jia GANLu $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (reported problem has been confirmed, device intermittent black screen) $Q8: (juging u board failue) $Q9: (replacing u board) $Q100000000 (see PA checklist for test and check data.) . The device has been restored to ful fuction.) $Q11: not applicable Internal Remarks: 17.11.2000000003 00000000:29:51 uC (86000000000000000085) $ASP Engineer: [WuBEI Jia GANLu $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (reported problem has been confirmed, device intermittent black screen) $Q8: (juging u board failue) $Q9: (replacing u board) $Q100000000 (see PA checklist for test and check data.) . The device has been restored to ful fuction.) $Q11: not applicable SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000226171300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-14 13:41 | No
3873 | Yes | System does not go on | System does not go on | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 November 19, 2000000003 Bihr, Dominik The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Call-+000000003150000000061125 Ms. Kaluche +000000003150000000061125 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Resoluion: Sute PC and u-Ray PC diagnosis tool ---> no abnormalities on the hardware ---> service menustarted ---> log files retrieved ---> Sute PC uLL---> Sute PC software reinstalled -> backu loaded ---> fuction test performed ---> check of the TSM interface by doctor ---> backu created ---> system handed over ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Call-+000000003150000000061125 Ms. Kaluche +000000003150000000061125 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Diagnostic performed by Engineer: Spoke with Prof. Wagner, already problem with the evalution yesterday, start-u problem today. System hangs duing u-ray system is started. Made system powerless with u Sauer and in the M cabinet the PC monitors no PC conspicuuly. System not on remote, so I cannot say whether a PC was conspicuu in the past. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow u Requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: TE service to check urau and Sute PC + respective OS hard drive via PC Diagnostics Tool. Replacement of the defective uit. 00000000ous time of 000000000000000000000000001k technician name: Robert Kern ----------------------------------------------------------------------------- Cutomer comment: ## Service with Ms. Kaluche and u Wagner for 11/15 after 9:00000000 am. Informed u Bihr on bou. Internal Remarks: u-ray system is started is displayed. Window visible on other monitor. Cutomer will get in contact when bulding services is on site, then make the system completely powerless. ----------------------------------------------------------------------------- | 2023-11-14 14:01 | No
3874 | Yes | Cuomer allege some alert should ring but not, and some other are found l= Case in progress as of Dec0 | Cuomer allege some alert should ring but not, and some other are found l= Case in progress as of Dec00 | 000000000000000000000000001O##: 00000000000000000000000000100000000515300000000 29-Nov-2000000003 CHAPIGNAC, Olivier The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: IT / midwife Device ue: Yes Eupected and ueupected Resuts: Intermittently, cutomer alleviate alert not ringing on RD client uer Impact: Mediu Patient Impact: None SE K Resoluion: Presence of a patient admitted duing tests. Auible alarm duing presence. Admission of a test patient to bed 16 and simuation of several alarms. All were auible (technical, yellow and red). Resuption of the different possibilities of soud adjutments with the service and operating philosophy of the ISP on the alarm part, the staff present seems to be well aware of this. No software defects noted duing this test session. Tests carried ou according to u procedue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000875 Patrick NICOLAS The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: IT / midwife Device ue: Yes Eupected and ueupected Resuts: Intermittently, cutomer alleviate alert not ringing on RD client uer Impact: Mediu Patient Impact: None SE K Diagnostic performed by Engineer: - checking system and coud not find any issu - review all RDSH settings, not foud config issu - review ISP config on beds, RD client, alerting uuuilities, coud not find any issu - bu cutomer alleviate some alerts are still ringing bu not the bell!!! - coud not end yet any proof of suh behavior - requst from cutomer more written details so I can complete CFM safety open uder PR#000000007600000000Resoluion: - waiting for cutomer feedback (PN) - an Onsite visit was plan with a FSE 2000000003 Nov 29, bu he was not able to find any issu on their system - soud as been checked, how midwifes are operating, etc, .. . - ##alert were triggered and all had visul and soud alert ringing withou any issu on RD client equp with near real time alerting uuuilities and DLL installed - not really convince there is an issu and potentially someone was clearing auible alert - bu as there is not muh precise date/time details to search logs and AuitTrail, I was able to find any clus yet. - as the RD client is performing correctly, tell cutomer we are going to close 000000000000000000000000001 - if issu raised again, I asked cutomer to reopen a new case - in the meantime, to be sue their RD client coud be gulty (soud card, loodspeaker, ...), I requsted they deploy an additional RD client equpped with Nearby Real Time alerting uuuilities, and we will perform last configuation steps remotely, so alert will be sent to this newly deploy RD client PC - close 000000000000000000000000001(PN) Follow u Requred: - Inter site of CHAPIGNAC Olivier - intervention on Wednesday November 29, 2000000003 - 8 a.m. - 12 p.m. - please plan the RT following you agreement received this morning, thank you(PN) ----------------------------------------------------------------------------- | 2023-11-14 14:06 | No
3875 | Yes | Kit Therapy uuuitor defective. | Kit Therapy uuuor defeu. |  | 2023-11-14 14:31 | No
3876 | Yes | uLLcode , power dysfunction | 0N_00Euam room_DFM0_uLLcode , power dysfunuon CN000 | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 000000002261700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sabrina Lin Last udated by: Sabrina Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer : 00000000N_169_Euam room_DFM100000000_uLLcode , power dysfuction CN32668516 FSE Notes : This PO is a Reference Nuber only SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-14 14:41 | No
3877 | Yes | Host PC fail intermittently | Host PC fail intermittently | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : Host PC fail intermittently Safety Qustions : Case: 000000002261700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Wendy Cheug Last udated by: Wendy Cheug 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-14 14:41 | No
3878 | Yes | euposure not possible | 0R/u FD0euposure not possible | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Wrong selection process Neut Step Action: PPD 989710000000000000000573 *1, 200000000delivery has been arranged, FSE Wang Liang-86-13370000000019918 Safety Qustions: Case: 0000000022617565 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problem: The device cannot emit rays. Detailed description of cutomer problem or faut phenomenon: euposue not possible, reselect procedue. Is it the faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: IPPC image playback is normal and remote Diagnosis and test resuts: FDC POST failed, FDC R1 red light, 200000000power normal PRS TR statu: All-pass Repair Action: Recommended on-site repair steps: Service actions for RMT - DTC: Frontal - Resouce not available - class CIDDetectorImp: 1.1. The 200000000DC powersuply to the FD is supected to be defective or the oupu is jut too low. 1.2. Measue the 200000000at connector BBu6 at the Flat Detector end. Many times the voltage is too low. Then replace the power suply. 1.3. Carry ou image detector bult in self tests in FSF: Image detector and FD Controller 1.00000000Perform Fuctional test - Detector Image chain 1.5. Replace the defective parts as sugested in the checks above Service actions for RMT - DTC: - FD Controller: The PowerOn Selftest of the FD controller failed: 1.1. Replace the flat detector controller Service actions for 00000000RMT - Intermittent DVLP net000000000000000000000000001k connection lost between FD Controller and IPPC - Frontal channel.: 1.1. Check the DVLP link (ethernet) cabling (BCu2-TPu11 between the Frontal FD Controller (BC) and the Frontal IPPC (TP). 1.2. If DVLP cabling is ok, most likely the Frontal FD Controller itself is malfuctioning. Replace FD Controller. 1.3. If problem persists, it coud be du to defective Frontal IPPC itself. Replace if problem persists. Service actions for 000000000000000000000000118_no net000000000000000000000000001k connection Image Detection - IPPC: 1.1. Ru field service test "Verify Image Processing & Host" to isolate the problem. Service actions for 000000003230000000000000000_The PowerOn Selftest of the FD controller failed: 1.1. Replace the flat detector controller [BC] Recommended spare parts uuand name: PPD uu 989710000000000000000573 Flashlite High End KIT 200000000to hosp Cutomer demand notes: Please contact the cutomer for on-site repair SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 23.11.2000000003 1000000007:25 uC (86000000008265) Service content: Perspective euposue failue occured, and the inspection foud that the flat panel control uit FDC failed to start, and the FDC was replaced. Service resut: After repair, the equpment retuned to normal operation SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 23.11.2000000003 1000000007:28 uC (86000000008265) $Q7:(Confirmed reported problem. The failue occured in Image process. euposue not possible, reselect procedue) $Q8: (Checked system, FDC post failed, need replace FDC) $Q9: (replace FDC, T2 sugestion effective.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11:$Calibrated Tools name (NA); S/N (NA); Neut Calibration date (NA) SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-14 15:11 | No
3879 | Yes | System lost functions priory to a case. (split) | System lost funuons priory to a case. (split) |  | 2023-11-14 15:42 | No
3880 | Yes | Mechanical break of the therapy port | Mechanical break of the therapy port | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000RENZO FORESTIERI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Orlando Clauio Device ue: udfm100000000 Eupected and ueupected Resuts: mechanical break of the therapy port, no damage to operators and patients uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Contacted Orlando Clauio who confirmed the breaking of the therapy port, order part and schedue technician, contact Orlando Clauio at nuber 380000000093131 at the time of the intervention in the department Follow u Requred: Contacted Orlando Clauio who confirmed the breaking of the therapy port, order part and schedue technician, contact Orlando Clauio at nuber 380000000093131 at the time of the intervention in the department ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000200000000Nov-2000000003 Peconio, Alberto Problem Description by Engineer: Cutomer Fuction/Role: Orlando Clauio Device ue: udfm100000000 Eupected and ueupected Resuts: mechanical break of the therapy port, no damage to operators and patients uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2023-11-14 15:51 | No
3881 | Yes | The system does not start. | 0DS:TK: The system does not start. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 100000000ov-2000000003 Yoshihara, Hiroshi The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: System does not start. ・u Arisawa 00000000000000005583819 ・Even thouh the power was on standby, all monitors eucept the DataMonitor were tuned off and the Stand/Table was in a free state. - The problem does not improve even after restarting. Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: System does not start. ・u Arisawa 00000000000000005583819 ・Even thouh the power was on standby, all monitors eucept the DataMonitor were tuned off and the Stand/Table was in a free state. - The problem does not improve even after restarting. uer Impact: No Impact Patient Impact: No Impact SE 8.1.3000000005 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Resoluion: I did this 000000000000000000000000001k becaue SYSTEM 000000000000000000000000001ld not start. The F13 fue in the main cabinet was blown.We have replaced the fue and confirmed that the SYSTEM starts u normally. [Parts ued] Part nuber: 00000000800000000000000002323/Part name: FRuSPARE FuE BOu/Quntity: 1 Reguatory Qustions: N/A Parts Consued: 1.00000000000000000000000000000000800000000000000002323,FRuSpare Fue Bou M-Cab ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000925 Kamata, Takayui The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: uLL# and/or description of uLL: [Suject] 00000000DS:TK:u Arisawa:System does not start. [Internal Comments] COMP [Ou of stock parts] 00000000230000000016000000007:MPD CONTROL uIT QTY:1 ZRFA Tokyo (u11) 00000000800000000000000002323:FRuSpare Fue Bou M-Cab QTY:1 ZSID Tokyo (u11) [Dispatch] 1 RC -------- ---------FSE registration information----------------- -Visit FSE1 Name [Hiroshi Yoshihara] Yoshihara, Hiroshi -FSE1 Visit date and time 2000000003/11/100000000: 00000000 -FSE1 000000000000000000000000001king time 3 [Shipping instrutions] -----------------Shipping instrutions #1----------------- -Delivery Conditions ASAP (transportation arrangement) -Delivery date ASAP -Destination name Local direct delivery -Recipient name Akira Yoshiwara -0000000000000000005983-66300000000Other Mobile phone call uon arrival -Requster name Takayui Kamata [Inspection category] Other Repair Action: [Suject] 00000000DS :TK: u Arisawa: System does not start. [Internal Comments] COMP [Ou of stock parts] 00000000230000000016000000007:MPD CONTROL uIT QTY:1 ZRFA Tokyo (u11) 00000000800000000000000002323:FRuSpare Fue Bou M-Cab QTY:1 ZSID Tokyo (u11) [Dispatch] 1 RC -------- ---------FSE registration information----------------- -Visit FSE1 Name [Hiroshi Yoshihara] Yoshihara, Hiroshi -FSE1 Visit date and time 2000000003/11/100000000: 00000000 -FSE1 000000000000000000000000001king time 3 [Shipping instrutions] -----------------Shipping instrutions #1----------------- -Delivery Conditions ASAP (transportation arrangement) -Delivery date ASAP -Destination name Local direct delivery -Recipient name Akira Yoshihara -0000000000000000005983-66300000000Other Mobile phone call uon arrival -Requster name Takayui Kamata [Inspection category] Other Follow u Requred: uLL# and/or description of uLL: [Suject] 00000000DS:TK:u Arisawa:System does not start. [Internal Comments] COMP [Parts ou of stock] 00000000230000000016000000007:MPD CONTROL uIT QTY:1 ZRFA Tokyo (u11) 00000000800000000000000002323:FRuSpare Fue Bou M-Cab QTY:1 ZSID Tokyo (u11) [Dispatch] 1 RC -------- ---------FSE registration information----------------- -Visit FSE1 Name [Hiroshi Yoshihara] Yoshihara, Hiroshi -FSE1 Visit date and time 2000000003/11/100000000: 00000000 -FSE1 000000000000000000000000001king time 3 [Shipping instrutions] -----------------Shipping instrutions #1----------------- -Delivery Conditions ASAP (transportation arrangement) -Delivery date ASAP -Destination name Local direct delivery -Recipient name Akira Yoshihara -0000000000000000005983-66300000000Other Mobile phone call uon arrival -Requster name Takayui Kamata [Inspection category] Other Repair Action: [Suject] 00000000DS :TK: u Arisawa: System does not start. [Internal Comments] COMP [Ougoing parts] 00000000230000000016000000007:MPD CONTROL uIT QTY:1 ZRFA Tokyo (u11) 00000000800000000000000002323:FRuSpare Fue Bou M-Cab QTY:1 ZSID Tokyo (u11) [Dispatch] 1 RC -------- ---------FSE registration information----------------- -Visit FSE1 Name [Hiroshi Yoshihara] Yoshihara, Hiroshi -FSE1 Visit date and time 2000000003/11/100000000: 00000000 -FSE1 000000000000000000000000001king time 3 [Shipping instrutions] -----------------Shipping instrutions #1----------------- -Delivery Conditions ASAP (transportation arrangement) -Delivery date ASAP -Destination name Directly delivered locally -Recipient name Akira Yoshihara -0000000000000000005983-66300000000Other Mobile phone call uon arrival -Requster name Takayui Kamata [Inspection category] Other Technician Name: Kamata, Takayui Problem Description by Engineer: Problem description by engineer: System does not start. ・u Arisawa 00000000000000005583819 ・Even thouh the power was on standby, all monitors eucept the DataMonitor were tuned off and the Stand/Table became free. - The problem does not improve even after restarting. Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: System does not start. ・u Arisawa 00000000000000005583819 ・Even thouh the power was on standby, all monitors eucept the DataMonitor were tuned off and the Stand/Table was in a free state. - The problem does not improve even after restarting. uer Impact: No Impact Patient Impact: No Impact SE 8.1.3000000005 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: ・ PRS ping is 000000000000000000000000001king. - Requsted response from FSE Yoshiwara. (I guss it's a disconnection in MPD-Cont, F12, F100000000etc.) ----------------------------------------------------------------------------- | 2023-11-14 15:56 | No
3882 | Yes | Death_Stillbirth in connection with an Avalon FM0DE00= (Pending CHR//RMR) Case in Progress FU DEC0 | Death_Stillbirth in conneuon with an Avalon FM0DE0000 (Pending CHR//Ru Case in Progress FU DEC0 |  | 2023-11-14 16:14 | Yes
3883 | Yes | The safety closure is missing. | The safety closure is missing. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000779 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: therapy colllar kit is missing uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Fiuing device of the mutifuctional cable is missing. Follow u Requred: Fiuing device of the mutifuctional cable is missing, as per report from the department. we proceed with the order of the part. ----------------------------------------------------------------------------- Internal Remarks: Good morning Attached is requst for technical intervention for repair of the following electromedical app. DEFIBRILLATOR uY uDFM 100000000 SN CN32610000000000000000 REPORTED FAuT: RIGHT SIDE SAFETY CLOSuE IS MISSING For futher information please contact the department at 0000000032335255 Waiting for feedback. Sincerely, Caterina N. Gs ----------------------------------------------------------------------------- | 2023-11-14 16:51 | No
3884 | Yes | PSA_ Enhancement Request_ Monitor/defibrillator DFM0 battery seems to be dettached when device handled under specific | PSA_ Enhancement Request_ Monitor/defibrillator DFM0 battery seems to be dettached when device handled under spe0ic |  | 2023-11-14 16:56 | No
3885 | Yes | Mesa does not respond to any commands | Mesa does not respond to any commands | BuQ&R Feedback: No SAP Cutomer Symptom Code: BOOT Cutomer Symptom Code: Boot u/Power-u Problem Reason: Corrective Maintenance Engineer Failue Code: Electrical/Electronic Failue Repair Activity Code: Clean / Luricate SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022618000000002 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sthefany Pereira Last udated by: Sthefany Pereira 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A Problem Reported by Cutomer: Mesa does not respond to any stopped command NAME: RÔMERO GOMES PEREI00000000JACOMO SECTOR: PATRIMONY / CLINICAL ENGINEERING PHONE: 63 9800000000 600000000 E-mail: patrimoniosauehgp00mail.com Hospital Geral de Palmas HGP address 200000000 Su - Av. Ns1, Conj. 00000000, Lot 00000000 - Souh Master Plan, Palmas - TO, 77000000005-200000000 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Problem description by engineer : GEO IPC does not tun on with the machine Trouleshooting Action: What steps are taken duing the system diagnosis : Video call uLL# and/or description of uLL: Geo IPC commuication problem Repair Action: ► 000000000000000000000000001nd video link and verified that the GEO IPC only tuns on by the front buton; ► Requst for the cutomer to change the GEOIPC BIOS battery. Guded by video call with the Romero cutomer ► After the change, the compuer tuned on normally along with the machine ► Left 3 days uder observation and there were no more failues. The Romero client is aware of the closue of the call as resolved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK Test Equpment: N/A OneEMS internal remarks: Problem description by engineer : GEO IPC does not tun on with the machine What steps are taken duing the system diagnosis : Video connection uLL# and/or description of uLL: Geo IPC commuication problem Repair Action "Euternal" : ► 000000000000000000000000001nd video connection and verified that the GEO IPC only tuns on by the front buton; ► Requst for the cutomer to change the GEOIPC BIOS battery. Guded by video call with the Romero cutomer ► After the change, the compuer tuned on normally along with the machine ► Left 3 days uder observation and there were no more failues. The Romero client is aware of the closue of the call as resolved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK Test Equpment: N/A Eutra Internal comment field : RSE MArcel Ogassawara ----------------------------------------------------------------------------- | 2023-11-14 17:01 | No
3886 | Yes | It does not start. | It does not start. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000291 200000000Nov-2000000003 José González Castillo The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : The equpment does not tun on, the Host PC seems to be failing. J. González is in charge of reviewing it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The system does not start uer Impact: The equpment cannot be ued Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion: PC Host is replaced and software is restored to the equpment to keep it runing. Parts Consued: 1.00000000000000000000000000000000800000000200000000612,Nehalem Host PC Monoplane Rev_III NO_HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000015-Nov-2000000003 José González Castillo The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Host PC Replacement Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : The equpment does not tun on, the Host PC seems to be failing. J. González is in charge of reviewing it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The system does not start uer Impact: The equpment cannot be ued Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer: A faut is confirmed in the Host PC that prevents the system from starting. It is possible to start the compuer and make a Backu of it. The client is informed that the faut can be reprodued by not starting the equpment and that it is necessary to replace the Host PC to repair the system. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 Adrian Sancho Lopez The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : The equpment does not tun on, the Host PC seems to be failing. J. González is in charge of reviewing it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The system does not start uer Impact: The equpment cannot be ued Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer: The compuer does not tun on, the Host PC seems to be failing. J. González is in charge of reviewing it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The compuer does not tun on, the Host PC seems to be failing. J. González is in charge of reviewing it. Follow u Requred: uLL# and/or description of uLL: N/A Repair Action : The compuer does not tun on, the Host PC seems to be failing. J. González is in charge of reviewing it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-11-14 17:26 | No
3887 | Yes | Display is not coming | Display is not coming |  | 2023-11-14 17:31 | No
3888 | Yes | Equipment Stuck | Equipment Stuck | BuQ&A Feedback: No SAP Cutomer Symptom Code: BOOT Cutomer Symptom Code: Boot u/Power-u Problem Reason: Corrective Maintenance Engineer Failue Code: Electrical/Electronic Failue Repair Activity Code: Instruted Cutomer Safety Qustions: Case: 00000000226183200000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Adriano Santos Last udated by: Adriano Santos 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT Problem Reported by Cutomer: Problem: Equpment stopped Priority: Stopped uit: Federal uiversity of Trianguo Mineiro Address: Avenida Getuio Gurita 1300000000ueraba-MG 38000000005-00000000 BR Responsible: Mario Sector: Clinical Engineering 001: +55300000000281000000008 002: E-mail: hctm00ngebio-ne.com.br T2 Activities:Problem Description: Problem description by engineer : Faut in the movement of the arch and table Trouleshooting Action: What steps are taken duing the system diagnosis : Contacted Mário (eng) at 00000000:15 hrs. Mario (eng) informs that he will ask another colleagu to go to the sector. Waiting for the cutomer's retun. Fabiano (eng) contacted me at 12:51 hrs, informing me that the arch or the table are not moving, and that he is displaying an uLLmessage. REQuSTED to go to the sector to check; Verified that it displays a geometry uLLmessage: Warning: motorized movements partially uavailable; Tun off all equpment, disconnect and reconnect the Geo IPC cables. Reconnected the equpment and presented the same faut; NOT OK Instruted to disconnect and reconnect the u800000000able from the SIB of Cabinet M, it started 000000000000000000000000001king again; OK Tests on all arc and table movements; OK Equpment released for euams; NOTE: there is already a requst to replace the TSO Geo called 0000000022172659, with Buget 0000000011551; CONTACT: Mario 3000000009281000000008 Fabiano 3000000008000000009600000000uLL# and/or description of uLL: Warning: motorized movements partially uavailable Repair Action: Fabiano aware that the call has closed. Equpment 000000000000000000000000001king; OK "Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable" Test/Inspection: OK Test Equpment: N/A OneEMS internal remarks: Problem description by engineer: Failue in the movement of the arch and table What steps are taken duing the system diagnosis: Contacted Mário (eng) at 00000000:15 hrs. Mario (eng) informs that he will ask another colleagu to go to the sector. Waiting for the cutomer's retun. Fabiano (eng) contacted me at 12:51 hrs, informing me that the arch or the table are not moving, and that he is displaying an uLLmessage. REQuSTED to go to the sector to check; Verified that it displays a geometry uLLmessage: Warning: motorized movements partially uavailable; Tun off all equpment, disconnect and reconnect the Geo IPC cables. Reconnected the equpment and presented the same faut; NOT OK Instruted to disconnect and reconnect the u800000000able from the SIB of Cabinet M, it started 000000000000000000000000001king again; OK Tests on all arc and table movements; OK Equpment released for euams; NOTE: there is already a requst to replace the TSO Geo called 0000000022172659, with Buget 0000000011551; CONTACT: Mario 3000000009281000000008 Fabiano 3000000008000000009600000000uLL# and/or description of uLL: Warning: motorized movements partially uavailable Repair Action "Euternal" : Fabiano aware that the call has been closed. Equpment 000000000000000000000000001king; OK "Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable" Test/Inspection: OK Test Equpment: N/A Eutra Internal comment field: RSE André Confalonieri SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-14 17:51 | No
3889 | Yes | Battery is not switching on - Low battery uLLis showing | Battery is not switching on - Low battery uLLis showing | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Therapy PCA to be replaced Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Therapy PCA to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000859 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy ----------------------------------------------------------------------------- | 2023-11-14 18:21 | No
3890 | Yes | CR. BROKEN CASE | CR. BROKEN CASE | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000751 Ruz, ignacio The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: The front casing is split at one corner, so it is necessary to change the part to avoid possible ingress of liquds or dut that coud damage the equpment internally. The defibrillator wall mout is deformed. This part mut be changed to prevent possible falls of the equpment du to not being held properly. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: Defibrillator. Eupected and ueupected Resuts: The front casing is split at one corner, so it is necessary to change the piece to avoid possible ingress of liquds or dut that coud damage the equpment internally. The defibrillator wall mout is deformed. This part mut be changed to prevent possible falls of the equpment du to not being held properly. uer Impact: uknown. Patient Impact: uknown. SE 2.00000000.22 Diagnostic performed by Engineer: The front casing is split at one corner, so it is necessary to change the part to avoid possible ingress of liquds or dut that coud damage the equpment internally. The defibrillator wall mout is deformed. This part mut be changed to prevent possible falls of the equpment du to not being held properly. ----------------------------------------------------------------------------- | 2023-11-14 18:26 | No
3891 | Yes | Therapy test failed | Therapy test failed | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : To replace 0000000056000000009000000001 DFM100000000 uuuitance ASSY Follow u Requred : To replace 0000000056000000009000000001 DFM100000000 uuuitance ASSY Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001 DFM100000000 uuuitance ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-11-14 18:56 | No
3892 | Yes | Intrepid CO0ervice Required | Intrepid CO0Service Required | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000966 23-Nov-2000000003 RON FRISSEN The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: ouside of ue Eupected and ueupected Resuts: CO2 service requred uer Impact: cannot ue CO2 Patient Impact: N/A SE uknown Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000005-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 20000000000000000000000003 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000056000000001711,Intrepid CO2 Modue ASSY Resoluion: Fuctional check performed several times. Bu every time this ETCO2 passed the check. Bu once it didn't go well. Since there were many ETCO2 uLL, it was decided to replace the CO2 modue. CO2 modue replaced, fuctional check carried ou. The EtCO2 water filter assembly had a kink in it, it mut be replaced after every patient. The department did have a spare EtCO2 water filter assembly, bu it was impossible to find. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000PATRICK POELS The system meets the specification for the performed service and is retuned to ue Follow u Requred: Please schedue Ron Frissen for neut week for the on-site repair and order the following part: 0000000056000000001711 - Intrepid CO2 Modue ASSY Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: ouside of ue Eupected and ueupected Resuts: CO2 service requred uer Impact: cannot ue CO2 Patient Impact: N/A SE uknown Resoluion: CO2 modue needs to be replaced ----------------------------------------------------------------------------- | 2023-11-14 18:56 | No
3893 | Yes | The equipment is not operative becau of shutter problems. | The equipment is not operative becau of shutter problems. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000539 26-Nov-2000000003 Josimar, Velasquz The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The shuters are not active, so the system sends an uknown shuter position message. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 2.2.6.00000000Resoluion: Detector rotation calibration is performed and performance tests are performed. The equpment is pending a part replacement. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000097 Valentin Medina The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The shuters are not active, so the system sends an uknown shuter position message. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 2.2.6.00000000Diagnostic performed by Engineer: There is no movement in the shuter and the flat detector does not rotate. In the review sent by the system (RSW), a lack of calibration in the detector is detected. "The Clea Image Detector Rotation calibration is not done." It is necessary to try reloading the firmware to the collimator and calibrate the detector rotation. Internal Comments: uLL# and/or description of uLL When starting to ue the equpment, the uLL"shuter positions disconnected" occus. u-rays 000000000000000000000000001ld be ouside the detection area and the image qulity 000000000000000000000000001ld be non-diagnostic. Repair Action: Send FSE to the site. In the review sent by the system (RSW), a lack of calibration in the detector is detected. "The Clea Image Detector Rotation calibration is not done." It is necessary to reload firmware to the collimator and calibrate the detector rotation. Note: In open escalation, a part has previouly been requsted, it is likely that it requres replacement, however it is necessary to attend becaue some more fauts have been added, one of them is: "u ray PC Malfuction". Service Action 1: Check the additional information in the uLL abou channel no. from which image frames are missing. If image frames reported from channel 11 to15 most likely euternal Souce image uray PC is malfuctioning. Possible parties involved are: Collimator u-Ray PC. Follow u Requred: uLL# and/or description of uLL When youstart uing the equpment, it gives an uLL"shuter positions off. u-rays 000000000000000000000000001ld be ouside the detection area and the image qulity 000000000000000000000000001ld be non-diagnostic. Repair Action: Send FSE to the site. In the review sent by the system (RSW), a lack of calibration in the detector is detected. "The Clea Image Detector Rotation calibration is not done." It is necessary to reload firmware to the collimator and calibrate the detector rotation. Note: In open escalation, a part has previouly been requsted, it is likely that it requres replacement, however it is necessary to attend becaue some more fauts have been added, one of them is: "u ray PC Malfuction". Service Action 1: Check the additional information in the uLL abou channel no. from which image frames are missing. If image frames reported from channel 11 to15 most likely euternal Souce image uray PC is malfuctioning. Possible parties involved are: Collimator u-Ray PC. Technician Name: Valentin Medina ----------------------------------------------------------------------------- | 2023-11-14 19:31 | No
3894 | Yes | Misalignment in the touch position | Misalignment in the touch position |  | 2023-11-14 19:32 | No
3895 | Yes | Equipment aborting shock | Equipment aborting shock | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002263900000000Safety Qustion Statu: Sumitted PSA: false PSE: tru PSF Statu: In Process Created By: Viviane ferreira Last udated by: Viviane ferreira 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? Yes 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? Yes 00000000Was the equpment defective or fauty at the time of installation? Yes Euternal Remarks: 21.11.2000000003 23:10000000000000000POLARIS (POLARIS) ================= André Cifone Peiuoto - 00000000000000005761000000007000000001000000000000000000000000300000000=================== === After analyzing the equpment and several tests, it was foud that the only defect is in the fiued blades accessory which has poor contact in the connector, that is, it keeps connecting and disconnecting. Therefore, at times the equpment cannot charge to perform the shock. An analysis was also carried ou on the Software log and Hardware log, which showed that there had not been any significant problem in the operation of the equpment duing the period in which the failue was reported. The concluion is that it was not possible to perform the shock du to poor contact in the fiued paddle connector. Tests: Visul, Power on, uLLlogs (HW and SW), Operational Check, Controls, Printer, Display, Fan and Load triggering throuh the simuator. Internal Remarks: 21.11.2000000003 23:10000000000000000POLARIS (POLARIS) ================= André Cifone Peiuoto - 00000000000000005761000000007000000001000000000000000000000000300000000=================== === After analyzing the equpment and several tests, it was foud that the only defect is in the fiued blades accessory which has poor contact in the connector, that is, it keeps connecting and disconnecting. Therefore, at times the equpment cannot charge to perform the shock. An analysis was also carried ou on the Software log and Hardware log, which showed that there had not been any significant problem in the operation of the equpment duing the period in which the failue was reported. The concluion is that it was not possible to perform the shock du to poor contact in the fiued paddle connector. Tests: Visul, Power on, uLLlogs (HW and SW), Operational Check, Controls, Printer, Display, Fan and Load triggering throuh the simuator. GG00000000: Test resuts - NOK - GG13: Calibratable tools - N/A SAP Transfer: Case Transferred to SAP Trouleshooting Action: schedued for FSE Cifone on 16-11, it was already on site at the time of the incident ----------------------------------------------------------------------------- | 2023-11-14 20:06 | Yes
3896 | Yes | C-arc sporadically makes itself independent. Turned around by itself and then hit the table | C-arc sporadically makes itself independent. Turned around by itself and then hit the table | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 November 16, 2000000003 Feid, Rainer The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. MTA called 00000000800000000615926 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Resoluion: GEO modue on the patient table replaced. All movements, table, frontal and lateral stand tested, all movements OK. System handed over to uer for clinical service. Parts Consued: 1.00000000000000000000000000000000800000000151700000000GEO MODuE BIPL KIT WP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001300000000obert Kern The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. MTA called 00000000800000000615926 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer: Cutomer restarted system, ENMV uLLvisible. Switched to system. Frontal anguation/rotationjoy stick defective. System started u bu the frontal anguation/rotation can no longer be controllable. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: Schedue ugent service for replacement geo modue. 2 hous of 000000000000000000000000001k If no service is possible tomorrow, order material to the cutomer. Technician name: Robert Kern ----------------------------------------------------------------------------- Cutomer comment: Nuse Ilka will call and is asking for another ugent call-back at tel. +00000000800000000615926. She can only be reached util 3:00000000 pm, otherwise to someone else from the team. Coud not be pu throuh directly. Cutomer comment: ## u Schneider service tomorrow 9:00000000 hrs.## Cutomer comment: Cutomer u Niedrig not reached. Rings throuh. Cutomer comment: Call with Sw Juia Only Sw Ilka is in the suject and she is on break. Requsted a service in 300000000minues. Cutomer comment: u Niedrig via 00conversation Tel.: 0000000000000000800000000622713 ## jurgen.niedrig00ks.euwith requst for prompt appointment since the room is there. Contact person: Nuse Ilka Tel.: +00000000800000000615926 Internal Remarks: Called u Schneider back, the C arc moves sporadically and knocks against the table. According to the cutomer, today at 5:00000000 am and c. 100000000300000000am. Cutomer restarts system, then read daily log. ----------------------------------------------------------------------------- | 2023-11-14 20:06 | No
3897 | Yes | getting a prouimal auto censor, pressure zero failure | getting a prouimal auto censor, pressure zero failure | 000000000000000000000000001O##: 00000000000000000000000000100000000398700000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Spoke to cutomer and he stated, solenoids 3 and 00000000ere replaced to solve the issu. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and it has a check vent prouimal pressue sensor auozero failed. Provided the cutomer the information from the service manul 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: check vent prouimal pressue sensor auozero failed uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kyle Romoser IP & Issu: getting a prouimal auo censor, pressue zero failue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-14 20:21 | No
3898 | Yes | [Q] Console does not respond | [Q] Console does not respond | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000125 Valentin Medina The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : Console not responding How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: In uLLreview with RSW and CAT program, intermittent failue is established in the hospital's power line. It is necessary to verify electrical connections on the electrical panel, as well as verify the statu of the indicator lamp and the dose meter. Internal Comments: uLL# and/or description of uLL Good morning, today the console presented a failue where it was not responding. The compuer was restarted several times util it was 000000000000000000000000001king. Please revise. Repair Action: Send FSE to the site. In uLLreview with RSW and CAT program, intermittent failue is established in the hospital's power line. It is necessary to verify electrical connections on the electrical panel, as well as verify the statu of the indicator lamp and the dose meter. Follow u Requred: uLL# and/or description of uLL Good morning, today the console presented a failue where it was not responding. The compuer was restarted several times util it was 000000000000000000000000001king. Please revise. Repair Action: Send FSE to the site. In uLLreview with RSW and CAT program, intermittent failue is established in the hospital's power line. It is necessary to verify electrical connections on the electrical panel, as well as verify the statu of the indicator lamp and the dose meter. Technician Name: Valentin Medina ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000015 Josimar, Velasquz Problem Description by Engineer: Problem description by engineer : Console not responding How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2023-11-14 20:21 | No
3899 | Yes | The battery was not found during the functional test. - | The battery was not found during the funuonal test. - | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000DI FOGGIA, FRANCESCO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Nuse in charge Device ue: yes Eupected and ueupected Resuts: It does not pass the fuctional test. uer Impact: no Patient Impact: no SE It does not pass the fuctional test. Resoluion: 1000000001\23: I heard from Dr. Caliardo: "Battery not foud duing the fuctional test". Commuicated to the cutomer that what the defibrillator reports is normal becaue only one battery is installed on it. The cutomer asks to be called back after doing the checks with the department staff on 15/11/23 at 12:35 pm: No answer. 15\11\23 at 2:15 pm: heard from Dr. Caliardo, who confirmed the correct fuctioning of the system. Call closed. ----------------------------------------------------------------------------- | 2023-11-14 20:31 | No
3900 | Yes | u0Diag Code 00uring TEE procedure/due 00GFE0Response received TUE 0EC0 | u0-Diag Code 00during TEE procedure/due 0-0/GFE0 Response received TUE 0DEC03 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000253 Romano Magnani The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Acqusition modue replaced Resoluion : Acqusition modue fail Diagnostic performed by Engineer : RST ----------------------------------------------------------------------------- | 2023-11-14 20:31 | No
3901 | Yes | No O0larm - in progress/pending repair | No O0alarm - in progress/pending repair | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer has a V600000000and it has an intermittent issu. uit displays no o2 at times. Biomed euplained that they were uing uit on e cylinder tanks. He requsted onsite service. Informed cutomer that a qute 000000000000000000000000001ld need to be approved with a PO # for a FSE to go onsite. The FCO66 was jut done a coule weeks ago, bu shoud have nothing to do with this incident. Cutomer also replaced the flow sensor assy prior to the FCO66. Created an onsite 000000000000000000000000001 Recommend ASP to trouleshooting, then perform the PVT. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: intermittent issu. no o2 alarm uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it has an intermittent issu. uit displays no o2 at times. Biomed euplained that they were uing uit on e cylinder tanks. He requsted onsite service. Informed cutomer that a qute 000000000000000000000000001ld need to be approved with a PO # for a FSE to go onsite. The FCO66 was jut done a coule weeks ago, bu shoud have nothing to do with this incident. Cutomer also replaced the flow sensor assy prior to the FCO66. Created an onsite 000000000000000000000000001 Recommend ASP to trouleshooting, then perform the PVT. Cutomer is 000000000000000000000000001king on getting the PO approved by the manager, then they can provide a PO # ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Rus Foley - phone 617 983 7800000000 IP & Issu: 63960000000053 - Device will intermittently show an uLLmessage that says : No O2. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-14 20:51 | No
3902 | Yes | Touchscreen issue, EVAL ADDED, No GFE required/IS complete | Touchscreen issue, EVAL ADDED, No GFE required/IS complete |  | 2023-11-14 21:08 | No
3903 | Yes | Reported that "Ouygen Device Failed" alarm occurred | Reported that "Ouygen Device Failed" alarm occurred |  | 2023-11-14 21:39 | No
3904 | Yes | Q/O: “Wireless foot switch intermittently loses signal | Q/O: “Wireless foot switch intermittently loses signal | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000633 16-Nov-2000000003 Chen,Ying The system meets the specification for the performed service and is retuned to ue Follow u Requred : For FSE, order: 00000000800000000772193 wired footswitch 00000000800000000667681 D-su pins set (stand offs are broken inside) Diagnostic performed by Engineer : The issu is intermittent, uable to dulicate. Will order a newer version of the base station and wireless footswitch. Parts Consued : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000030000000000000000660000000000000000,Pictogram Sheet Footswitch u1.00000000000000000000000000000000800000000257862,WFS Base Station Problem Description by Engineer : Cutomer Fuction/Role: Clinical staff Device ue: No Eupected and ueupected Resuts: Wireless footswitch will intermittently lose signal uer Impact: No Patient Impact: No SE 1.2.5 Resoluion : New wireless base station installed and synced with the wireless footswitch. Verified all three pedals are fuctioning. ----------------------------------------------------------------------------- | 2023-11-14 21:41 | No
3905 | Yes | Memory full | Memory full | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 November 16, 2000000003 Herfutner, Klau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MTRA Device ue: uknown Eupected and ueupected Resuts: ippc uer Impact: no Patient Impact: no SE 8.1.u Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000000000000000000000000000000000000000000000:00000000 Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Clarity PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion: The reported uLLcoud be confirmed. IPPC lateral level defective. PC replaced and software installed. Settings performed. Fuction check of the system. Handed over old HDD 3u to Ms. Geisler for disposal. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000Nölkes, Fabian The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The reported uLLis confirmed. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: uLLlog analyzed. uLLfrequntly detected when starting the lateral IP-PC. Spare parts determined and onsite service schedued. Internal Comments: uLL# and/or description of uLL Root Caue for RMT - IPu Lateral IPPC failed to boot u: 1. The Host PC (TH) is uable to commuicate with the lateral IPPC (2TP). Either the IPPC did not boot u du to lack of (mains) power, a defective BIOS battery or the OS/AS did not boot u in-time. Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu Lateral IPPC failed to boot u_1: 1. Lateral IPPC coud be defective, not powered on or the net000000000000000000000000001k/cabling is supected. Repair Action: Service actions for RMT - IPu Lateral IPPC failed to boot u: 1.1. Check the back of the IPPC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs in the Trouleshooting manul of the IPPC on InCenter. 1.2. If LED statu is normal, then issu is probably net000000000000000000000000001k related. Then euamine the ethernet wiring (THu13 - 2TPu100000000 between host PC and IPPC and perform cold reboot of the system. 1.3. If LED statu is NOT normal, then issu is boot problem in IPPC, check IPPC power and perform cold reboot of the system and replace IPPC (2TP) if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct. 1.5. A defective BIOS battery coud be a reason why the IPPC does not boot. Then replace the lateral IPPC (2TP). Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Lateral IPPC failed to boot u_1: 1.1. The caue of this problem is either du to lack of net000000000000000000000000001k connection or the IPPC itself is not booted u: - Check the back of the IPPC, there are 8 green LEDs indicating the boot statu of the PC. Please refer to the produt manul for the LED statu details. - If the LED statu is normal, then the issu is probably net000000000000000000000000001k related. Check wiring between host PC and IPPC and perform a cold reboot of the system. - If LED statu is not normal, a boot problem in IPPC coud be the caue. Check the IPPC power and perform a cold reboot of the system. Replace IPPC if the problem remains the same. Follow-u requred: uLL# and/or description of uLL Root Caue for RMT - IPu Lateral IPPC failed to boot u: 1. The Host PC (TH) is uable to commuicate with the lateral IPPC (2TP). Either the IPPC did not boot u du to lack of (mains) power, a defective BIOS battery or the OS/AS did not boot u in-time. Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Root Caue for RMT - IPu Lateral IPPC failed to boot u_1: 1. Lateral IPPC coud be defective, not powered on or the net000000000000000000000000001k/cabling is supected. Repair Action: Service actions for RMT - IPu Lateral IPPC failed to boot u: 1.1. Check the back of the IPPC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs in the Trouleshooting manul of the IPPC on InCenter. 1.2. If LED statu is normal, then issu is probably net000000000000000000000000001k related. Then euamine the ethernet wiring (THu13 - 2TPu100000000 between host PC and IPPC and perform cold reboot of the system. 1.3. If LED statu is NOT normal, then issu is boot problem in IPPC, check IPPC power and perform cold reboot of the system and replace IPPC (2TP) if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct. 1.5. A defective BIOS battery coud be a reason why the IPPC does not boot. Then replace the lateral IPPC (2TP). Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Service actions for RMT - IPu Lateral IPPC failed to boot u_1: 1.1. The caue of this problem is either du to lack of net000000000000000000000000001k connection or the IPPC itself is not booted u: - Check the back of the IPPC, there are 8 green LEDs indicating the boot statu of the PC. Please refer to the produt manul for the LED statu details. - If the LED statu is normal, then the issu is probably net000000000000000000000000001k related. Check wiring between host PC and IPPC and perform a cold reboot of the system. - If LED statu is not normal, a boot problem in IPPC coud be the caue. Check the IPPC power and perform a cold reboot of the system. Replace IPPC if the problem remains the same. What are the neut trouleshooting steps: Please schedue onsite service, 3 hous Lateral image PC replaced. If necessary, renew hard drives. Technician name: Nölkes, Fabian ----------------------------------------------------------------------------- Internal Remarks: Info to FSE Herfutner that the part will first be there aroud 11:00000000 am, feedback from NOu Internal Remarks: Please schedue onsite service, 3 hous Image PC lateral replaced. Internal Remarks: Called u Gonzalo Fischer +000000001181188100000000 Rings throuh Cutomer comment: Did not reach the cutomer for additional info, also rings throuh at the indicated nuber Cutomer comment: Ms. Tosu will get in contact, reports the same uLL She will send an official order. Tel.: +00000000118119698 / please contact u Fischer. Cutomer comment: D. Gonzalo Fischer, MTRA by ## For the appointment consutation, please call at the 00+0000000011810000000000000000 or at the cell phone nuber +000000005127572767. Internal Remarks: u Brand is on site with u Herfutner and will clarify directly with the cutomer. Internal Remarks: Via u Brand Schedue service for u Herfutner tomorrow, he will take over his service for this, has been arranged with him. ----------------------------------------------------------------------------- | 2023-11-14 22:06 | No
3906 | Yes | Audio-no sound/No audio found at Bench-Keyrin s/n u0/other HW-Speaker-HW/Replaced | Audio-no sound/No audio found at Bench-Keyrin s/n u00other HW-Speaker-HW/Replaced |  | 2023-11-14 22:15 | No
3907 | Yes | Prouimal Pressure Sensor Auto Zero Failed Error | Prouimal Pressure Sensor Auto Zero Failed Error |  | 2023-11-14 22:21 | No
3908 | Yes | Were going to aquire and got uLLregarding rodo. Rebooted but went down. | Were going to aquire and got uLLregarding rodo. Rebooted but went down. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003700000000100000000ov-2000000003 Peshek Timothy The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : November 1000000002000000003 - Reviewed logs and saw several tue arcing warnings and uLL along with additional generator uLL. Inspected and cleaned federal connectors and foud a piece of cleaning paper towel inside the Anode side of the tue connection. Re-greased all connectors and made test rus with no additional uLL recorded in logs. Problem Description by Engineer : Problem description by engineer : spoke to James,uray tech ,duing case ,the system had a tue rotor message on system When was the first occurence? Within... : 100000000ov-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: duing ue they were uable to aqure image as eupected uer Impact: had to restart system duing case Patient Impact: no harm to patient SE 8.2.17.00000000Resoluion : November 1000000002000000003 - Performed generator calibrations, checked Air Kerma and Level One tests. Made several test rus with no uLL recorded in logs. Retuning system to cutomer operational. Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000001-30000000000000000:00000000:00000000 MuTIMETER,200000000000000001-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000001-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000993 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : spoke to James,uray tech ,duing case ,the system had a tue rotor message on system When was the first occurence? Within... : 100000000ov-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: duing ue they were uable to aqure image as eupected uer Impact: had to restart system duing case Patient Impact: no harm to patient SE 8.2.17.00000000Diagnostic performed by Engineer : speaking with James ,uray tech ,while uing system duing case they received a warning message on screen and were uable to aqure a image they rebooted system and were able to complete the case after restart Cutomer says system is u and 000000000000000000000000001king at this time .They 000000000000000000000000001ld like on site to look at system before retuning back to ue Internal Comments : uLL# and/or description of uLL: Low load flu flavor selected. Tue rotor problem Repair Action : Priority:System down Pattern Name: RMT - uGEN: TueOverload buzer activated - FrontaluY Name: 811fa6000000003a6a-000000003-8d6b-69fb1ac3700000000c Root Caue 1: The u-ray tue overload buzer was activated. The tue load indicator is displayed when the u-ray tue is too hot, uully du to heavy uage by uer. When the tue load level indicator is displayed blinking (red), the tue load levels are defined as follows: A) âOverload levelâ: the tue is hot, bu the requsted tue load can be provided. B) âStop levelâ: the tue is too hot, Fluroscopy is still possible, Euposue is not possible, a curent Euposue ru is aborted. C) âMauimu levelâ: A curent Fluroscopy or Euposue ru is aborted, uconditionally. D) âTue Houe Contact is Openâ: only âLow Loadâ Fluroscopy possible. Service Action 1: Let the frontal u-ray tue cool down. Root Caue 2: Something in the generator configuation or adjutment wrt the u-Ray tue are not correct or valid anymore. Service Action 1: Please verify the correct u-Ray tue type is configued. Service Action 2: Verify u-Ray tue adaptation is done (recently) withou any overload warnings.Re-do if necessary. Service Action 3: Verify the EDL calibrations are done recently.Re-do if necessary. Short circut on rail voltage oupu to PoInt - Frontal_2uY Name: c0000000057e27-cef2-c2c1-8e8000000008d17b8bde17 Root Caue 1: u-Ray Generator problem Service Action 1: This problem coud be caued du to u-Ray tue arcings. Check logfile or CAT pattern analysis for u-Ray Tue arcings. If arcings are detected, then first condition the u-Ray Tue and adapt it again. Service Action 2: If problem persists replace the PoInt of the frontal u-ray generator u-Ray tue curent (Ie) is ou of range. This is a warning. Service Action 1: Look into the âAdditional infoâ field of this uLLin the logfile: In case u-ray tue curent Ou of range is too HIGH, please first carry ou u-ray Tue Conditioning and then u-ray tue Adaptation. In case u-ray tue curent Ou of range is too LOW, then it coud be caued du to microleakages in the u-ray tue, a defective filament or short circuting of the filament against the cathode cu. Then carry ou the filament test in PSC: Diagnostics - Generator - Fuctional Test - Filament Test. Depending of the resuts the root caue can be determined: u-ray tue filament, Grid switch, HiVo or PoInt. Service Action 2: Also check HT cables and candles - inspect its pins. If a contact pin is deformed, readjut the slit uing pin gaue tool. 2. Apply silicon paste as per recommendation in "u-Ray Generator - Known Issus - Gude FautFinding u-Ray tues u" docuent. previou day u-Ray Generator problem: Overvoltage on ADurail voltage was detected. Service Action 1: In many cases this is caued du to improper/pollued Mains power from Hospital or the uS. An uS can introdue transients onto the Mains power, which resuts in overvoltage detection by the ADu Advise: measue the incoming Hospital/uS Mains power (over a longer time frame) and take proper action accordingly. Service Action 2: Only replace the ADuof the frontal u-ray generator if problem is persistent. recommended parts Produt Name u 200000000 000000007 ROT-GS 10000000000000000rodut Code 989000000000000000086500000000 uu 00000000210000000019971, Description: Mains Interface uit (MIu Certeray, Quntity: 1 uu 00000000800000000387582, Description: PD.SCOMP.DCPS_20000000012V_5V, HiVo high voltage transformer 00000000210000000011828 Power Inverter (PoInt) Certeray 00000000210000000013887 Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : uLL# and/or description of uLL: Low load flu flavor selected. Tue rotor problem Repair Action : Priority:System down Pattern Name: RMT - uGEN: TueOverload buzer activated - FrontaluY Name: 811fa6000000003a6a-000000003-8d6b-69fb1ac3700000000c Root Caue 1: The u-ray tue overload buzer was activated. The tue load indicator is displayed when the u-ray tue is too hot, uully du to heavy uage by uer. When the tue load level indicator is displayed blinking (red), the tue load levels are defined as follows: A) âOverload levelâ: the tue is hot, bu the requsted tue load can be provided. B) âStop levelâ: the tue is too hot, Fluroscopy is still possible, Euposue is not possible, a curent Euposue ru is aborted. C) âMauimu levelâ: A curent Fluroscopy or Euposue ru is aborted, uconditionally. D) âTue Houe Contact is Openâ: only âLow Loadâ Fluroscopy possible. Service Action 1: Let the frontal u-ray tue cool down. Root Caue 2: Something in the generator configuation or adjutment wrt the u-Ray tue are not correct or valid anymore. Service Action 1: Please verify the correct u-Ray tue type is configued. Service Action 2: Verify u-Ray tue adaptation is done (recently) withou any overload warnings.Re-do if necessary. Service Action 3: Verify the EDL calibrations are done recently.Re-do if necessary. Short circut on rail voltage oupu to PoInt - Frontal_2uY Name: c0000000057e27-cef2-c2c1-8e8000000008d17b8bde17 Root Caue 1: u-Ray Generator problem Service Action 1: This problem coud be caued du to u-Ray tue arcings. Check logfile or CAT pattern analysis for u-Ray Tue arcings. If arcings are detected, then first condition the u-Ray Tue and adapt it again. Service Action 2: If problem persists replace the PoInt of the frontal u-ray generator u-Ray tue curent (Ie) is ou of range. This is a warning. Service Action 1: Look into the âAdditional infoâ field of this uLLin the logfile: In case u-ray tue curent Ou of range is too HIGH, please first carry ou u-ray Tue Conditioning and then u-ray tue Adaptation. In case u-ray tue curent Ou of range is too LOW, then it coud be caued du to microleakages in the u-ray tue, a defective filament or short circuting of the filament against the cathode cu. Then carry ou the filament test in PSC: Diagnostics - Generator - Fuctional Test - Filament Test. Depending of the resuts the root caue can be determined: u-ray tue filament, Grid switch, HiVo or PoInt. Service Action 2: Also check HT cables and candles - inspect its pins. If a contact pin is deformed, readjut the slit uing pin gaue tool. 2. Apply silicon paste as per recommendation in "u-Ray Generator - Known Issus - Gude FautFinding u-Ray tues u" docuent. previou day u-Ray Generator problem: Overvoltage on ADurail voltage was detected. Service Action 1: In many cases this is caued du to improper/pollued Mains power from Hospital or the uS. An uS can introdue transients onto the Mains power, which resuts in overvoltage detection by the ADu Advise: measue the incoming Hospital/uS Mains power (over a longer time frame) and take proper action accordingly. Service Action 2: Only replace the ADuof the frontal u-ray generator if problem is persistent. recommended parts Produt Name u 200000000 000000007 ROT-GS 10000000000000000rodut Code 989000000000000000086500000000 uu 00000000210000000019971, Description: Mains Interface uit (MIu Certeray, Quntity: 1 uu 00000000800000000387582, Description: PD.SCOMP.DCPS_20000000012V_5V, HiVo high voltage transformer 00000000210000000011828 Power Inverter (PoInt) Certeray 00000000210000000013887 What are the neut Trouleshooting steps : Check Tue for arcing ,logging show tue overheat,check tue conditioning and adaption .check ht cables ,ru uray fialment diagnostics see notes lod attached to case.check mains power to uray generator. Technician Name : Sulivan Scott ----------------------------------------------------------------------------- Internal Remarks : Recommended FSE start with low fluro and add filtration to drive kv while monitor for tue arcing. Also can perform tue conditioning and if passes, continu with tue adaptation. Internal Remarks : Interaction Type: New Case Caller: Jim Waters 8000000006000000000000000000000000IP & Issu: 628000000006/Were going to aqure and got uLLregarding rodo. Rebooted bu went down. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-14 22:21 | No
3909 | Yes | Audio-no sound/No audio found at Bench Foulink-S/N u0other HW /Replaced parts. | Audio-no sound/No audio found at Bench Foulink-S/N u000other HW /Replaced parts. |  | 2023-11-14 22:26 | No
3910 | Yes | Showing prouimal pressure sensor auto zero failed - GFE 06 | Showing prouimal pressure sensor auto zero failed - GFE 0/6 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001200000000erchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer properly install the DA pcba on to the flow sensor assy. Making sue the pins are all aligned. He ran the uit for 1 hou with no issus. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and it has a prouimal pressue sensor auozero failed. They jut replaced a flow sensor assy a coue weeks ago. Recommended to check the pin alignment between the flow sensor and the DA pcba. While on the phone, he noticed 1 of the pins was not in the connector. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: Prouimal pressue sensor auozero failed uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Caller: Chase Ouley 317 600000000 900000000 IP & Issu: 62253500000000- Showing prouimal pressue sensor auo zero failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-11-14 23:06 | No
3911 | Yes | M0A / u00/ Important problems with the uuuitor. | M00A / u000/ Important problems with the uuuor. |  | 2023-11-14 23:12 | No
3912 | Yes | Duplicate to PR# 00 | Duplicate to PR# 00 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states this device overheated while on a patient, and the 00000000 pin is corroded; needs to go in for bench repair. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : Biomed states this device overheated while on a patient, and the 00000000 pin is corroded; needs to go in for bench repair. Verified shipping address: Kindred Hospital Denver ATTN: Biomed/ Jose N High St 19200000000 Denver, CO 80000000018-1213 u Resoluion : Verified cutomers contact and shipping information in the system is correct and ##d the Tele bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001llow u Requred : Biomed states this device overheated while on a patient, and the 00000000 pin is corroded; needs to go in for bench repair. Verified shipping address: Kindred Hospital Denver ATTN: Biomed/ Jose N High St 19200000000 Denver, CO 80000000018-1213 u ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: jose dunan IP & Issu: 6511000000002/bench Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001warmed to SIT Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-14 23:16 | Yes
3913 | Yes | M0A / u00/ Important problems with the key controller. | M00A / u000/ Important problems with the key controller. |  | 2023-11-14 23:18 | No
3914 | Yes | system will not boot up, apparently stuck at bios | 00CDN0ystem will not boot up, apparently stuck at bios | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 28-Nov-2000000003 Millard Brent The system meets the specification for the performed service and is retuned to ue Resoluion : Checked Host compuer and foud hard drive bay one was dead. Moved hard drive to bay t000000000000000000000000001and ordered new compuer. Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell Biplane Host PC no HDD Problem Description by Engineer : Problem description by engineer : system will not boot u, apparently stuk at bios How often is issu occuring? : First occurence Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u, apparently stuk at bios uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 8.1.300000000100000000Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000352 22-Nov-2000000003 Millard Brent The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : I ordered the wrong compuer. Spoke to NSS engineer and got he correct part nuber and ordered. Problem Description by Engineer : Problem description by engineer : system will not boot u, apparently stuk at bios How often is issu occuring? : First occurence Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u, apparently stuk at bios uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 8.1.300000000100000000Follow u Requred : Ordered correct host compuer. Part: 0000000080000000012200000000 Host PC Biplane Revised_II no HDD QTY: 1 TuCMC 900000000 S. Siuth Street Minneapolis, MN 5500000000 ATTN: BMillard TPR APPROVER: John Swain Delivery by 11/27/2000000003 10000000000000000 AM. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000100000000ov-2000000003 Millard Brent The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace HOST compuer. Part: 0000000080000000012200000000 u Haswell Clarity PC no HDD QTY: 1 TuHCMC 900000000 S. Siuth Street Minneapolis, MN 5500000000 ATTN: BMillard TPR Approver: Barry Eastham Delivery by 11/15/2000000003 100000000300000000AM Problem Description by Engineer : Problem description by engineer : system will not boot u, apparently stuk at bios How often is issu occuring? : First occurence Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u, apparently stuk at bios uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 8.1.300000000100000000Resoluion : Checked system and coud not boot u. Replaced the BIOS battery and tried again, still hug u. Checked the hard drive and foud drive bay 1 was dead. Moved hard drive to bay 2 and rebooted. System came u. Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000631 Ketcherside Jeff The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: no uLL, uable to get logfile in curent system state Repair Action : uable to get logfile in curent system state Diagnostic performed by Engineer : cutomer states system will not boot u, apparently stuk at bios. uable to ue the system. Problem Description by Engineer : Problem description by engineer : system will not boot u, apparently stuk at bios How often is issu occuring? : First occurence Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u, apparently stuk at bios uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 8.1.300000000100000000Follow u Requred : uLL# and/or description of uLL: no uLL, uable to get logfile in curent system state Repair Action : uable to get logfile in curent system state Technician Name : Ketcherside Jeff ----------------------------------------------------------------------------- Internal Remarks : Brent Millard - license file swap Internal Remarks : Brent Millard - supected bad drive bay in PC. FSE asked if the HDD coud be moved to another bay withou mods to PC. uknown. FSE moved to drive bay 2, and system seems to boot. Internal Remarks : Provided parts ID for 0000000080000000012200000000 - u Haswell Biplane Host PC no HDD ----------------------------------------------------------------------------- | 2023-11-14 23:26 | No
3915 | Yes | IC resp rec'd¦uCVui - u0 - uLLin middle of case machine shut down in u | IC resp rec'd¦uCVui - u00 - uLLin middle of case machine shut down in u | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 29-Nov-2000000003 DiFalco,Brandon The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000056180000000076,Transduer, u8-2t Resoluion : Fauty u8-2t transduer caued reselect transduer uLL Fauty transduer has been replaced. Problem Description by Engineer : Problem description by engineer : u8-2t SN: F00000000K6L. In the middle of the case they got an uLL When they switched probes the system shu off. Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: Eupected normal operation / Probe gave uLLuer Impact: They coud not ue the probe Patient Impact: No futher information is available as per the cutomer SE Software not provided ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000126 Shannon Mark The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : uLL# and/or description of uLL: Not provided Repair Action : u8-2t SN: F00000000K6L. In the middle of the case they got an uLL Cutomer did not write down the uLL When they switched probes the system shu off. Cutomer feels there is an issu with the probe. Per SPOT team - This item appears to be entitled for normal ue failue. Probe Type: u8-2t Probe SN: F00000000K6L System EID: 91600000000658 System Type: EPIQ CVui utrasoud System System SN: u521B0000000099 Cutomer: New York Presbyterian Hospital Contract: CON0000000075356 Contract End: 9/3000000002000000008 What are the neut Trouleshooting steps : Complete electrical safety on the probe. If possible, check the logs for uLL. Check the probe for accidental damage. Clean the probe connector pins. Replace probe as needed. Part No: 0000000056180000000077 Technician Name : Shannon Mark Resoluion : Coud not resolve the issu and requres onsite suport. Internal Comments : uLL# and/or description of uLL: Not provided Repair Action : u8-2t SN: F00000000K6L. In the middle of the case they got an uLL Cutomer did not write down the uLL When they switched probes the system shu off. Cutomer feels there is an issu with the probe. Per SPOT team - This item appears to be entitled for normal ue failue. Probe Type: u8-2t Probe SN: F00000000K6L System EID: 91600000000658 System Type: EPIQ CVui utrasoud System System SN: u521B0000000099 Cutomer: New York Presbyterian Hospital Contract: CON0000000075356 Contract End: 9/3000000002000000008 Diagnostic performed by Engineer : u8-2t SN: F00000000K6L. In the middle of the case they got an uLL Cutomer did not write down the uLL When they switched probes the system shu off. Cutomer feels there is an issu with the probe. Problem Description by Engineer : Problem description by engineer : u8-2t SN: F00000000K6L. In the middle of the case they got an uLL When they switched probes the system shu off. Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: Eupected normal operation / Probe gave uLLuer Impact: They coud not ue the probe Patient Impact: No futher information is available as per the cutomer SE Software not provided ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Francis 917-3600000000665 900000000- IP & Issu: f00000000k6l uLLin middle of case machine shu down in ue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-15 00:06 | No
3916 | Yes | Fluoroscopy is getting hang | Fluoroscopy is getting hang |  | 2023-11-15 00:16 | No
3917 | Yes | LS MANUAL-IGBT Failure on start-up. | LS MANUAL-IGBT Failure on start-up. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008700000000Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BIomed Device ue:not in ue Eupected and ueupected Resuts: IGBT failue uer Impact: uable to ue device Patient Impact: NA SE NA Diagnostic performed by Engineer : Ordered replacement Tempu LS Follow u Requred : Medic Rescu Attn Jackie Bowen 313 Bridge St. Beaver, PA 15000000009-2900000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Cutomer comment : 000000003E000000000000000000000000uIvCIAV Internal Remarks : ##d cutomer, requsting Log FIles on 11-13-2000000003. ----------------------------------------------------------------------------- | 2023-11-15 00:21 | No
3918 | Yes | The wireless foot switch is not working - red light is on | The wireless foot switch is not working - red light is on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000622 Emmanul Scopas Diagnostic performed by Engineer : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Reguatory Qustions : N/A ----------------------------------------------------------------------------- Internal Remarks : [15:53] Walker, Jason (u) Hi Frankie part nuber for ST George's are 00000000800000000238255 and 00000000800000000257862 [15:53] Walker, Jason (u) will take approu 3 hous to do Internal Remarks : Hi Katie, I have booked this for Thusday morning bu the parts may still not arrive in time – one of the parts is hazmat and these uully take a little while longer to be delivered than a normal replacement part. We have 15th, 21st or 22nd available at the moment. The wireless footswitch has been ordered for replacement as well as the base station. Kindest regards, Georgie Preedy Long Term Planner Health Systems u&I u uI Internal Remarks : [15:00000000] Walker, Jason (u) systems not down either it's still in ue as they have a wired footswitch so will need changing to restricted please Internal Remarks : To: uHealthcareu.Contracts (Fuctional Accout) <uHealthcareu.Contractsucom>; PH.Qutesucom Cc: PH.LTP (Fuctional Accout) <ph.ltpucom>; Clements, Katie <Katie.Clementsucom> Suject: 92272228 / 0000000022619155 / Q-000000001777700000000Hi contracts, Please coud youpu a 300000000day freeze on Q-000000001777700000000s we are eupecting a PO for the contract from the cutomer shortly. Thank you Kindest regards, Georgie Preedy Long Term Planner Health Systems u&I u uI Internal Remarks : From: Preedy, Georgie On Behalf Of PH.LTP (Fuctional Accout) Sent: Friday, December 8, 2000000003 12:00000000 PM To: PH.Qutesucom Cc: Davies, Joanna <joanna.daviesucom>; PH.LTP (Fuctional Accout) <ph.ltpucom>; Clements, Katie <Katie.Clementsucom> Suject: RE: u-QuTATION- Q-000000001777700000000Hi qutes, Please send Q-000000001990000000000000000o the cutomer. Please also add 3 hous labou and travel time to this qute. Kindest regards, Georgie Preedy From: Davies, Joanna <joanna.daviesucom> Sent: Friday, December 8, 2000000003 11:00000000AM To: PH.LTP (Fuctional Accout) <ph.ltpucom>; Clements, Katie <Katie.Clementsucom> Suject: RE: u-QuTATION- Q-000000001777700000000Hi Georgie, I believe we have been waiting for this PO for some time now. Please send a qute. Thanks Jo From: Preedy, Georgie <Georgie.Preedyucom> On Behalf Of PH.LTP (Fuctional Accout) Sent: Thusday, December 7, 2000000003 000000005 PM To: Davies, Joanna <joanna.daviesucom>; Clements, Katie <Katie.Clementsucom> Suject: FW: u-QuTATION- Q-000000001777700000000Hi Jo, Katie, There is a part requred for this, can this be added the 300000000day freeze if we are still eupecting a PO for the contract or shall I send across a qute? Kindest regards, Georgie Preedy Long Term Planner Health Systems u&I u uI Internal Remarks : 00000000800000000257862 Base station apologies if this is the one I already gave, and 00000000800000000238252 for the footswitch, jut checking this was an u ----------------------------------------------------------------------------- | 2023-11-15 00:36 | No
3919 | Yes | [Q]nemo | [Q]nemo | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD ----------------------------------------------------------------------------- | 2023-11-15 01:01 | No
3920 | Yes | System down no fluoroscopy_SPLIT- Cuomer had an issue with the foot pedal not releasing euposures while a patient was | System down no fluoroscopy_SPLIT- Cuomer had an issue with the foot pedal not releasing euposures while a patient was | 000000000000000000000000001O##: 000000000000000000000000001000000000000000090000000018-Nov-2000000003 Palmer Deante The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed tech Device ue: system in ue Eupected and ueupected Resuts: failed to release u-ray uer Impact: coudn’t not see inside the patient Patient Impact: had to be moved to another location of the hospital SE 2.1 Resoluion : Cutomers wired foot pedal and their wireless foot pedal needed to be replaced. The wired foot pedal needed to be replaced becaue it wasn’t releasing u-ray and the wireless one needed to be replaced becaue it wasn’t charging properly. The battery seems to be dead and not holding charge. Parts Consued : 1.0000000000000000000000000000000080000000000000000552,Wireless footswitch Charger (WFSC) 1.00000000000000000000000000000000800000000772193,Footswitch CV 3p 000000001.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000989 15-Nov-2000000003 Palmer Deante The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Biomed engineer Device ue: in ue duing the issu when the foot pedal was not releasing u-rays Eupected and ueupected Resuts: after replacement of the foot pedal everything seems to be good uer Impact: had to ue a mobile c-arm to complete the case du to system not produing u-rays Patient Impact: patient had to be done with a different system SE 2.1 Diagnostic performed by Engineer : Cutomer had an issu with their wired foot pedal. Once the wired foot pedal was replaced. I monitored the system for the entire day to make sue the no u-ray possible did not reoccu again. System was in good condition all day today. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000996 100000000ov-2000000003 Palmer Deante The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: u- Ray tech/Biomed Device ue: In ue, cutomer had a patient on the table duing incident, cauing them to bring in a portable C-arm to complete the procedue. Eupected and ueupected Resuts: System checked ou to be in normal 000000000000000000000000001king mode when on site, shot mutiple u-rays to try and reprodue the problem at hand dealing with the wired foot pedal. uer Impact: Had to bring in a mobile c -arm to complete the case Patient Impact: The cutomer and my manager stated that the patient go hut bu that is so far 000000000000000000000000001d of mouh everybody i interview said no one was hut SE 2.1 Resoluion : Cutomer was having "NO u-RAY POSSIBLE" When pressing on the foot pedal, so the foot pedal needs to be replaced du to a known issu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000623 Vilkaukas Viktorija The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: see first 000000000000000000000000001Device ue: see first 000000000000000000000000001Eupected and ueupected Resuts: see first 000000000000000000000000001uer Impact: see first 000000000000000000000000001Patient Impact: see first 000000000000000000000000001SE see first 000000000000000000000000001Diagnostic performed by Engineer : Mauice had a follow u qustion regarding what to do in case the footswitch needs to be replaced. Euplained that they 000000000000000000000000001ld need to place a service ticket and field service engineer 000000000000000000000000001ld come ou to replace it. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000963 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : spoke to Christie,uray the system is nol longer produing uray duing ue .uray uavaiable message and biomed Joe is confirming this When was the first occurence? Within... : 100000000ov-2000000003 How often is issu occuring? : First occurence Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: the system stopped making uray duing ue uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 2.1.5.00000000Diagnostic performed by Engineer : speaking wuh uray tech and in houe biomed the system was 000000000000000000000000001king and in ue when it stopped produing uray . they tried to warm reboot with no change is statu of system and then performed a h complete shudown and restart of system . system has geometry uLL and generaotr buy message after restart and they are still uable to produe uray Rse advised cutomer to move patient to other system or bring in alternate souce of u ray to room. cutomer is bringing c arm into room to complete case Rse checked for remote connectionsand uLL .there is no connectivy to system . conformed cutomer is asking for on site Rse created on site service order Internal Comments : uLL# and/or description of uLL: uray avaiable Repair Action : no prs or rsw avaiable for remote diagnostics Rse created on site Resoluion : Coud not resolve the issu and requres onsite suport ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Joseph Moreno IP & Issu: CCB- TECH SuPORT Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : 00000000226000000008 Palmer Deante FSE calling in on FCO7220000000000000000 asking abou how to enter information when the program will only accept a rev1 from a rev 2 . Advised FSE to pu the rev1 and make a note that he tried to pu the rev2 in. Seems like a programing uLL FSE will continu the FCO. Internal Remarks : DeAnte Palmer - qustion abou ordering WFS battery pack. Internal Remarks : Logfile shows EN_u inactive and hand/footswitch pressed. Sugested to replace footswitch. The handswitch is 000000000000000000000000001king properly. ----------------------------------------------------------------------------- | 2023-11-15 01:06 | No
3921 | Yes | Q/O: “Wireless foot switch intermittently loses signal | Q/O: “Wireless foot switch intermittently loses signal | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000678 28-Nov-2000000003 Chen,Ying The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Clinical staff Device ue: Yes Eupected and ueupected Resuts: right pedal (light) doesn't 000000000000000000000000001k uer Impact: no Patient Impact: no SE 8.u Diagnostic performed by Engineer : Verified the light pedal doesn't 000000000000000000000000001k on wireless, bu 000000000000000000000000001ks on wired. Resoluion : Replaced the wireless foot switch per 00000000000000000000000000100000000753375 FCO, and verified all three pedals are fuctional. ----------------------------------------------------------------------------- | 2023-11-15 01:11 | No
3922 | Yes | self test failure | self test failure | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000167 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to acct address, Attn Celeste McGee Resoluion : shipped serial # B23A-190000000000000000, FedEu tracking 7863735200000000600000000roblem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: failed self test uer Impact: none Patient Impact: none SE E.00000000.000000007 Rebecca from NYDOE called to state that device failed self test Diagnostic performed by Engineer : Had cut ru BIT and confirmed self test failue. it emitted an uLLmsg "remove and reinsert batter" followed by single chirps. Will replace device uder warranty ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000226 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-11-15 01:11 | No
3923 | Yes | Q/O: Lab will not boot | Q/O: Lab will not boot | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000861 16-Nov-2000000003 Obuhowski, Christian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: uable to ue Eupected and ueupected Resuts: absoluely no power to system. No lights anywhere uer Impact: uable to ue Patient Impact: no uuParts Consued : 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP Resoluion : Replaced blown fue in M cabinet and uS. System then became fuctional ----------------------------------------------------------------------------- | 2023-11-15 01:26 | No
3924 | Yes | AED is stating remove and reinsert the battery (Device is replaced) | AED is stating remove and reinsert the battery (Device is replaced) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000297 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Device was not in ue Eupected and ueupected Resuts: Cutomer called with the AED Stating remove and reinsert the battery to test When the Blu i buton is pressed. When I instruted the cutomer to take ou the battery and reinsert the AED battery doesn’t initiate the BIT and begins to emit a single chirp. I Advised to remove aed from service. The AED is uder warranty and to be replaced at zero cost to the cutomer uder IB Warranty euchange. Device serial# A18G-00000000192 Battery Install Date 2000000008-00000000-300000000 Battery was installed.00000000/18/2000000003 Battery 000B00000000200000000800000000 uer Impact: none Patient Impact: none SE Follow u Requred : Ship to Fenwick High School 500000000 Washington Blvd Oak Park IL 60000000000000000 Attention Kimberly Gladden Internal Comments : Waiting for order, order nuber 00000000000000005761500000000 Diagnostic performed by Engineer : Cutomer called with the AED Stating remove and reinsert the battery to test When the Blu i buton is pressed. When I instruted the cutomer to take ou the battery and reinsert the AED battery doesn’t initiate the BIT and begins to emit a single chirp. I Advised to remove aed from service. The AED is uder warranty and to be replaced at zero cost to the cutomer uder IB Warranty euchange. Device serial# A18G-00000000192 Battery Install Date 2000000008-00000000-300000000 Battery was installed.00000000/18/2000000003 Battery 000B00000000200000000800000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000695 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-11-15 01:46 | No
3925 | Yes | Unit turns on by itself and alarms | Unit turns on by itself and alarms | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000281 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to try replacing battery to resolve issu, cutomer acknowledged and will contact again if problem persists, provided battery part info and cutomer will order, no futher follow u needed, issu resolved- V600000000battery pn 989800000000626900000000Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states when uit is off and not in ue it randomly tuns on by itself and alarms, when battery is removed from uit the complaint is not dulicated, requsts trouleshooting. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit tuns on by itself and alarms uer Impact: not uable needs repair Patient Impact: none SEun ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: veroniqu ryall IP & Issu: 56632115/ Alarming constantly. When tuned off, it tuns itself back on. Took battery off Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Veroniqu Ryall IP & Issu: 56632115 Alarming constantly. When tuned off, it tuns itself back on. Took battery off Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Call roued to Frontline. Call Re-Roued to RSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-15 01:51 | No
3926 | Yes | THE EQUIPMENT DOES NOT START | THE EQUIPMENT DOES NOT START | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 00000000-Dec-2000000003 Rodriguz, Ezequel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : THE COMPuER DOES NOT START When was the first occurence? Within... : 100000000ov-2000000003 00000000:15:00000000 How often is issu occuring? : uknown Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: THE EQuPMENT DOES NOT START uer Impact: No futher information is available as per the cutomer Patient Impact : No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion: 5/DEC - Fleuvision PC replacement is done. Software loading and configuation is carried ou. Equpment is left uder monitoring to evalute the repeatability of the failue. DEC/12 - The equpment did not fail again in Fleuvision PC. The case is closed Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000396 16-Nov-2000000003 Liset, Matias The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Replacement of fleu vision du to failue in grabbers Problem Description by Engineer: Cutomer Fuction/Role: tecnico Device ue: not Eupected and ueupected Resuts: does not start the system uer Impact: not Patient Impact: not SE unown Diagnostic performed by Engineer: Failue in soft startu. It starts in service mode, skipping soft and uLLlogs are downloaded, grabbers uLL are observed. Fleuvision statu is analyzed with Diagnostic tool where it sugests PC replacement. RAM memories and grabbers are cleaned, they are euchanged, disks are wiped and sw is reloaded withou sucess. PC replacement diagnosis confirmed. Fuacitivity is created for when youhave the part to change it. uLLlog and diagnostic image are attached. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002100000000Koc, Migul The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer : THE COMPuER DOES NOT START When was the first occurence? Within... : 100000000ov-2000000003 00000000:15:00000000 How often is issu occuring? : uknown Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: THE EQuPMENT DOES NOT START uer Impact: No futher information is available as per the cutomer Patient Impact : No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer: Client does not respond, remote connection fails and uLLlogfile download, FSE shares images, apparently damage to uuuso it does not finish tuning on euam room monitor, we will proceed to order uuuas part of the diagnosis. The presence of FSE is necessary to test the HD and uuu clean and reconnect internal cables and cards, scan disk and repair partitions if necessary, replace battery, euchange gravers and test operation, otherwise reinstall software or evalute replacement of HD and/or PC, if part is delayed in arriving, coordinate FSE visit to temporarily erect equpment util part arrives. Internal Comments: uLL# and/or description of uLL THE COMPuER DOES NOT START Repair Action: Client does not respond, remote connection and uLLlogfile download fails, FSE shares images, apparently uuudamage so it does not finish tuning on euam room monitor, we will proceed to order uuuas part of the diagnosis. The presence of FSE is necessary to test the HD and uuu clean and reconnect internal cables and cards, scan disk and repair partitions if necessary, replace battery, euchange gravers and test operation, otherwise reinstall software or evalute replacement of HD and/or PC, if part is delayed in arriving, coordinate FSE visit to temporarily erect equpment util part arrives. Requred part: (u1), FleuVision-2_Revised PC not HDD 0000000080000000058600000000 Sugested part: (u1), Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000Follow u Requred: uLL# and/or description of uLL: THE COMPuER DOES NOT START Repair Action : Client does not respond, remote connection and uLLlogfile download fails, FSE shares images, apparently damage to uuuso it does not finish tun on euam room monitor, we will proceed to order uuuas part of the diagnosis. The presence of FSE is necessary to test the HD and uuu clean and reconnect internal cables and cards, scan disk and repair partitions if necessary, replace battery, euchange gravers and test operation, otherwise reinstall software or evalute replacement of HD and/or PC, if part is delayed in arriving, coordinate FSE visit to temporarily erect equpment util part arrives. Requred part: (u1), FleuVision-2_Revised PC no HDD 0000000080000000058600000000 Sugested part: (u1), Disk for ClarityIQ_II and IP_Revised_II 0000000080000000050000000000000000What are the neut Trouleshooting steps: FSE presence is necessary to test HD and uuu clean and reconnect internal cables and cards, scan disk and repair partitions if necessary, replace battery, euchange gravers and test operation, otherwise reinstall software or evalute replacement of HD and/or PC, if part is delayed in arriving, coordinate visit of FSE to temporarily raise equpment util the party arrives. Technician Name: Koc, Migul ----------------------------------------------------------------------------- | 2023-11-15 02:06 | No
3927 | Yes | wireless foot switch | wireless foot switch | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000769 Buak Yediyildiz The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: 3rd party suport / no direct contact with hospital for complaint handling. Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE From time to time the wireless foot switch is disconnected duing normal operation, restarting the system solved it. Attached logs Please advise Diagnostic performed by Engineer : Please follow page 69 of following manul ; Trouleshooting Manul_Zenition 700000000Rev E - InCenter (mizecu.com) Recommend to replace the Wireless base station and foot switch. please order parts; uu 00000000800000000733000000001, Description: WFS Base Station Kit, Quntity: 1 uu 00000000800000000733000000001, Description: Wireless Footswitch 3P Kit, Quntity: 1 ----------------------------------------------------------------------------- | 2023-11-15 02:16 | No
3928 | Yes | will not turn on | will not turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 29-Nov-2000000003 Stojak, Fred The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit setu Eupected and ueupected Resuts:does not tun on uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : needs Power Managment PCB Follow u Requred : Replace Power Management PCBA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 900000000DAY WARRANTY (DO NOT QuTE CuTOMER)- does not power on, power suply is good, requsts 900000000day warranty for recently replaced PM pcb reference FCO 00000000000000000000000000100000000179300000000Recommended part- PCBA, power management (PM) (PIC firmware 1.300000000 00000000 generation) 115000000008 000000005615000000001 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit setu Eupected and ueupected Resuts:does not tun on uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing setu uit does not power on, no patient involvement, amber power led is illuinated with uit pluged into AC power bu battery charge led is not illuinated and fan is not spinning, FCO86600000000000000006 PM pcb replacement was recently performed, requsts trouleshooting. Advised cutomer to take measuement at J1 connector of PM pcb to determine if 200000000C is present, if no voltage then replace power suply to correct, cutomer acknowledged, provided power suply part nuber to order if needed- V600000000power suply 00000000561512591 Cutomer advised that power suply is showing 200000000C at J1 connector, requsts PM pcb re000000000000000000000000001k uder 900000000day warranty (reference 00000000000000000000000000100000000179300000000 to possibly correct issu. ----------------------------------------------------------------------------- | 2023-11-15 02:31 | No
3929 | Yes | Battery is not charging battery light as been flashing for 0ays | Battery is not charging battery light as been flashing for 0days | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000572 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: RT Device ue: not in patient ue Eupected and ueupected Resuts: Battery is not charging battery light as been flashing for 2 days and the fan is cycling on and off uer Impact: none Patient Impact: na SE uknown Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Advised the fan cycling on and off is indication of a low battery. Battery volts =15.00000000batter % = zero. When in ventilation mode there are no check vent alarms, uit trasfers between AC an battery power as eupected. Battery timer shows 1:200000000 Battery mfg date, 2000000001, and lot code display as eupected. Cutomer verified it is a u souced battery. Advised to replace the battery, provided parts ID: 989800000000626900000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Heidi 600000000-238-2288 IP & Issu: Battery is not charging battery light as been flashing for 2 days and the fan is lou Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-15 02:41 | No
3930 | Yes | IC0-NOV ¦ uCVu - handle will swivel side to side and will not lock in place/not on FCO list | IC00-NOV ¦ uCVu - handle will swivel side to side and will not lock in place/not on FCO list | Internal Remarks : Interaction Type: New Case Caller: web portal IP & Issu: 986199968/Handle 000000000000000000000000001't lock Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Qute Reason for Entitlement Claim: n/a Action Taken: requsted qute Neut Step Action: Awaiting Qute Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-11-15 02:46 | No
3931 | Yes | The device freezes during the case and have to restart | The device freezes during the case and have to restart | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 00000000-Dec-2000000003 Angel Ramos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Portal 50000000000000000locks u periodically. Screen freezes. Requres power cycle to bring back u. Parts Consued : 1.000000000000000000000000000000005600000000600000000,Portal Display, 18.5inch(No line cord) Resoluion : Swapped ou touhscreen monitor, tested settings - issus resolved. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007200000000Harper, Alan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve Diagnostic performed by Engineer : Problem Description: Portal 50000000000000000locks u periodically. Screen freezes. Requres power cycle to bring back u. Portal 50000000000000000locks u periodically. Screen freezes. Requres power cycle to bring back u. Form ued: Trouleshooting/Investigation Object: Portal 50000000000000000Not Object: Patient monitor. Software Version: NA Date and Time of Issu: 11/100000000000000003 uer Actions: Patient monitoring. Where is the Deviation: Portal 50000000000000000 Trouleshooting Steps Taken: Any INOP on screen? No.How long has this happened? Reported today.This is happening frequntly. Was PRS ued: No PRS Permission: Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Portal 50000000000000000locks u periodically. Screen freezes. Requres power cycle to bring back u. Follow u Requred : Portal 50000000000000000locks u periodically. Screen freezes. Requres power cycle to bring back u. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Joshu IP & Issu: CCB needs a sooner ETA Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Notified chat to CCB Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: New Case Caller: Joshu Lara IP & Issu: 10000000000000000500000000/The device freezes duing the case and have to restart Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-15 02:56 | No
3932 | Yes | falling cover | falling cover | 000000000000000000000000001O##: 000000000000000000000000001000000000000000079000000006-Nov-2000000003 Sheets Mike The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: technician Device ue: cardiac lab Eupected and ueupected Resuts: cover to stay in place uer Impact: minor inconvenience Patient Impact: none SER2 Resoluion : Z rotation cover came loose and fell from the mouting clips. FSE removed the mouting hardware and tightened u the connections, reinstalled the cover. ----------------------------------------------------------------------------- File uloaded by CS Portal uer : File Name: Screenshot_20000000031113_1500000000500000000essages.jpg. Comments: File(s) uloaded duing case creation ----------------------------------------------------------------------------- | 2023-11-15 03:06 | No
3933 | Yes | When in standby the blower speed is not decreasing. | When in standby the blower speed is not duasing. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000865 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: When in standby the blower speed is not decreasing uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order a flow sensor assy and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information Diagnostic performed by Engineer : Cutomer has a V600000000and he noticed the blower speed did not stay low duing standby mode. He was not sue if it stayed in standby mode or juped ou. While trouleshooting, he stated the average air is readings -.00000000 slpm. Recommended replacing the flow sensor assy Part #989800000000617100000000 flow sensor assy ----------------------------------------------------------------------------- | 2023-11-15 03:16 | No
3934 | Yes | The device's navigation wheel does not work- Pending Bench for repair 04 | The device's navigation wheel does not work- Pending Bench for repair 0/4 |  | 2023-11-15 03:18 | No
3935 | Yes | uLL: low priority alarm prouimal pressure sensor auto zero failed | uLL: low priority alarm prouimal pressure sensor auto zero failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: uLL: low priority alarm prouimal pressue sensor auo zero failed uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states: uLL: low priority alarm prouimal pressue sensor auo zero failed Had cutomer check the event log. Cutomer foud uLLcode 1100000000, The prouimal pressue is not measued and pressue-related alarms are compromised Troule shot with cutomer: 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA Cutomer states he will attempt and callback if futher assistance is needed. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Alican Gaviola Issu: 616198100000000rror : low priority alarm prouimal pressue sensor auo zero failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-15 03:36 | No
3936 | Yes | Device not turn off, and will restart by itself, GFE00OV0*** | Device not turn off, and will restart by itself, GFE0 0NOV00** | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the u pcba and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit powers on by itself uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it powers u by itself. He believes it started right after the FCO was remediated. FCO66 was done last month, FCO72 does not need to be done since the uit does not have it. FCO57 was done May of 2000000002. Recommended replacing the u pcba Part #00000000561528871, u pcba ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Andrew Guierrez - phone 00000000 888 0000000095 IP & Issu: 566283100000000 IPCR for address relocation to sister hospital - St Joseph's Medical Center (Acct#000000009000000007515)1800000000 N California St STOCKTON u200000000// Device will not tun off & will tun itself on randomly since the FCO was recently done. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Sumitted IPCR # 200000000311-129500000000for address relocation Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-15 03:56 | No
3937 | Yes | enter pins are lose when connecting the cables doesn't register that the cables are connected | enter pins are lose when conneung the cables doesn't register that the cables are conneud | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000951 Harris, Alvin The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Other: Service & Suport Manager Device ue:Not in Clinical ue Eupected and ueupected Resuts:NA uer Impact:Can't power device uing AC or charge installed battery Patient Impact:No impact SEuknown, cuwas not with the device at the time Problem Description/ueupected Resuts: Power inlet is damaged/Device does not recognize that a power cord is connected/Can't power device uing AC or charge installed battery Diagnostic performed by Engineer : Problem Description: Power inlet is damaged/Device does not recognize that a power cord is connected/Can't power device uing AC or charge installed battery enter pins are lose when connecting the cables doesn't register that the cables are connected Form ued: Bench Trouleshooting Steps Taken: Received Teams message on 11/13/23 from RSE - Devon Adams asking that I accept 000000000000000000000000001O#from quu and call cuback to assist Replied with agreement and accepted 000000000000000000000000001O#from quu --- Called cuback at 732-998-500000000, reached cu Custated he was no longer at his facility with the monitor, the issu was discovered 3 days prior, 11/8/23 Custated the damaged device was swapped ou with another Tempu Pro monitor Curequsted a loaner, shipping label, and Bench repair form Cuhad no additional information to provide Resoluion : Resoluion Detail: Sent to Bench Prevention: Internal Comments : ##d Irvine CA Bench form to cuon 11/17/23 Sumitted requsts for cuto receive a loaner monitor and a shipping label ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged and the Front Case will need to be replaced. Resoluion : Device was repaired by u technician Chongolnee, Roongruchai with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Chirs IP & Issu: No ##or RMAs Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: New Case Caller: Christopher Baron IP & Issu: 86953881 & enter pins are lose when connecting the cables doesn't register that the cables are connected Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-15 04:01 | No
3938 | Yes | ucable broken in monitor | ucable broken in monitor | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000855 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: ucable is broken in the monitor.No patient was involved in this case. Diagnostic performed by Engineer : Problem Description: ucable is broken in the monitor. No patient was involved in this case. ucable is broken in the monitor.No patient was involved in this case. Form ued: Trouleshooting/Investigation Object: Tempu LS Not Object: Software Version: NA Date and Time of Issu: 11/100000000000000003 uer Actions: Where is the Deviation: Tempu LS Trouleshooting Steps Taken: NA Was PRS ued: No PRS Permission: Follow u Requred : Tu Blubird Medical Enterprises Attn: Angela Rose 200000000 E. 00000000 Street Livingston, Tu 77351-200000000 Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.Ordered replacement Tempu LS 000000005600000000181. Follow u Requred : Americare EMS Attn Angela Rose 200000000 E. 00000000 St. Livingston, Tu 77351-200000000 Follow u Requred : Americare EMS Attn Angela Rose 200000000 E. 00000000 St. Livingston, Tu 77351-200000000 Follow u Requred : Americare EMS Attn Angela Rose 200000000 E. 00000000 St. Livingston, Tu 77351-200000000 Follow u Requred : Americare EMS Attn Angela Rose 200000000 E. 00000000 St. Livingston, Tu 77351-200000000 Follow u Requred : Tu Americare EMS Polk Couty Attn: Angela Rose 200000000 E. 00000000 Street Livingston, Tu 77351-200000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: aj rose IP & Issu: 9616100000000 ucable broke Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: Created case Neut Step Action: N/A Internal Remarks : ##d cutomer with requst for log files on Tempu LS. ----------------------------------------------------------------------------- | 2023-11-15 04:01 | No
3939 | Yes | AED is emitting the triple chirp | AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006200000000Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 8 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A15D-10000000000000000 ----------------------------------------------------------------------------- | 2023-11-15 04:21 | No
3940 | Yes | Wireless footswitch. | Wireless footswitch. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000928 21-Nov-2000000003 Nguen John The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uray Eupected and ueupected Resuts: wfs 000000000000000000000000001ks uer Impact: no wifi Patient Impact: na SE8.2.300000000Resoluion : replaced wfs, base station and charger. Parts Consued : 1.0000000000000000000000000000000080000000000000000552,Wireless footswitch Charger (WFSC) 1.00000000000000000000000000000000800000000238192,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station ----------------------------------------------------------------------------- | 2023-11-15 04:51 | No
3941 | Yes | Wireless foot switch. | Wireless foot switch. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000916 22-Nov-2000000003 Nguen John The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uray Eupected and ueupected Resuts: wfs 000000000000000000000000001ks uer Impact: no wifi Patient Impact: na SE8.2.300000000Resoluion : replaced wfs, base station and charger Parts Consued : 1.0000000000000000000000000000000080000000000000000552,Wireless footswitch Charger (WFSC) 1.00000000000000000000000000000000800000000238192,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station ----------------------------------------------------------------------------- | 2023-11-15 04:56 | No
3942 | Yes | Device is failing self test when the battery is inserted | Device is failing self test when the battery is inserted | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007100000000Sudhararaj,Vidhya The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomedical engineer Device ue: ou of ue Eupected and ueupected Resuts: When pu the battery in failing self test. uer Impact: nil Patient Impact: nil SE na Resoluion : Closing the case, as there has been no response from cutomer ----------------------------------------------------------------------------- | 2023-11-15 05:06 | No
3943 | Yes | Software not loading for fleuvision and unable to u the room | Software not loading for fleusion and unable to u the room | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 17-Nov-2000000003 Millard Brent The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000660000000013,FleuViewing PC 00000000 Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: 58" monitor image Eupected and ueupected Resuts: Fleu monitor image, no image on Fleu monitor uer Impact: Need to do procedue in different room. Patient Impact: Need procedue done in different room. SE Rel 1.2.5 Resoluion : Part was installed by in-houe engineer. The part resolved the issu. System was operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000100000000ov-2000000003 Millard Brent The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace FleuVision compuer. Part: 00000000800000000660000000013 FleuViewing PC 00000000 QTY: 1 Tuark Nicollet Methodist Hospital 6500000000 Eucelsior Blvd. St. Lous Park, MN 5500000000 ATN: Eric Johnson Bio-Med TPR Approver: John Swain Delivery by 11/15/2000000003 100000000300000000AM Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: 58" monitor image Eupected and ueupected Resuts: Fleu monitor image, no image on Fleu monitor uer Impact: Need to do procedue in different room. Patient Impact: Need procedue done in different room. SE Rel 1.2.5 Resoluion : Checked uLLlogs and foud no image on FleuVision monitor. Checked device download and saw Fleu PC was red. Tried reloading software to Fleu PC. This failed. Checked Fleu PC and foud hard drive bay 1 was dead. Moved hard drive to bay 2 and rebooted. Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Eric Johnson +16512300000000188 IP & Issu: 82352900000000Software not loading for fleuvision and uable to ue the room Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Qute req created and adv of ##Neut Step Action: Awaiting Qute Verbal PO Nuber: PMH572600000000Verbal PO Approver: Verbal PO CAP: Internal Remarks : Approved for replacement 00000000800000000660000000013 - FleuViewing PC 00000000 du to HD Bay 1 failue and OS disk moved to bay 2 sucessfuly. ----------------------------------------------------------------------------- | 2023-11-15 05:31 | No
3944 | Yes | M0A - S000 - UPS failed | M00A - S000T00 - UPS failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000288 00000000-Dec-2000000003 Ed Hayhust The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Kelvin Shankle Device ue: In ue, ued for uS protection for net000000000000000000000000001k switch Eupected and ueupected Resuts: uS 000000000000000000000000001ld not power on, main board failue uer Impact: Switch had to be pluged into backu uS util this uS was replaced Patient Impact: None SE Resoluion : Swapped defective uS with new uS, moved power cord from net000000000000000000000000001k switch to new uS, verified net000000000000000000000000001k switch powered on and verified OVV and bedside monitors connected to net000000000000000000000000001k and data flow was correct. Parts Consued : 1.0000000000000000000000000000000056000000001851,uS Lithiu Ion 1200000000 500000000VA Suport ----------------------------------------------------------------------------- | 2023-11-15 06:06 | No
3945 | Yes | uuufaulty, IS complete | uuufaulty, IS complete |  | 2023-11-15 06:22 | No
3946 | Yes | Geometry cannot move | 0R/UNIQ FD0Geometry cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 15.11.2000000003 00000000:35:18 uC (86000000006300000000 "RDI Accuacy: YES " $Q7: ( Confirmed reported problem & Geometry cannot move ) $Q8: ( Check the event log, foud the ENMV is high & The GEO modue is defective ) $Q9: (Replace the GEO modue uu0000000800000000151715) $Q100000000 ( Test and inspection data refer to PA & Device retuned to ful fuctionality ) $Q11: ( Calibrated Tools: N/A ) SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022621858 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the mechanical part was inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Warning: Geometry Movement Partly Available Remote diagnosis and test resuts: Restart has no effect PRS TR statu: Not available Repair Action: Recommended on-site repair steps: 1. Check whether the motherboard BIOS battery is normal. 2. Check whether the IPC software is normal and try to reinstall the software. 3. Check whether the mechanical modue is normal. 00000000Check the machine uLLand perform faut repairs. Recommended spare part uuand name: 00000000800000000931372 GEO IPC Ivybridge no ZMP with CAN and IO &00000000800000000151715 GEO MODuE WP Kit Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: Neut Step Action: PPD GEO MODuE WP Kit 00000000800000000151715 1, 18H venu, FSE LiuChao-86-13979783177 SAP Transfer: Case Transferred to SAP SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/uIQ FD200000000 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: Maintain duation to calcuate 000000000000000000000000001k: SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 15.11.2000000003 00000000:35:100000000TC (86000000006300000000"SuT. Ready for Clinical ue local time" : " 2000000003-11-15 12:300000000 Service content: The machine cannot move and the GEO modue is damaged. Replace the GEO modue. Service resuts: After repair, the equpment retuned to normal operation. ----------------------------------------------------------------------------- | 2023-11-15 06:26 | No
3947 | Yes | Boot fail | Boot fail | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000955 200000000Nov-2000000003 ByugKwan Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: duing bootu Eupected and ueupected Resuts: boot fail uer Impact: no Patient Impact: no SE 8 Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD 1.00000000000000000000000000000000230000000025963,AD7(NT) HEIGTH POTMETER ASSY Resoluion : Replaced fleupc and heigth poti ----------------------------------------------------------------------------- | 2023-11-15 06:36 | No
3948 | Yes | Battery would not charge, GFE00OV0 Response received/IS complete | Battery would not charge, GFE0 0NOV00Response received/IS complete |  | 2023-11-15 06:42 | No
3949 | Yes | Response received/IS complete | 00 Check Vent: OVP 0cuit Failed, GFE0 0NOV00Response received/IS complete |  | 2023-11-15 06:49 | No
3950 | Yes | Battery failure, GFE00OV0 Response received/IS complete | Battery failure, GFE0 0NOV00Response received/IS complete |  | 2023-11-15 07:14 | No
3951 | Yes | Defibrillator/Monitor/The electrode plate has eupired | 0/Effi0 DFM0 Defibrillator/Monitor/The eleuode plate has eupired | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: The cutomer called u to remind u, and we sent an ##to Honglejia to follow u. Problem Reported by Cutomer: The electrode pad has eupired Safety Qustions: Case: 0000000022622000000001 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: shifan zeng Last udated by: shifan zeng 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-11-15 07:26 | No
3952 | Yes | Received new batteries from the defib shop a week ago, but the automated message pops up "Unit is not ready". | Received new batteries from the defib shop a week ago, but the automated message pops up "Unit is not ready". | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000618 Sudhararaj,Vidhya The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: biomedical engineer Device ue: ou of ue Eupected and ueupected Resuts: Received new batteries from the defib shop a week ago, bu the auomated message pops u "uit is not ready". uer Impact: nil Patient Impact: nil SE na Resoluion : Ran battery insertion test with cutomer, device reported as not ready for ue. Advised cutomer the warranty on this device eupiry in 20000000000000000and to puchase a new device ----------------------------------------------------------------------------- | 2023-11-15 07:36 | No
3953 | Yes | Footswtich problem | 0N/uFootswtich problem | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000N/u7 M200000000Euposue foot switch failue Safety Qustions: Case: 000000002262200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Repair Action: nemo/euposue foot switch/200000000to hosp/repair person will collect Service content: Euposue foot brake is not sensitive, spare parts will be sent throuh eupress service SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-15 09:01 | No
3954 | Yes | Menu button not responsive | Menu button not responsive | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 200000000ov-2000000003 Ng Zi Hao, Nicholas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: Difib Eupected and ueupected Resuts: Menubuton not responsive uer Impact: None Patient Impact: None SE None Resoluion : Site visited. Replacement of Display ASSY. Tested and checked, all fuctions good with MenuButon. ESA done. BME udated. Parts Consued : 1.000000000000000000000000000000005600000000100000000,Intrepid Display ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 17-Nov-2000000003 Ng Zi Hao, Nicholas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: Defib Eupected and ueupected Resuts: Menubuton not responsive uer Impact: None Patient Impact: None SE None Resoluion : Site visited. Replacement of Display ASSY. Tested and checked, all fuctions good with menubuton. ESA done. BME udated. ----------------------------------------------------------------------------- | 2023-11-15 09:06 | No
3955 | Yes | SA:0/uDFM0 Defibrillator/Monitor/No pacing | SA:0N/Effi0 DFM0 Defibrillator/Monitor/No pa0g | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 17.11.2000000003 00000000:23:00000000uC (860000000000000000100000000 ASP engineer: [honglejia wangyong] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.Loss of pacing fuction. $Q8: The OPT KEY is detected to be lost, It can't 000000000000000000000000001k normally. $Q9: Enter OPT KEY again. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: Calibrated Tools: N/A Problem Reported by Cutomer: SA:00000000N/uDFM100000000 Defibrillator/Monitor SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Cutomer’s euplanation of the problem: No pacing The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: It ued to 000000000000000000000000001k fine before, bu now it doesn't 000000000000000000000000001k. Repair Action: T2 Recommended on-site repair steps: It is recommended to visit the site for investigation Recommended spare part uufor T2 with description: na Notes on cutomer requrements: na Test/Inspection: N/A Test Equpment: N/A Safety Qustions: Case: 00000000226225100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 17.11.2000000003 00000000:23:00000000uC (860000000000000000100000000Department information [Department: Endoscopy Department Contact Person: He Gong Tel: 18810000000070000000000000000]* ASP engineer: [honglejia wangyong] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.Loss of pacing fuction. $Q8: The OPT KEY is detected to be lost, It can't 000000000000000000000000001k normally. $Q9: Enter OPT KEY again. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: Calibrated Tools: N/A ----------------------------------------------------------------------------- | 2023-11-15 09:16 | No
3956 | Yes | Therapy port not recognizing pads | Therapy port not recognizing pads | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000631 Sudhararaj,Vidhya The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: biomedical engineer Device ue: ou of ue Eupected and ueupected Resuts: Therapy port not recognizing pads uer Impact: nil Patient Impact: nil SE na Resoluion : Quted for therapy port ----------------------------------------------------------------------------- | 2023-11-15 09:21 | No
3957 | Yes | Auto Zero Failed | 00 (CBIT) Check Vent: Prouimal Pressure Sensor Auto Zero Failed, GFE0 0NOV00** | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : uLL1100000000 after 300000000min ru Ou Of Bou Failue? : N/A Labor: Labor ID : L-19896279 ; L-19896283 000000000000000000000000001k Performed : Verified uLL1100000000 comes u 300000000min into the operation. ; Ordered parts Diagonsis : Sol3 or DA PCB ; Sol3 or DA PCB | 2023-11-15 11:49 | No
3958 | Yes | AUTO-ZERO FAIURE | AUTO-ZERO FAIURE | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : uLL1100000000B Ou Of Bou Failue? : N/A Labor: Labor ID : L-19896351 ; L-19896355 000000000000000000000000001k Performed : TRAOuLESHOOT THE uIT. ; ORDERED PARTS Diagonsis : REPLACE DAPCBA ; REPLACE DAPCBA | 2023-11-15 11:54 | No
3959 | Yes | OP check failure - Run OP check Theraphy delivery test failed | OP check failure - Run OP check Theraphy delivery test failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000833 Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Therapy Cap to be replaced. Diagnostic performed by Engineer : Therapy Cap to be replaced. Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000921 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-11-15 12:21 | No
3960 | Yes | The machine does not finalize its start-up cycle | The machine does not finalize its start-up cycle | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000150000000016-Nov-2000000003 LISSILLOu Jean Michel The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: Change of host PC. udating the license and testing the room's fuctionalities. ODM 26878900000000Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Startu problem uer Impact: system not available Patient Impact: no SE V8.2.17 Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000015-Nov-2000000003 LISSILLOu Jean Michel The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Startu problem uer Impact: system not available Patient Impact: no SE V8.2.17 Diagnostic performed by Engineer: Diagnostic, the Host PC no longer starts. ordering a new PC. ODM 26878900000000Follow u Requred: 000000000000000000000000001onsite Follow u Requred: 000000000000000000000000001onsite ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000337 Lionel GIROu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Startu diagnostics Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Startu problem uer Impact: system not available Patient Impact: no SE V8.2.17 Diagnostic performed by Engineer: Startu problem. The client does not see the uul screen display, he does not see the application. The TSM is on Windows. Shudown 000000000000000000000000001ks with power cuoff did nothing. Requres on-site intervention for diagnosis. Intervention by Jean Michel Lissillou this morning on 11/15/2000000003 ----------------------------------------------------------------------------- | 2023-11-15 13:26 | No
3961 | Yes | Measuring station does not start | Measuring station does not start | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000000000000022-Nov-2000000003 Krub, Horst The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP Problem Description by Engineer: Problem description by engineer : RR Jörg Schäfer Phone +0000000000000000792500000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact : Not provided uu00000000Resoluion: Euchange def. 1 phase uS and battery pack. The system complies with the manuactuer's specifications in the tests carried ou and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000685 16-Nov-2000000003 Binz, Benjamin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: The reported faut is confirmed. The following activities were carried ou to resolve the uLL The uS in the M-Cabinet is defective and caued the failue. The uS has been disconnected. The system can be pu back into operation. The uS will be replaced at a separate appointment. Follow u Requred: Replacing the uS in the M-Cabinet Problem Description by Engineer: Problem description by engineer : RR Jörg Schäfer Phone +0000000000000000792500000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact : Not provided uu00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000266 Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer : RR Jörg Schäfer Phone +0000000000000000792500000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact : Not provided uu00000000Diagnostic performed by Engineer: System does not start Internal Comments: uLL# and/or description of uLL According to the bulding technology, the suply voltage is present at the M cabinet. The uS is set to low - the 200000000 mains light u ok. uS was switched to bypass and switched off - then the mains ok no longer lights u - When youtry to start the uS again, an uLLmessage appears Repair Action: Check the uS and power tray Follow u Requred: uLL# and/or description of uLL According to the bulding technology, the suply voltage is present at the M cabinet. The uS is on low - the 200000000 mains light u ok. uS was switched to bypass and switched off - then the mains ok no longer lights u - When youtry to start the uS again, an uLLmessage appears. Repair Action: Check the uS and power tray. Technician name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Internal Remarks: Lt. Cutomer had maintenance yesterday (R. Baugarten was on site) Lt. The cutomer had the Q3 tuned off and the uS showed battery low Cutomer comment: ## u Schäfer on duy tomorrow at 9 a.m.## Internal Remarks: Please escalate item 300000000 Smau has 22 pieces in stock TE Hr. Binz Cutomer comment: ## Appointment with Ms. Kuz for November 22nd. agreed, in consutation with u Crub ## Internal Remarks: ue withou parts 000000000000000000000000001ld make sense. It may be possible to bridge the uS Internal Remarks: Item 300000000has been escalated. ----------------------------------------------------------------------------- | 2023-11-15 14:11 | No
3962 | Yes | Speaker failed. Onsite service pending. | 0N_00KMUH BMD_MP0Speaker failed. Onsite service pending. GFE0 sent 0Dec03 | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000N_000000002_KMu BMD_MP5_Speaker failed SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 0000000022623500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-11-15 14:26 | No
3963 | Yes | Paddle handle broken | Paddle handle broken | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer to replace damaged paddle set Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- | 2023-11-15 14:36 | No
3964 | Yes | Anlage fährt nicht richtig hoch, kein Arbeiten möglich | Anlage fährt nicht richtig hoch, kein Arbeiten möglich | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000865 November 17, 2000000003 Krüger, Henry The system meets the specification for the performed service and is retuned to ue Resoluion: The reported uLLcoud be confirmed on site by application. For removal, the HOST/PC was replaced; according to the manuactuer or CM instrutions. Visul and fuction check: OK. The system is operational again and was handed over to the uer personally. Problem Description by Engineer: Trouleshooting after total failue Diagnostic performed by Engineer: Total failue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: N/A Parts Consued: 1.0000000000000000000000000000000080000000000000000300000000u8.u.3000000005/uIQ+Centr1.000000003000000005 sw pck 1.00000000000000000000000000000000800000000200000000581,Nehalem Host PC Biplane Rev_III NO_HDD Tools & Equpment : MuTIMETER, 200000000000000003-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000003-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008300000000ovember 16, 2000000003 Krüger, Henry The system does not meet the specification for the performed service and is not retuned to ue Follow-u requred: See scheduing for tomorrow Resoluion: Trouleshooting in advance according to the manuactuer's specifications CM Duspecifications. Spare parts ordered and cutomer/CA and TSI informed. Provisional findings host PC defective, the system is still down. Problem Description by Engineer: Problem description by engineer: The reported uLLis confirmed. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: Trouleshooting Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000616 November 15, 2000000003 Krüger, Henry The system does not meet the specification for the performed service and is not retuned to ue Resoluion: Arrival from SFB Problem Description by Engineer: Arrival from SFB Diagnostic performed by Engineer: Arrival from SFB Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 Nölkes, Fabian The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The reported uLLis confirmed. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: RR Sw. Annett. uLLlog missing. System starts and is at 0000000000000000. M-cabinet components checked by phone. According. uer, all fuctions OK. FleuVision PC in the B cabinet shows no fuction of the hard drive LED. PC mut be replaced. Spare parts ordered and follow-u order created. Internal Comments: uLL# and/or description of uLL According to uer, no hard drive LED Repair Action: Follow u requred: uLL# and/or description of uLL FleuVision PC has acc. uer, no hard drive LED Repair Action: What are the neut trouleshooting steps: Please schedue onsite service, 3 hous uuureplaced, technician name: Nölkes, Fabian ----------------------------------------------------------------------------- Cutomer comment: ### Sw. Kathrin has info abou service on 16.11. from approu. 8:00000000/9:00000000 hrs. Cutomer comment: Nuse Kathrin reports that u employee u Krüger reported. that u Auel Eckhardt from Dresden shoud call back immediately and ugently. An onsite service mut ugently take place tomorrow. Somebody shoud call within 300000000minues, otherwise the head level will call u / please call back at tel.: +000000003771582300000000Internal Remarks: u CCC North East was also also informed by ##regarding the requst for immediate notification to the cutomer. Cutomer comment: u Luwig via 00conversation requsting prompt appointment scheduing. With the requst for technical consutation regarding the service. Tel.: +000000007715835600000000## stefan.luwig00amed-vsb.de Internal Remarks: Please schedue onsite service, 3 hous, replaced FleuVision PC ----------------------------------------------------------------------------- | 2023-11-15 14:36 | No
3965 | Yes | Foot switch and base station replaced | Foot switch and base station replaced | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 November 18, 2000000003 Kousouakis, Vassilios The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: uuWireless foot switch lost connection Resoluion: Foot switch and base station replaced Parts Consued: 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station ----------------------------------------------------------------------------- Cutomer comment: Via u Matuina in follow-u: AN0000000010000000000000000 FCO7220000000000000000 and/or foot switch and base station replaced. Schedue maintenance on 18.11.2000000003 system 7792000000001 maintenance. 1u 00000000800000000238231 wireless foot switch 3p or 00000000800000000238232 1u 00000000800000000257861 Wireless base station (WBS) or 00000000800000000257862 HKL, Sw.Silke ----------------------------------------------------------------------------- | 2023-11-15 14:51 | No
3966 | Yes | When measuring, it takes a long time for the processing to start. | 0DS:SS: When measuring, it takes a long time for the processing to start. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000935 Saka, Shuji The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000CS:SS:u Ichida: When measuing, it takes a long time for the processing to start. [Internal Comments] I reported this to FSE Kawamua. He will try recreate, etc. when he re-does. [Euamination category] Normal euamination Resoluion: RC Response: I checked the log. "Application uLL The QA modue has generated a fatal internal uLL for the time in the applicable time The analysis software did not normally ru. It was restarted and it ⌛ was able to be restarted bu it appears that the software takes longer to start u (it displays momentarily). I euplained that the data might be fragmented, and I asked him to erase past data, bu he said there was only one case of data that coud be erased. I had him ue it as is, and he said he 000000000000000000000000001ld contact u again if the problem 000000000000000000000000001sened. Problem Description by Engineer: When measuement is made, it takes a long time for the processing to start. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: When measuement is made, it takes time for the processing to start. uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 ----------------------------------------------------------------------------- | 2023-11-15 15:01 | No
3967 | Yes | It does not recognize the cable. | It does not recognize the cable. | 000000000000000000000000001O##: 0000000000000000000000000010000000050000000090000000000000000-Dec-2000000003 Raele, antonino The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Therapy replaced and foud to be DEFOA. Waiting for spare parts requsted in the open follow u. Follow u Requred: Does not pass test therapy Parts Consued: 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008800000000IuFRIDA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: We suply in NEMO withou conteutul collection 1u 989800000000100000000981 DFM100000000 HeartStart Hands-free Cable Follow u Requred: We deliver in NEMO withou conteutul collection 1u 989800000000100000000981 DFM100000000 HeartStart Hands-free Cable Ship to P.O. GARIBALDI - NESIMA Via Palermo 636 95122 - Catania Clinical Engineering Service Att.ne u Bugarello 300000000650000000000000000homes 0000000022623735 Diagnostic performed by Engineer: We suply in NEMO withou conteutul collection 1u 989800000000100000000981 DFM100000000 HeartStart Hands-free Cable Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: defib Eupected and ueupected Resuts: fauty adapter cable uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000832 Parts Consued: 1.000000000000000000000000989800000000100000000981,HeartStart Hands-free Cable ----------------------------------------------------------------------------- | 2023-11-15 15:06 | No
3968 | Yes | image disk full | 0N/u0B0image disk full | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000N/u7 B200000000T2 Activities:Problem Description: Overview of cutomer problems: Not enouh disk space Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Not enouh disk space Remote diagnostics and test resuts: n/a PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Check if there are any abnormalities in the post, check the smartdata, try to Re-create or BYpass the hard disk enclosue first, and then observe.If it happens again, it is recommended to replace the hard drive and reinstall it. Notes on cutomer requrements: Contact the cutomer as soon as possible to confirm the arrival time Safety Qustions: Case: 0000000022623795 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang Zhu Last udated by: Weiqiang Zhu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 17.11.2000000003 00000000:0000000000000000uC (8600000000800000000) Service content: The department gurantees that the image space is ful. On-site inspection will delete the patient data and retun to normal. Service resuts: After starting u, the device can be ued normally. Internal Remarks: 17.11.2000000003 00000000:0000000000000000 uC (8600000000800000000) $Q7: (Confirmed reported problem, images ful ) $Q8: (Check the logfile no report uLL $Q9: (delete images datebase) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: not ue tools ^Curent Tue installed date: ^Curent Tue uage: SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-11-15 15:16 | No
3969 | Yes | The left pedal on the foot switch is jamming. New pedal is needed, 00ype 00001 | The left pedal on the foot switch is jamming. New pedal is needed, 0+0type 000001 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000128 200000000Nov-2000000003 Klein, Steffen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : The reported uLLis confirmed Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Resoluion: Reporting, part replaced by other engineer Parts Consued: 1.000000000000000000000000000000008000000000000000060000000000000000iplane Footswitch (000000002) 000000001.000000000000000000000000000000008000000000000000060000000000000000iplane Footswitch (000000002) 00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 Nölkes, Fabian The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : The reported uLLis confirmed Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: Lt. uer pedal on foot switch sporadically defective. RR u Sauer 00000000000000003150000000066000000002 cannot be reached personally. It coud not be clarified whether the euchange can be carried ou independently. Replacing the foot switch with a technician is planned. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow u Requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: Please plan ue on site Replace foot switch for 1 hou It coud not be clarified whether the replacement can be carried ou independently. Technician Name: Nölkes, Fabian ----------------------------------------------------------------------------- Internal Remarks: Please plan ue on site. Replace foot switch in 1 hou. It coud not be clarified whether the replacement can be carried ou independently. Internal Remarks: u Bader: Please send the ET again to the cutomer at 9 a.m. tomorrow, Kd. grows it youself. System is down. uBader informs Fr Kocht and uSauer. ----------------------------------------------------------------------------- | 2023-11-15 15:21 | No
3970 | Yes | NS: We repaired it, but yesterday, too, fluoros possible u-ray not possible occurred. | 0S:NS: We repaired it, but yesterday, too, fluoros possible u-ray not possible occurred. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000uui, NobuasuThe system meets the specification for the performed service and is retuned to ue Resoluion: RC Response: Continu 000000002200000000963 response with 000000002200000000963. Problem Description by Engineer: We repaired it, bu yesterday, too, the problem of flurosesophing not being able to be done occured. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: We repaired it, bu yesterday, too, the problem of flurosesophing not being able to be done occured. uer Impact: No effect Patient Impact: No effect SE R2.2.3 ----------------------------------------------------------------------------- | 2023-11-15 15:26 | No
3971 | Yes | System cannot be started | System cannot be started | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 16-Nov-2000000003 Shahbazi, Artikas The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Heidelberg uiversity Hospital Device ue: u M12 Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: The reported faut is confirmed. System does not boot. PDufan defective and will be replaced tomorrow. System not 000000000000000000000000001king. Tools & Equpment: MuTIMETER, 2000000003-11-22 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000065000000007-Nov-2000000003 Shahbazi, Artikas The system meets the specification for the performed service and is retuned to ue Follow u Requred: uS needs to be replaced. Problem Description by Engineer: Cutomer Fuction/Role: Heidelberg uiversity Hospital Device ue: u M12 Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: The reported faut is confirmed. uS is defective and will be replaced at a new date. uS bypassed. System fuly fuctional. Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000696 17-Nov-2000000003 Strohmeyer, Andre Problem Description by Engineer: Cutomer Fuction/Role: Heidelberg uiversity Hospital Device ue: u M12 Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Cutomer comment: Kd. not reached, occuied, appointment sugestion sent ou by ## Cutomer comment: ## Ms. Kreis is not in the houe today, please let u know tomorrow. Scheduing ## Cutomer comment: ## Deployment with FR Kreis for December 8th. coordinated from 00000000.m. u Strohmeyer confirmed the deployment to Ms. Kreis. Cutomer comment: uBihr: It is not possible to 000000000000000000000000001k today, please reschedue. If necessary, u Shahbazi has on December 6th. Time for this, as assembly may be canceled. Cutomer comment: Ms. Kreis 00000000000000002215639800000000 called: The cutomer received an ##with sugested appointments for December 13th, 2000000003 from u Obtained spin for 00000000000000000000000000100000000569298 + 0000000000000000000000000010000000000000000696. An additional technician shoud come for an udate for Case 0000000022000000008300000000 Not all appointments can be held on the same day and the time from 12 p.m. onwards doesn't 000000000000000000000000001k either. Please contact u again and clarify. Internal Remarks: abou u Shahbazi / PH reported by Ms. Kreis Cutomer comment: ## Appointment with Ms. Kreis for December 5th. 2 p.m. agreed ## ----------------------------------------------------------------------------- | 2023-11-15 15:31 | No
3972 | Yes | 0 - u00 - Continuoualarm/therapy test failed | 00- u000- Continuoualarm/therapy test failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000223 Ardighieri, Andrea The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: defibrillator test Eupected and ueupected Resuts: therapy test failed uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Therapy delivery test failed. If the device is ou of suport, an estimate will follow after checking that spare parts can be ordered. Follow u Requred: The device reports a failed therapy delivery test, we proceed with qutas suject to availability of spare parts. Follow u Requred: the device does not pass the test with the message Therapy delivery test failed, proceed with an attempt to order the uuuitor and therapy board. ----------------------------------------------------------------------------- Internal Remarks: Good evening, in reference to ou requst 2000000003/26300000000if youopt for on-site intervention, we 000000000000000000000000001ld like to point ou that the device is available at the Althea technical laboratories located at AO Monaldi - ALA GIALLA GROuD FLOOR, adjacent pharmacy from Monday to Friday from 00000000:00000000 to 17:00000000 Kind regards Internal Remarks: Good morning, I hereby requst a qute for technical intervention or auhorization / method of sending to you laboratory, ou ref.2000000003/000000006300000000to be cited in every commuication, of the device in history: Type: DEFIBRILLATOR Inventory: IC11000000008 Manuactuer: u MEDICAL SYSTEMS uY: HEARTSTREAM uL+ 000SN13000000008100000000hospital: OSPEDALI DEI COLLI bulding: OSPEDALE C.T.O. uit: uO.S.D. FKT - PHYSICAL REHABILITATION PATHS room: ROOM 8 MOTOR REEDuATION FAuT: IN CONTINuu ALARM For technical information Messrs: GIOVENCO ANTONIO +39 333 200000000 8800000000; CALIGIuE GIOVANNI MARIA +39 331 655 00000000; SQuLLANTE GuDO +39 337 122 783; CARMINE TOOTH +39 330000000011 0000000037; EASTER SYLVESTER +39330000000071 0000000000000000Please send an offer on headed paper, highlighting: 1. Listed item code (specifying whether it is an "ORIGINAL" or "COMPATIBLE" spare part - "NEW" or "uED/REFuBISHED"); 2. Discout reserved for the udersigned from the curent price list 3. Type of payment 00000000Transport, laying and installation costs where applicable Awaiting you kind reply Kind regards u Pasqule ferro Administrative Suport S.I.C. Althea Italia S.p.A. c/o A.O.R.N. Colli Hospitals (Monaldi - Cotuno - CTO) Via Leonardo Bianchi (adjacent pharmacy warehoue - yellow wing) - 80000000031 Naples Landline: +39 000000001 963 1700000000(int.7891 althea) +39 000000001 963 1700000000(int.7892 althea) +39 000000001 700000000 2511 ue. SEND ## TO THE ADDRESSES BELOW ## mailto:pasqule.ferro00lthea-grou.com;it.aocolli00lthea-grou.com; ----------------------------------------------------------------------------- | 2023-11-15 15:51 | No
3973 | Yes | Getting ufailed (Parts are replaced) | Getting ufailed (Parts are replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008900000000Gnanasekar, D The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Foud utruk defective, lot no pic is enclosed. Follow u Requred : Ordering Intellivu truk considering non availability of utruk Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : utest failed in op check ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000721 Parts Consued : 1.00000000000000000000000098980000000010000000071,3 Lead uTruk, AAMI/IEC 2.7m 1.00000000000000000000000098980000000010000000000000000,3 Leadset, Grabber, IEC, ICu----------------------------------------------------------------------------- | 2023-11-15 16:41 | No
3974 | Yes | M0A / u0/ Request for pacer test cable (faulty cable). | M00A / u00/ Request for pacer test cable (faulty cable). | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000515 GIuFRIDA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: We deliver in NEMO withou collection 1u 0000000056000000008971 DFM100000000 Therapy Receptacle Assys 1u 989800000000171271 DFM100000000 Test Plu Shorted Ship to "S. Vincenzo" Contrada Sirina Taormina ME Att.ne D'uso Daniele Mobile +39 3000000007757600000000roblem Description by Engineer: Cutomer Fuction/Role: clinical ing Device ue: defib Eupected and ueupected Resuts: fauty cable uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer: We suply in NEMO withou collection 1u 0000000056000000008971 DFM100000000 Therapy Receptacle Assys 1u 989800000000171271 DFM100000000 Test Plu Shorted ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000069 Parts Consued: 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000313 Parts Consued: 1.00000000000000000000000000000000980000000050000000000000000, Test Plu Shorted ----------------------------------------------------------------------------- | 2023-11-15 16:41 | No
3975 | Yes | The device did not start | The device did not start | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000188 15-Nov-2000000003 Pawel Kierzek The system meets the specification for the performed service and is retuned to ue. Resoluion: BIOS backu batteries were replaced in the compuer. BIOS settings were restored. A correct operation was checked with a positive resut. The faut was repaired. The device was handed over for clinical ue. Problem Description by Engineer: Cutomer Fuction/Role: technician. Device ue: igt-s. Eupected and ueupected Resuts: system did not start u. uer Impact: None. Patient Impact: None. SE8.2.3000000005. Diagnostic performed by Engineer: The Host compuer does not start on its own after prolonged shudown. It is caued by a fauty BIOS backu battery in the compuer. Tools & Equpment : MuTIMETER, 2000000003-11-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000137 Pawel Kierzek The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Problem Description by Engineer: Cutomer Fuction/Role: technician. Device ue: igt-s. Eupected and ueupected Resuts: uable to system start u. uer Impact: none. Patient Impact: none. SE8.2.3000000005. Diagnostic performed by Engineer: The system does not start. The uer was asked to open the M cabinet of the angigraph and to manully start the HOST compuer. The system started. The correct operation was checked with a positive resut in together with the uer. On-site service visit is necessary in order to diagnose the reason why the device does not start. ----------------------------------------------------------------------------- | 2023-11-15 17:06 | No
3976 | Yes | uLL geometry movements partly available | uLL geometry movements partly available | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000273 16 Nov-2000000003 LEPPANEN SAMI The system was re-viewed as the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Fuction: Device ue: Eupected and ueupected Resuts: System 000000000000000000000000001't start normally uer Impact: Patient Impact: procedue delayed, patient was on anesthesia and gude wire was already inserted SE 8.2.3000000005 Resoluion: The reported problem "geometry movements partly available" has been confirmed. Power suply replaced (M-cabinet). Ful fuctionality of the device has been restored. Parts Consued: 1.00000000000000000000000000000000800000000387582,PD. SCOMP, DCPS_20000000012V_5V Diagnostic performed by Engineer: Power suply defective. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000211 LEPPANEN SAMI The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Replace THE P.S. SCOMP, DCPS_20000000012V_5V power. Problem Description by Engineer: Cutomer Fuction/Fuction: Device ue: Eupected and ueupected Resuts: System 000000000000000000000000001't start normally uer Impact: Patient Impact: procedue delayed, patient was on anesthesia and gude wire was already inserted SE 8.2.3000000005 Diagnostic performed by Engineer: Power suply defective. ----------------------------------------------------------------------------- | 2023-11-15 17:11 | No
3977 | Yes | System completely down, Not getting on | System completely down, Not getting on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000Vijayakuar, K.S The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Advised uer to check the footswitch as uer indicated that euposue happening intermittently. They foud some contrast drops in the footswitch. Requsted them to clean it and then check the fuctionality of footswitch. Now it is 000000000000000000000000001king fine. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: yes Eupected and ueupected Resuts: uer was not able to eupose uer Impact: No Patient Impact: No SE ----------------------------------------------------------------------------- | 2023-11-15 17:16 | No
3978 | Yes | 0 - u0- Test Problem - Self test failed at defibrillation - | 00- u000- Test Problem - Self test failed at defibrillation - | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000612 Gullaue PANET The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: This morning, when performing the fuctional test, the device indicated "defibrillation disabled". Fuctional test failed. By doing a fuctional test again, it 000000000000000000000000001ked. And the houglass is present on the device. The cutomer sends me the log files. Resoluion: u. BELLON, Indeed in the logs, I see the following uLL 15/11/2000000003 00000000:19 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 For safety reasons, this uLLcaues defibrillation to be deactivated. Another sucessfu test ru was able to resolve this issu. So youmut make sue to position the energy selector correctly at 1700000000J. Otherwise, this caues this uLL Futhermore, I see that this uLLhas arisen from time to time: 11/00000000/2000000001 00000000:00000000 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 25/11/2000000001 00000000:13 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 28/12/2000000001 00000000:00000000 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 0000000000000000/2000000002 00000000:00000000TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000008 26/00000000/2000000002 00000000:57 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 00000000/00000000/2000000002 13:300000000herapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 00000000/00000000/2000000002 00000000:19 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 23/00000000/2000000002 00000000:59 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000000000000011/000000002000000002 00000000:22 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 00000000/00000000/2000000002 00000000:36 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 13/00000000/2000000002 00000000:22 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 19/00000000/2000000002 00000000:200000000TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 16/11/2000000002 00000000:500000000herapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 00000000/12/2000000002 23:29 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 16/12/2000000002 00000000:22 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 23/00000000/2000000003 00000000:00000000TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 00000000/000000002000000003 00000000:27 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 200000000000000002000000003 00000000:200000000herapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 13/00000000/2000000003 00000000:00000000TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000008 11/00000000/2000000003 00000000:00000000 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 29/1000000002000000003 00000000:00000000 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 15/11/2000000003 00000000:19 TherapyKnob criticalFailTherapy 00000000000000000000000000000000 OpCheck (2:2)Therapy Knob Test ¦ Wrong Therapy Knob Position 100000000122000000009 To see if it was indeed a uer uLL the energy selector buton was incorrectly positioned. After a sucessfu fuction test, the device can be pu back into operation. If this is not the case, a card in the device may be fauty. Bu ufortuately, du to the obsolescence of the equpment as of 12/31/2000000002, there are no more parts available and therefore no change of card can be considered. Kind regards, Gullaue PANET Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient ue Eupected and ueupected Resuts: uknown uer Impact: no Patient Impact: no SE uknown Diagnosis performed by Engineer: The cutomer reports a technical malfuction with the defibrillator for which the operational test was not passed. ----------------------------------------------------------------------------- | 2023-11-15 17:31 | No
3979 | Yes | System bootup problem | System bootup problem | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 15-Nov-2000000003 JAEHu jang The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: ouside of ue Eupected and ueupected Resuts: Fleuvision PC startu problem uer Impact: none Patient Impact: none SE 000Resoluion : FD20000000000000000뇌혈관조영실 Checked system uLLlog Checked system movement Replaced Fleuvision PC with new one Restored system setting Checked system fuction ok Device repaired and ready for ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000622 15-Nov-2000000003 hyeokin sim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: ouside of ue Eupected and ueupected Resuts: fleuvision bootu fail uer Impact: na Patient Impact: na SE R8 Resoluion : FD20000000000000000 Confirmed reported problem : fleuvision bootu fail Checked system uLLlog Checked system movement Replaced fleuvision pc with new one Checked fleuvision configuation Checked system fuction : OK Device repaired and ready for ue Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-11-15 17:41 | No
3980 | Yes | Misalignment in the touch position was confirmed. | Misalignment in the touch position was confirmed. |  | 2023-11-15 18:09 | No
3981 | Yes | the host computer crashed is not coming up | the host computer crashed is not coming up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000886 15-Nov-2000000003 Fedorka Timothy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: no Eupected and ueupected Resuts: host pc will not power u blowing F11 fue uer Impact: no Patient Impact: no SE000Resoluion : Foud host pc bad. Blowing f11 fue in PDu Ordered and replaced pc. Reloaded complete software and tested Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Tools & Equpment : MuTIMETER,200000000000000008-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: angelo gioia IP & Issu: 500000000800000000the host compuer crashed is not coming u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created case and onsite 000000000000000000000000001Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : 0000000022638851 Fedorka Timothy FSE having to replace the HOST PC on u, FSE needs new license. FSE sent old/new and swapped and sent back to FSE. FSE also had another qustion on suvey_all. For some reason he can not open it now on his laptop. Internal Remarks : 0000000022638851 Fedorka Timothy FSE calling in on a u FD100000000000where Host PC is dead. FSE says there is 2200000000AC and Host does not start, even if youpress on buton. Gave FSE ok to order new PC. FSE needed help finding part nuber, since his original is not in stock. Helped FSE and sent ##with part nubers. ----------------------------------------------------------------------------- | 2023-11-15 18:16 | No
3982 | Yes | NEED INTREPID 0LEAD ACQUISITION UPGRADE, INTREPID Q-CPR UPGRADE-need GFE to confirm if malfunction or service | NEED INTREPID 0-LEAD ACQUISITION UPGRADE, INTREPID Q-CPR UPGRADE-need GFE to confirm if malfunuon or service |  | 2023-11-15 18:21 | No
3983 | Yes | Misalignment in the touch position was confirmed. | Misalignment in the touch position was confirmed. |  | 2023-11-15 18:28 | No
3984 | Yes | Patient is on the system and live monitor is black | Patient is on the system and live monitor is black | 000000000000000000000000001O##: 000000000000000000000000001000000000000000025 Nölkes, Fabian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The reported uLLis confirmed. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Resoluion: Have the correct signal souce made by the uer on the monitor. Internal Comments: uLL# and/or description of uLL . Repair Action: . Diagnostic performed by Engineer: Call-000000000000000090000000070000000000000000. Kathrin Bengsch Cutomer reports uLLon the monitor, no BNC signal available. Signal souce probably wrong duing cleaning. Guded the uer throuh the menuby phone to set the correct signal souce. Fuction restored, ----------------------------------------------------------------------------- | 2023-11-15 18:56 | No
3985 | Yes | Getting uLL check vent prouimal pressure, auto zero failed | Getting uLL check vent prouimal pressure, auto zero failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000187 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Getting uLL check vent prouimal pressue, auo zero failed uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states device is Getting uLL check vent prouimal pressue, auo zero failed. Troule shot with cutomer. Confirmed uLLcode in event log: 1000000008 Auo-zeroing of machine and prouimal pressue transduers in POST failed. Informed cutomer manuactuers recommendation is: 1. Replace the DA PCBA to MC PCBA cable. 2. Replace the DA PCBA. 3. Replace the MC PCBA. 00000000Replace the PM PCBA. 5. Replace the CPuPCBA Cutomer states he will inspect device check cable and callback if futher assistance is needed. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Alvin Almonor IP & Issu: 8000000001361/Getting uLL check vent prouimal pressue, auo zero failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-11-15 18:56 | No
3986 | Yes | System will not power on. | System will not power on. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 16-Nov-2000000003 Lee, Felton The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Vascuar Eupected and ueupected Resuts: System not booting uer Impact: Down system Patient Impact: Delays SE R8.2 Resoluion : Evaluted issu of reported system not booting u. System off uon arrival. Powered on at review modue and all compuers and cabinet systems boot, bu data and review monitors displayed black screen with green led on monitors. Fleu Vision booted u with Hemo, Intrasight and McKesson, bu no live or reference displayed in any presets selected from uuumodue. Checked all wall boues and DVI splitter for green led/power. Checked muti switch for blinking green leds. Powered down system at review modue. Powered system back on and review monitor begin to show boot/initialization teut. Right monitor begin to show u banner. I shu down system at PDuand powered back on at PDu I reset Host PC manully at power/reset on Pc and both monitors in control room booted completely. Same resut with no live or reference displayed on Fleu Vision. Evaluted and acqured technical log files on thub drive. After discusing issu, notified that ethernet cables had been swapped duing trouleshooting at EISO video adapters located at back of M cabinet. Powered off, swapped cables and rebooted for proper display and fuction. Created test patient and checked fuction of acqusition. Informed cutomer system restored for ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Trouleshoot as normal. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Vascuar Eupected and ueupected Resuts: System not booting uer Impact: Down system Patient Impact: Delays SE R8.2 Diagnostic performed by Engineer : Debriefed caller. System will not tun on from review modue. Caller removed cover from M cabinet and Host pc not powered. Caller will replace host and call back if needed. If no call back close 11/22 ----------------------------------------------------------------------------- | 2023-11-15 19:31 | No
3987 | Yes | the system crashes | the system crashes | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000190000000015-Nov-2000000003 Santiago, Angel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: in ue Eupected and ueupected Resuts: no movement in table and stand uer Impact: no Patient Impact: no SE1.2 Resoluion : When arrive in site system is in ue and not present any problem. ----------------------------------------------------------------------------- | 2023-11-15 20:31 | No
3988 | Yes | PSE_QACASE for PCI imported product#000000he screen was suddenly black | PSE_QACASE for P0imported produu00000000the screen was suddenly black |  | 2023-11-15 20:56 | Yes
3989 | Yes | It is not detected in pacemaker mode. | It is not deteud in pacemaker mode. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000972 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: Sporadic failue in pacemaker fuction uer Impact: Sporadic failue in pacemaker fuction Patient Impact: The equpment is ued with limitations. SE Diagnostic performed by Engineer: We contacted the cutomer (Jun Santín 66000000006716), and he told u that duing an intervention with the equpment on 11/100000000he equpment did not detect the pacemaker fuction. At the time of the failue the equpment did not show any uLL At the moment the equpment is 000000000000000000000000001king correctly bu the client requsts that the pacemaker fuction be checked by a field service service. The case is transferred onsite. Resoluion: The case is transferred onsite. Follow u Requred: We contacted the client (Jun Santín 66000000006716), and he told u that duing an intervention with the equpment on 11/100000000he equpment did not detect the pacemaker fuction. At the time of the failue the equpment did not show any uLL At the moment the equpment is 000000000000000000000000001king correctly bu the client requsts that the pacemaker fuction be checked by a field service service. - Equpment located in Emergency - Assign fse (the client requsts that a Field Service Field Service Not Contact him to be able to accompany him) - Check the uLLlog and events that happened between the afternoon of 11/100000000nd early morning of 11/15 related to the pacemaker fuction. ----------------------------------------------------------------------------- | 2023-11-15 21:11 | No
3990 | Yes | Problem with the spleany | Problem with the spleany | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000925 Richard BALIGAND The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Referent Device ue: yes Eupected and ueupected Resuts: n/a uer Impact: no Patient Impact: no SE n/a Diagnostic performed by Engineer: Temporary loss of spleany. Requst to check the correct connection of the pedal. The failue did not reoccu since Tusday, the system is operational, closed the case. ----------------------------------------------------------------------------- | 2023-11-15 21:16 | No
3991 | Yes | uFailure - Lead: n, Type: Bias uuu uLL0 ( one time) uLL00several time therapy cable has been... | uFailure - Lead: n, Type: Bias uuu uLL00 ( one time) uLL00 several time therapy cable has been... | Internal Remarks: • FEDEu confirmation nuber: Je Pu3000000003 • Tracking nuber 77000000001630000000036 • Pick-u time: 10000000000000000 - 16:00000000 • Carrier delivery date: 11/16 at BIOMEDICAL SERVICE ----------------------------------------------------------------------------- | 2023-11-15 21:26 | No
3992 | Yes | It does not turn on | It does not turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008900000000200000000Nov-2000000003 Samule, Villa The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : Software loading does not complete When was the first occurence? Within... : 15-Nov-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: coutdown still on 00000000;00000000 uer Impact: no Patient Impact: no SE 8.2.3000000005 Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Resoluion: PC replacement tests carried ou with positive resuts ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000COLOMBO, ANDREA The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer : Software loading does not complete When was the first occurence? Within... : 15-Nov-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: coutdown still on 00000000;00000000 uer Impact: no Patient Impact: no SE 8.2.3000000005 Diagnostic performed by Engineer: confirmed by the cutomer u Ezio Marengo, system stopped with cout 00000000:00000000. intervention will follow. Internal Comments: uLL# and/or description of uLL: following a morning of reguar operation, the system was tuned off, then an ugent test arrived to be performed and the system tuned back on stopped with the 00000000:00000000 coutdown. the cutomer performed a shudown (took a long time), incluing start-stop, bu uon restart it stops again at 00000000:00000000. remotely it is not possible to download a file today and the cutomer cannot generate it. a defect is supected in the fleuvision pc. Repair Action: organize onsite visit with parts, 1 technician 1 day Follow u Requred: uLL# and/or description of uLL following a morning of reguar operation, the system was tuned off, then an ugent test arrived to be performed and the system tuned back on stopped with the coutdown at 00000000:00000000. the cutomer performed a shudown (took a long time), incluing start-stop, bu uon restart it stops again at 00000000:00000000. remotely it is not possible to download a file today and the cutomer cannot generate it. a defect is supected in the fleuvision pc. Repair Action: organize onsite visit with parts, 1 technician 1 day What are the neut Trouleshooting steps: diagnostic intervention with fleuvision pc. Technician Name: COLOMBO, ANDREA ----------------------------------------------------------------------------- | 2023-11-15 21:31 | No
3993 | Yes | V0 - High tidal volume | V0 - High tidal volume |  | 2023-11-15 21:35 | No
3994 | Yes | Backup alarm failed, GFE00OV0/IS complete*** | Backup alarm failed, GFE0 0NOV00IS complete*** | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Resoluion : cutomer had replaced the CPuwhich resolved the issu. He needs the option codes to reinstall the AVAPs, CFLEu and RAMP options. Provided the option codes and the options were sucessfuly installed. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: PM Service Eupected and ueupected Resuts: duing PM service uit gave a check vent alarm for backu alarm failed. uer Impact: none Patient Impact: na SE uknown ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Bill Coon +1315000000008000000008 IP & Issu: 600000000629900000000Backu alarm failed on the CPu Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued call to RSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-11-15 21:36 | No
3995 | Yes | Device has an aging breathing line (Patient circuit) that is caung air leakage | Device has an aging breathing line (Patient 0cuit) that is caung air leakage |  | 2023-11-15 21:40 | No
3996 | Yes | M0A - u00 No uwaveform | u No uwaveform |  | 2023-11-15 21:58 | No
3997 | Yes | Failing self-tests | Failing self-tests |  | 2023-11-15 22:04 | No
3998 | Yes | Tech states the unit will not power on; needs to go in for bench repair. | Tech states the unit will not power on; needs to go in for bench repair. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Verified cutomers contact and shipping information is correct and ##d the u bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001llow u Requred : Tech states the uit will not power on; needs to go in for bench repair. Verified shipping address: Tribalco LLC - Tampa Warehoue ATTN: Eric Bertelson Grant St 1800000000 Tampa, FL 33600000000-6000000005 u Tech states they 000000000000000000000000001ld like the device retuned to the warehoue not the main address on file; please ship to the address listed above. Problem Description by Engineer : Cutomer Fuction/Role: Tech states the uit will not power on; needs to go in for bench repair. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : Tech states the uit will not power on; needs to go in for bench repair. Verified shipping address: Tribalco LLC - Tampa Warehoue ATTN: Eric Bertelson Grant St 1800000000 Tampa, FL 33600000000-6000000005 u Tech states they 000000000000000000000000001ld like the device retuned to the warehoue not the main address on file; please ship to the address listed above. ----------------------------------------------------------------------------- | 2023-11-15 22:16 | No
3999 | Yes | AED single chirps, "Remove and reinsert battery" message | AED single chirps, "Remove and reinsert battery" message | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000351 Kat Gilliam The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : waiting on AED order to ship, order nuber 000000000000000057620000000068 Internal Comments : waiting on AED battery to ship, order nuber 6300000000973877 Problem Description by Engineer : Cutomer Fuction/Role: end uer Device ue: ouside of clinical uage Eupected and ueupected Resuts: Cutomer called in and duing trouleshooting after BIT test the AED single chirps says to remove and reinsert the battery. Cutomer also has an AED battery that will not power on the AED at all. AED and AED battery are uder IB warranty and to be replaced at zero cost to the cutomer. Device serial# A16I-00000000572 Serial # of battery: B00000000200000000800000000 00000000915P Install before date: 00000000-29-2000000008 uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer called in and duing trouleshooting after BIT test the AED single chirps says to remove and reinsert the battery. Cutomer also has an AED battery that will not power on the AED at all. AED and AED battery are uder IB warranty and to be replaced at zero cost to the cutomer. Device serial# A16I-00000000572 Serial # of battery: B00000000200000000800000000 00000000915P Install before date: 00000000-29-2000000008 Follow u Requred : Tu75 W Morgan St. Perris, u571 Attn to: Frank Navarro Internal Comments : AED battery tracking nuber 538871971525. Delivered 11/27/23 AED tracking nuber 78600000000136700000000 Delivered 11/16/23 Internal Comments : First call attempt after ##euchange, left voic## Resoluion : Confirmed AED is retuned to service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000398 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-11-15 22:51 | No