## Question
These all are complaints data reeived from customers
Questions
Analyze Description 1 , Description 2 and Notes to provide insights on the following
what are the main concerns customers have
what are the different topics/topic clusters that are emerging-
Build a Classification model to predict if it is an important event
Any other inight that you want to share,
Please submit approach/code used for analysis and arriving at the result

## Data
ID | Indicator Flag | Description 1 | Description 2 | Notes | Date | Important Event
1 | Yes | Battery request | Battery request | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000033 Assad MAJEED The system does not meet full specifications for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled. Resolution: Recommendation: 0000000056000000009791 Part requested for replacement. Follow u Required: Delivery of the part to the biomedical department if contract, otherwise, request to AVF. Regulatory Questions: No/No Problem Description by Engineer: Customer Faction/Role: Biomedical Device use: unknown Expected and unexpected Results: unknown our Impact: None Patient Impact: None SE unknown Diagnostic performed by Engineer: The customer wants to order a battery. ----------------------------------------------------------------------------- Internal Remarks: The customer called again and did not get his order. I am escalating the case. ----------------------------------------------------------------------------- | 2023-09-01 14:11 | No
2 | Yes | u - Paddle Plate is Not Working | Paddle Plate is Not Working | 000000000000000000000000001O##: 000000000000000000000000001000000009000000000000000012-Sep-2000000003 Baraca, Tiziani The system meets the specification for the performed service and is retuned to use. Problem Description by Engineer: Customer Faction/Role: Clinical Engineer Device use: defibrillator test Expected and unexpected Results: external plates are damaged on the handles our Impact: no impact Patient Impact: no impact SE an Parts Consumed: 1.0000000000000000000000000000000056300000000991, Paddle Repl. Kit Water Resistant Resolution: Replaced plates. Carried our checks with a positive outcome. The device is u and running. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000090000000000000000 FALASCA, IVO The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u scheduled 000000000000000000000000001k Problem Description by Engineer: Customer Faction/Role: Clinical Engineer Device use: defibrillator test Expected and unexpected Results: external plates are damaged on the handles our Impact: no impact Patient Impact: no impact SE an Follow u Required: Sterne plates are damaged, replacement is necessary, I order part, schedule on-site service for replacement and functional tests. If they are not available, they need to be tested. To request from Sided. The uluas not supported. If the first code is not available, try with this one: 98980000000012900000000 Diagnostic performed by Engineer: External plates are damaged. ----------------------------------------------------------------------------- Internal Remarks: Good morning, This is to request supply of plate for the following equipment: T000000000000000000000000001PHASE Manual DEFIBRILLATOR Inventory number: A-0000000022 buy: M 3535 A HEART START ulu Manufacture: u MEDICAL SYSTEMS 000S00000000538795 u3 - RED CODE, Site: A000000001, A.O. PAPARDO Medical Department: A000000001-37, EMERGENCY Reported faut: plate is not 000000000000000000000000001king, du to accidental damage. Attached photo. For further information contact the reference technician: Grammatist was in use 380000000015189 Details of the intervention Number: 2000000003000000001900000000 Cordially Falsetto Elisa Service Administration Specialist Clinical Engineering Service. Pepard Messina A.O. Conrad Sherone Building A 98158 Messina (Me) ----------------------------------------------------------------------------- | 2023-09-01 15:51 | No
3 | Yes | Alarms/No sound: No sound alarm anymore: No Trouble Found: Ur Configuration: Provided Information. | Alarms/No sound: No sound alarm anymore: No Trouble Found: Ur Configuration: Provided Information. | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000011 00000000-Sep-2000000003 Olivier MACARIO The system meets the specifications for the performed service and is retuned to use. Regulatory Questions: No/No Resolution: Problem occurred on 8/300000000in room 3300000000: checked the correct operation OK, no problem observed. Retrieved the logs and the configuration of the monitor and the central stations. It is normal that the alarms of the ullaged not reported on the monitors and on the central station. Problem Description by Engineer: Customer Faction/Role: Biomed Device use: unknown Expected and unexpected Results: No alarm sound. our Impact: No. Patient Impact: No. SE N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000090000000058 MESNIL, Maxime The system meets the specifications for the performed service and is retuned to use. Diagnostic performed by Engineer: The department indicates that on an occasional basis the monitor lost the audio alarms. The customer could not observe or reproduce the failure but wants on-site service because an undesirable event sheet related to this event was opened. Problem Description by Engineer: Customer Faction/Role: Biomed Device use: unknown Expected and unexpected Results: No alarm sound. our Impact: No. Patient Impact: No. SE N/A Regulatory Questions: No/No Resolution: Request for on-site service within the frame000000000000000000000000001k of the contract. REF: 0000000056000000006971 Loudspeaker Assy QTY: 1 Follow u Required: The department indicates that on an occasional basis the monitor lost the audio alarms. The customer could not observe or reproduce the failure but wants on-site service because an undesirable event sheet related to this event was opened. LOOKING FOR Resources Delivery of the part to the biomedical department. ----------------------------------------------------------------------------- | 2023-09-01 16:46 | No
4 | Yes | The u cable does not work - | The usable does not work - | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000091 Kimberton Masipa, Francisco The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Regulatory Questions: N/A Problem Description by Engineer: Customer Faction/Role: Electromedicine Device use: Defibrillator Expected and unexpected Results: usable failure. our Impact: Impossible to register u Patient Impact: The equipment is used with limitations. SE Diagnostic performed by Engineer: The client (Raquel Carrasco 9158680000000000000000 says that the usable is damaged and requests its replacement under warranty. Material order is handled under warranty. Resolution: Material order is handled. Follow u Required: The client (Raquel Carrasco 9158680000000000000000 says that the usable is damaged and requests its replacement under warranty. Manage material order under warranty. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000090000000079 ----------------------------------------------------------------------------- | 2023-09-01 18:31 | No
5 | Yes | Exposer in not possible reselect application uLLis appearing | Exposer in not possible reseleuapplication uLLis appearing | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000000000000 00000000-Sep-2000000003 Souav, Poddar The system meets the specification for the performed service and is retuned to use Parts Consumed : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resolution : Image disk replaced and system 000000000000000000000000001king properly Problem Description by Engineer : Problem description by engineer : Exposer in not possible reselect application uLLis appearing When was the first occurrence? Within... : 00000000-Sep-2000000003 18:00000000:00000000 How often is issue occurring? : First occurrence Malfunction area : HDD Information to support the complaint handling process Customer Function/Role: BioMed Device use: No patient involved Expected and unexpected Results: Exposer in not possible reselect application uLLis appearing user Impact: No user impact Patient Impact: No patient impact SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000090000000091 Vijayakuar, K.S The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : Exposer in not possible reselect application uLLis appearing When was the first occurrence? Within... : 00000000-Sep-2000000003 18:00000000:00000000 How often is issue occurring? : First occurrence Malfunction area : HDD Information to support the complaint handling process Customer Function/Role: BioMed Device use: No patient involved Expected and unexpected Results: Exposer in not possible reselect application uLLis appearing user Impact: No user impact Patient Impact: No patient impact SE 000Diagnostic performed by Engineer : One of the image disk is throwing read full This needs replacement. It is recommended to replace both image disks of 1 TB to make it equal unity instead of replacing only one. Image disks needs replacement for permanent solution. Here the problem is intermittent. Internal Comments : full# and/or description of full: Exposer in not possible reselect application uLLis appearing Repair Action : One of the image disk is throwing read full This needs replacement. It is recommended to replace both image disks of 1 TB to make it equal unity instead of replacing only one. Image disks needs replacement for permanent solution. Here the problem is intermittent. Resolution : One of the image disk is throwing read full This needs replacement. It is recommended to replace both image disks of 1 TB to make it equal unity instead of replacing only one. Image disks needs replacement for permanent solution. Here the problem is intermittent. Follow u Required : full# and/or description of full: Exposer in not possible reselect application uLLis appearing Repair Action : One of the image disk is throwing read full This needs replacement. It is recommended to replace both image disks of 1 TB to make it equal unity instead of replacing only one. Image disks needs replacement for permanent solution. Here the problem is intermittent. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-01 18:36 | No
6 | Yes | fluoroscopy was not working. | fluoroscopy was not working. | 000000000000000000000000001O##: 000000000000000000000000001000000009000000002 00000000-Sep-2000000003 Shula Chintan The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal 00000000-Sep-2000000003 00000000:3000000003 RMT - urgent: HostPC could not open or lost connection with Velara generator - Frontal 00000000-Sep-2000000003 00000000:3000000003 How often is issue occurring? : First occurrence Malfunction area : u-ray generation Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device use: No further information is available as per the customer Expected and unexpected Results: Not provided user Impact: usable to use system Patient Impact: patient had to be moved to another room to complete procedure SE 8.1.300000000100000000Resoluion : Generator was found not tuned on. Checked connection and found loose “L1” connection on K2 relay. Tightened the wire. Tested system to confirm proper operation. Retuned system back to clinical use. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000090000000057 Van Pelt Daniel The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Internal Comments : full# and/or description of full: Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. Relay (tap) in MPDucontrol board is defective Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. Cue is defective Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. The 200000000cÂ power supply (NT200000000located in the u-ray generator is malfunctioning. Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. Problem with the hospital Mains power Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. This could be caused du a timing issue between Host(PC) and Generator. Please ignore this alert, if it only occurs sporadically. Treat it as a 'one time only'. Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. Problem with EN100000000 PCB Root Cause for RMT - urgent: HostPC could not open or lost connection with Velara generator - Frontal: 1. The HostPC (TH) was not able to connect or suddenly lost connection to the Frontal Velara u-ray generator. Repair Action : Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. If the problem is permanent: Switch off the system (but make sue F1 to F5 MCBs in MPDuare on) and connect D-connector measuring board at NPu: 300000000 in M-cabinet. Measure the continuity between pins 00000000nd 11. This should be normally closed, If it is open then replace MPDucontrol board. If the problem is intermittent: Switch off the system. Disconnect the cable from NPu: 300000000 (frontal) in M-cabinet(Frontal) and connect it to NPu:300000000 (lateral). Observe the system. If the problem is resolved, MPDucontrol board is defective. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Check the LED status of Cue. If no LED's are lit, please check 200000000power supply first.(see also 'Cause 5') If Power supply is ok but LEDs are off, replace Cue. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Measure the 200000000c power supply (NT200000000located on the EN unit in the u-ray generator. Replace when necessary. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Check the hospital Mains inputs power to the Velara Generator. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Perform a reboot the CV system. If the problem still persists after a reboot, carry our recommended actions of the other Solution here below. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Check whether the green power LED is on at EN100000000 pcb. If the LED is not lit, replace this PCB Service actions for RMT - urgent: HostPC could not open or lost connection with Velara generator - Frontal: 1.1. If this occurs only sporadically youkan ignore it. Please do a cold reboot of the whole system in order to recover. (This can happen if the HostPC was booted in the morning but the u-ray generator was not switched off the last time so the generator is still on. Then there is no initialization sequence hence the HostPC is usable to connect.) 1.2. ENF1 inside the frontal Velara generator could be tripped. Reset breaker and restart system. 1.3. If it occurs more frequently check MPDu Could be that it does not switches off the mains tap 100000000for the Velara generator. 1.00000000Check tap 100000000of the MPDu Could be a malfunctioning MPD Control unit. (NP100000000 1.5. Check the 200000000c Low Voltage Power Supply (NT200000000in the Velara generator. Is the voltage stable and Ok? 1.6. Check condition of the safety resistors ENR1, ENR2, ENR3 (00000000hm) which are located in the Velara generator. If defective please also check the condition of the converters. (EQ) 1.7. Could also be du a defective EN100000000 PCB. (=Power On circuit) 1.8. Check connections on the SIBBOu (MDY). Especially check connectors: MDY-u71 (= CAN Velara Frontal) 1.9. Could be du to a defective Cue. (EZ) Possible required parts: 00000000210000000020000000091 - Cue CVFD-5 ASSY. 00000000800000000000000002855 - SIBBOu unit 00000000210000000083771 - Power Supply 200000000100000000 00000000210000000015282 - Power on circuit 000000002100000000000000002 - Safety resistor 00000000, 300000000133mm 00000000230000000016000000007 - MPD CONTROL unit Follow u Required : full# and/or description of full: Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. Relay (tap) in MPDucontrol board is defective Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. Cue is defective Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. The 200000000cÂ power supply (NT200000000located in the u-ray generator is malfunctioning. Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. Problem with the hospital Mains power Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. This could be caused du a timing Issu between Host(PC) and Generator. Please ignore this alert, if it only occurs sporadically. Treat it as a 'one time only'. Root Cause for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1. Problem with EN100000000 PCB Root Cause for RMT - urgent: HostPC could not open or lost connection with Velara generator - Frontal: 1. The HostPC (TH) was not able to connect or suddenly lost connection to the Frontal Velara u-ray generator. Repair Action : Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. If the problem is permanent: Switch off the system (but make sue F1 to F5 MCBs in MPDuare on) and connect D-connector measuring board at NPu: 300000000 in M-cabinet. Measure the continuity between pins 00000000nd 11. This should be normally closed, If it is open then replace MPDucontrol board. If the problem is intermittent: Switch off the system. Disconnect the cable from NPu: 300000000 (frontal) in M-cabinet(Frontal) and connect it to NPu:300000000 (lateral). Observe the system. If the problem is resolved, MPDucontrol board is defective. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Check the LED status of Cue. If no LED's are lit, please check 200000000power supply first.(see also 'Cause 5') If Power supply is ok but LEDs are off, replace Cue. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Measure the 200000000c power supply (NT200000000located on the EN unit in the u-ray generator. Replace when necessary. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Check the hospital Mains inputs power to the Velara Generator. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Perform a reboot the CV system. If the problem still persists after a reboot, carry our recommended actions of the other Solution here below. Service actions for RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal: 1.1. Check whether the green power LED is on at EN100000000 pcb. If the LED is not lit, replace this PCB Service actions for RMT - urgent: HostPC could not open or lost connection with Velara generator - Frontal: 1.1. If this occurs only sporadically youkan ignore it. Please do a cold reboot of the whole system in order to recover. (This can happen if the HostPC was booted in the morning but the u-ray generator was not switched off the last time so the generator is still on. Then there is no initialization sequence hence the HostPC is usable to connect.) 1.2. ENF1 inside the frontal Velara generator could be tripped. Reset breaker and restart system. 1.3. If it occurs more frequently check MPDu Could be that it does not switches off the mains tap 100000000for the Velara generator. 1.00000000Check tap 100000000of the MPDu Could be a malfunctioning MPD Control unit. (NP100000000 1.5. Check the 200000000c Low Voltage Power Supply (NT200000000in the Velara generator. Is the voltage stable and Ok? 1.6. Check condition of the safety resistors ENR1, ENR2, ENR3 (00000000hm) which are located in the Velara generator. If defective please also check the condition of the converters. (EQ) 1.7. Could also be du a defective EN100000000 PCB. (=Power On circuit) 1.8. Check connections on the SIBBOu (MDY). Especially check connectors: MDY-u71 (= CAN Velara Frontal) 1.9. Could be du to a defective Cue. (EZ) Possible required parts: 00000000210000000020000000091 - Cue CVFD-5 ASSY. 00000000800000000000000002855 - SIBBOu unit 00000000210000000083771 - Power Supply 200000000100000000 00000000210000000015282 - Power on circuit 000000002100000000000000002 - Safety resistor 00000000, 300000000133mm 00000000230000000016000000007 - MPD CONTROL unit Technician Name : Van Pelt Daniel Problem Description by Engineer : Problem description by engineer : RMT - urgent: Host(PC) is usable to communicate via CAN with Velara Generator - Frontal 00000000-Sep-2000000003 00000000:3000000003 RMT - urgent: HostPC could not open or lost connection with Velara generator - Frontal 00000000-Sep-2000000003 00000000:3000000003 How often is Issu occurring? : First occurrence Malfunction area : u-ray generation Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device use: No further information is available as per the customer Expected and unexpected Results: Not provided user Impact: usable to use system Patient Impact: patient had to be moved to another room to complete procedure SE 8.1.300000000100000000Diagnostic performed by Engineer : Reviewed log files ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: David Cridland IP & Issu: 60000000000000000587-patient on table fluo not 000000000000000000000000001king Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Create R000000000000000000000000001and warm trans to rise Neat Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-01 18:46 | No
7 | Yes | The system did not startup. | The system did not startup. | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000012 00000000Sep-2000000003 ANDREA BOCCAFOGLI The system meets the specification for the performed service and is retuned to use. Problem Description by Engineer: Problem description by engineer: Technician opens, problems at the start-u the M Cabinet does not start. The system is down. uCotignola Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device use: Diagnostic Expected and unexpected Results: Technician opens, problems at the start-u the M Cabinet does not start. The system is down. Location; Cotignola user Impact: yes Patient Impact: yes SE 7.2.16.00000000Resoluion: Replaced power tray cabinet M. The fuction tests were regular. Parts Consumed: 1.00000000000000000000000000000000230000000028523,PFS272 POWERTRAY FRu----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009000000005 TESTOLIN, DARIO The system does not meet the specification for the performed service and is not retuned to use. Internal Comments: full# and/or description of full . Repair Action: As per information, the part is necessary. Follow u Required: full# and/or description of full . Repair Action: As per information, the part is necessary. What are the neat Troubleshooting steps: . Technician Name: TESTOLIN, DARIO Problem Description by Engineer: Problem description by engineer: Technician opens, problems at the start-u the M Cabinet does not start. The system is down. uCotignola Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device use: Diagnostic Expected and unexpected Results: Technician opens, problems at the start-u the M Cabinet does not start. The system is down. Location; Cotignola user Impact: yes Patient Impact: yes SE 7.2.16.00000000Diagnostic performed by Engineer: as per information, part is necessary ----------------------------------------------------------------------------- Internal Remarks: Good morning, Please open a call on configuration 51000000002523 for "Problems at the start-u the M Cabinet does not start" Reference technician Tedeschi 0000000000000000217127 For RSE it is necessary to order PFS272 POWERTRAY FRu00000000800000000000000001500000000system down, no need to contact the customer. They are waiting for a plan as early as possible. Ettore Cattoli Field Service Engineer u HealthNet 1000000006 Cadui Del Pilastro Street 0000000055 Castenaso (BO) - Italy Phone: +39 800000000 232100000000, Fau: +39 000000001600000000200000000 e-mail: ettore.cattoliucom www.u.com/healthcare Logo New u ----------------------------------------------------------------------------- | 2023-09-01 20:01 | No
8 | Yes | Blower not running, alarming for patient disconnect | Blower not running, alarming for patient disconnect | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000011 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer : Customer Function/Role: biomed Device use: not in patient use Expected and unexpected Results: blower not running, alarming for patient disconnect user Impact: none Patient Impact: an SE unknown Resolution : Customer is taking responsibility to correct/repair the device. The customer has been notified to contact u if this does not address their device Issu at which point a new service order will be created. No further information will be obtained as this conclude u remote support to the customer for this event. Diagnostic performed by Engineer : with no patient circuit connected, on the pneumatics page with pressure set to 100000000 blower RPM = 30000000000000000 blower amps = 2. Advised to replace the blower. Provided parts ID. 00000000561512761 ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Thomas O'Leary Issu: 566310000000000000000Will not run and alarm patients Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001 Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-01 20:06 | No
9 | Yes | The touchscreen is not working. | The touchscreen is not working. | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000000000000Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to use Resolution : Customer is taking responsibility to correct/repair the device. The customer has been notified to contact u if this does not address their device Issu at which point a new service order will be created. No further information will be obtained as this conclude u remote support to the customer for this event. Diagnostic performed by Engineer : Tried touchscreen calibration, failed. provided parts ID for the touchscreen for them to provide to in-house biomed. 00000000561539000000001 Problem Description by Engineer : Customer Function/Role: RT Device use: not in patient use Expected and unexpected Results: touchscreen unresponsive user Impact: none Patient Impact: an SE unknown ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jason Byun IP & Issu: 678982900000000- The touchscreen is not 000000000000000000000000001king. Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Regular Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001 Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-01 20:21 | No
10 | Yes | Failed Self Test/Pads iMessage | Failed Self Test/Pads iMessage | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000032 Pirarucu, Ida The system meets the specification for the performed service and is retuned to use Internal Comments : Order 630000000080000000000000000Diagnostic performed by Engineer : A few days ago, one of my ports of entry used the defibrillator. Plugging in the new pair (#23000000006-0000000000000000, the machine wrote: TEST FAILED Several other ports tested their pads and 8 of them failed and I’m still waiting for the results from other ports. They tried to take our the battery and up it back in and tested the new pads again; no Luk, still TEST FAILED They also tested with old pairs (up 00000000-31-2000000003) and they all passed the test. # REF 98980000000010000000081 8 defectives # LOT 23000000006-0000000000000000 #REF 98980000000010000000081 1 defective # LOT 220000000021-000000002 ASFC-CBSA ATT: ELAINE Cuso Battement B (Vieu Commercial) 500000000 AutoRoute 15 us St-Bernard-de-Lucille, QC J00000000 1V00000000 Will replace pads under warranty but usable to retrieve pads back for evolution du to its location in Canada. Problem Description by Engineer : Customer Function/Role: Admin Device use: no Expected and unexpected Results: pads alms user Impact: none Patient Impact: none SE PR2.1.00000000 See ref case 000000002220000000000000000 Test failed with new pads Resolution : Pads delivered. However upon confirmation upon receiving shipment and installing the pads with the customer, it was determined that the cut did not have the pads in the sentry doing self tests. Advised cut that pads need to be in the sentry. Closing case. ----------------------------------------------------------------------------- | 2023-09-01 20:51 | No
11 | Yes | Failed Self Test | Failed Self Test | 000000000000000000000000001O##: 00000000000000000000000000100000000900000000600000000Apiruvaraku, Ida The system meets the specification for the performed service and is retuned to use Internal Comments : Order 630000000080000000000000000Diagnostic performed by Engineer : A few days ago, one of my ports of entry used the defibrillator. Plugging in the new pair (#23000000006-0000000000000000, the machine wrote: TEST FAILED Several other ports tested their pads and 8 of them failed and I’m still waiting for the results from other ports. They tried to take our the battery and up it back in and tested the new pads again; no Luk, still TEST FAILED They also tested with old pairs (up 00000000-31-2000000003) and they all passed the test. # REF 98980000000010000000081 8 defectives # LOT 23000000006-0000000000000000 #REF 98980000000010000000081 1 defective # LOT 220000000021-000000002 ASFC-CBSA ATT: ELAINE Cuso Battement B (Vieu Commercial) 500000000 AutoRoute 15 us St-Bernard-de-Lucille, QC J00000000 1V00000000 Will replace pads under warranty but usable to retrieve pads back for evolution du to its location in Canada. Problem Description by Engineer : Customer Function/Role: Admin Device use: no Expected and unexpected Results: pads alms user Impact: none Patient Impact: none SE PR2.1.00000000 See ref case 000000002220000000000000000 Test failed with new pads Resolution : Pads delivered. However upon confirmation upon receiving shipment and installing the pads with the customer, it was determined that the cut did not have the pads in the sentry doing self tests. Advised cut that pads need to be in the sentry. Closing case. ----------------------------------------------------------------------------- | 2023-09-01 20:56 | No
12 | Yes | u- The defibrillator plate corrosion was found in regular testing, | The defibrillator plate corrosion was found in regular testing, |  | 2023-09-01 21:07 | No
13 | Yes | Audio-No Sound / Speaker Malt-Intermittent Sound-Foulink s/nu/ Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / Speaker Malt-Intermittent Sound-Foulink s/n u000/ Other HW-Speaker-HW / Replaced Parts | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000551 The system meets the specification for the performed service and is retuned to use ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000950000000085 Lefker, Harry The system does not meet the specification for the performed service and is not retuned to use Resolution : Resolution Detail: Sending to bench Follow u Required : Bench has speaker Malt sound on Mu00000000is intermittent Problem Description by Engineer : Customer Function/Role: Biomed Device use: Not in Clinical use Expected and unexpected Results: user Impact: Patient Impact: SE Problem Description/unexpected Results: Bench has speaker Malt sound on Mu00000000is intermittent Diagnostic performed by Engineer : Problem Description: Bench has speaker Malt sound on Mu00000000is intermittent Bench has speaker Malt sound on Mu00000000is intermittent Form used: Other Troubleshooting Steps Taken: Bench has speaker Malt sound on Mu00000000is intermittent ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-09-01 21:51 | No
14 | Yes | Oxygen failed to be flowed when the u manifold kit | Oxygen failed to be flowed when the O0manifold kit |  | 2023-09-01 22:54 | No
15 | Yes | No exposure | No exposure | 000000000000000000000000001O##: 00000000000000000000000000100000000950000000000000000 11-Sep-2000000003 Maua George The system meets the specification for the performed service and is retuned to use Parts Consumed : 1.00000000000000000000000000000000210000000013887,Power Inverter (Point) Certeray Resolution : Inverter full replaced power inverter/calibrated true Problem Description by Engineer : Customer Function/Role: user Device use: u Expected and unexpected Results: no up. user Impact: n/a Patient Impact: SE Tools & Equipment : DOSE-KVP-LIGHT METER,2000000007-00000000-28 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000950000000026 00000000-Sep-2000000003 Maua George The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Follow u Required : part Problem Description by Engineer : Customer Function/Role: user Device use: u Expected and unexpected Results: no up. user Impact: n/a Patient Impact: SE Diagnostic performed by Engineer : defective point/ part on order ----------------------------------------------------------------------------- Internal Remarks : FSE George Maua. FSE sent in logs and see many full of From Generator device: Point: Resonance frequency too high. Issu covered in training and webinars. Suggested to replace Point. Approved for replacement and provided p/n: 00000000210000000013887 - Power Inverter (Point) Certeray. ----------------------------------------------------------------------------- | 2023-09-01 23:01 | No
16 | Yes | Power cord failure | Power cord failure |  | 2023-09-02 | No
17 | Yes | u Repeating Self Test .. | AED Repeating Self Test .. | 000000000000000000000000001O##: 00000000000000000000000000100000000950000000057 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer : Customer Function/Role: End Customer Device use: Device was not in use Expected and unexpected Results: Customer called and stated that the AED was found to be chirping 2 weeks ago and was not brought to the attention of customer contact util most recent. The AED seems to not accept the AED battery when first inserted. The customer reinserted the battery, the AED began the self test and then restarted the self test in the middle of testing. The AED needs to be replaced at zero cost to the customer under IB Warranty Exchange. Device serial# A18A-00000000716 Serial # of battery B00000000200000000800000000 00000000229P Install before date: 00000000-3000000002000000008 user Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Customer called and stated that the AED was found to be chirping 2 weeks ago and was not brought to the attention of customer contact util most recent. The AED seems to not accept the AED battery when first inserted. The customer reinserted the battery, the AED began the self test and then restarted the self test in the middle of testing. The AED needs to be replaced at zero cost to the customer under IB Warranty Exchange. Device serial# A18A-00000000716 Serial # of battery B00000000200000000800000000 00000000229P Install before date: 00000000-3000000002000000008 Internal Comments : waiting for order to ship. order number 000000000000000057296675 Follow u Required : Ship to Cartersville Fire Department Address: 195 Cassville Rd Cartersville GA 300000000200000000 Attention: Tracey Mcfadden Resolution : Customer Tracey called 00000000/00000000/2000000003 and has confirmed She received the replacement AED and the AED is 000000000000000000000000001king as it should. In addition she is also mailing back the defected AED ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009510000000000000000 Parts Consumed : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Each ----------------------------------------------------------------------------- | 2023-09-02 00:11 | No
18 | Yes | u is emitting the triple chirp | FRu AED is emitting the triple chirp | 000000000000000000000000001O##: 000000000000000000000000001000000009500000000300000000Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to use Resolution : Closing case AED no longer under warranty. Instructed the customer to remove the AED from service and not use the device on anyone . Problem Description by Engineer : Customer Function/Role: Admin Device use: N/A Expected and unexpected Results: Customer reported that AED is emitting chirps and will report uLLcode Not ready to use "Remove and reinsert battery to test. user Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Customer reported that AED is emitting chirps and will report uLLcode Not ready to use "Remove and reinsert battery to test. Diagnostic performed by Engineer : Customer reported that AED is emitting chirps and will report uLLcode Not ready to use "Remove and reinsert battery to test. Instruct customer to reinsert battery the AED Failed iMessage of Not ready to use "Remove and reinsert battery to test. Instructed the Customer to do du another self-test to remove battery and reinsert, AED had the same iMessage Not ready to use "Remove and reinsert battery to test.. Advise customer that the AED has a 8 to 100000000year life span and you current machine is 12 years old, the Issu is with the device. Instructing the customer they will need to remove from service and not use the device on anyone. AED Serial# B11F00000000268 Battery Install Before# 8/31/2000000008 Lot#00000000975P Pads Up Date: 8/19/2000000003 Pads Lot #2100000000190000000058 ----------------------------------------------------------------------------- | 2023-09-02 00:16 | No
19 | Yes | Incident investigation for patient event-off-No product malfunction- ur training-u error | In0ent investigation for patient event-off-No produumalfunuon- ur training-u error | 000000000000000000000000001O##: 00000000000000000000000000100000000951200000000Miller, Dell The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Resolution : Is Cause Known: No Cause: Is Resolution Known: No Resolution Detail: Onsite service required, FSE to resolvases to pull IVPM configuration files, logs, and perform a fuction ck. Also, pull PICIu clinical audit logs. Please attach logs, configuration files, any us trips, etc. to the case for Q/R review. Sent to planners for onsite support. Follow u Required : PSE du to delay in treatment. FSE to pull IVPM configuration files, logs, and perform a fuction ck. Also, pull PICIu clinical audit logs. Please attach logs, configuration files, any us trips, etc. to the case for Q/R review. Problem description by Clinical Support Specialist : Customer Fuction/Role: Biomed Device use: In Clinical use Expected and unexpected Results: u to evaluate the issue. user Impact: No impact Patient Impact: undelays in treatment/no pt info available. SE uu Problem Description/unexpected Results: Need onsite support to pull logs from IVPM/PICIu. Diagnostic performed by Engineer : Problem Description: Need onsite support to pull logs from IVPM/PICIu. Cureports delay in treatment on 00000000/2000000003 at approach. 0000000000000000 hrs., Rm #3119/3 Souh. SPO2 alarms were allegedly tuned off by clinical staff on MP700000000IVPM. Requesting onsite support for u to investigate the incident by puling logs from IVPM/PICIu and evaluated by Q/R. Form used: Troubleshooting/Investigation Object: un. Not Object: Other functions. Software Version: uu Date and Time of Issu: 9/1/2000000003 user Actions: Pt monitoring. Where is the Deviation: 3 Souh PICIu. Troubleshooting Steps Taken: Confirmed contact information. Was PRS used: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000951300000000 00000000-Sep-2000000003 Yui Ruakov The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: Biomed Device use: In Clinical use Expected and unexpected Results: u to evaluate the issue. user Impact: No impact Patient Impact: undelays in treatment/no pt info available. SE uu Problem Description/unexpected Results: Need onsite support to pull logs from IVPM/PICIu. Diagnostic performed by Engineer : SPo2 alarm was not going off for patient and patient in room 3S3119 u number 228000000001 adult male went into low oxygen alarm and no alarm sound from monitor SN# - DE8000000006325 MP700000000 Later found that staff were tuning off SPo2 alarms on certain monitors. Resolution : SPo2 alarm was not going off for patient and patient in room 3S3119 u number 228000000001 adult male went into low oxygen alarm and no alarm sound from monitor SN# - DE8000000006325 MP700000000 Later found that staff were tuning off SPo2 alarms on certain monitors. Staff were retrained. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Jeff Bassett Issu: 9850000000000000000 Incident investigation for patient event Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to clinical support Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-02 00:56 | Yes
20 | Yes | Bad display | Bad display | 000000000000000000000000001O##: 00000000000000000000000000100000000951165 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer : Customer Fuction/Role: purchasing Device use: not in patient use Expected and unexpected Results: needs u says, bad display user Impact: none Patient Impact: an SE unknown Diagnostic performed by Engineer : customer actually wants to order several u Assys, different colors. When trying to order 0000000056150000000051 or 0000000056150000000091 they are being told by parts that they are not orderable, not yet released form the vendor for ordering. I called parts and verified that is the case. Advised customer that we expect them to be available in December du to supply chain issues and to try ordering closer to then to place the backorder. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tamya Goud IP & Issu: 5700000000900000000Parts Id Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and round to RSE Neat Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-02 01:16 | No
21 | Yes | Part u display goes out. | Part ID. LCD display goes out. | 000000000000000000000000001O##: 00000000000000000000000000100000000951200000000 Miller Jason The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: biomed Device use: outside of use Expected and unexpected Results: Parts ID user Impact: system restricted Patient Impact: none reported SE unknown Resolution : Provided customer with part number and price: 0000000056153600000000 - RP-Cable, LCD, NEC, 2nd Gen, V600000000- 138.00000000 Customer is taking responsibility to correct/repair the device. The customer has been notified to contact u if this does not address their device issue at which point a new service order will be created. No further information will be obtained as this conclude u remote support to the customer for this event. Diagnostic performed by Engineer : Customer states the LCD cable is faulty and needs replaced. Customer is requesting replacement LCD cable part number. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000951800000000 Parts Consumed : 1.0000000000000000000000000000000056153600000000,RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Lester Embry 93732157300000000 IP & Issu: 73000000002736/Part ID. LCD display goes our. Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and round to RSE Neat Step Action: N/A Verbal PO Number: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-02 02:11 | No
22 | Yes | not getting any ac power | not getting any ac power | 000000000000000000000000001O##: 00000000000000000000000000100000000950000000071 Miller Jason The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: biomed Device use: outside of use Expected and unexpected Results: will not 000000000000000000000000001k with AC power user Impact: system restricted Patient Impact: none reported SE unknown Resolution : Provided customer with part number: 00000000561512591 - power supply Customer is taking responsibility to correct/repair the device. The customer has been notified to contact u if this does not address their device issue at which point a new service order will be created. No further information will be obtained as this conclude u remote support to the customer for this event. Diagnostic performed by Engineer : Customer states unit will only operate on battery power. Trouble shot with customer: 1. Verify AC outlet for proper voltage. 2. Verify that power cord is functional. 3. Verify power supply (200000000: a. With AC power disconnected, remove J1 from PM PCBA. b. Reconnect AC power. c. Verify that 200000000is present between the orange and brown wires on the power supply harness connector. d. If 200000000is present, replace the PM PCBA. e. If 200000000is not present, go to neat step. 00000000Verify that AC power is present: a. Disconnect AC power. b. Remove EMI shrou. c. Reconnect AC power. d. Verify that AC voltage is present between the blue and brown wires coming from the AC inlet. e. If AC power is present, replace the power supply Customer seeing 1200000000coming in. Customer is requesting replacement part number for power supply. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: matt IP & Issu: 60000000072271 not running on ac Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Regular Entitlement Options: Free Remote Reason for Entitlement Claim: an Action Taken: created case Neat Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-02 02:11 | No
23 | Yes | Unable to turn off the brightness and looks very dim | Unable to turn off the brightness and looks very dim | 000000000000000000000000001O##: 0000000000000000000000000010000000095100000000 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to use Resolution : Advised customer that unit may have original HYDIS lcd and will need to be updated to 2nd gen NEC lcd, provided parts for either replacing u says or rebuilding current display, customer acknowledged, no further follow u needed, issue resolved- u assembly, without nav-ring, V600000000(gray) pn 0000000056150000000091 or can order internal display parts- 00000000561528891 [RP-LCD, Assy,2nd GEN, V600000000 0000000056153600000000 [RP-Cable, LCD, NEC, 2nd Gen, V600000000 00000000561528911 [RP-Cable, u PCBA to LCD, V600000000 00000000561528871 [RP-PCBA,uer Interface 2nd GEN,V600000000 00000000561531651 [RP-Tray,LCD,2nd GEN V600000000 00000000561528900000000 [RP-Screw,M2u3 Socket had, PKG/100000000V600000000 Customer is taking responsibility to correct/repair the device. The customer has been notified to contact u if this does not address their device issue at which point a new service order will be created. No further information will be obtained as this conclude u remote support to the customer for this event. Diagnostic performed by Engineer : Customer states doing set u the display is dim at highest Brightness setting, requests part id to replace to resolve issue, no patient involvement. Problem Description by Engineer : Customer Fuction/Role: biomed Device use: unit set u Expected and unexpected Resuts:display is dim user Impact: not unable needs repair Patient Impact: none SE3.200000000----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Eriberto Surez Issu: 566317300000000nable to tun off the brightness and looks very dim Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech support Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-02 02:16 | No
24 | Yes | Alarms-No Sound/ Speaker Malt & No Sound-s/n u / Unable to Confirm-Not Received for Eval-Cau Unk/Replace Sent | Alarms-No Sound/ Speaker Malt & No Sound-s/n u000 / Unable to Confirm-Not Received for Eval-Cau Unk/Replace Sent | 000000000000000000000000001O##: 00000000000000000000000000100000000951638 Massey Ryan The system meets the specification for the performed service and is retuned to use Resolution : Is Cause Known: No Cause: Is Resolution Known: Yes Resolution Summary: Malfunction Resolution Detail: Customer is taking responsibility to correct/repair the device. The customer has been notified to contact u if this does not address their device issue at which point a new service order will be created. No further information will be obtained as this conclude u remote support to the customer for this event. Parts/Supplies Information Provided Part ID and Price: 0000000056000000002551 Mu00000000- PWM 2.00000000z ITS (and SpO2) - Not used in us. A. CAD 2,619.00000000 Problem Description by Engineer : Customer Fuction/Role: Biomed Device use: In Clinical use Expected and unexpected Results: Normal monitoring user Impact: unable to monitor Patient Impact: unable to monitor SE uu Problem Description/unexpected Results: Speaker malfunction message Diagnostic performed by Engineer : Problem Description: Speaker malfunction message Giving speaker malfunction message Form used: Parts Troubleshooting Steps Taken: Provided part number and price ----------------------------------------------------------------------------- | 2023-09-02 02:41 | No
25 | Yes | Will not go pass countdown menu after power up | Will not go pass countdown menu after power up | 000000000000000000000000001O##: 000000000000000000000000001000000009000000009000000001-Sep-2000000003 Bonus Kenneth The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: user Device use: unknown Expected and unexpected Results: start u user Impact: unknown Patient Impact: unknown SE 8.1.17. Resolution : Replace battery in PC ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gene IP & Issu: 58330000000022- Will not go pass countdown menu after power u Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and advised customer of callback. Called on call for IuR Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-02 03:06 | No
26 | Yes | System would not boot up | System would not boot up | 000000000000000000000000001O##: 00000000000000000000000000100000000951951 Clark, Byron The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer, no follow-u requested Problem Description by Engineer : Problem description by engineer : detector uLLno unable image tuned off unit and tuned off breaker in M-rack .. when resetting breaker, breaker in transfer switch is now tripping on us When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : Continuously Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: No patient involved Expected and unexpected Results: detector uLLno unable image tuned off unit and tuned off breaker in M-rack .. when resetting breaker, breaker in transfer switch is now tripping on us user Impact: room down Patient Impact: room down SE 7.2.16.00000000Resoluion : advised of part for control for MPDu OK to close 000000000000000000000000001Internal Comments : full# and/or description of full: no power Repair Action : MPD CONTROL unit 00000000230000000016000000007 Diagnostic performed by Engineer : issues when troubleshooting, cannot get unit to boot, upon resetting breakers in m-rack the breaker in transfer switch for us trips, when that is reset the mains trips again in m-rack.. looking for steps and part for controller ----------------------------------------------------------------------------- | 2023-09-02 04:36 | No
27 | Yes | The Fleuvision monitor is out, and patient is on the table. | The Fleusion monitor is out, and patient is on the table. | 000000000000000000000000001O##: 00000000000000000000000000100000000951621 00000000-Sep-2000000003 Haynes Mike The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: Physician Device use: unknown. Screen blanked our Expected and unexpected Results: Expected system to have a Fleuvision viewing / unexpected blank screen user Impact: Could not use system util resolved Patient Impact: none SE 8.1.3000000005 Resolution : This service case was closed by the customer over the weekend. Fleuvision PC and Hard drive replaced. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000951955 Kuis Peter The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Customer Fuction/Role: biomed Device use: not at time of call Expected and unexpected Results: an user Impact: an Patient Impact: an SE Diagnostic performed by Engineer : Niki states that patient is off the table and 000000000000000000000000001ld like to further test and troubleshoot Fleuvision. RSW is also confirming that the Fleuvision hard drive may be what is causing the issue. Niki says she has a HDD available. Confirmed that it is the correct part number and she will continue with locating the replacement procedure on In Center. No other action is required. An onsite 000000000000000000000000001has been created earlier. ----------------------------------------------------------------------------- | 2023-09-02 04:36 | No
28 | Yes | u interface is broken (Problem solved by Plug and unplug the u interface board cable) | u- interface is broken (Problem solved by Plug and unplug the interface board cable) | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Internal Remarks: 18.00000000.2000000003 00000000:30000000039 cut (86000000001000000005) $ASP engineer: [qiangyi wuianku] $Q1: (N) $Q2: (N) $Q7: (Confirmed reported problem & interface is broken) $Q8: (Restart the device, check the device, It was found that the function could not be used. After plugging and unplugging the cable of the interface board, the fuction was used normally, so it was jugged that it was and the cable of the interface board was faulty） $Q9: (Plug and uplu the interface board cable) $Q100000000 (Test and inspection data refer to :PA checklist. Device retuned to full functionally.) $Q11:Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neat Calibration Date (00000000/11/20000000000000000SAP Transfer: Case Transferred to SAP External Remarks: 18.00000000.2000000003 00000000:30000000036 cut (86000000001000000005)*Department information: [Department: Equipment Department Contact person: Shen Gong Tel: 18200000000258718]* $ASP engineer: [qiangyi wuianku] $Q1: (N) $Q2: (N) $Q7: (Confirmed reported problem & interface is broken) $Q8: (Restart the device, check the device, It was found that the function could not be used. After plugging and unplugging the cable of the interface board, the fuction was used normally, so it was jugged that it was and the cable of the interface board was faulty） $Q9: (Plug and uplu the interface board cable) $Q100000000 (Test and inspection data refer to :PA checklist. Device retuned to full functionally.) $Q11:Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neat Calibration date (00000000/11/20000000000000000 No open Field Change Order found for this service. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Problem Reported by Customer: ECR:00000000R/Heart Start uuu Safety Questions: Case: 000000002221700000000 Safety Question State: Submitted PSA: false PSE: false PSF State: NA Created By: Zhujian Lin Last updated by: Zhujian Lin 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A Neat Step Action: PH Chen Shengju 1581000000005000000001 called The customer who has retuned to visit has a problem with the interface. Neat Step Action: FSE Chen Shengju requested to send orders to Quang Yi T2 Activities: Problem Description: Customer's explanation of the problem: The interface is broken The customer observed the iMessage displayed on the monitor: an Troubleshooting Action: The customer observed the iMessage displayed on the monitor: an T2 Remote Diagnostics and Testing: an Repair Action: T2 Recommended on-site repair steps: It is recommended to replace the interface on site Recommended spare part uufor T2 with description: an Notes on customer requirements: an Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 2023-09-02 08:06 | No
29 | Yes | Shock is not delivering properly - | Shock is not delivering properly - | 000000000000000000000000001O##: 00000000000000000000000000100000000952977 00000000-Sep-2000000003 Muesh, Kuar The system meets the specification for the performed service and is retuned to use Tools & Equipment : ELECTRICAL SAFETY TESTER,2000000003-00000000-00000000 00000000:00000000:00000000 Problem Description by Engineer : Customer Fuction/Role: Biomed Device use: No Expected and unexpected Results: N/A user Impact: No Patient Impact: No SE N/A Resolution : unit is already checked with Defib Analyzer on dated 00000000.00000000.2000000003 and found 000000000000000000000000001king ok. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000952867 Gnanasekar, D The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Resolution : FSE Visit required. Follow u Required : FSE Visit required. Problem Description by Engineer : Customer Fuction/Role: Biomed Device use: No Expected and unexpected Results: N/A user Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Shock is not delivering properly ----------------------------------------------------------------------------- | 2023-09-02 11:31 | No
30 | Yes | "System" is being up and it won't start | "System Starting 0" is being up and it won't start | 000000000000000000000000001O##: 00000000000000000000000000100000000953000000002 00000000-Sep-2000000003 Ono, Shinji The system meets the specification for the performed service and is retuned to use Resolution: [Parts used] Part number: 0000000080000000012200000000/Part name: u Haswell FleuVision 2 PC no HDD/units: 1 [Parts used] Part number: 0000000080000000050000000000000000Part name: Hard disk spare part/Number of parts: 1 Regulatory Questions: N/A Problem Description by Engineer: "System Starting 00000000" is being used and it 000000000000000000000000001't start Customer Fuction/Role: Radiology Department Device use: Not used Expected and unexpected Results: "System Starting 00000000" is displayed and the system does not start. user Impact: No effect Patient Impact: No effect SE 8.1.30000000015 Resolution: A diagnostic test confirmed that the FleuVision control PC was faulty. I replaced the PC and installed the software. After the 000000000000000000000000001k, I confirmed that there was no problem with system operation. Follow u Required: N/A Follow u Required: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000956220000000000000000ep-2000000003 Ono, Shinji The system meets the specification for the performed service and is retuned to use Regulatory Questions: N/A Problem Description by Engineer: "System Starting 00000000" is being used and it 000000000000000000000000001't start Customer Fuction/Role: Radiology Department Device use: Not used Expected and unexpected Results: "System Starting 00000000" is displayed and the system does not start. user Impact: No effect Patient Impact: No effect SE 8.1.30000000015 Resolution: 000000000000000000000000001for parts processing I reported the details of the 000000000000000000000000001k on September 2, 2000000003 and he agreed. Parts Consumed: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000095700000000 00000000-Sep-2000000003 Ono, Shinji The system meets the specification for the performed service and is retuned to use Regulatory Questions: N/A Problem Description by Engineer: "System Starting 00000000" is being used and it 000000000000000000000000001't start Customer Fuction/Role: Radiology Department Device use: Not used Expected and unexpected Results: "System Starting 00000000" is displayed and the system does not start. user Impact: No effect Patient Impact: No effect SE 8.1.30000000015 Resolution: 000000000000000000000000001for parts processing I reported the details of the 000000000000000000000000001k on September 2, 2000000003 and he agreed. Parts Consumed: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000952992 Noda, Kotaro The system does not meet the specification for the performed service and is not retuned to use Internal Comments: [Subject] 00000000DS:KN:u Ota: "System starting 00000000" is being used and the system does not start. [Parts from the equipment] 0000000080000000012200000000:u Haswell Fleuvision 2 PC no HDD QTY:1 ZREP Tokyo (u11) 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Shinji Ono] Ono, Shinji - FSE1 Visit date and time 2000000003/00000000/0000000018:15 - FSE1 000000000000000000000000001k time 2 [Shipping instructions] ----------------- shipping instructions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Shinji Ono - 000000000000000000900000000-000000002 - Others Call mobile upon arrival Requester name Kotaro Noda [Examination category] Normal examination Follow u Required: * Problem Description by Engineer: "System Starting 00000000" is being used and it 000000000000000000000000001't start Customer Fuction/Role: Radiology Department Device use: Not used Expected and unexpected Results: "System Starting 00000000" is displayed and the system does not start. user Impact: No effect Patient Impact: No effect SE 8.1.30000000015 Diagnostic performed by Engineer: At startup, "System starting 00000000" and it 000000000000000000000000001't start. It also takes time to tun the power off. I had him tun the circuit breaker off/on, but that does not improve the situation. "Switching not possible" occurred in the TSM.FSE response. ----------------------------------------------------------------------------- | 2023-09-02 12:01 | No
31 | Yes | u - Meterage test: low energy delivery | u Meterage test: low energy delivery | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Questions: Case: 000000002221800000000 Safety Question State: Submitted PSA: false PSE: false PSF State: NA Created By: Tian Ju Last updated by: Tian Ju 1. Was the device in clinical use at the time the issue was discovered? no 2. Was any patient or user harmed? no 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A Problem Reported by Customer: Measurement problem ----------------------------------------------------------------------------- | 2023-09-02 12:41 | No
32 | Yes | Alarm: upfault | Alarm: upfault | BuQ&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Incompatible use Repair Activity Code: No Faut Food: unconfirmed Problem / Failure Problem Reported by Customer: failure SAP Transfer: Case Transferred to SAP Neat Step Action: The equipment is currently in shanghaiing company T2 Activities: Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST Occurrence? WITHIN. . . Same Day HOW OFTEN IS Issu Occurring? First Occurrence WHERE IN THE END-user 000000000000000000000000001KFLOW IS THE Issu Occurring? Pre-Acquisition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE Occurred? First Power on day HAS Customer CREATED A SPECIFIC SYSTEM EVENT REPORT OF Issu? None Troubleshooting Action: ---------TA-----Troubleshooting ACTION--------------- NOTE full# AND / OR DESCRIPTION OF full Alarm: device faut WHAT STEPS WERE TAKEN During THE DIAGNOSIS? Instruct dealer to run the operation check without leadwork connection and all items are passed. DFM100000000 retunes to full functionality Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: NA Test/Inspection: Power on test and visual test are passed, device return to full functionality Test Equipment: NA Safety Questions: Case: 0000000022218621 Safety Question State: Submitted PSA: false PSE: false PSF State: NA Created By: Tian Ju Last updated by: Tian Ju 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? Yes 00000000Was this an our of bou failure? No ----------------------------------------------------------------------------- | 2023-09-02 13:06 | No
33 | Yes | It does not start. | It does not start. | 000000000000000000000000001O##: 000000000000000000000000001000000009000000003 00000000Sep-2000000003 José Manul Mora Pastor The system meets the specification for the performed service and is retuned to use. Regulatory Questions: N/A Resolution: After performing several diagnostic tests, we proceed to replace the Hard Disk of the Host PC (Technical Room) + reinstallation of ghost image of the system. After restarting the equipment, several functional tests are performed. Right. It is determined that the equipment is repaired and 000000000000000000000000001king normally according to the manufacture's specifications, restoring the operating conditions prior to the incident. Problem Description by Engineer: Customer Fuction/Role: Nursing Device use: DIAGNOSTIC BY IMAGE Expected and unexpected Results: DIAGNOSTICS BY IMAGE user Impact: NONE Patient Impact: NONE SE 000----------------------------------------------------------------------------- | 2023-09-02 13:36 | No
34 | Yes | Self test loop (Device is replaced) | Self test loop (Device is replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000900000000900000000Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to use Resolution: "I replaced the AED and confirmed that it passed the battery fit test. [Simultaneous FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is required." October 000000002000000003 Fuua Denshi Kanagawa Hanbai I sent the equipment along with a report to the on-site person in charge and contacted him. Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Result: Pass Regulatory Questions: N/A Problem Description by Engineer: Customer Fuction/Role: Person in charge of equipment management in the facility. Device use: In standby mode Expected and unexpected Results: The AED failed a self test. user Impact: Cannot use the AED in an emergency. Patient Impact: No patient used. SE Equipment without version notation Diagnostic performed by Engineer: I performed a battery test on the AED and confirmed that the button test was skipped, the "Cannot use" message was made, and the three-beep alarm sounded repeatedly. "I verified the operation record of the AED and found that the internal high voltage element was faulty, so it is necessary to replace the AED. I apologize for the inconvenience this has caused." ----------------------------------------------------------------------------- Internal Remarks: Failure content 2 When I checked on-site and pressed the i-buton, the message "Cannot use. To test, remove and re-install the battery. "Message was made. When the battery was removed and re-installed⇒ "Self-test⇒ testing in progress.....testing⇒ Cannot be used. To test, remove and re-install the battery. Message was made.⇒ substitute. Lot No. Adult pad Y000000000000000021-00000000 Adult pad Y000000000000000021-00000000 Battery 000000000000000000000000----------------------------------------------------------------------------- | 2023-09-02 13:51 | No
35 | Yes | u- External Paddles are not recognized correctly | u- External Paddles are not recognized correuy | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000000000000Hübler Christina The system meets the specification for the performed service and is retuned to use Resolution: The replacement of the external paddles will be performed by the dealers. Can they send the order to SMS Medipool? Location Name SMS medipool AG Street Robert-Koch-Str. 2a City Friedrichsthal Zip 66299 Country Germany Problem Description by Engineer: Customer Fuction/Role: Biomed Device use: no Expected and unexpected Results: FSN question user Impact: no Patient Impact: no SE ----------------------------------------------------------------------------- Customer comment: The uLLaffects all our defis in the building. Bu this 000also has the following 000: CN7390000000089, CN7390000000088, CN7390000000087, CN7390000000017 please send new free paddles. ----------------------------------------------------------------------------- | 2023-09-02 14:36 | No
36 | Yes | Device Fails Self-Test - log analysis requested | Device Fails Self-Test - log analysis requested | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000039 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to use Resolution: "I replaced the AED and confirmed that the AED passed a battery fitted test. [Simultaneous FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is required." "September 13, 2000000003 Fuua Denshi Seibuinami Hanbai Report to the on-site Person in Charge and I also sent the equpment.September 22, 2000000003 Fuua Denshi Seibuinami Hanbai I sent the equpment along with a report to the on-site person in charge." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Result: Pass Parts Consumed: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Each Problem Description by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device use: In standby mode Expected and unexpected Results: The AED failed a self test. user Impact: Cannot use the AED in an emergency. Patient Impact: No patient used. SE Equpment without version notation Diagnostic performed by Engineer: Received. I used a test pad and battery on the AED and performed a battery test, and I confirmed that the AED passed the test.After that, I monitored the AED in standby mode, but the alarm did not re-sound. I analyzed the AED operating record and found that the alarm that occurred at the time was that the high voltage calibration item in the self test was outside the passed range, and that is why the AED beeped once as an full The alarm has now been resolved, and the above uLLitems are within the passed range, so I will replace the AED in order to ensue safety. I apologize for the inconvenience this has caused. Regulatory Questions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failure content 2 The AED did not pass a self test, and the message "To run a test, remove the battery..." message was generated. The same thing happens when he replaces the pad or the battery. He asked u to inspect it. Lot No. Not listed. ----------------------------------------------------------------------------- | 2023-09-02 15:01 | No
37 | Yes | u connector broken | DC concur broken | 000000000000000000000000001O##: 000000000000000000000000001000000000000000063800000000Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Fuction/Role: Medical Devices Coordinator & Senior Technician Device use: not in use Expected and unexpected Results: During routine inspection it was noted that there was a pin from the DC cable stuck in the socket meaning that the monitor 000000000000000000000000001ld not charge. user Impact: n/a Patient Impact: n/a SE n/a Resolution : unit assembled and tested OK. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000961661 Robert Jon de Ste Croiu The system does not meet the specification for the performed service and is not retuned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Fuction/Role: Medical Devices Coordinator & Senior Technician Device use: not in use Expected and unexpected Results: During routine inspection it was noted that there was a pin from the DC cable stuck in the socket meaning that the monitor 000000000000000000000000001ld not charge. user Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : unit disassembled as requested by RDT, awaiting repair by Bradley. Follow u Required : unit disassembled as requested by RDT, awaiting repair by Bradley. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000090000000081 Marc Chambers The system does not meet the specification for the performed service and is not retuned to use Follow u Required : Device to be collected and sent to u Bench for Investigation & Repair 1u Front Enclosure Kit with all connectors 0000000056000000005000000001 Regulatory Questions : N/A Problem Description by Engineer : Customer Fuction/Role: Medical Devices Coordinator & Senior Technician Device use: not in use Expected and unexpected Results: During routine inspection it was noted that there was a pin from the DC cable stuck in the socket meaning that the monitor 000000000000000000000000001ld not charge. user Impact: n/a Patient Impact: n/a SE n/a Resolution : Device to be collected and sent to u Bench for Investigation & Repair 1u Front Enclosure Kit with all connectors 0000000056000000005000000001 ----------------------------------------------------------------------------- Internal Remarks : Tracking Nuber:1Z3R6000000006693221863 COLLECT TOMORROW ----------------------------------------------------------------------------- | 2023-09-02 15:31 | No
38 | Yes | Gas analyzer module failure giving no readings | Gas analyzer module failure giving no readings | 000000000000000000000000001O##: 0000000000000000000000000010000000090000000000000000 Parts Consumed : 1.0000000000000000000000000000000056000000007391,M_G7m Gas Module Teut Problem Description by Engineer : gas analyzer module failure giving no readings Resolution : Replaced G7 module ----------------------------------------------------------------------------- | 2023-09-02 17:31 | No
39 | Yes | Machine is automatically getting shutdown | Machine is automatically getting shutdown | 000000000000000000000000001O##: 00000000000000000000000000100000000953517 00000000Sep-2000000003 Chaudhary, Deepak The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : Host PC not Switching ON intermediately . After switching ON manually System is 000000000000000000000000001king fine Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: No further information is available as per the customer SE 1.000000002.1 Tools & Equpment : MuTIMETER,200000000000000006-27 00000000:00000000:00000000 Resolution : Replaced the CMOS battery with new, reseated the hard disk and other connections of Host PC after that problem has been resolved. Handed over the machine in 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000952600000000Kuar, Atu The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : Host PC not Switching ON intermediately . After switching ON manually System is 000000000000000000000000001king fine Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: No further information is available as per the customer SE 1.000000002.1 Follow u Required : full# and/or description of full: No full Repair Action : After switching ON manually System is 000000000000000000000000001king fine Technician Name : Kuar, Atu Internal Comments : full# and/or description of full: No full Repair Action : After switching ON manually System is 000000000000000000000000001king fine Diagnostic performed by Engineer : After switching ON manually System is 000000000000000000000000001king fine ----------------------------------------------------------------------------- | 2023-09-02 18:21 | No
40 | Yes | The equipment displays an "u" on the display indicating a Statues u. | u The equipment displays an "u" on the display indicating a Statues u. | BuQ&R Feedback: No Reason: Corrective Maintenance. Problem Reported by Customer: Equpment shows "u" on the Display indicative of Status of use. When switching on it displays the message "deactivated therapy" and "equipment defect". When trying to run the operational test, the equpment shows a failure in the shock discharge, apparently it does not charge" and failure in the u even with the cable connected. The paddles and the usable were tested in another DFM100000000 and they 000000000000000000000000001k correctly. The equpment was tested with the paddles and usable from another device (good accessories) and it continued with the failure. Safety Questions: Case: 000000002222100000000 Safety Question Status: Submitted PSA: false PSE: false PSF Status: Not at all. Created By: uyni Edwirges. Last updated by: uyni Edwirges. 1. Was the equpment in clinical use when the problem was detected? No 2. Was the patient or user injured? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? unknown 00000000Did the equpment show a defect or failure when it was installed? unknown ----------------------------------------------------------------------------- | 2023-09-02 18:46 | No
41 | Yes | Temperature is rising for detector. | Temperature is rising for deteur. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000388000000001-Sep-2000000003 Colon, Efrain The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: Device use: Expected and unexpected Results: user Impact: Patient Impact: SE Diagnostic performed by Engineer : retuned used part for credit, will create follow u case for pending software and geometry issues ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000982999 100000000ep-2000000003 Colon, Efrain The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: tech Device use: unknown Expected and unexpected Results: detector 000000000000000000000000001king, doesn't user Impact: low Patient Impact: low SE 7.2.9 Diagnostic performed by Engineer : Checked unit our, looks like there are some software bus. 000000000000000000000000001king with NSS for upgrade quote ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096700000000 00000000-Sep-2000000003 Colon, Efrain The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Follow u Required : replace rotor control board Problem Description by Engineer : Customer Fuction/Role: tech Device use: unknown Expected and unexpected Results: detector 000000000000000000000000001king, doesn't user Impact: low Patient Impact: low SE 7.2.9 Parts Consumed : 1.00000000000000000000000000000000200000000000000001700000000FIRE u BOARD Diagnostic performed by Engineer : rotor controller needs to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000951616 00000000-Sep-2000000003 Colon, Efrain The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Follow u Required : replace FDC board Problem Description by Engineer : Customer Fuction/Role: tech Device use: unknown Expected and unexpected Results: detector 000000000000000000000000001king, doesn't user Impact: low Patient Impact: low SE 7.2.9 Diagnostic performed by Engineer : checked bulldogs FDC board needs replacing ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Customer Call Back Caller: Sai IP & Issu: ccb wants to be contacted with eta for onsite Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? No, not entitled Preferred Service Time: Regular Entitlement Options: Quote Reason for Entitlement Claim: N/A Action Taken: created aa Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Ephriam - FSE inquiring about uufor 7.2.16 SW. 00000000800000000550000000052 - u uuuFam. SW PKT REL. 7.2.16 Internal Remarks : Efrain Colon - detector cooler problem. FDuD chiller replaced. Warning "Chiller communication full. Suggested replacing FDC board. Internal Remarks : Interaction Type: Customer Call Back Caller: Nazar Fadhil (600000000 659-00000000 IP & Issu: 536178 Need ETA FSE isn't onsite Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: Entitlement Options: Quote Reason for Entitlement Claim: Action Taken: ##d service manager to escalate the cu req to contact the FSE and CCB with updates to the ETA Neat Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: Internal Remarks : FSE is replacing the FDC board and the bulldogs don't seem to be happy with the correction dataset restore. While on the phone, geometry stopped 000000000000000000000000001king intermittently and there is also a message about true anode defect. -FSE will get new logging for the geometry issue. Suggested that the RoCo could be the problem since we were able to make ray once by running a diagnostic, but after a reboot it defaulted back to the anode issue. -Approved: RoCo Velara: 000000002100000000125700000000nternal Remarks : Interaction Type: Customer Call Back Caller: Nazar 600000000659-00000000 IP & Issu: case 00000000222228300000000SE has not been onsite Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: Entitlement Options: Quote Reason for Entitlement Claim: Action Taken: created AA 0000000000000000689800000000FSE was scheduled to be onsite today at 8am est but has not been onsite please reach our to Nazar at 600000000659-00000000 to advise of FSE ETA asap.. thanks // created AA 0000000000000000689897 Expedite service was requested on 9/1/23 but FSE was scheduled for 9/5/23 please remove expedite fee from cu's quote.. thanks Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Fuse didn't follow replacement procedure for FDC. Had FSE use procedure and call back if faut still present. Possible DEFOA. Internal Remarks : Interaction Type: New Case Caller: Nazar Fadhil IP & Issu: 536178/Temperature is rising for detector. Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Quote Reason for Entitlement Claim: N/A Action Taken: Created case, created quote and advised of callback for scheduling Neat Step Action: Awaiting Quote Verbal PO Number: N/A Verbal PO Approver: Nazar Fadhil Verbal PO CAP: N/A Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: Customer Callback Caller: Nazar Fadhil IP & Issu: 536178/Temperature is rising for detector. Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Quote Reason for Entitlement Claim: N/A Action Taken: Verified Onsite Date & Time Neat Step Action: Awaiting Quote Verbal PO Number: N/A Verbal PO Approver: Nazar Fadhil Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-02 21:51 | No
42 | Yes | unit not charging | Tempuunit not charging | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003217 Robert Jon de Ste Croiu The system does not meet the specification for the performed service and is not retuned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Fuction/Role: Clinical Support Assistant Device use: not in use Expected and unexpected Results: During routine checks it was noted that the Monitor was not charging - multiple chargers were tried but issue remained. user Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : Parts retuned ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009500000000100000000Marc Chambers The system does not meet the specification for the performed service and is not retuned to use Resolution : Device to be collected and sent to u Bench for Investigation & Repair 1u Single Board Computer 000000005600000000100000000 Follow u Required : Device to be collected and sent to u Bench for Investigation & Repair 1u Single Board Computer 000000005600000000100000000 Regulatory Questions : N/A Problem Description by Engineer : Customer Fuction/Role: Clinical Support Assistant Device use: not in use Expected and unexpected Results: During routine checks it was noted that the Monitor was not charging - multiple chargers were tried but issue remained. user Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : End user has been asked to advise if they is any sign of damage to the DC socket or PINS being stuck in the socket from a charger. ----------------------------------------------------------------------------- Internal Remarks : End user has been asked to provide the 000of the affected LS so that a case can be logged. Internal Remarks : Tempo LS issue logged on case ref 00000000222923600000000nternal Remarks : Hi Tracey, I think the 000on the system is wrong. As far as I can see we don’t have 6100000000 I will check our lists and come back to you Thanks, Alice Internal Remarks : Hi Sam , I am guessing the wrong device SN 615900000000has been shipped instead of the below that we have a case for SN 6100000000 Case 0000000022222922 W00000000000000003217 I will ##the customer Internal Remarks : Hi Alice , Hope you are well I have been contacted by the bench the wrong device SN 615900000000has been shipped instead of the below that we have a case for the numbers are very similar SN 6100000000 Case 0000000022222922 W0000000000000000000000003217 Please can you check and I will arrange a collection for SN 6100000000 and then get the bench to return SN 615900000000uless that has a faut to Internal Remarks : The end user has reported an issue with an LS but is yet to provide any details of the LS. Once we have this then a separate case can be logged. Internal Remarks : The 00061000000003 should not be registered to EHAA, this device was sold to another end user. Please see case ref 000000002200000000000000002 as this will be the case that any required 000000000000000000000000001ks will be actioned against. Internal Remarks : Hi Sam , Still awaiting customer as they were checking the 000 Will follow u again today Internal Remarks : Tracking Nuber:1Z3R6000000006699000000002500000000OLLECT TOMORROW Esau AND HERTS AIR Ambulance ALICE GARDNER EHAAT Hangar 12 North Weald Air Field North Weald CM166HR 0000000000595270000000000000000Internal Remarks : Collection address: EHAAT Hangar 12 North Weald Air Field North Weald CM16 6HR ----------------------------------------------------------------------------- | 2023-09-02 22:21 | No
43 | Yes | Sky Plate dropped ( calibration unsuccessful replacement needed) | Sky Plate dropped ( calibration unsuccessful replacement needed) | 000000000000000000000000001O##: 0000000000000000000000000010000000097165000000008-Sep-2000000003 Jones, Jutin The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: an Device use: an Expected and unexpected Results: an user Impact: an Patient Impact: an SEna Resolution : Sky plate has bad artifact/ distorted image. One of the corners was damaged and shock log was showing 3 heavy shocks. One of the days was June 7th 2000000003 which the customer mentioned there was a drop that day. Failed pixel calibration as well as other detector calls. Beginning process of getting approval for replacement and quotes for customer. ----------------------------------------------------------------------------- | 2023-09-02 23:06 | No
44 | Yes | Speaker Malfunction-Exchange request/ audible sound found at Bench keyrin-s/nu/other HW /Replaced parts. | Speaker Malfunuon-Euchange request/ audible sound found at u/Replaced parts. | 000000000000000000000000001O##: 0000000000000000000000000010000000015700000000 The system meets the specification for the performed service and is retuned to use ----------------------------------------------------------------------------- | 2023-09-03 00:16 | No
45 | Yes | Bad display | Bad display | Internal Remarks : Interaction Type: New Case Caller: Tamya Goud IP & Issu: 56631000000002 Parts Id Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and routed to RSE Neat Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-03 00:31 | No
46 | Yes | Wireless foot switch is not working | Wireless foot switch is not working | 000000000000000000000000001O##: 00000000000000000000000000100000000951557 26-Sep-2000000003 Dun Peter The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: u Device use: hybrid OR Expected and unexpected Results: wireless footswitch not functional even after FCO user Impact: cannot use wireless footswitch Patient Impact: N/A SE 1.2.5 Resolution : Replaced wireless footswitch, paired to wireless base station. Performed functional test. Fully functional. After FCO7220000000000000000. Parts Consumed : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P ----------------------------------------------------------------------------- Internal Remarks : Called FSE Peter Dun they advised dispatching today Sept 5 at anytime, they are on their way to site. Internal Remarks : Replan 000000000000000000000000001 - Changed to in process, FSE advised moving the 000000000000000000000000001to Monday Sept 25 Internal Remarks : Potential FCO issue stated by customer, as they have recalls regarding footswitch. Customer states FSE will be onsite to address other issues Tuesday and 000000000000000000000000001ld like the FSE to investigate this issue at that time. ----------------------------------------------------------------------------- | 2023-09-03 02:01 | No
47 | Yes | Touchscreen malfunction during u | Touchscreen malfunuon during FCO0 | 000000000000000000000000001O##: 0000000000000000000000000010000000095168000000002-Sep-2000000003 Siku Fanfan The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: BIOMED Device use: NO Expected and unexpected Results: user Impact: unable TO CONTROL SETTING Patient Impact: SE Diagnostic performed by Engineer : PROBLEM: Touch SCREEN Failure/ NOT RESPONDING Solution: REPLACED A NEW Touch SCREEN, AND COMPLETED ALL TEST AND VERIFICATION Parts Consumed : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-03 02:36 | No
48 | Yes | system locks up and need to reboot the system | system locks up and need to reboot the system | 000000000000000000000000001O##: 0000000000000000000000000010000000096720000000012-Sep-2000000003 Muoz Michael The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: an Device use: cv Expected and unexpected Results: an user Impact: an Patient Impact: an SE 000Resoluion : Part received and replaced. Verified system operation, tuned system over to customer. Parts Consumed : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA ----------------------------------------------------------------------------- Internal Remarks : Michael Muoz - Image size of Live image on FleuVision monitor is wrong. Intermittent issue. Suggested reloading IPPC software. Internal Remarks : Interaction Type: New Case Caller: Walter (5200000000 272-300000000 IP & Issu: system locks u and need to reboot the system Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? No, not entitled Preferred Service Time: ASP Entitlement Options: Quote Reason for Entitlement Claim: an Action Taken: created case and quote Internal Remarks : Michael Muoz - Grabber card full for FV2PC on this SW R000system. Suggested replacing FV2PC. ----------------------------------------------------------------------------- | 2023-09-03 02:41 | No
49 | Yes | No charge | No charge | 000000000000000000000000001O##: 0000000000000000000000000010000000095200000000 Parts Consumed : 1.000000000000000000000000989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000----------------------------------------------------------------------------- | 2023-09-03 05:06 | No
50 | Yes | u/Footswitch problem | Footswitch problem | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Neat Step Action: PPD 00000000800000000772193*1 200000000WuJiajie-86-13192679800000000 Normal shipment SAP Transfer: Dispatched for Onsite Support SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 11.00000000.2000000003 00000000:0000000000000000 cut (86000000000000000000000000 RDI Accuracy: Yes. $Q7:(Confirm reported problem, footswitch problem. ) $Q8:(Performed visual inspection, footswitch problem.) $Q9: (replace the footswitch do the adjustment and do the fuction test. ) $Q100000000 (Test and inspection data refer to:(PA). Device retuned to full functionality) $Q11: Calibrated Tools name ( NA ) External Remarks: 11.00000000.2000000003 00000000:0000000059 cut (86000000000000000000000000[Service Content] The foot pedal caused intermittent fluoroscopy failure. After replacing spare parts, the machine retuned to normal use. [Service Result] After repairing according to the manual, the machine retuned to normal use. T2 Activities: Problem Description: Overview of customer problems: The exposure pedal is intermittently malfunctioning and needs to be replaced. Engineers are required to be present Whether it is a faut of the device itself: Yes Troubleshooting Action: Remote diagnosis and test results: The exposure pedal is intermittently malfunctioning and needs to be replaced. Engineers are required to be present PRS TR statue: No PRS Repair Action: Recommended on-site repair steps: The exposure pedal is intermittently malfunctioning and needs to be replaced. Engineers are required to be present Recommended spare parts uuand designation: " 00000000800000000772193 Footswitch CV 3p 00000000 Notes on customer requirements: Onsite Safety Questions: Case: 000000002222000000007 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jiankang Jiang Last updated by: Jiankang Jiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? Yes 00000000Was this an our of bou failure? No Problem Reported by Customer: channel manager line ruuin1392000000000000000062 Call for repair, usable to expose ----------------------------------------------------------------------------- | 2023-09-03 08:06 | No
51 | Yes | u/startup fault | 0R/u FD0startup fault | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Support Neat Step Action: Zhang uiaouin 186000000000000000095 Problem Reported by Customer: 00000000R/u uuuFD200000000SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neat Step Action: Shutdown is slow, Neat Step Action: PPD spare parts 0000000080000000050000000000000000*1 Arrange 200000000to the hospital, Zhang uiaouin-86-186000000000000000095 T2 Activities: Problem Description: Overview of customer problems: Startup failure Detailed description of customer problem or faut phenomenon: C-arm movement failure Whether it is a faut of the equpment itself: Yes Troubleshooting Action: Key faut codes or phenomena: usable to power on, the screen remains black after powering on Remote diagnostics and test results: N/A PRS TR statue: All Access Repair Action: Recommended on-site repair steps: Check the startup statue of hosp PC Check BIOS battery Manul switch PC If it is invalid, replace the host PC. Recommended spare parts uuand designation: PPD uu0000000080000000050000000000000000 Host PC Monoplane Revised no HDD200000000to Hosp Notes on customer requirements: Please contact the customer for on-site repair Safety Questions: Case: 000000002222000000006 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Min Yang Last updated by: Min Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A External Remarks: 10000000000000000.2000000003 00000000:39:200000000uC (86000000008266) Service content: [Intermittent failure to boot, monitor black screen, no BIOS display, check BIOS battery and find insufficient power, replace BIOS battery, restore Ghost, reinstall IPPC system and software] Service results: [After service, the overall equpment operates normally] Other important matters: [None] Internal Remarks: 10000000000000000.2000000003 00000000:52:53 cut (86000000008266) $Q7: (Confirmation & CAN'T POWER on sometimes and no message reported on the monitor) $Q8: (Reviewed the uLLlog, found no uLLlog ;Performed visual inspection and check disk and it's ok then check the host PC and IPPC BIOS & the BIOS battery is defect ) $Q9: (replace BIOS battery of Host PC and IPPC then reinstall IPPC OSand SW, T2 suggestion effective.) $Q100000000 (Test and inspection data refer to PA. Device retuned to full functionality.) $Q11: (No Calibrated Tools used) For Mainland China CV / our, CT engineers only. Please fill in: ^Current Tue installed date: N/A ^Current Tue usage: N/A ----------------------------------------------------------------------------- | 2023-09-03 09:06 | No
52 | Yes | u/exposure failure | 0N/u/exposure failure | BuQ&R Feedback: No Reason: Corrective Maintenance Neat Step Action: PPD spare parts 00000000800000000660000000032*1 Arrange 200000000to the hospital, Zhang Chen-86-177179577100000000uternal Remarks: 11.1000000002000000003 15:0000000059 cut (86000000006533) Service content (required): Do not expose doing use. Check the grid voltage and the input voltage is low intermittently. Service results (required): After repairs, the equpment retuned to normal operation. Other important matters: none. Internal Remarks: 11.1000000002000000003 15:55:00000000uC (86000000006533) "RDI Accuracies" $Q7: (Confirmed reported problem, The device in use cannot be exposed. ) $Q8: (Check log and found the hospital power grid input voltage is 397V, lower than 00000000V. ) $Q9: (Current voltage is normal. Keep watching. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to full functionality. ) $Q11: Calibrated Tools name (NA) S/N (NA) Neat Calibration date (NA) SAP Transfer: Case Transferred to SAP Problem Reported by Customer: Hard drive failure SAP Transfer: Dispatched for Onsite Support Neat Step Action: Zhang Chen-86-177179577100000000AP Transfer: Case Transferred to SAP Safety Questions: Case: 000000002222000000007 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Min Yang Last updated by: Min Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities: Problem Description: Overview of customer problems: Espouse failure Detailed description of customer problem or faut phenomenon: Can only see through, cannot expose and save images, intermittent Whether it is a faut of the equpment itself: Yes Troubleshooting Action: Key faut codes or phenomena: n/a Remote diagnostics and test results: N/A PRS TR statue: No PRS Repair Action: Recommended on-site repair steps: 1. Check the logs 2. Reinstall ray PC software Recommended spare parts uuand designation: PPD uu00000000800000000660000000032 SSD data Haswell 200000000to Hosp Notes on customer requirements: Please contact the customer for on-site repair ----------------------------------------------------------------------------- | 2023-09-03 11:11 | No
53 | Yes | Video issues, 0ifferent functions aren't available | Video issues, 0different funuons aren't available | 000000000000000000000000001O##: 00000000000000000000000000100000000953159 00000000-Sep-2000000003 Black Edward The system meets the specification for the performed service and is retuned to use Resolution : Reset power supply Problem Description by Engineer : Customer Fuction/Role: Hospital/Biomed Device use: Cath lab Expected and unexpected Results: Video sources not displayed user Impact: usable to review Patient Impact: None SERel1.2 Diagnostic performed by Engineer : Checked log files, cable connections and diagnostic HW tests. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000095375000000003-Sep-2000000003 Oshiro Kenneth The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: Biomed Engineer Device use: Cath lab Expected and unexpected Results: usable to display multiple video sources on u7 M200000000 user Impact: usable to display on monitors Patient Impact: usable to display on monitors SE R1.2.5 Resolution : Troubleshooting of B-Cabinet indicated a fuel in u37 Depower supply was blown resulting in a bank of DVI splitters to stop 000000000000000000000000001king. Replaced Fee with a new one, ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jason IP & Issu: 8328200000000-Video issues, 00000000ifferent functions aren't available Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and Notified on call RSE. Advised customer of callback Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Ken - usable to display some video sources - found DVI splitter bank powered off. Approve B cabinet 3NY unit part order. ----------------------------------------------------------------------------- | 2023-09-03 14:41 | No
54 | Yes | Exposure not possible error | Exposure not possible error |  | 2023-09-03 15:31 | No
55 | Yes | Some geometry movements are not available. | Some geometry movements are not available. | 000000000000000000000000001O##: 0000000000000000000000000010000000095000000009 00000000-Sep-2000000003 Jetlag, Pharurat The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : The system show message "Some geometry movements are not available." user try to restart main system it back to 000000000000000000000000001king again. This problem is intermittent. How often is issue occurring? : Multiple times per week Malfunction area : Geometry Information to support the complaint handling process Customer Fuction/Role: Not provided Device use: Not provided Expected and unexpected Results: Not provided user Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000Resoluion : Replace part ,Reinstall Clea_ceil_Carc_bou ,adjustment Geometry Currents and test Parts Consumed : 1.00000000000000000000000000000000800000000990000000052,AMC Triple Servo Drive HP-LP-LP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000095000000001 CHATCHAI SANTANA The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : The system show message "Some geometry movements are not available." user try to restart main system it back to 000000000000000000000000001king again. This problem is intermittent. How often is issue occurring? : Multiple times per week Malfunction area : Geometry Information to support the complaint handling process Customer Fuction/Role: Not provided Device use: Not provided Expected and unexpected Results: Not provided user Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000Diagnostic performed by Engineer : FSE try to check all Ethercat cable and connection : All OK. Try to swap GIB from another room : Can't solve the problem. Try to reinstall software of Sute PC : Can't solve the problem. After reinstall software of Sute PC can't download software of "Clea_Ceil_Carc_Bou" so FSE try to swap 2SPC from another room : Problem is solved. Internal Comments : full# and/or description of full: Application full EtherCat POST full[item=TopologyTest] [Exceptions Info=[detail=Cannot start set master state to SAFEOP (Result = full Slave full(00000000981100000000200000000;]] Application full Frontal stand statue information not sent by GSCRT, check POST ethercat POST Failed [Detail=!TestName:EtherCat HW;!TestReport:involved FRus: Power Supply 200000000code#: check the electronic spare-part-finder SutePC code#: check the electronic spare-part-finder All EtherCat cabling code#: check the electronic spare-part-finder The EtherCat cabling between the PC and the first EtherCat slave) code#: check the electronic spare-part-finder The first EtherCat slave connected to the PC code#: check the electronic spare-part-finder All EtherCat slaves code#: check the electronic spare-part-finder POSTItem: Ethercat topology test Done/Failed [Description]: Tests whether the expected EtherCat slaves in the net000000000000000000000000001k are detected, [Details]: Cannot start set master state to SAFEOP (Result = full Slave full(00000000981100000000200000000; POSTItem: Ethercat hardware test Done/Passed [Description]: Tests whether the EtherCat net000000000000000000000000001k master initialization succeeded, [Details]: No details available. ;!Description: POST EtherCat HW Failed;started:9/1/2000000003 9:25:17 PM ended:9/1/2000000003 9:25:17 PM POST for EtherCat HW;] (GibPresent-Rest)Data[7}: pain=00000000600000000b00000000, pdinSz78, PdOu=0000000060000000016100000000 PdOusz=600000000 (GibPresent-Rest)Data[7]: EcatAli=00000000 Provide=00000000 VenId=00000000 Stated=1000000009, AdrAuoInc=8 EtherCat Hardware Test Passed EtherCat TopologyTest Failed GIB_Bou MemoryAccessTest Skipped GIB_Bou VersionTest Skipped Table_Height_Bou MemoryAccessTest Skipped Table_Height_Bou VersionTest Skipped Table_Tilt_Bou MemoryAccessTest Skipped Table_Tilt_Bou VersionTest Skipped Table_Lat_Bou MemoryAccessTest Skipped Table_Lat_Bou VersionTest Skipped Table_Long_Bou MemoryAccessTest Skipped Table_Long_Bou VersionTest Skipped Clea_Ceil_Prop_Bou MemoryAccessTest Skipped Clea_Ceil_Prop_Bou VersionTest Skipped Clea_Ceil_Carc_Bou MemoryAccessTest Skipped Clea_Ceil_Carc_Bou VersionTest Skipped BIB_Bou MemoryAccessTest Skipped BIB_Bou VersionTest Skipped SAM_TblHeight DriveAmpStateTest Skipped SAM_TblHeight SafetyBrakeTest Skipped SAM_TblHeight PositionPotmeterTest Skipped SAM_TblHeight ForceSensorTest Skipped SAM_TblHeight MotorThermistorTest Skipped SAM_TblPivot FiuationBrakeTest Skipped SAM_TblPivot PositionPotmeterTest Skipped SAM_TblTilt DriveAmpStateTest Skipped SAM_TblTilt SafetyBrakeTest Skipped SAM_TblTilt PositionPotmeterTest Skipped SAM_TblTilt MotorThermistorTest Skipped SAM_TblLat DriveAmpStateTest Skipped SAM_TblLat FiuationBrakeTest Skipped SAM_TblLat VABrakeTest Skipped SAM_TblLat PositionPotmeterTest Skipped SAM_TblLat MotorThermistorTest Skipped SAM_TblLong DriveAmpStateTest Skipped SAM_TblLong SafetyBrakeTest Skipped SAM_TblLong FiuationBrakeTest Skipped SAM_TblLong CluchTest Skipped SAM_TblLong PositionPotmeterTest Skipped SAM_TblLong MotorThermistorTest Skipped SAM_CleaLNG DriveAmpStateTest Skipped SAM_CleaLNG CluchTest Skipped SAM_CleaLNG PositionPotmeterTest Skipped SAM_CleaLNG MotorThermistorTest Skipped SAM_CleaPROP DriveAmpStateTest Skipped SAM_CleaPROP SafetyBrakeTest Skipped SAM_CleaPROP MotorThermistorTest Skipped SAM_CleaZCeil DriveAmpStateTest Skipped SAM_CleaZCeil FiuationBrakeTest Skipped SAM_CleaZCeil MotorThermistorTest Skipped SAM_CleaIDS DriveAmpStateTest Skipped SAM_CleaIDS SafetyBrakeTest Skipped SAM_CleaIDS ForceSensorTest Skipped SAM_CleaIDS MotorThermistorTest Skipped SAM_CleaCARC DriveAmpStateTest Skipped SAM_CleaCARC SafetyBrakeTest Skipped SAM_CleaCARC MotorThermistorTest Skipped SAM_CleaIDR DriveAmpStateTest Skipped SAM_CleaIDR SafetyBrakeTest Skipped SAM_CleaIDR MotorThermistorTest Skipped 2000000003-Sep-00000000, 21:21:200000000MCI Initialize call failed, recode<1>. 2000000003-Sep-00000000, 21:21:200000000GEOCAN endnote switch is OFF Repair Action : Replace 2SPC 00000000800000000990000000052 - AMC triple servo drive HP-LP-LP Resolution : Replace 2SPC 00000000800000000990000000052 - AMC triple servo drive HP-LP-LP Follow u Required : full# and/or description of full: Application full EtherCat POST full[item=TopologyTest] [Exceptions Info=[detail=Cannot start set master state to SAFEOP (Result = full Slave full(00000000981100000000200000000;]] Application full Frontal stand statue information not sent by GSCRT, check POST ethercat POST Failed [Detail=!TestName:EtherCat HW;!TestReport:involved FRus: Power Supply 200000000code#: check the electronic spare-part-finder SutePC code#: check the electronic spare-part-finder All EtherCat cabling code#: check the electronic spare-part-finder The EtherCat cabling between the PC and the first EtherCat slave) code#: check the electronic spare-part-finder The first EtherCat slave connected to the PC code#: check the electronic spare-part-finder All EtherCat slaves code#: check the electronic spare-part-finder POSTItem: Ethercat topology test Done/Failed [Description]: Tests whether the expected EtherCat slaves in the net000000000000000000000000001k are detected, [Details]: Cannot start set master state to SAFEOP (Result = full Slave full(00000000981100000000200000000; POSTItem: Ethercat hardware test Done/Passed [Description]: Tests whether the EtherCat net000000000000000000000000001k master initialization succeeded, [Details]: No details available. ;!Description: POST EtherCat HW Failed;started:9/1/2000000003 9:25:17 PM ended:9/1/2000000003 9:25:17 PM POST for EtherCat HW;] (GibPresent-Rest)Data[7}: pain=00000000600000000b00000000, pdinSz78, PdOu=0000000060000000016100000000 PdOusz=600000000 (GibPresent-Rest)Data[7]: EcatAli=00000000 Provide=00000000 VenId=00000000 Stated=1000000009, AdrAuoInc=8 EtherCat Hardware Test Passed EtherCat TopologyTest Failed GIB_Bou MemoryAccessTest Skipped GIB_Bou VersionTest Skipped Table_Height_Bou MemoryAccessTest Skipped Table_Height_Bou VersionTest Skipped Table_Tilt_Bou MemoryAccessTest Skipped Table_Tilt_Bou VersionTest Skipped Table_Lat_Bou MemoryAccessTest Skipped Table_Lat_Bou VersionTest Skipped Table_Long_Bou MemoryAccessTest Skipped Table_Long_Bou VersionTest Skipped Clea_Ceil_Prop_Bou MemoryAccessTest Skipped Clea_Ceil_Prop_Bou VersionTest Skipped Clea_Ceil_Carc_Bou MemoryAccessTest Skipped Clea_Ceil_Carc_Bou VersionTest Skipped BIB_Bou MemoryAccessTest Skipped BIB_Bou VersionTest Skipped SAM_TblHeight DriveAmpStateTest Skipped SAM_TblHeight SafetyBrakeTest Skipped SAM_TblHeight PositionPotmeterTest Skipped SAM_TblHeight ForceSensorTest Skipped SAM_TblHeight MotorThermistorTest Skipped SAM_TblPivot FiuationBrakeTest Skipped SAM_TblPivot PositionPotmeterTest Skipped SAM_TblTilt DriveAmpStateTest Skipped SAM_TblTilt SafetyBrakeTest Skipped SAM_TblTilt PositionPotmeterTest Skipped SAM_TblTilt MotorThermistorTest Skipped SAM_TblLat DriveAmpStateTest Skipped SAM_TblLat FiuationBrakeTest Skipped SAM_TblLat VABrakeTest Skipped SAM_TblLat PositionPotmeterTest Skipped SAM_TblLat MotorThermistorTest Skipped SAM_TblLong DriveAmpStateTest Skipped SAM_TblLong SafetyBrakeTest Skipped SAM_TblLong FiuationBrakeTest Skipped SAM_TblLong CluchTest Skipped SAM_TblLong PositionPotmeterTest Skipped SAM_TblLong MotorThermistorTest Skipped SAM_CleaLNG DriveAmpStateTest Skipped SAM_CleaLNG CluchTest Skipped SAM_CleaLNG PositionPotmeterTest Skipped SAM_CleaLNG MotorThermistorTest Skipped SAM_CleaPROP DriveAmpStateTest Skipped SAM_CleaPROP SafetyBrakeTest Skipped SAM_CleaPROP MotorThermistorTest Skipped SAM_CleaZCeil DriveAmpStateTest Skipped SAM_CleaZCeil FiuationBrakeTest Skipped SAM_CleaZCeil MotorThermistorTest Skipped SAM_CleaIDS DriveAmpStateTest Skipped SAM_CleaIDS SafetyBrakeTest Skipped SAM_CleaIDS ForceSensorTest Skipped SAM_CleaIDS MotorThermistorTest Skipped SAM_CleaCARC DriveAmpStateTest Skipped SAM_CleaCARC SafetyBrakeTest Skipped SAM_CleaCARC MotorThermistorTest Skipped SAM_CleaIDR DriveAmpStateTest Skipped SAM_CleaIDR SafetyBrakeTest Skipped SAM_CleaIDR MotorThermistorTest Skipped 2000000003-Sep-00000000, 21:21:200000000MCI Initialize call failed, recode<1>. 2000000003-Sep-00000000, 21:21:200000000GEOCAN endnote switch is OFF Repair Action : Replace 2SPC 00000000800000000990000000052 - AMC triple servo drive HP-LP-LP Technician Name : CHATCHAI SANTANA ----------------------------------------------------------------------------- | 2023-09-03 20:51 | No
56 | Yes | No audible alarms/seems like speaker is out | 000no audible alarms/seems like speaker is out | 000000000000000000000000001O##: 0000000000000000000000000010000000095000000009 00000000Sep-2000000003 David Cedotal The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: Device use: Expected and unexpected Results: user Impact: Patient Impact: SE Customer Fuction/Role: Biomed Device use: In Clinical use Expected and unexpected Results: user Impact: No impact Patient Impact: No impact SE uu Problem Description/unexpected Results: biomed states no audible sound- external speaker Resolution : audio device settings were incorrect #remote ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000095000000005 Al Hammad Lamiae The system meets the specification for the performed service and is retuned to use Follow u Required : Problem Description: biomed states no audible sound- external speaker biomed states no audible sound- external speaker Form used: Troubleshooting/Investigation Object: Aicu Not Object: Software Version: uu Date and Time of Issu: 9/3/2000000003 morning user Actions: Where is the Deviation: adult Icon Troubleshooting Steps Taken: customer replaced external speaker that did not flu issue monitor does not have external speaker Was PRS used: No PRS Permission: Diagnostic performed by Engineer : Problem Description: biomed states no audible sound- external speaker biomed states no audible sound- external speaker Form used: Troubleshooting/Investigation Object: Aicu Not Object: Software Version: uu Date and Time of Issu: 9/3/2000000003 morning user Actions: Where is the Deviation: adult Icon Troubleshooting Steps Taken: customer replaced external speaker that did not flu issue monitor does not have external speaker Was PRS used: No PRS Permission: Problem Description by Engineer : Customer Fuction/Role: Device use: Expected and unexpected Results: user Impact: Patient Impact: SE Customer Fuction/Role: Biomed Device use: In Clinical use Expected and unexpected Results: user Impact: No impact Patient Impact: No impact SE uu Problem Description/unexpected Results: biomed states no audible sound- external speaker ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jimmy Davis IP & Issu: no audible alarms/seems like speaker is our 7220000000000000000 Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and routed to RSE Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-03 21:36 | No
57 | Yes | Not getting any sound to central station-P-0 affecting patient care-Insufficient information-unknown reason | Not getting any sound to central station-P-0/ affeung patient care-Insuffi0nt information-unknown reason | Internal Remarks : Interaction Type: New Case Caller: Jason 650000000000000000700000000 IP & Issu: SN 26526KKDC / P-1 / affecting patient care- not getting any sound to central station. Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Routed RTECH Neat Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-03 21:36 | No
58 | Yes | u / It fails operational check, can't u it anymore routine call. | M00A / u00/ It fails operational check, can't u it anymore routine call. | 000000000000000000000000001O##: 00000000000000000000000000100000000953720000000000000000ep-2000000003 Naraine, Kenric The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: run Device use: in use Expected and unexpected Results: device replaced with backup user Impact: no Patient Impact: no SE n/a Resolution : Device failing daily Operational Check. FSE attempted troubleshooting over the phone. Drove on site. Device IP 39200000000300000000 running to additional tests "invasive BP and Bluetooth" neither which are installed options. Device should not be running these tests. Defib replaced with Backup IP 39200000000267. Replacement does not run additional tests and passes Op checks. FSE will open new 000000000000000000000000001uder IP for reconfiguring and further investigation. Staff satisfied. Defib crash cart operational in CVH ICuPod A. ----------------------------------------------------------------------------- | 2023-09-03 21:46 | No
59 | Yes | Live view monitor goes blank | Live view monitor goes blank | 000000000000000000000000001O##: 00000000000000000000000000100000000953800000000 13-Sep-2000000003 Buler Thomas The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: Tech Device use: Clinical Expected and unexpected Results: Monitor failed on boom user Impact: None Patient Impact: None SE 7.2 Resolution : Monitor replaced by bio-med on site with another u monitor. Monitor on 000000000000000000000000001will be retuned. ----------------------------------------------------------------------------- | 2023-09-04 02:21 | No
60 | Yes | QAAu/startup failure | QAA0L/startup failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Support External Remarks: 11.00000000.2000000003 00000000:30000000000000000uC (8600000000600000000) Faut description: The machine cannot move, the screen in the control room is black and the screen does not start normally. Service content: Software log prompts software failure, reinstall the system and software Service results: After repair, the equpment retuned to normal operation. Internal Remarks: 11.00000000.2000000003 00000000:30000000000000000uC (8600000000600000000) RDI Accuracies $Q7: (Confirm reported problem &The machine does not move, the screen in the control room is black and does not start normally) $Q8: (Software logs indicate that the software is faulty. Reinstall the system and software) $Q9: (Reinstall the system and software after maintaining the system normal) $Q100000000 (Test and inspection data refer to PA, Device return to full functionality) $Q11: Calibrated Tools name:N/A For Mainland China CV / our, CT engineer SAP Transfer: Case Transferred to SAP Neat Step Action: LiuMin 18627163378 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Customer: QAA00000000L/u7 M200000000 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Questions: Case: 0000000022225859 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Min Yang Last updated by: Min Yang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A T2 Activities: Problem Description: Overview of customer problems: Startup failure Detailed description of customer problem or faut phenomenon: The machine cannot move, the screen in the control room is black and the screen does not start normally. Whether it is a faut of the equpment itself: Yes Troubleshooting Action: Key faut codes or phenomena: The customer requires the engineer to arrive today Remote diagnostics and test results: N/A PRS TR statue: No PRS Repair Action: Recommended on-site repair steps: Check the site PC, plug and uplu the board, and try to reinstall the software if necessary Notes on customer requirements: Please contact the customer for on-site repair ----------------------------------------------------------------------------- | 2023-09-04 07:16 | No
61 | Yes | u-ray is lagging when trying to fluro | u-ray is lagging when trying to fluro | 000000000000000000000000001O##: 00000000000000000000000000100000000955000000007 00000000-Sep-2000000003 RouseauDavid The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: tech Device use: catch Expected and unexpected Results: IPPC user Impact: n Patient Impact: n SE 8.2 Parts Consumed : 1.0000000000000000000000000000000080000000050000000092,IP PC Revised no HDD Resolution : ordered and install IPPC downloaded software and tested ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000095000000006 00000000-Sep-2000000003 RouseauDavid The system meets the specification for the performed service and is retuned to use Follow u Required : install and configure part Resolution : trouble shot and found IPPC failing Problem Description by Engineer : Customer Fuction/Role: tech Device use: catch Expected and unexpected Results: IPPC user Impact: n Patient Impact: n SE 8.2 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Customer Call Back Caller: Robert Folger IP & Issu: 63850000000000000000-ray is lagging when trying to uuu Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: 9/3/2000000003 Entitlement Options: Quote Reason for Entitlement Claim: no entitled Action Taken: cu called back stating that he was initially given the phone number to pay back credit card but that dept doesn’t 000000000000000000000000001k on the wknd. Advised him to provide a PO but he said he could only do that if the quote was ##d to him. Pinged QA on day to ##cu the quote so that he could get a PO so that he could get onsite support for today Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: Robert Folger Verbal PO CAP: $8,600000000 Internal Remarks : Interaction Type: Customer Call Back Caller: Erin IP & Issu: 63850000000000000000 u-ray is lagging when trying to uuu Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Remote Bypass Preferred Service Time: Regular Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created quote Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : No answer, requesting quote and left VM to reply to quote team ASAP if onsite required over weekend/holiday Internal Remarks : Interaction Type: New Case Caller: Erin IP & Issu: 63850000000000000000 u-ray is lagging when trying to uuu Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Remote Bypass Preferred Service Time: Regular Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case r000000000000000000000000001and routed to rise Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Customer Call Back Caller: Robert Folger IP & Issu: customer callback --- adv waiting for the quote to be sent which can not be done util credit approval is completed case should be updated by Tuesday Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No, not entitled Preferred Service Time: N/A Entitlement Options: Quote Reason for Entitlement Claim: N/A Action Taken: spoke to lead ship who adv credit check needs to be done 1st ----------------------------------------------------------------------------- | 2023-09-04 07:41 | No
62 | Yes | Battery not charging fully, continues flashing - Work performed under u | Battery not charging fully, continues flashing - Work performed under SO:0000 | Test Equpment Info : Null Device Alarm/Alert as intended? : Yes Problem Observed : Not Answered Ou Of Bou Failure? : No Labor: Labor ID : L-1970000000036 000000000000000000000000001k Performed : Found the Battery indicator light continues flashing even after being plugged in. Tried different batteries and same issue. Needs to be sent in to u for repair. Diagnosis : Not Answered | 2023-09-04 10:22 | No
63 | Yes | Device shuts off when port is blocked, u/Repair *** | Device shuts off when port is blocked, 0** | Test Equpment Info : Null Device Alarm/Alert as intended? : N/A Problem Observed : Not Answered Ou Of Bou Failure? : No Labor: Labor ID : L-1970000000037 000000000000000000000000001k Performed : unit shuts off when circuit is attached. Needs to be sent in to u for Depot Repair. Diagnosis : Not Answered | 2023-09-04 10:28 | No
64 | Yes | u/Geometry movements partly available | 0R/u FD0Geometry movements partly available | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Questions: Case: 00000000222260000000000000000 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jiankang Jiang Last updated by: Jiankang Jiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A SAP Transfer: Case Transferred to SAP External Remarks: 12.1000000002000000003 21:15:52 cut (86000000006533) Service content (required): Mechanical movement is unavailable, the control power fuel is damaged, replace the fuel. Service results (required): After repair, the equpment retuned to normal operation. Other important matters: none. T2 Activities: Problem Description: Overview of customer problems: After the equpment is tuned on, all mechanical movements are powered on and locked. Only the bed surface and L arm can be unlocked manually, and other electromechanical movements cannot be ulocked.Engineers are required to be present Whether it is a faut of the equpment itself: Yes Troubleshooting Action: Key faut codes or phenomena: Enable Movement (ENMV) is active doing system startup. Remote diagnostics and test results: Mechanical motion control module button check failure PRS TR statue: No PRS Repair Action: Recommended on-site repair steps: Root Cause 1: Enable Movement (ENMV) request signal is active doing system startup. Since this is not allowed doing startup phase, geometry will not become available for the user. Service Action 1: - Check whether an object is perhaps activating one or more buttons/joysticks on one of the GEO modules . - Switch off the system. Remove suspected geometry module and connect spare Emergency stop plug instead of the removed module (or use D connector board with pins 100000000and 22 shorted). - Switch the system 'ON' and observe the POST results (IO Geometry safety line check (Passed?) . If it is passed in this bypassed condition, Check for damaged, broken or loose cabling And connectors for the Geo module. Replace module if necessary Recommended spare parts uuand designation: 00000000800000000151715 GEO Module WP Kit Notes on customer requirements: Onsite Internal Remarks: 12.1000000002000000003 21:15:55 cut (86000000006533) "RDI Accuracies" $Q7: (Confirmed reported problem, geometry movement not available. ) $Q8: (Check and test the fuel of NC unit is defective. ) $Q9: (Replace the device spare fuel. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to full functionality. ) $Q11: Calibrated Tools name (NA) S/N (NA) Neat Calibration date NA) SAP Transfer: Case Transferred to SAP Problem Reported by Customer: The machine cannot move SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 2023-09-04 12:01 | No
65 | Yes | Intermittently machine is not switching on issue | Intermittently machine is not switching on issue | 000000000000000000000000001O##: 00000000000000000000000000100000000956952 00000000-Sep-2000000003 Kuar, Vipin The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : Intermittently machine is not switching on issue Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: No further information is available as per the customer SE 1.000000002.1 Resolution : Checked and found system started 000000000000000000000000001king.observed for a week no issues with start uphanded over system in good 000000000000000000000000001king condition to user. Tools & Equpment : MuTIMETER,2000000003-00000000-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000950000000000000000emesh, rajan The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : Intermittently machine is not switching on issue Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: No further information is available as per the customer SE 1.000000002.1 Diagnostic performed by Engineer : Intermittently machine is not switching on issue Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Internal Comments : full# and/or description of full: Intermittently machine is not switching on issue Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Repair Action : Intermittently machine is not switching on issue Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Resolution : Intermittently machine is not switching on issue Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Follow u Required : full# and/or description of full: Intermittently machine is not switching on issue Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Repair Action : Intermittently machine is not switching on issue Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Technician Name : Remesh, rajan ----------------------------------------------------------------------------- | 2023-09-04 12:16 | No
66 | Yes | The equipment startup stopped with the message "u". | The equipment startup stopped with the message "0:0". | 000000000000000000000000001O##: 00000000000000000000000000100000000955890000000000000000-Sep-2000000003 Saito, Yoshiki The system meets the specification for the performed service and is retuned to use Regulatory Questions: N/A Problem Description by Engineer: Problem description by engineer : The equpment startup stopped with the message "00000000:00000000". u Morita Restarting the system improved the situation. - This was responded to several months ago. He 000000000000000000000000001ld like u to check the log. Customer Fuction/Role: Radiology Department Device use: Not used Expected and unexpected Results: The equpment startup stopped with the display "00000000:00000000". u Morita Restarting the system improved the situation. - This was responded to several months ago. He 000000000000000000000000001ld like u to check the log. user Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Information to support the complaint handling process Customer Fuction/Role: Not provided Device use: Not provided Expected and unexpected Results: Not provided user Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Resolution: I checked it on site. I restarted the system a number of times, but I was not able to detect any abnormal full. I checked the log for the past week, and I did not think there was a problem. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000955319 Kamata, Takayui The system meets the specification for the performed service and is retuned to use Problem Description by Engineer: Problem description by engineer : The equpment startup stopped with the message "00000000:00000000". u Morita Restarting the system improved the situation. - This was responded to several months ago. He 000000000000000000000000001ld like u to check the log. Customer Fuction/Role: Radiology Department Device use: Not used Expected and unexpected Results: The equpment startup stopped with the display "00000000:00000000". u Morita Restarting the system improved the situation. - This was responded to several months ago. He 000000000000000000000000001ld like u to check the log. user Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Information to support the complaint handling process Customer Fuction/Role: Not provided Device use: Not provided Expected and unexpected Results: Not provided user Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: I checked the log. The power was tuned on at 7:59:00000000 and then shut down 5 minutes later at 8:0000000011. - During this 5-minue 8:00000000:52,8:00000000:19 and 2 "Application Restart event received from MPD" commands were recorded. - The above command is not listed in the Daily Log for the past few days. There is no record of an uLLon the FleuVision PC, and there is no record of "connect lost". - I think it was a power on switch malfunction or a MPD Control malfunction. Case.0000000021723831 5/25 (0000000080000000058600000000: FleuVision-2 Revised PC no HDD replacement) I reported the above to u Morita and he understood, but he said he was concerned and wanted u to come and check it on Wednesday when he does not have an examination scheduled. I asked FSE Saito to respond. Internal Comments: [Subject] 00000000DS:TK:u Morita: The equpment startup stopped with the message "00000000:00000000". [Internal Comments] Comp [Parts from the equpment] 00000000230000000016000000007:MPD CONTROL unit QTY:1 ZSID Osaka (u15) [u-ed parts] 00000000800000000000000005100000000Service part REVIEW Module QTY:1 ZRFA Osaka (u15) Reason for u-collection: Consulted with engineer. Result of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Yoshio Saito] Saito, Yoshiki - FSE1 Visit date and time 2000000003/00000000/0000000015:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instructions] ----------------- shipping instructions #1----------------- - Delivery conditions NBD (mined loading, after the neat day) - Delivery Mut arrived by 10000000000000000 on September 5, 2000000003. - Delivery address name Kobe Office - Received person name Yoshio Saito - 00 - Others Requester name Takayui Kamata [Examination category] Other: Follow u Required: * ----------------------------------------------------------------------------- | 2023-09-04 12:16 | No
67 | Yes | u does not pass function test (no shock delivered) - Parts are replaced. | DFM0 does not pass union test (no shock delivered) - Parts are replaced. | 000000000000000000000000001O##: 000000000000000000000000001000000009559600000000 Parts Consumed: 1.000000000000000000000000989800000000190000000071,Rechargeable Li-ion Battery Pack Problem Description by Engineer: Visual inspection: OK - > fuction test: Defibrillator does not pass fuction test (no shock delivered), read our hardware uLLlog - > unlocalized: therapy board defective - replace > therapy board -> -> battery deeply de-de-charge - replace > battery -> perform technical safety test and technical safety test - > final test: OK Resolution: Visual inspection: OK - > fuction test: Defibrillator does not pass fuction test (no shock delivered), read our hardware uLLlog - > unlocalized: Therapy board defective - replaced > therapy board - > battery deep de-charge - > battery replaced -> technical safety test and mechanical safety test performed - > final test: OK ----------------------------------------------------------------------------- | 2023-09-04 12:36 | No
68 | Yes | Machine is not switching On | Machine is not switching On | 000000000000000000000000001O##: 00000000000000000000000000100000000956369 00000000Sep-2000000003 VEMBARASAN K The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : Machine is not switching On Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: No further information is available as per the customer SE 000Resoluion : Host pc SMPS defective, visit site and check FD200000000system host pc SMPS got defective and replaced same SMPS system made functional. And return the system to costumer for clinical use ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000955292 Remesh, rajan The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer : Problem description by engineer : Machine is not switching On Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: No further information is available as per the customer SE 000Diagnostic performed by Engineer : Machine is not switching On Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Internal Comments : full# and/or description of full: Machine is not switching On Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Repair Action : Machine is not switching On Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Resolution : Machine is not switching On Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Follow u Required : full# and/or description of full: Machine is not switching On Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Repair Action : Machine is not switching On Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Technician Name : Remesh, rajan ----------------------------------------------------------------------------- | 2023-09-04 13:31 | No
69 | Yes | system do not startup, restart do not resolve, system down | system do not startup, restart do not resolve, system down | 000000000000000000000000001O##: 0000000000000000000000000010000000095660000000000000000-Sep-2000000003 BRATT MATHIAS The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: MT Device use: not in use Expected and unexpected Results: u-ray system 000000000000000000000000001t startup. user Impact: cant use system Patient Impact: NA resolution : Re-routed Power supply to host-disc on Sute-PC. Test ok ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000956000000008 JALAR JIM The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Follow u Required : u-ray system 000000000000000000000000001t startup. After faultfinding with customer we have found that the issue seems to be with the u-ray PC, no power to OS-disk. CCC: Order a new u-ray PC and send FSE to replace part. system is down, system code is 722000000008. Problem Description by Engineer : Customer Fuction/Role: MT Device use: not in use Expected and unexpected Results: u-ray system 000000000000000000000000001t startup. user Impact: cant use system Patient Impact: NA uuDiagnostic performed by Engineer : u-ray system 000000000000000000000000001t startup. After faultfinding with customer we have found that the issue seems to be with the u-ray PC, no power to OS-disk. CCC: Order a new u-ray PC and send FSE to replace part. system is down, system code is 722000000008. ----------------------------------------------------------------------------- | 2023-09-04 14:01 | No
70 | Yes | cant start up. get failure u-ray system not awailable. restart do not resolve | cant start up. get failure u-ray system not awailable. restart do not resolve | 000000000000000000000000001O##: 00000000000000000000000000100000000956675 00000000 September 2000000003 Fridh, Johan The system meets the specifications for the performed service and is retuned to use Parts Consumed: 1.00000000000000000000000000000000800000000660000000093,u-ray PC Monoplane Resolution: Customer Fuction/Role: MT Device use: Not in use Expected and unexpected Results: Customer exchanged u-ray PC to a new one, system now 000000000000000000000000001ks according to specifications. user Impact: NA Patient Impact: NA problem Description by Engineer: Customer Fuction/Role: MT Device use: Not in use Expected and unexpected Results: u-ray system 000000000000000000000000001't start u. user Impact: Can't use system Patient Impact: NA uu----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000955812 JALAR JIM The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow-u 000000000000000000000000001k scheduled Follow-u Required: The u-ray system is not starting u. After faultfinding with the customer, we found that the issue seems to be with the u-ray PC, no power to the OS disk. CCC: Order a new u-ray PC and send an FSE to replace part. System is down, system code is 722000000008. Problem Description by Engineer: Customer Fuction/Role: MT Device use: Not in use Expected and unexpected Results: u-ray system 000000000000000000000000001't start u. user Impact: Can't use system Patient Impact: NA uuDiagnostic performed by Engineer: u-ray system 000000000000000000000000001't start u. After faultfinding with the customer, we found that the issue seems to be with the u-ray PC, no power to the OS disk. CCC: Order a new u-ray PC and send an FSE to replace part. System is down, system code is 722000000008. ----------------------------------------------------------------------------- | 2023-09-04 14:06 | No
71 | Yes | Unable to screen or expose. | Unable to screen or expose. | 000000000000000000000000001O##: 00000000000000000000000000100000000961931 21-Sep-2000000003 David Chatterton The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: Clinical engineer Device use: in use Expected and unexpected Results: usable to screen or expose user Impact: abort procedure Patient Impact: delay - move to another system resolution : Replaced system disk in Sute PC as recommended by helpdesk. Reinstalled software an configuration. Parts Consumed : 1.00000000000000000000000000000000800000000660000000022,SSD OS Haswell ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000956968 00000000Sep-2000000003 David Chatterton The system meets the specification for the performed service and is retuned to use Resolution : Performed complete software reinstall. system appears stable after re-install. continue to test system for 1 hou without incident. no full recorded in log file. Further information from tier 3 recommended replacement of system disk of site PC. ordered disk as follow u 000000000000000000000000001k. Follow u Required : Replace hard disk in site PC Problem Description by Engineer : Customer Fuction/Role: Clinical engineer Device use: in use Expected and unexpected Results: usable to screen or expose user Impact: abort procedure Patient Impact: delay - move to another system uuDiagnostic performed by Engineer : customer called 100000000300000000AM usable to screen or expose. arrived at 10000000000000000 confirmed that no fluro or exposures were possible. No iMessage on monitors, green ready symbol displayed. usable to open PSC window - no response from PC. Shu down system and restarted, after restart system was 000000000000000000000000001king. downloaded log file for escalation. Tier 2 recommended complete software reinstall. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000956985 David Chatterton The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Customer Fuction/Role: Clinical engineer Device use: no fluro or expose Expected and unexpected Results: no fluro or expose possible user Impact: delay Patient Impact: delay SE Diagnostic performed by Engineer : Asked customer to restart system and check for messages on monitor. after 2 restarts system was still not 000000000000000000000000001king. ----------------------------------------------------------------------------- | 2023-09-04 15:06 | No
72 | Yes | Loudspeaker uLLis displayed-Sound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker uLLis displayed-Sound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 000000000000000000000000001000000009576200000000Problem Description by Engineer: Visual inspection: OK. - > fuction test: Loudspeaker full- > unlocalized: Loudspeaker defect - replace > loudspeaker - perform > technical safety test and > technical safety test Resolution: Visual inspection: OK. - > fuction test: Loudspeaker full- > unlocalized: Loudspeaker defect - > loudspeaker replaced - performed technical safety test > MKT Parts Consumed: 1.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- | 2023-09-04 16:21 | No
73 | Yes | u / The equipment gives a failure in the defibrillation test. | 00/ u00 / The equipment gives a failure in the defibrillation test. | 000000000000000000000000001O##: 00000000000000000000000000100000000957630000000000000000-Sep-2000000003 Run Álvarez Villa The system does not meet the specification for the performed service and is not retuned to use. Regulatory Questions: N/A Problem Description by Engineer: Customer Fuction/Role: Electromedicine Device use: Defibrillator Expected and unexpected Results: System failure. user Impact: Impossible to use the equpment. Patient Impact: A different equpment is used. SE Diagnostic performed by Engineer: After the diagnosis performed, it is determined that the equpment cannot be repaired because spare parts are not available since it arrived on December 12, 2000000002 at the end of its life and service. Obsolescence letter is attached. The equpment is stored at the moment in the u 000000000000000000000000001k room, waiting for instructions on how to proceed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000957517 Imbernon Masip, Francisco The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Resolution: The case is transferred onsite. Regulatory Questions: N/A Problem Description by Engineer: Customer Fuction/Role: Electromedicine Device use: Defibrillator Expected and unexpected Results: System failure. user Impact: Impossible to use the equpment. Patient Impact: A different equpment is used. SE Diagnostic performed by Engineer: The client (José Lus García 666370000000065) says that the equpment shows a defibrillation test failure. We replaced the cable and used a different test cable and the failure persists. The case is transferred onsite so that the equpment can be checked by a field service. Follow u Required: The client (José Lus García 666370000000065) says that the equpment shows a defibrillation test failure. We replaced the cable and used a different test cable and the failure persists. - Equpment located in N+1 Z00000000ugical Block, neat to supervisor's office - Assign fse - Take into account that this equpment is EOL: Heart Start uL+ Defibrillator/Monitor 2000000002-12-31. ----------------------------------------------------------------------------- | 2023-09-04 16:21 | No
74 | Yes | u malfunctioning | DFM0 umalfunuoning | 000000000000000000000000001O##: 00000000000000000000000000100000000957585 Nandha, Rahu The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: user Device use: yes Expected and unexpected Results' user Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : issue could not be confirmed. Opchk passed ----------------------------------------------------------------------------- | 2023-09-04 16:41 | No
75 | Yes | Machine is not booting | Machine is not booting | 000000000000000000000000001O##: 00000000000000000000000000100000000957672 00000000Sep-2000000003 Guta, Pranjal The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : Machine is not booting Check power to GPDuin M cabinet. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: No further information is available as per the customer SE 8.2.17.00000000Resoluion : Check and found fleuvision PC not getting ON, change the battery and check the machine completely. Handed over the machine in good 000000000000000000000000001king condition to lab technician for clinical use. Tools & Equpment : MuTIMETER,200000000000000007-17 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000950000000000000000emesh, rajan The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : Machine is not booting Check power to GPDuin M cabinet. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: No further information is available as per the customer SE 8.2.17.00000000Diagnostic performed by Engineer : Machine is not booting Check power to GPDuin M cabinet. Check Host PC is booting. Internal Comments : full# and/or description of full: Machine is not booting Check power to GPDuin M cabinet. Check Host PC is booting. Repair Action : Machine is not booting Check power to GPDuin M cabinet. Check Host PC is booting. Resolution : Machine is not booting Check power to GPDuin M cabinet. Check Host PC is booting. Follow u Required : full# and/or description of full: Machine is not booting Check power to GPDuin M cabinet. Check Host PC is booting. Repair Action : Machine is not booting Check power to GPDuin M cabinet. Check Host PC is booting. Technician Name : Remesh, rajan ----------------------------------------------------------------------------- | 2023-09-04 16:46 | No
76 | Yes | u calibration due (Quote provided). | M00A / u00/ CO0calibration due (Quote provided). | 000000000000000000000000001O##: 000000000000000000000000001000000009552700000000wight Dave Ragadio Carbon The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer : Customer Fuction/Role: ENDuER Device use: MONITORING Expected and unexpected Results: FOR CALIBRATION OF CO2 user Impact: NONE Patient Impact: NONE SE Resolution : -Recommended for onsite service to calibrate CO2. -For quotation. -need to verify if unit is end of life of end of service. ----------------------------------------------------------------------------- | 2023-09-04 16:46 | No
77 | Yes | The Wireless Footswitch is not switching on | The Wireless Footswitch is not switching on | 000000000000000000000000001O##: 00000000000000000000000000100000000960000000073 00000000-Nov-2000000003 Dawid Strydom The system meets the specification for the performed service and is retuned to use Follow u Required : Replacement of Base Station required. 1 u 00000000800000000257862 - WFS Base Station Resolution : As per R000000000000000000000000001 doing FCO it was found that the WFS is not charging and not switching on. Replacement of WFS, WBS & charger. Did all relevant checks and no problems found. Parts Consumed : 1.0000000000000000000000000000000080000000000000000552,Wireless footswitch Charger (WFSC) 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station Problem Description by Engineer : Customer Fuction/Role: Onsite CE Device use: Outside Expected and unexpected Results: WFS not switching on user Impact: NONE Patient Impact: NONE SE R1.2.u ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000958000000002 Dawid Strydom The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Diagnostic performed by Engineer : During another site visit is was found that the WFS is not coming on. New spare required. Problem Description by Engineer : Customer Fuction/Role: Onsite CE Device use: Outside Expected and unexpected Results: WFS not switching on user Impact: NONE Patient Impact: NONE SE R1.2.u Follow u Required : Please order: 1 u 00000000800000000238231 - wireless footswitch 3p (wfs 3p) 1 u 0000000080000000000000000552 - Wireless footswitch Charger (WFSC) ----------------------------------------------------------------------------- | 2023-09-04 18:06 | No
78 | Yes | Alarms/No Sound: The alarms are not heard on the central unit: Speaker not connected properly. No Trouble Found. | Alarms/No Sound: The alarms are not heard on the central unit: Speaker not coined properly. No Trouble Found. | 000000000000000000000000001O##: 00000000000000000000000000100000000958000000001 Andrea Ardighieri The system meets the specification for the performed service and is retuned to use. Problem Description by Engineer: Customer Fuction/Role: Engineer Clinical Device use: parameter monitor Expected and unexpected Results: speaker is not connected correctly user Impact: no impact Patient Impact: no impact SE an Resolution: The speaker was not connected correctly, restored audio connection with the aid of the head nose. ----------------------------------------------------------------------------- | 2023-09-04 18:51 | No
79 | Yes | still having issues with machine not flouring intermittently and sometimes shutters unavailable and table stuck | still having issues with machine not flouring intermittently and sometimes shutters unavailable and table stuck | 000000000000000000000000001O##: 0000000000000000000000000010000000095000000007 00000000-Nov-2000000003 Muphy, Kyle The system meets the specification for the performed service and is retuned to use Resolution : conducted visual and functional inspection of system. Could not duplicate issues and no issues were discovered in visual inspection of system. Problem Description by Engineer : Customer Fuction/Role: engineer Device use: it Expected and unexpected Results: geo movement fuction user Impact: none Patient Impact: none SE unknown ----------------------------------------------------------------------------- Internal Remarks : Field Complete this column Interaction Type: New Case Caller: Carly Everett IP & Issu: 816600000000 still having issues with machine not flouring intermittently and sometimes shutters unavailable and table stuck Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: 9/5/2000000003 Entitlement Options: Entitled Reason for Entitlement Claim: contract Action Taken: created O000000000000000000000000001 pinged RP chat to dispatch for Tuesday; called cu back to confirm tomorrow was a holiday and FSE 000000000000000000000000001ld be dispatched on Tuesday Neat Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-04 19:01 | No
80 | Yes | shock not delivered | shock not delivered | 000000000000000000000000001O##: 000000000000000000000000001000000009577700000000Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to use Follow u Required : issue with module Opchk not passing. Problem Description by Engineer : Customer Fuction/Role: Be Device use: no Expected and unexpected Results' user Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : issue with module Opchk not passing. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000957900000000 Parts Consumed : 1.0000000000000000000000000000000056000000009131,DFM100000000 Measurement Module u----------------------------------------------------------------------------- | 2023-09-04 19:36 | No
81 | Yes | Security Issue Legal Issue_0atients in the last month sequenced significant alopecia after perform | Sensuality Issue Legal Issue_0patients in the last month sequenced significant alope0 after perform |  | 2023-09-04 19:37 | Yes
82 | Yes | u Battery backup issue and will not boot up | 00CDN0Battery backup issue and will not boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000095878000000005-Sep-2000000003 Jorgensen, Rolf The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : system not booting u When was the first occurrence? Within... : 00000000Sep-2000000003 How often is issue occurring? : Continuously Malfunction area : Power Supply Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: system down reboot does not resolve user Impact: system down Patient Impact: no impact SE 8.2.17.00000000Resoluion : Troubleshot, reset and verified us configuration. Consult DC Group/3rd party for support. Verified configuration, reset circuit breakers and us initialized and powered u properly, resuming monitoring/charging and protection for system. Reset M cab mains and observed procedure for consistent operation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000095800000000 Chhlam, Sy The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : system not booting u When was the first occurrence? Within... : 00000000Sep-2000000003 How often is issue occurring? : Continuously Malfunction area : Power Supply Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: system down reboot does not resolve user Impact: system down Patient Impact: no impact SE 8.2.17.00000000Diagnostic performed by Engineer : Marc reported that the us was tripped and now the system is not tuning on. Per caller he resettled the us breaker that he found tripped and the system 000000000000000000000000001ld still not boot u. Internal Comments : full# and/or description of full: system not booting u Repair Action : Check power to pdupdm check if there are any tripped, check dc power supply in the back of the M cabinet. Replace defective parts. 00000000800000000387582 - PD.SCOMP.DCPS_20000000012V_5V Follow u Required : full# and/or description of full: system not booting u Repair Action : Check power to pdupdm check if there are any tripped, check dc power supply in the back of the M cabinet. Replace defective parts. 00000000800000000387582 - PD.SCOMP.DCPS_20000000012V_5V Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Mark Bergeron +1600000000851000000006 IP & Issu: 68000000006100000000Battery backup issue and will not boot u (input failure for monitoring system) Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: R000000000000000000000000001created and adv of cb du to Holiday and contacted on-call Neat Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: Internal Remarks : 0000000000000000uS in bypass with full Call On Power for service. ----------------------------------------------------------------------------- | 2023-09-04 19:46 | No
83 | Yes | u LS unit not charging -Split from PR u | Tempo LS unit not charging -Split from PR 000 |  | 2023-09-04 19:50 | No
84 | Yes | "Barometer Sensor Range full alarm occurred ju before u on a patient. | "Barometer Sensor Range full alarm occurred ju before u on a patient. |  | 2023-09-04 20:00 | No
85 | Yes | "Oxygen Not Available" alarm occurred during u on a patient | "Oxygen Not Available" alarm occurred during u on a patient |  | 2023-09-04 20:08 | No
86 | Yes | Shooting is being made, it is not recorded. | Shooting is being made, it is not recorded. | 000000000000000000000000001O##: 00000000000000000000000000100000000958851 18-Sep-2000000003 Faith Teller The system meets the specification for the performed service and is retuned to use. Problem Description by Engineer: Problem description by engineer: RMT - IPu Espouse is not possible. Image disk is full 00000000Sep-2000000003 00000000:56:100000000 RMT - ACQ: Image Store: Inconsistent Image Store - Frontal IPPC 00000000Sep-2000000003 00000000:55:00000000 RMT - IPu Image Disk write problem. Images are possibly lost. 00000000Sep-2000000003 16:50000000033 How often is the issue occurring? : Intermittent Malfunction area: HDD Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: Diagnostic Expected and unexpected Results: No further information is available as per the customer. user Impact: No further information is available as per the customer. Patient Impact: No further information is available as per the customer. SE 1.000000003000000005 Diagnostic performed by Engineer: The image hard discs that had been found to be faulty before were used to replace the old one. The images were seen to be recorded on the hard drive. The system was made sound and 000000000000000000000000001king. Parts Disks: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000958672 Buak Yediyildiz The system does not meet the full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow-u 000000000000000000000000001k scheduled. Internal Comments: full# and/or description of the full Root Cause for RMT - IPu Espouse is not possible. Image disk is full: 1. Espouse is not possible. Image disk is full. Root Cause for RMT - ACQ: Image Store: Inconsistent Image Store - Frontal IPPC: 1. Software issue: there is an inconsistency in the patient Database wrt the stored images on the Frontal IPPC. Root Cause for RMT - IPu Image Disk write problem. Images are possibly lost. 1. Image Disk write problem. Images are possibly lost. Repair Action: Service actions for RMT - IPu Espouse not possible. Image disk is full: 1.1. Instruct the customer to free u disk space. 1.2. Could be caused du to malfunction Image disk. Check whether any other CAT pattern hits related to malfunctioning HDD. Replace the disk when necessary. Service actions for RMT - ACQ: Image Store: Inconsistent Image Store - Frontal IPPC: 1.1. Perform the Verify Image Store Data Consistency test in FSF. 1.2. If problem persists, Recreate Image Store in FSF. Don't forget to archive patient data if we wanted before euecuing this step. 1.3. If the database consistency repair fails, replace the image disk. Service actions for RMT - IPu Image Disk write problem. Images are possibly lost. 1.1. Replace the image disk. Follow u Required: full# and/or description of the full Root Cause for RMT - IPu Espouse is not possible. Image disk is full: 1. Espouse is not possible. Image disk is full. Root Cause for RMT - ACQ: Image Store: Inconsistent Image Store - Frontal IPPC: 1. Software issue: there is an inconsistency in the patient Database wrt the stored images on the Frontal IPPC. Root Cause for RMT - IPu Image Disk write problem. Images are possibly lost. 1. Image Disk write problem. Images are possibly lost. Repair Action: Service actions for RMT - IPu Espouse not possible. Image disk is full: 1.1. Instruct the customer to free u disk space. 1.2. Could be caused du to malfunction Image disk. Check whether any other CAT pattern hits related to malfunctioning HDD. Replace the disk when necessary. Service actions for RMT - ACQ: Image Store: Inconsistent Image Store - Frontal IPPC: 1.1. Perform the Verify Image Store Data Consistency test in FSF. 1.2. If problem persists, Recreate Image Store in FSF. Don't forget to archive patient data if we wanted before euecuing this step. 1.3. If the database consistency repair fails, replace the image disk. Service actions for RMT - IPu Image Disk write problem. Images are possibly lost. 1.1. Replace the image disk. Technician Name: Buak Yediyildiz Problem Description by Engineer: Problem description by engineer: RMT - IPu Espouse is not possible. Image disk is full 00000000Sep-2000000003 00000000:56:100000000 RMT - ACQ: Image Store: Inconsistent Image Store - Frontal IPPC 00000000Sep-2000000003 00000000:55:00000000 RMT - IPu Image Disk write problem. Images are possibly lost. 00000000Sep-2000000003 16:50000000033 How often is the issue occurring? : Intermittent Malfunction area: HDD Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: Diagnostic Expected and unexpected Results: No further information is available as per the customer. user Impact: No further information is available as per the customer. Patient Impact: No further information is available as per the customer. SE 1.000000003000000005 Diagnostic performed by Engineer: I talked to u Tayfu. I reviewed the failure logs. THERE ARE IPPC hard disk full. The hard disks will be changed. ----------------------------------------------------------------------------- | 2023-09-04 20:11 | No
87 | Yes | SPLIT- there seems to be an issue with one of the ‘pads’ on the dock connector-See Reference case PR u | SPLIT- there seems to be an issue with one of the ‘pads’ on the dock concur-See Reference case PR 0002 |  | 2023-09-04 20:16 | No
88 | Yes | Loudspeaker fault-Sound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker fault-Sound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000958719 QuROY, Nathan The system does not meet full specifications for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled. Regulatory Questions: No/No Problem Description by Engineer: Customer Fuction/Role: BIOMED Device use: NA Expected and unexpected Results: NA user Impact: NA Patient Impact: NA SE NA Diagnostic performed by Engineer: quote requested. Resolution: Quote requested. QR - 0000000000000000 QTY:1 0000000056000000003831 Speaker Assembly Follow u Required: Quote requested. QR - 0000000000000000 QTY:1 0000000056000000003831 Speaker Assembly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009590000000000000000 Parts Consumed: 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- | 2023-09-04 20:41 | No
89 | Yes | Alarm or Messages/u_No valid alarm during cardiac arrest/ur Error | Alarm or Messages/PSA_QA00000000No valid alarm during cardiac arrest/ur Error |  | 2023-09-04 20:44 | No
90 | Yes | Touch buttons do not work and are invalid | Touch buttons do not work and are invalid |  | 2023-09-04 20:55 | No
91 | Yes | external paddles are not recognized correctly | SPLIT: u-external paddles are not recognized correuy ( See reference PR 0000) |  | 2023-09-04 22:59 | No
92 | Yes | external paddles are not recognized correctly | SPLIT: u-external paddles are not recognized correuy( See reference PR 0000) |  | 2023-09-04 23:01 | No
93 | Yes | u/ external paddles are not recognized correctly. | 000/ CN0000/external paddles are not recognized correuy (split, ref. PR 0000. |  | 2023-09-04 23:01 | No
94 | Yes | u / external paddles are not recognized correctly . | 000/ u external paddles are not recognized correuy (split, ref. PR 0000. |  | 2023-09-04 23:02 | No
95 | Yes | CO0ensor is not working | CO0sensor is not working | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000001 Bronte Westaway The system meets the specification for the performed service and is retuned to use Parts Consumed : 1.0000000000000000000000000000000056000000001711,Intrepid CO2 Module ASSY Tools & Equpment : FLOW METER,20000000000000000000000002 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000009-13 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER,200000000000000005-21 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000003-15 00000000:00000000:00000000 Problem Description by Engineer : Customer Fuction/Role: biomed tech Device use: defibrillation Expected and unexpected Results: co2 equips malf user Impact: cant use Patient Impact: none SE 1.00000000.53 Resolution : EtCO2 module replaced du to service required message. Performance verification, calibration and safety testing completed - all passed within specification. Config updated as per Clinical Apps ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 Bronte Westaway The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer : Customer Fuction/Role: biomed tech Device use: defibrillation Expected and unexpected Results: co2 equips malf user Impact: cant use Patient Impact: none SE 1.00000000.53 Follow u Required : hrf repair and t&V Diagnostic performed by Engineer : Confirmed CO2 service required full Attempted SW upgrade - did not remedy, reseated cables, issue persists. EtCO2 module requires replacement. 0000000056000000001711 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000983600000000Alhad Apte The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Customer Fuction/Role: Biomedical Device use: Ou of use Expected and unexpected Results: CO2 sensor is not 000000000000000000000000001king user Impact: No impact on user Patient Impact: no impact on patient SE Not provided Diagnostic performed by Engineer : unit is being sent to repair facility for assessment and repair. ----------------------------------------------------------------------------- | 2023-09-05 05:21 | No
96 | Yes | u/boot up failed | 0/u FD0boot up failed | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 00000000/u uuuFD200000000Safety Questions: Case: 0000000022226671 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weiqiang Zhu Last updated by: Weiqiang Zhu 1. Was the device in clinical use at the time the issue was discovered? no 2. Was any patient or user harmed? no 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A ----------------------------------------------------------------------------- | 2023-09-05 06:01 | No
97 | Yes | Live stopped being displayed on FleuVision. | Live stopped being displayed on Fleusion. | 000000000000000000000000001O##: 00000000000000000000000000100000000960000000033 00000000-Sep-2000000003 Nakashima, Shogo The system meets the specification for the performed service and is retuned to use Regulatory Questions: N/A Problem Description by Engineer: Live is no longer displayed on FleuVision. Customer Fuction/Role: Radiology Department Device use: During examination Expected and unexpected Results: Live stopped being displayed on FleuVision. user Impact: Examination delay Patient Impact: Examination delay SE unknown Resolution: Du to a faulty grabber card on the FleuVision PC (large monitor control PC), It was not possible to output images. Yesterday, I reinstalled the software for the same problem, but the problem recured today, so I replaced the PC. After I replaced it, I installed the software and restored the backup. When I had completed the 000000000000000000000000001k, I checked the operation and confirmed normal operation. [Parts used] Part No.: 0000000080000000050000000000000000/Part Name's Haswell Fleuvision 2 PC on HDD /Serial No.: 722300000000/Batch No.: an /Number of units: 1 Parts Consumed: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009599800000000Ohgua, Yui The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Internal Comments: [Subject] 00000000DM:YO2:u Kouake: Live is no longer displayed on FleuVision. [Parts from the equpment] 0000000080000000058600000000:FleuVision-2_Revised PC no HDD QTY:1[GATP] ZREP/Traceable Tokyo (u11) 0000000080000000012200000000:u Haswell Fleuvision 2 PC no HDD QTY:1[GATP] ZREP/Traceable Nagoya (u100000000 [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Nakashima, Shogo [Shogo Nakajima] - FSE1 Visit date and time 2000000003/00000000/00000000--None-- - FSE1 000000000000000000000000001k time 2 [Shipping instructions] ----------------- shipping instructions #1----------------- - Delivery conditions [Delivery conditions not specified] - Delivery - Delivery address name - Received person name [Receiver not specified] - 00***-****-**** - Others Call mobile uon arrival Requester name Yui Ogua [Examination category] Normal examination Follow u Required: /* Problem Description by Engineer: Live is no longer displayed on FleuVision. Customer Fuction/Role: Radiology Department Device use: During examination Expected and unexpected Results: Live stopped being displayed on FleuVision. user Impact: Examination delay Patient Impact: Examination delay SE unknown Diagnostic performed by Engineer: RC Response: I explained how to identify the failure. We plan to respond on site. ----------------------------------------------------------------------------- | 2023-09-05 06:41 | No
98 | Yes | u/System cannot make exposure | SA:0R/u FD00System cannot make exposure | BuQ&R Feedback: No Reason: Corrective Maintenance External Remarks: 00000000.00000000.2000000003 12:1000000008 cut (8600000000800000000) Service content: The upright position cannot be exposed, the IPPC hard disk cannot be read, and the hard disk can be reinserted. Service result: The equpment is in normal use. External Remarks: 00000000.00000000.2000000003 1000000001:27 cut (8600000000900000000) Service content: The upright position cannot be exposed, and the IPPC hard disk cannot be read. Re-plug the hard disk. Service result: The equpment is in normal use. SAP Transfer: Case Transferred to SAP Neat Step Action: Chai Yanfei 1310000000050000000061 calls and adds OP to FSE Siyuuin SAP Transfer: Dispatched for Onsite Support SAP Transfer: Dispatched for Onsite Support Internal Remarks: 00000000.00000000.2000000003 12:10000000000000000uC (8600000000800000000)$Q7: (Confirmed reported problem, can't read disk ) $Q8: (Check the logfile no report full images disk uLLread patient information full） $Q9: (reconnect images diks) $Q100000000 (Test and inspection data refer to: PA. Device retuned to full functionality.) $Q11: Calibrated Tools name: not use tools ^Current Tue installed date: ^Current Tue usage: Internal Remarks: 00000000.00000000.2000000003 1000000001:300000000uC (8600000000900000000)Q7: (Confirmed reported problem, can't read disk ) $Q8: (Check the logfile no report full images disk uLLread patient information full） $Q9:ages (reconnect diks) $Q100000000 (Test and inspection data refer to: PA. Device retuned to full functionality.) $Q11: Calibrated Tools name: not use tools ^Current Tue installed date: ^Current Tue usage: Internal Remarks: 0000000000000000.2000000003 00000000:200000000300000000TC (8600000000800000000) $Q7: (Confirmed reported problem, can't read disk ) $Q8: (Check the logfile no report full images disk uLLread patient information full $Q9: (reconnect images diks ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to full functionality.) $Q11: Calibrated Tools name: not use tools ^Current Tue installed date: ^Current Tue usage: Safety Questions: Case: 0000000022226826 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jiankang Jiang Last updated by: Jiankang Jiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? no 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A T2 Activities: Problem Description: Overview of customer problems: The device often reports that the available image storage space is low, and many patients have been deleted, or that the disk space is low, and now there are not many patients. Sometimes, only fluoroscopy, exposure, and image storage cannot be performed. You can look back at previous patient images. An engineer is required to check the cause of the faut on site. Whether it is the faut of the equpment itself: Yes Troubleshooting Action: Critical Faut Code or Symptom: Image disk free space low Image disk full Remote diagnosis and test results: Image storage management failure for the device PRS TR statue: No PRS Repair Action: Suggested field repair steps: 1. Do database consistency test, 2. Do database consistency repair, 3. Do Host and IPPC connection verification, 00000000Do Re-creation image store, 5. Check whether the image net000000000000000000000000001k connection cable between host PC and IP PC is damaged and oxidized, 6. Check the smart data of the image disk. If necessary, replace the image hard disk. 7. Reinstall the IP PC software. It is recommended to format the system hard disk before reinstalling the software. 8. Bypass the hard disk bou and do troubleshooting. 9. Replace the IP PC Customer Requirements Remarks: Onsite SAP Transfer: Case Transferred to SAP External Remarks: 0000000000000000.2000000003 00000000:20000000032 cut (8600000000800000000) Service content: the upright position cannot be exposed, the IPPC hard disk cannot be read, re-plug the hard disk. Service result: The equpment is in normal use. SAP Transfer: Dispatched for Onsite Support Problem Reported by Customer: SA:00000000R/u uuuFD200000000200000000no rays SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-05 07:01 | No
99 | Yes | u-auditor failure. | 00/ u000/ defibrillator/Monitor / BUY PARTS-usurer failure. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: auditor failure Governor Jiang Guming 13700000000182000000005 call to make a quote FSE Notes: Block Invoice Reason: NA FSE Notes: This PO is a Reference Number only SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support SAP Transfer: Case Transferred to SAP Safety Questions: Case: 0000000022227100000000 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: uiyuyang Last updated by: uiyuyang 1. Was the device in clinical use at the time the issue was discovered? no 2. Was any patient or user harmed? no 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A ----------------------------------------------------------------------------- | 2023-09-05 08:26 | No
100 | Yes | The three-beep sounded - | The three-beep sounded - | 000000000000000000000000001O##: 00000000000000000000000000100000000955389 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to use Resolution: "I replaced the AED and confirmed that it passed the battery fit test. [Simultaneous FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is required." September 7, 2000000003 Fuua Denshi Hyogo Hanbai I sent the equpment along with a report to the on-site person in charge. September 15, 2000000003 Fuua Denshi Hyogo Hanbai Received by the on-site Person in Charge. Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Result: Pass Regulatory Questions: N/A Problem Description by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device use: In standby mode Expected and unexpected Results: The AED failed a self test. user Impact: Cannot use the AED in an emergency. Patient Impact: No patient used. SE Equpment without version notation Diagnostic performed by Engineer: Received. I performed a battery test on the AED and confirmed that it passed the test.I monitored the AED in standby mode, but no alarm, etc. occurred. As a result of verifying the AED main unit's operation record, there was a record of the 3-beep alarm sounding 2 times in the recent automatic self-test.An AED high voltage circuit uLLwas detected, so the AED is faulty. "I will replace the faulty unit. I apologize for the inconvenience this has caused. " Parts Consumed: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Each ----------------------------------------------------------------------------- Internal Remarks: Failure content 2 3. The AED beeped continuuly.8/29 When the battery was re-re-pu and the self-test was run it was without problem. Please inspect. Lot No. Adult pad C100000000522-00000000 Adult pad C100000000522-00000000 Battery 97828P ----------------------------------------------------------------------------- | 2023-09-05 09:41 | No
101 | Yes | Failing self-test. | Failing self-test. | 000000000000000000000000001O##: 0000000000000000000000000010000000095500000000 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to use Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Result: Pass Resolution: "I replaced the AED and confirmed that it passed the battery fit test. [Simultaneous FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is required." September 7, 2000000003 Fuua Denshi Kiki Hanbai Co.,Ltd. I sent the equpment along with a report to the on-site person in charge. Regulatory Questions: N/A Problem Description by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device use: In standby mode Expected and unexpected Results: The AED failed a self test. user Impact: Cannot use the AED in an emergency. Patient Impact: No patient used. SE Equpment without version notation Diagnostic performed by Engineer: I performed a battery test on the AED and confirmed that the single-beep alarm sounded and the message "To perform a test, remove and re-install the battery". I verified the operation record of the AED and found that a self test had detected an uLLin the audio circuit element inside the AED, which led to the test not passing and the alarm sounding. "The AED is faulty, so it needs to be replaced. I apologize for the inconvenience this has caused. " Parts Consumed: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Each ----------------------------------------------------------------------------- Internal Remarks: Failure content 2 The AED beeped about t000000000000000000000000001weeks ago. The AED retuned to normal temporarily when a self test was run, but the AED started beeping again on the 25th of last week. The message "To test, remove and re-install the battery" was announced. Lot No. Adult pad C112322-00000000 Adult pad Battery 87532P ----------------------------------------------------------------------------- | 2023-09-05 09:51 | No
102 | Yes | Machine is not functioning in battery mode | Machine is not funuoning in battery mode (Replaced ASSY Therapy) | 000000000000000000000000001O##: 00000000000000000000000000100000000956296 Madhan, kuar The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Follow u Required : To replace 0000000056000000009000000001, DFM100000000 Therapy PCA ASSY Problem Description by Engineer : Customer Fuction/Role: BME Device use: No Expected and unexpected Results: N/A user Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001, DFM100000000 Therapy PCA ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000960000000000000000 Parts Consumed : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy ----------------------------------------------------------------------------- | 2023-09-05 11:46 | No
103 | Yes | u/Defibrillator/Monitor/equipment malfunction | 0N/Effi0 DFM0 Defibrillator/Monitor/equipment malfunuon (Replaced ASSY Processor) | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP External Remarks: 11.00000000.2000000003 00000000:21:00000000 cut (860000000000000000100000000Department information: [Department: Equpment Department Contact Person Liang Gong Tel: 18652000000002268]* $ASP engineer:[honglejia Zhoukoudian] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem An allocated in the self-test $Q8: Disassembly check, the processing board was faulty $Q9:Replaced the processing board $Q100000000Test and inspection data refer to PA checklist, Device retuned to full functionality $Q11: $ Calibrated Tools name (ProSim8) S/N (26712200000000 Neat Calibration date (10000000028/2000000003), (ESA6200000000 S/N (300000000000000003) Neat Calibration date (11/3000000002000000003), (70000000000000000P) /SN:3953000000006/Neat Calibration (10000000029/2000000003) Internal Remarks: 11.00000000.2000000003 00000000:21:100000000uC (860000000000000000100000000 $ASP engineer:[honglejia Zhoukoudian] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem An allocated in the self-test $Q8: Disassembly check, the processing board was faulty $Q9:Replaced the processing board $Q100000000Test and inspection data refer to PA checklist, Device retuned to full functionality $Q11: $ Calibrated Tools name (ProSim8) S/N (26712200000000 Neat Calibration date (10000000028/2000000003), (ESA6200000000 S/N (300000000000000003) Neat Calibration date (11/3000000002000000003), (70000000000000000P) /SN:3953000000006/Neat Calibration (10000000029/2000000003) SAP Transfer: Case Transferred to SAP T2 Activities: Problem Description: Customer's explanation of the problem: equipment failure The customer observed the iMessage displayed on the monitor: NA Troubleshooting Action: The customer observed the iMessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine but it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: It is recommended to check the usable and processor Recommended spare part uufor T2 with description: #0000000056000000009000000001 Notes on customer requirements: NA Test/Inspection: NA Test Equpment: NA Problem Reported by Customer: failure Safety Questions: Case: 00000000222277500000000 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gujju Yu Last updated by: Gujju Yu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 2023-09-05 11:46 | No
104 | Yes | Incorrect saturation readings; No Trouble Found. Cau Unknown. | Incorreusaturation readings; No Trouble Found. Cau Unknown. | 000000000000000000000000001O##: 00000000000000000000000000100000000961000000008 00000000-Sep-2000000003 Piotr Zinko The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled. Diagnostic performed by Engineer: Verification of situation measurement in t000000000000000000000000001cardiac monitors: Mu700000000 SN: DE58563200000000+u3 SN: DE6900000000000000006 Mu700000000 SN: DE585632300000000u3 SN: DE6900000000000000009. No irregularities in situation measurements were found. A message about a malfunction of the SPO2 sensor was observed in the history of alarm messages for the set Mu700000000 SN: DE585632300000000u3 SN: DE6900000000000000009. LED malfunction was found in server u3 DE6900000000000000009 - further diagnostics is necessary and possibly a replacement of the motherboard. The measurement server was taken for 000000000000000000000000001kshop diagnostics. Tools & Equpment : PATIENT Simulator, 2000000003-10000000020000000000000000:00000000:00000000. Problem Description by Engineer: Customer Fuction/Role: biomed. Device use: monitored monitor. Expected and unexpected Results: incorrect SPO2 measurement - not confirmed. user Impact: none. Patient Impact: none. SE unknown. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000958279 Piotr Zinko The system does not meet the specification for the performed service and is not retuned to use. Problem Description by Engineer: Customer Fuction/Role: biomed. Device use: monitored monitor. Expected and unexpected Results: incorrect SPO2 measurement - not confirmed. user Impact: none. Patient Impact: none. SE unknown. Diagnostic performed by Engineer: Further diagnostics with the service tester is necessary du to the reported problem with situation measurement. Follow u Required: The client requests an on-site visit in order to verify the situation measurement operation - to be scheduled. ----------------------------------------------------------------------------- | 2023-09-05 12:16 | No
105 | Yes | u / uL Defibrillator/u wave display abnormal. | M00A / u00 / Heart Start uL Defibrillator/Monitor / wave display abnormal. | BuQ&R Feedback: No Reason: Corrective Maintenance Neat Step Action: Contact ASP Yigong for business follow-u Problem Reported by Customer: ECR:00000000/Heart Start uL Defibrillator/Monitor Safety Questions: Case: 0000000022227927 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: DiFeng Liang Last updated by: DiFeng Liang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? no 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A Neat Step Action: The equpment is in Shanghai Puong Hospital Transfer to sales handling, this case is considered invalid and without any further steps if quotation is not accepted, will be closed without final solution after customer refuse our service One EMS internal remarks: There are waveforms and heart rate even when the usable is not connected. It is the same when using AC power or not connecting to AC power but only connected to the battery. Business follow-u is required. ----------------------------------------------------------------------------- | 2023-09-05 12:26 | No
106 | Yes | u arm issue | OCAP:0R/u FD0C arm issue | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 000000003818291737 call customer feedback contact teacher already contacted fse refused to transfer to rise Internal Remarks: 0000000000000000.2000000003 13:100000000 cut (8600000000600000000) $Q7: (Confirmation & Description of problem/C arm issue) $Q8: (customer report device C arm issue, C-arm can't move, check and found foreign matter stuck movement) $Q9: (clean foreign matter and test movement ok, test system fuction ok then system OK) $Q100000000 (Test system run ok and inspection data refer to: PA. Device retuned to full functionality. ) $Q11: N/A T2 Activities: Problem Description: Detailed description of customer problems or symptoms: The C-arm does not move, the sheet is rolled into the C-arm Whether it is the faut of the equpment itself: Yes Troubleshooting Action: Critical Faut Code or Symptom: The C-arm does not move, the sheet is rolled into the C-arm Remote diagnosis and test results: N/A PRS TR statue: No PRS Repair Action: Suggested field repair steps: " It is recommended that engineers come to the scene to repair and deal with it" Safety Questions: Case: 0000000022227998 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qingfeng Gu Last updated by: Qingfeng Gu 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? no 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A SAP Transfer: Dispatched for Onsite Support External Remarks: 0000000000000000.2000000003 13:1000000008 cut (8600000000600000000) Service content (required): The movement of the C-arm of the equpment is faulty. It is found that a foreign object is stuck in the movement of the C-arm. After removing the foreign object, it can be used normally. Service result (required): [After repairs, the device retuned to normal operation.] Other important things: [none] Neat Step Action: 15821257166 Incoming call from SS wuian, transferred to u Gu SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-05 12:36 | No
107 | Yes | Device does not completely go up - child is under anesthesia. | Device does not completely go up - child is under anesthesia. | 000000000000000000000000001O##: 00000000000000000000000000100000000961358 September 6, 2000000003 Beer, Dietmar The system meets the specification for the performed service and is retuned to use Problem Description by Engineer: Problem description by engineer: System start hangs at 00000000:00000000 Information to support the complaint handling process Customer Fuction/Role: Not provided Device use: Not provided Expected and unexpected Results: Not provided user Impact: Not provided Patient Impact: Not provided SE 000Resoluion: Fleuvision PC replaced, fuction check, system is ready for clinical use Parts Consumed: 1.0000000000000000000000000000000080000000058600000000,FleuVision 2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000961100000000Baugarten, Rüdiger The system does not meet the specification for the performed service and is not retuned to use Internal Comments: full# and/or description of full No FleuVision PC start, PC or HD defective Repair Action: The with part Follow-u required: full# and/or description of full No FleuVision PC start, PC or HD defective Repair Action: The with part Technician name: Baugarten, Rüdiger Problem Description by Engineer: Problem description by engineer: System start hangs at 00000000:00000000 Information to support the complaint handling process Customer Fuction/Role: Not provided Device use: Not provided Expected and unexpected Results: Not provided user Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: Call-back, uuuoff/on ----------------------------------------------------------------------------- Customer comment: Request only for 00contact - customer is on his way in the building. Please check via remote. Customer comment: ### u Beer will inform u Clement about service on 00000000.00000000. ----------------------------------------------------------------------------- | 2023-09-05 13:16 | No
108 | Yes | Speaker technical uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Speaker technical uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000962000000003 Fhager, Andreas The system meets the specifications for the performed service and is retuned to use Resolution: Replaced the speaker assembly. Problem Description by Engineer: Customer Fuction/Role: Biomed Device use: Not in use Expected and unexpected Results: INOP message "speaker malfunction" appears on the monitor sporadically, no other device connected when INOP appears. The sound is not 000000000000000000000000001king from the monitor when INOP appears. Sound OK central. user Impact: INOP "speaker malfunction" Patient Impact: None SE M.0000000000000000 Diagnostic performed by Engineer: Problem Acknowledged Parts Consumed: 1.0000000000000000000000000000000056000000006631,IV2-FLEu MECHASY Loudspeaker assembly Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000003-12-12 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000095600000000 Ersbrand, Gutav The system does not meet the specifications for the performed service and is not retuned to use Follow-u required: The customer will send the device for bench repair. Problem Description by Engineer: Customer Fuction/Role: Biomed Device use: Not in use Expected and unexpected Results: INOP message "speaker malfunction" appears on the monitor sporadically, no other device connected when INOP appears. The sound is not 000000000000000000000000001king from the monitor when INOP appears. Sound OK central. user Impact: INOP "speaker malfunction" Patient Impact: None SE M.0000000000000000 Diagnostic performed by Engineer: Biomed reports INOP message "speaker malfunction" appears on the monitor sporadically, no other device connected when INOP appears. No sound from the monitor when INOP appears. The customer will send the device for bench repair. ----------------------------------------------------------------------------- | 2023-09-05 13:41 | No
109 | Yes | PC boot sequence problem | u-ray pc boot sequence problem | 000000000000000000000000001O##: 0000000000000000000000000010000000095600000000 21-Sep-2000000003 seokhyeon jeong The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: RT Device use: randomly occued Expected and unexpected Results: u-ray pc booting full user Impact: n/a Patient Impact: n/a SE R1.2.5 Resolution : Completed part monitoring. Tools & Equpment : MuTIMETER,200000000000000001-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000961975 00000000-Sep-2000000003 seokhyeon jeong The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: RT Device use: when system boot u Expected and unexpected Results: u-ray pc problem user Impact: n/a Patient Impact: n/a SE R1.2.5 Resolution : Checked bulldogs. Checked u-ray pc statue. -> had & memory full Reconnected had cables & Cleaned memory cards. Reinstalled u-ray pc software. Performed system booting tests. Passed Tools & Equpment : MuTIMETER,200000000000000001-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- | 2023-09-05 13:41 | No
110 | Yes | The equipment does not start | The equipment does not start | 000000000000000000000000001O##: 0000000000000000000000000010000000096100000000 Dragoscosmin Bogdan The system meets the specification for the performed service and is retuned to use Problem Description by Engineer: Customer Fuction/Role: as chief Device use: no Expected and unexpected Results: fleuvision not booting u user Impact: medium Patient Impact: no SE 8.2.17 Diagnostic performed by Engineer: It is reported that the equpment does not start (stuck at the start screen 00000000:00000000). The start mode and computer statue are checked. It is recommended to manually stop the FleuVision computer and restart it. After the procedures, the equpment is functional. ----------------------------------------------------------------------------- | 2023-09-05 14:06 | No
111 | Yes | u The Wi-Fi pedal does not work anymore. | FSE CDP The Wi-Fi pedal does not work anymore. | 000000000000000000000000001O##: 00000000000000000000000000100000000956655 Parts Consumed: 1.00000000000000000000000000000000800000000238192, Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862, WFS Base Station ----------------------------------------------------------------------------- | 2023-09-05 14:06 | No
112 | Yes | System won’t boot up. | System won’t boot up. | 000000000000000000000000001O##: 0000000000000000000000000010000000096170000000022-Sep-2000000003 From, Titian Nicolae The system meets the specification for the performed service and is retuned to use Problem Description by Engineer: Customer Fuction/Role: operator Device use: startup Expected and unexpected Results: no boot user Impact: none Patient Impact: none SE 2.2 L5 Parts Consumed: 1.00000000000000000000000000000000800000000117161, us 1Phase Data Integrity Controller SP Diagnostic performed by Engineer: us Controller replacement operations are performed, functionality tests are performed. The angiograph is taught fictionally with tests. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000961188 00000000-Sep-2000000003 From, Titian Nicolae The system meets the specification for the performed service and is retuned to use Follow u Required: us Controller replacement. Problem Description by Engineer: Customer Fuction/Role: operator Device use: startup Expected and unexpected Results: no boot user Impact: none Patient Impact: none SE 2.2 L5 Diagnostic performed by Engineer: A faut is detected at the level of us in M's cabinet. It is necessary to replace the defective us Controller part. The functional angiograph is handed over and kept under observation. Tools & Equpment : MuTIMETER, 2000000003-10000000028 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-05 14:21 | No
113 | Yes | Oxygen plant failure | 0K/V0uLLOuygen plant failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Support Safety Questions: Case: 0000000022229000000006 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhujian Lin Last updated by: Zhujian Lin 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A Problem Reported by Customer: V600000000uLLFSE Notes: This PO is a Reference Number only SAP Transfer: Case Transferred to SAP T2 Activities: Problem Description: Customer's explanation of the problem: Oxygen device failure The customer observed the iMessage displayed on the monitor: an Troubleshooting Action: The customer observed the iMessage displayed on the monitor: an T2 Remote Diagnostics and Testing: Other machines can be used normally in the same location. Repair Action: T2 Recommended on-site repair steps: It is recommended to visit the site for investigation Recommended spare part uufor T2 with description: " 00000000561512191 RP-Gas Delivery System (GDS), V600000000 Notes on customer requirements: an Test/Inspection: N/A Test Equpment: N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 18.1000000002000000003 00000000:00000000:00000000uC (86000000000000000000000000) $ASP engineer:Yingwei[suuhu$Q1:no $Q2：no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.The machine RP-Flow Sensor Assy and RP-Solenoid Valve are fauty.$Q8: Download "DRPT" report, uLLCode have not, Should be problems of RP-Flow Sensor Assy and RP-Solenoid Valve Q9: Replace the RP-Flow Sensor Assy and RP-Solenoid Valve .Q100000000Test and inspection data refer to: PA checklist. Device retuned to full fuctionality.$Q11:Calibrated Tools name (FLuE16500000000 S/N (000000006000000002) Neat Calibration date (2000000003/11/11) External Remarks: 18.1000000002000000003 00000000:00000000:00000000uC (86000000000000000000000000) *Department information: [Department: Department of Respiratory and Critical Care Medicine Contact person: Chen Yan Tel: 137767000000000000000000000000 $ASP engineer:Yingwei[suuhu$Q1:no $Q2：no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.The machine RP-Flow Sensor Assy and RP-Solenoid Valve are fauty.$Q8: Download "DRPT" report, uLLCode have not, Should be problems of RP-Flow Sensor Assy and RP-Solenoid Valve Q9: Replace the RP-Flow Sensor Assy and RP-Solenoid Valve .Q100000000Test and inspection data refer to: PA checklist. Device retuned to full fuctionality.$Q11:Calibrated Tools name (FLuE16500000000 S/N (000000006000000002) Neat Calibration date (2000000003/11/11) SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-05 14:51 | No
114 | Yes | No u-ray possible | No u-ray possible | 000000000000000000000000001O##: 00000000000000000000000000100000000961796 Baugarten, Rüdiger The system meets the specification for the performed service and is retuned to use Problem Description by Engineer: Problem description by engineer: No u-raying possible Information to support the complaint handling process Customer Fuction/Role: Not provided Device use: Not provided Expected and unexpected Results: Not provided user Impact: Not provided Patient Impact: Not provided uu00000000Diagnostic performed by Engineer: No iMessage, customer was not able to be Internal Comments: full# and/or description of full Foot switch does not respond. Repair Action: Foot switch cable on the table not completely plugged in Resolution: Plug re-attached, system 000000000000000000000000001ks correctly again. ----------------------------------------------------------------------------- | 2023-09-05 15:26 | No
115 | Yes | System is too slow | System is too slow | 000000000000000000000000001O##: 00000000000000000000000000100000000961995 Madhan, kuar The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Follow u Required : To replace 0000000056000000009000000001, DFM100000000 SOM PCA ASSY Install Problem Description by Engineer : Customer Fuction/Role: BME Device use: No Expected and unexpected Results: N/A user Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001, DFM100000000 SOM PCA ASSY Install ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000962197 Parts Consumed : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY SOM ----------------------------------------------------------------------------- | 2023-09-05 15:31 | No
116 | Yes | Machine is not booting up | Machine is not booting up | 000000000000000000000000001O##: 0000000000000000000000000010000000096200000000 00000000-Sep-2000000003 Chaniyara, Manan The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : Machine is not booting u Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: Not provided SE 000Resoluion : Reloaded ghost and replaced host pc HDD now system is 000000000000000000000000001king as intended Diagnostic performed by Engineer : Found host pc HDD faulty Tools & Equpment : MuTIMETER,2000000003-11-0000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096000000001 Remesh, rajan The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : Machine is not booting u Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: No further information is available as per the customer Patient Impact: Not provided SE 000Diagnostic performed by Engineer : Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Internal Comments : full# and/or description of full: Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Repair Action : Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Resolution : Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Follow u Required : full# and/or description of full: Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Repair Action : Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. Technician Name : Remesh, rajan ----------------------------------------------------------------------------- | 2023-09-05 16:06 | No
117 | Yes | u/Stand cannot move | 0/u FD0Stand cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities: Problem Description: Detailed description of customer problems or symptoms: The customer reported that some parts of the rack cannot be operated, and engineers need to be present for repairs Is the device itself faulty: Yes Troubleshooting Action: Critical Faut Code or Symptom: Part of the rack is inoperable Remote diagnosis and test results: restarting has no effect PRS TR statue: unavailable Repair Action: Suggested field repair steps: 1. Check whether the GIB EtherCAT indicator light is normal. 2. Check whether the AMC is normal. 3. Check whether each potentiometer is normal. 00000000Check the machine for uLLreporting, and perform faut maintenance. If it cannot be repaired, replace the corresponding spare parts Suggested spare part uuand designation: 00000000800000000265772 GEO IO BOu Remarks for customer needs: contact the customer in time to confirm the Onsite time Safety Questions: Case: 0000000022229576 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yangyang Gu Last updated by: Yangyang Gu 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? no 3. If the device has alarm/alert utility, did it alarm/alert as it should have at the time the issue was discovered? N/A 00000000Was this an our of bou failure? N/A Problem Reported by Customer: Rack cannot rotate Neat Step Action: ##ASP rujia for confirmation Neat Step Action: The specific dispatch time will be notified by Wuan Rujia Company (permission 137970000000012900000000. Neat Step Action: Shiyan Hospital of Integrated Traditional Chinese and Western Medicine One FD200000000 the device number cannot be provided The rack cannot be rotated, and an uLLis reported: some functions are available Wang 13986912819 Transferred to RSE Gu Yangyang Best Regards, Zheng Zhang Parts Team ----------------------------------------------------------------------------- | 2023-09-05 16:31 | No
118 | Yes | Not booting up stuck on initializing connection to host | Not booting up stuck on initializing conneuon to host | 000000000000000000000000001O##: 00000000000000000000000000100000000989792 11-Sep-2000000003 Davis Barry The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: interventional radiology tech Device use: system not in use Expected and unexpected Results: system will not boot user Impact: no user impact Patient Impact: no, patient impact SE8.1 Resolution : Return part to depot. Parts Consumed : 1.00000000000000000000000000000000800000000200000000611,Nehalem Host PC Monoplane Rev_III NO_HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000965868 00000000-Sep-2000000003 Davis Barry The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : System is not booting u When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: system not booting u and reboot does not resolve user Impact: system down Patient Impact: No further information is available as per the customer SE 8.1.3000000005 Resolution : Host PC is defective. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000962586 00000000-Sep-2000000003 Davis Barry The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer : Problem description by engineer : System is not booting u When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: system not booting u and reboot does not resolve user Impact: system down Patient Impact: No further information is available as per the customer SE 8.1.3000000005 Diagnostic performed by Engineer : System will not boot Follow u Required : Overnight parts order required. Request parts First overnight delivery on: Auto 6, 2000000003, 8:00000000 AM Ship parts to the site. Part = 00000000800000000200000000611 Qty = 1 Desc = host PC Part = 0000000080000000050000000000000000 Qty = 2 Desc = disk ￼ ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096200000000 Chhlam, Sy The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Internal Comments : full# and/or description of full: System is not booting u Repair Action : Check M cabinet for power issues and reset breakers or fuel if necessary. Check weather the host pc has power and tuns on. Check also the tun on board NP100000000and swap connection on it to see if it helps. 0000000080000000050000000000000000 - Host PC Monoplane Revised no HDD 00000000230000000016000000007 - MPD CONTROL unit Problem Description by Engineer : Problem description by engineer : System is not booting u When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: system not booting u and reboot does not resolve user Impact: system down Patient Impact: No further information is available as per the customer SE 8.1.3000000005 Diagnostic performed by Engineer : Renee reported that the system will not boot u and reboot does not resolved. Per caller, the tsm displayed message, "initializing connection to host," both data monitor and review monitor has black screens, rebooted twice and issue remains. Follow u Required : full# and/or description of full: System is not booting u Repair Action : Check M cabinet for power issues and reset breakers or fuel if necessary. Check weather the host pc has power and tuns on. Check also the tun on board NP100000000and swap connection on it to see if it helps. 0000000080000000050000000000000000 - Host PC Monoplane Revised no HDD 00000000230000000016000000007 - MPD CONTROL unit Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- Neat Step Action : Called on call left message on vm to call customer back Internal Remarks : Interaction Type: New Case Caller: Renee 813-929-50000000000000000 IP & Issu: Not booting u stuck on initializing connection to host Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and routed to RSE Neat Step Action: Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-05 16:41 | No
119 | Yes | Audio-no sound/No audio found at uu. | Audio-no sound/No audio found at Bench-Keyrin s/n u000other HW-Speaker-HW/Replaced. |  | 2023-09-05 16:49 | No
120 | Yes | suddenly stop | suddenly stop | 000000000000000000000000001O##: 00000000000000000000000000100000000962800000000 18-Sep-2000000003 AuAbdelrazek The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : system stopped on 00000000:00000000 How often is issue occurring? : First occurrence Malfunction area : PC's Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: Diagnostic Expected and unexpected Results: No further information is available as per the customer user Impact: no Patient Impact: no SE 8.2.17.00000000Resoluion : Replaced the Fleuvision PC in B cabinet and Hard Disk. Reloaded the Ful software and did the configuration on New Fleuvision PC. now the system is fully functioning and operation as customer satisfied Parts Consumed : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009611800000000mjad Aburqou The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : system stopped on 00000000:00000000 How often is issue occurring? : First occurrence Malfunction area : PC's Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: Diagnostic Expected and unexpected Results: No further information is available as per the customer user Impact: no Patient Impact: no SE 8.2.17.00000000Diagnostic performed by Engineer : system stopped on 00000000:00000000 doing start u need to order new fleuvioisn PC. Internal Comments : full# and/or description of full: system stopped on 00000000:00000000 doing start u need to order new fleuvioisn PC. Repair Action : system stopped on 00000000:00000000 doing start u need to order new fleuvioisn PC. Follow u Required : full# and/or description of full: system stopped on 00000000:00000000 doing start u need to order new fleuvioisn PC. Repair Action : system stopped on 00000000:00000000 doing start u need to order new fleuvioisn PC. Technician Name : Amjad Aburqou ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005800000000 18-Sep-2000000003 ----------------------------------------------------------------------------- | 2023-09-05 17:21 | No
121 | Yes | LSR problem | LSR problem | 000000000000000000000000001O##: 0000000000000000000000000010000000096100000000 00000000-Sep-2000000003 jinhyeok choi The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Fuction/Role: operator Device use: outside of use Expected and unexpected Results: LSR bootfail user Impact: no Patient Impact: no SER2.2.3 Resolution : LSR problem site pc reinstall system fuction check Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-05 17:26 | No
122 | Yes | Assistance to be snag. | Assistance to be snag. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 18-Sep-2000000003 use SERRA The system meets the specifications for the performed service and is retuned to use. Diagnostic performed by Engineer: reconfiguration of automatic sending and cardio report. Tests. Ok. Regulatory Questions: No/No Problem Description by Engineer: Customer Fuction/Role: Field service Hapdey Device use: No. Expected and unexpected Results: automatic configuration user Impact: No. Patient Impact: SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000992300000000 12-Sep-2000000003 Franck HAPDEY The system meets the specifications for the performed service and is retuned to use. Regulatory Questions: No/No Problem Description by Engineer: Customer Fuction/Role: service Device use: yes Expected and unexpected Results: NA user Impact: NA Patient Impact: late procedure SE 1.2.5 Resolution: Finalization of parameters. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009970000000000000000September 11, 2000000003 use SERRA The system meets the specifications for the performed service and is retuned to use. Regulatory Questions: No/No Problem Description by Engineer: Customer Fuction/Role: field engineer Device use: No. Expected and unexpected Results: reinstallation of software. user Impact: No. Patient Impact: SE Diagnostic performed by Engineer: Assistance to u Hapdey. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096000000001 12-Sep-2000000003 Franck HAPDEY The system does not meet full specifications for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled. Follow u Required: Finalization of replacement following PC Regulatory Questions: No/No Problem Description by Engineer: Customer Fuction/Role: service Device use: no Expected and unexpected Results: no user Impact: no Patient Impact: No. SE no. Resolution: Replacement of follow-u pc and software reload. Parts Consumed: 1.0000000000000000000000000000000080000000000000000900000000, PDuFues Oupus 1.00000000000000000000000000000000800000000660000000083, Following PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000962872 100000000Sep-2000000003 use SERRA The system does not meet the specifications for the performed service and is not retuned to use. Follow u Required: Further to the intervention, u Hapdey followed u. Regulatory Questions: No/No Problem Description by Engineer: Customer Fuction/Role: service Device use: yes Expected and unexpected Results: NA user Impact: NA Patient Impact: late procedure SE 1.2.5 Diagnostic performed by Engineer: Problem immediately: PC does not tun on. The power supply of the pc site is our of 000000000000000000000000001k. The us of the system is short-circuited. Ordered a new pc. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000962725 Regis LAMBING The system does not meet the specifications for the performed service and is not retuned to use. Internal Comments: (no impact for now) Resolution: Intervention by P. Serra Follow u Required: intervention by P. Serra Regulatory Questions: No/No Problem Description by Engineer: Customer Fuction/Role: service Device use: yes Expected and unexpected Results: NA user Impact: NA Patient Impact: late procedure SE 1.2.5 Diagnostic performed by Engineer: the machine tuned off at the beginning of the procedure, tried to restart. No log possible. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000001 12-Sep-2000000003 Franck HAPDEY Problem Description by Engineer: Customer Fuction/Role: service Device use: no Expected and unexpected Results: no user Impact: no Patient Impact: No. SE no. Regulatory Questions: No/No Diagnostic performed by Engineer: No/No ----------------------------------------------------------------------------- Internal Remarks: The customer wants to talk to Franck HAPDEY. ----------------------------------------------------------------------------- | 2023-09-05 17:31 | No
123 | Yes | The foot pedal as lost its Wi-Fi signal | The foot pedal as lost its Wi-Fi signal | 000000000000000000000000001O##: 000000000000000000000000001000000009629100000000arry White The system meets the specification for the performed service and is retuned to use Regulatory Questions : N/A Diagnostic performed by Engineer : 0000000021200000000737 / FCO open for this issue. Will speak to planning to make sue this gets booked in Problem Description by Engineer : Customer Fuction/Role: Rad Device use: Clinical Expected and unexpected Results: Wireless pedal issue user Impact: NA Patient Impact: NA SE1.u ----------------------------------------------------------------------------- Internal Remarks : 0000000021200000000737/FCO open for wireless footswitch upgrade ----------------------------------------------------------------------------- | 2023-09-05 17:51 | No
124 | Yes | Monitors did not red alarm at central - No product malfunction- unknown reason - insufficient information | Monitors did not red alarm at central - No produumalfunuon- unknown reason - insuffi0nt information | 000000000000000000000000001O##: 00000000000000000000000000100000000962931 Austin, Mike The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer, no follow-u requested Problem Description by Engineer : Customer Fuction/Role: Biomed Device use: Monitor Patients Expected and unexpected Results: Monitor Patients or Not Monitor Patients user Impact: None Patient Impact: None SE N/A Diagnostic performed by Engineer : Problem Description: reported red alarm did not alert for room 571-T2. -Iu C.00000000.00000000 -Affected unit- CICu -Affected Device-571T2 & Mu00000000- 571 2 & CICuuVA -biomed did not have any details to search the clinical audit trail outside of the room number. -biomed to interface with person reporting issue to gather detailed information. -biomed agreed case could be closed end of the day if no callback. -Permission Granted by: Ed Daniels -PRS-https://pww.portal.rs.healthcare.u.com/PRs Portal/defaut.aspu?PrsID=300000000662 -recommended to following the documented performance assurance tests to determine if device is functional. -u cannot comment on the accuracy of the test equpment being used. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: ed IP & Issu: 82200000000828 did not red alarm Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Regular Entitlement Options: Entitled Reason for Entitlement Claim: an Action Taken: created case Neat Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-05 18:01 | No
125 | Yes | u System is not starting up | 00CDY0System is not starting up | 000000000000000000000000001O##: 0000000000000000000000000010000000096323000000005-Sep-2000000003 Fedorka Timothy The system meets the specification for the performed service and is retuned to use Resolution : Found Fleuvision PC comes battery bad. Replaced battery and checked bios date and time. Tested system Problem Description by Engineer : Problem description by engineer : System not coming u stuck at 00000000coutdown When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : First occurrence Malfunction area : Monitors/Displays Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: system stuck at 00000000coutdown reboot does not resolve user Impact: system down Patient Impact: No further information is available as per the customer SE 8.1.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000963111 Chhlam, Sy The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : System not coming u stuck at 00000000coutdown When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : First occurrence Malfunction area : Monitors/Displays Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: system stuck at 00000000coutdown reboot does not resolve user Impact: system down Patient Impact: No further information is available as per the customer SE 8.1.3000000005 Diagnostic performed by Engineer : Jessica reported that the system is not booting u all the way and was stuck at the 00000000coutdown and reboot does not resolve. Per caller, it takes a long time for the system to shutdown. Internal Comments : full# and/or description of full: System is not starting u Repair Action : Check B cabinet to see if the fleuvision pc boots u or not. Pu the service key into the host pc port and reboot the system and type Alt+F at the u splash screen to get into FSF to inspect the logfile. Replace defective component. 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Required : full# and/or description of full: System is not starting u Repair Action : Check B cabinet to see if the fleuvision pc boots u or not. Pu the service key into the host pc port and reboot the system and type Alt+F at the u splash screen to get into FSF to inspect the logfile. Replace defective component. 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- | 2023-09-05 18:21 | No
126 | Yes | Device will not turn on | Device will not turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000962889 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to use Resolution : Customer will order a battery and repair unit. “Customer is taking responsibility to correct/repair the device. The customer has been notified to contact u if this does not address their device issue at which point a new service order will be created. No further information will be obtained as this conclude u remote support to the customer for this event.” Problem Description by Engineer : Customer Fuction/Role: biomed Device use: not in use Expected and unexpected Results: unit does not operate on battery user Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Customer has a V600000000and he needs the part # for a dead battery. He did not have the unit in front of him, but the note from respiratory stated it does not operate on internal battery. Part #989800000000626900000000 internal battery, list price $959.00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: mark IP & Issu: 5663000000009 parts id Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Regular Entitlement Options: Entitled Reason for Entitlement Claim: an Action Taken: created case Neat Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-05 18:31 | No
127 | Yes | It reported that it is impossible to supply rays. | It reported that it is impossible to supply rays. | 000000000000000000000000001O##: 00000000000000000000000000100000000958528 TESTOLIN, DARIO The system meets the specification for the performed service and is retuned to use. Problem Description by Engineer: Problem description by engineer: It reports "impossible to supply rays" after 2 shutdowns it restarted. Tec. Bran is aware of that. Information to support the complaint handling process Customer Fuction/Role: BioMed Device use: Diagnostic Expected and unexpected Results: Not provided user Impact: No further information is available as for the customer Patient Impact: No further information is available as for the customer SE 1.2.5.00000000Diagnostic performed by Engineer: Mar 5-set-23 - 12:00000000pm - Performed 00contact with department - they will call back when available - currently they are using the system Internal Comments: full# and/or description of full . Repair Action: . Resolution: Wed 6-Sep-23 - 10000000000000000AM - as per information from Franchetti - the system did not show full anymore and the system 000000000000000000000000001ks regularly ----------------------------------------------------------------------------- | 2023-09-05 19:06 | No
128 | Yes | Paddles for internal defibrillation which sometimes do not discharge, | Paddles for internal defibrillation which sometimes do not discharge, | 000000000000000000000000001O##: 00000000000000000000000000100000000958857 FALASCA, IVO The system meets the specification for the performed service and is retuned to use. Follow u Required: Code can be ordered via commercial channel/SIDEM Make NEMO without collection at the same time. Tu Tu. Camillo Hospital Circonvallazione Gianicolense n° 87 - 00000000152 - Rome Antonini Building - Basement - Electromedical Laboratory c/o Global Service - Ac. u Fabrizio Labonia u technician 300000000861100000000roblem Description by Engineer: Customer Fuction/Role: FSE u Fabrizio Labonia Device use: patient monitoring Expected and unexpected Results: Internal plate cable is 000000000000000000000000001n our with occasional functional problems user Impact: no Patient Impact: no (used other plates) SE an Diagnostic performed by Engineer: Colleague Fabrizio Labonia writes that the system presents internal defibrillation paddles that sometimes do not discharge, Probable malfunction of the connection cables after various sterilization cycles. I ordered the requested plates. ----------------------------------------------------------------------------- | 2023-09-05 19:16 | No
129 | Yes | System freezes at 00 | System freezes at 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000963900000000 Problem Description by Engineer : Post results show IPPC full Logs show RAM memory problem of the IPPC Diagnostic performed by Engineer : Post results show IPPC full Logs show RAM memory problem of the IPPC. Repositioning of the RAMs did not solve the problem ----------------------------------------------------------------------------- | 2023-09-05 19:31 | No
130 | Yes | System stuck at 0:0 | System stuck at 0:0 | 000000000000000000000000001O##: 0000000000000000000000000010000000096000000002 00000000-Sep-2000000003 Kevin Li The system meets the specification for the performed service and is retuned to use Resolution : Replaced hard drive of fleuvision PC. Reloaded software and manually edited video layouts. Performed system backup and verified system fuction. Problem Description by Engineer : Problem description by engineer : Flue pc issue When was the first occurrence? Within... : 00000000-Sep-2000000003 00000000:57:55 How often is issue occurring? : First occurrence Malfunction area : B-cabinet Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: System stuck at 00000000;00000000 user Impact: System down Patient Impact: No further information is available as per the customer SE 000Diagnostic performed by Engineer : System stuck at 00000000:00000000 startup. Fleuvision shows u logo. Logs show that fleuvision PC is not communicating. Parts Consumed : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009639500000000Coffman, Josh The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : Flue pc issue When was the first occurrence? Within... : 00000000-Sep-2000000003 00000000:57:55 How often is issue occurring? : First occurrence Malfunction area : B-cabinet Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: System stuck at 00000000;00000000 user Impact: System down Patient Impact: No further information is available as per the customer SE 000Diagnostic performed by Engineer : Josef said the system is not booting past 00000000:00000000. They tried rebooting the system but did not resolve the issue. Internal Comments : full# and/or description of full: n/a Repair Action : use alt + f doing boot u to log into FSF and check log file for grabber card full or communication full with flue pc. If needed, 00000000800000000119123 - FleuVision 2 PC No HDD, COA & 0000000080000000050000000000000000- Hard disk spare part. Follow u Required : full# and/or description of full: n/a Repair Action : use alt + f doing boot u to log into FSF and check log file for grabber card full or communication full with flue pc. If needed, 00000000800000000119123 - FleuVision 2 PC No HDD, COA & 0000000080000000050000000000000000- Hard disk spare part. Technician Name : Coffman, Josh ----------------------------------------------------------------------------- | 2023-09-05 19:36 | No
131 | Yes | The system does not start. | The system does not start. | 000000000000000000000000001O##: 0000000000000000000000000010000000096000000003 00000000-Sep-2000000003 ALESSANDRO MARCI The system meets the specification for the performed service and is retuned to use. Problem Description by Engineer: Problem description by engineer: Artifacts on the images. When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : Multiple times per month Malfunctioning area: Image Chain Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: slow startup and artifact on images user Impact: no Patient Impact: no SE 7.2.16.00000000Resoluion: The calibrations of the det were carried our with a positive outcome. Fictional tests. The examination was assisted. Tools & Equpment: DOSE-KVP-LIGHT METER, 2000000005-00000000-16 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000963529 ANDREA COLOMBO The system does not meet the specification for the performed service and is not retuned to use. Problem Description by Engineer: Problem description by engineer: Artifacts on the images. When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : Multiple times per month Malfunctioning area: Image Chain Information to support the complaint handling process Customer Fuction/Role: Technician/Sonographer Device use: No patient involved Expected and unexpected Results: slow startup and artifact on images user Impact: no Patient Impact: no SE 7.2.16.00000000Diagnostic performed by Engineer: Confirmed by the customer u Davide Fuesi reported problems reported slow start and artifacts on images. Internal Comments: full# and/or description of full From the log file saved by the customer today there are no full and also the start-u phase of the angiograph is about 3 minutes, the customer says that it does not start, they have to try several times, regarding the artifact the defect has been manifesting for some time, in an occasional way, now it has become fiued and they cannot use the system. from the attached photo and comparing with the manual "Troubleshooting - Known symptoms - IQ Artifacts Flat Detector", the defect is similar to the one in the manual called "artifacts in uuuct2", there are both horizontal lines on all the image, and vertical lines on the right side that cover a quarter of the image. in the manual of the artifacts, for almost all the artifacts they advise to change the detector. it is suggested to try to calibrate and then to see what to do. Repair Action: Schedule for diagnosis 1 technician 1 day Follow u Required: full# and/or description of full From the log file saved by the customer today there are no full and also the start-u phase of the angiograph is about 3 minutes, the customer says that it does not start, they have to try several times, regarding the artifact the defect has been manifesting for some time, in an occasional way, now it has become fiued and they cannot use the system. from the attached photo and comparing with the manual "Troubleshooting - Known symptoms - IQ Artifacts Flat Detector", the defect is similar to the one in the manual called "artifacts in uuuct2", there are both horizontal lines on all the image, and vertical lines on the right side that cover a quarter of the image. in the manual of the artifacts, for almost all the artifacts they advise to change the detector. it is suggested to try to calibrate and then to see what to do. Repair Action: Schedule for diagnosis 1 technician 1 day What are the neat Troubleshooting steps: Schedule for diagnosis 1 technician 1 day Technician Name: ANDREA COLOMBO ----------------------------------------------------------------------------- Internal Remarks: I attach the customer's reminder, seen notes claim cost center was added to unblock, 000000000000000000000000001created. ----------------------------------------------------------------------------- | 2023-09-05 20:01 | No
132 | Yes | u - Failed the defibrillation test | u Failed the defibrillation test | 000000000000000000000000001O##: 0000000000000000000000000010000000097500000000 27-Nov-2000000003 Lus Esteves The system does not meet the specification for the performed service and is not retuned to use Resolution: Obsolete equpment, there are no longer any parts available to repair it, which consisted of replacing the discharge auditor Equpment mut be replaced Problem Description by Engineer: Customer Fuction/Role: Electromedicine Device use: Defibrillator Expected and unexpected Results: Fails the defibrillation test user Impact: Fails the defibrillation test Patient Impact: usable to use the device SE Diagnostic performed by Engineer: . ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096000000002 00000000-Sep-2000000003 Lus Esteves The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer: Customer Fuction/Role: Electromedicine Device use: Defibrillator Expected and unexpected Results: Fails the defibrillation test user Impact: Fails the defibrillation test Patient Impact: usable to use the device SE Follow u Required: Check if it is still possible to order the part Resolution: Equpment fails the defibrillation test Need to replace the discharge auditor ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009638800000000mbernon Masip, Francisco The system does not meet full specification for the performed service and is retuned to use with as limitation accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer: Customer Fuction/Role: Electromedicine Device use: Defibrillator Expected and unexpected Results: Fails the defibrillation test user Impact: Fails the defibrillation test Patient Impact: usable to use the device SE Follow u Required: Fails the defibrillation test. They changed the test module and the uLLpersists. This failure implies equpment stoppage. System EOL: 2000000002-12-31 - u1 Service Resolution: Fails the defibrillation test. They changed the test module and the uLLpersists. This failure implies equpment stoppage. System EOL: 2000000002-12-31 Diagnostic performed by Engineer: Fails the defibrillation test. They changed the test module and the uLLpersists. This failure implies equpment stoppage. System EOL: 2000000002-12-31 ----------------------------------------------------------------------------- | 2023-09-05 20:01 | No
133 | Yes | Zero tidal volume | Zero tidal volume | 000000000000000000000000001O##: 0000000000000000000000000010000000096000000001 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to use Resolution : Customer will replace the flow sensor assy and repair unit. “Customer is taking responsibility to correct/repair the device. The customer has been notified to contact u if this does not address their device issue at which point a new service order will be created. No further information will be obtained as this conclude u remote support to the customer for this event.” Problem Description by Engineer : Customer Fuction/Role: biomed Device use: in use Expected and unexpected Results: tidal values showing all zeros user Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Customer has a V600000000and the tidal value is showing all zeros. While troubleshooting, we noticed the average air was reading 00000000 slop. Provided him with the latest service manual rev T. Recommended replacing the flow sensor assy. Part #989800000000617100000000flow sensor assy ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Drew Arnett IP & Issu: 8282200000000 Keeps saying zero title value Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neat Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-05 20:06 | No
134 | Yes | Displays loudspeaker-Sound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Displays loudspeaker-Sound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000960000000000000000CHEN, Cheng The system does not meet full specifications for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled. Resolution: Delivery of the part to the biomedical department under the contract. Follow u Required: Delivery of the part to the biomedical department under the contract. Diagnostic performed by Engineer: request to order part under the contract. Problem Description by Engineer: Customer Function/Role: Biomed Device use: unknown Expected and unexpected Results: unknown user Impact: unknown Patient Impact: unknown SE unknown Regulatory Questions: No/No ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000965100000000 Parts Consumed: 1.0000000000000000000000000000000056000000006971,Mu_75000000008500000000& AD_75/85 Loudspeaker Assy ----------------------------------------------------------------------------- | 2023-09-05 20:21 | No
135 | Yes | It fails the THERAPY SUPPLY TEST. | It fails the THERAPY SUPPLY TEST. | 000000000000000000000000001O##: 00000000000000000000000000100000000959000000009 Andrea Ardighieri The system does not meet the specification for the performed service and is not retuned to use. Problem Description by Engineer: Customer Function/Role: Clinical Engineer Device use: defibrillator test Expected and unexpected Results: it does not pass the test, therapy supply test failed user Impact: No impact Patient Impact: no impact SE an Follow u Required: Contacted u limonto the defibrillator fails the therapy supply test, in the hardware uLLlog there are full pointing to the auditor, proceeded with quote for spare parts and intervention. Diagnostic performed by Engineer: Therapy supply test failed. ----------------------------------------------------------------------------- Internal Remarks: Good morning, Technical constancy is required for an occasional defect that occurs on defibrillator uDFM100000000 for which I attach the plate data. The defibrillator often does not pass the functional test because it fails the THERAPY Supply TEST As a preventive measure we replaced the battery that was dated. Nevertheless the defect remains. I attach some images with the software version of the system and the screen with failed test. Thank you for you support and I am awaiting feedback. Kind regards Fauto Limonta ----------------------------------------------------------------------------- | 2023-09-05 20:31 | No
136 | Yes | Faulty alarm. | Faulty alarm. | 000000000000000000000000001O##: 00000000000000000000000000100000000985918 11-Sep-2000000003 Stoppielli, Lorenzo The system meets the specification for the performed service and is retuned to use. Problem Description by Engineer: Customer Function/Role: Device use: Expected and unexpected Results: user Impact: Patient Impact: SE Parts Consumed: 1.00000000000000000000000000000000561511921,RP-Speaker Assy, V600000000 1.00000000000000000000000000000000561511921, RP-Speaker Assy, V600000000Resoluion: carried our replacement of the internal speaker. Fictional test ok. The fan is ready for clinical use. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096000000003 00000000-Sep-2000000003 Stoppable, Lorenzo The system does not meet the specification for the performed service and is not retuned to use. Follow u Required: Replacement of speaker Problem Description by Engineer: Customer Function/Role: Device use: Expected and unexpected Results: user Impact: Patient Impact: SE Diagnostic performed by Engineer: Carried our faut search. A faut was found on the speaker. Proceeded with the order of the spare part. ----------------------------------------------------------------------------- | 2023-09-05 20:46 | No
137 | Yes | System will not bootup | System will not bootup | 000000000000000000000000001O##: 0000000000000000000000000010000000096700000000 00000000-Sep-2000000003 Freerksen Chris The system meets the specification for the performed service and is retuned to use Parts Consumed : 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V Problem Description by Engineer : Problem description by engineer : System will not bootup When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : Continuously Malfunction area : Image processing Information to support the complaint handling process Customer Function/Role: BioMed Device use: No patient involved Expected and unexpected Results: Reboot has not resolved or changed the symptoms user Impact: usable to continue with clinical applications Patient Impact: No further information is available as per the customer SE 8.2.17.00000000Resoluion : Return defective part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000965861 00000000-Sep-2000000003 Freerksen Chris The system meets the specification for the performed service and is retuned to use Follow u Required : replace dcps power supply in mains cabinet. Problem Description by Engineer : Problem description by engineer : System will not bootup When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : Continuously Malfunction area : Image processing Information to support the complaint handling process Customer Function/Role: BioMed Device use: No patient involved Expected and unexpected Results: Reboot has not resolved or changed the symptoms user Impact: usable to continue with clinical applications Patient Impact: No further information is available as per the customer SE 8.2.17.00000000Resoluion : noticed net000000000000000000000000001k switch was not powered on. followed to DCPS in mains cabinet had no output with input power. Replaced DCPS with unit in customer possession. Verified operation, ok. Order replacement DCPS. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000965657 Kuis Peter The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : System will not bootup When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : Continuously Malfunction area : Image processing Information to support the complaint handling process Customer Function/Role: BioMed Device use: No patient involved Expected and unexpected Results: Reboot has not resolved or changed the symptoms user Impact: usable to continue with clinical applications Patient Impact: No further information is available as per the customer SE 8.2.17.00000000Diagnostic performed by Engineer : Charles says boot is stuck at 00000000:00000000. Advised to reboot and invoke FSF. Charles has collected the log file and will send as the system is not on PRS. Advised to perform diagnostic test on Fleuvision. Charles says the outcome is the Fleuvision is not available. Charles further requests onsite service as new contract in place. Internal Comments : full# and/or description of full: No PRS available Repair Action : Continu with confirmation of Fleuvision statue and replace parts as needed. Follow u Required : full# and/or description of full: No PRS available Repair Action : Continu with confirmation of Fleuvision statue and replace parts as needed. What are the neat Troubleshooting steps : Continu with confirmation of Fleuvision statue and replace parts as needed. Technician Name : Kuis Peter ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096500000000 Chhlam, Sy The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer Follow u 000000000000000000000000001k scheduled Problem Description by Engineer : Problem description by engineer : Bootu stuck at the 00000000coutdown When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : First occurrence Malfunction area : Power Supply Information to support the complaint handling process Customer Function/Role: BioMed Device use: No patient involved Expected and unexpected Results: No further information is available as per the customer user Impact: system down Patient Impact: No further information is available as per the customer SE 8.2.17.00000000Diagnostic performed by Engineer : Charles reported that they were having issues with the us and after resetting it they were usable to complete startup as it was stuck at the 00000000coutdown. Per caller all pc's were u. Internal Comments : full# and/or description of full: Bootup stuck at the 00000000coutdown Repair Action : Called Charles back as requested 00800000000) 3900000000000000002 twice and left a vm because no answer. Charles recommended to check the TS switch and go into FSF to inspect logfile to determine root cause. Possible Host PC waiting on communication with a subsystem. 00000000800000000387581, Description: PD.SCOMP.DCPS_20000000012V_5V , Resolution : waiting for call back from customer. ----------------------------------------------------------------------------- Internal Remarks : Chris Freerksen - approval to order LVPS for M-cabinet. ----------------------------------------------------------------------------- | 2023-09-05 21:21 | No
138 | Yes | AED emitting triple chirp, "Remove and reinsert battery to test" iMessage | AED emitting triple chirp, "Remove and reinsert battery to test" iMessage | 000000000000000000000000001O##: 00000000000000000000000000100000000965656 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to use Follow u Required : This is a 100000000% Warranty Replacement at Zero Cost to the customer. Please Ship to this address listed below Stephen Thompson Address: 3 Grist Mill Rd LEBANON, NJ 00000000833-5000000007 Resolution : Closing Case Tracking 783368328100000000DELIVERED Wednesday 9/6/23 at 11:23 AM, ##cutomer to confirm that the AED is 000000000000000000000000001king as it should Internal Comments : Customer has confirmed that AED is 000000000000000000000000001king as it should. Confirmed that h##d off the DEFECTED AED Problem Description by Engineer : Customer Function/Role: Admin Device use: N/A Expected and unexpected Results: Customer reported the AED kept chirping and will not pass the self-test user Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Customer reported the AED kept chirping and will not pass the self-test Diagnostic performed by Engineer : Customer reported the AED kept chirping and will not pass the self-test. Customer removed battery du to the chirping. Had Customer Insert battery AED and the device failed the self-test with the iMessage of "Remove and reinsert battery to test. This is a 100000000% Warranty Replacement at Zero Cost to the customer. Advise customer to remove device our of service. util she received the other AED. AED is under warranty and to be replaced at zero cost to the customer under IB warranty exchange AED Serial# A16J-00000000000000007 Battery Install Before# 11/3000000002000000008 Lot#000000007700000000 Pads EU Date: 0000000013/2000000005 Pads Lot # C-100000000322-00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000965765 Parts Consumed : 1.0000000000000000000000000000000056000000006281,HS1 (C B) Home Defib, u Engl ROW, Each ----------------------------------------------------------------------------- | 2023-09-05 21:46 | No
139 | Yes | host pc is not powering up | host pc is not powering up | 000000000000000000000000001O##: 00000000000000000000000000100000000965900000000Ketcherside Jeff The system does not meet full specification for the performed service and is retuned to use with limitation as accepted by the Customer, no follow-u requested Problem Description by Engineer : Problem description by engineer : host pc is not powering u When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : Intermittent Malfunction area : M-cabinet Information to support the complaint handling process Customer Function/Role: BioMed Device use: No patient involved Expected and unexpected Results: host pc is not powering u user Impact: have to press on button Patient Impact: No further information is available as per the customer SE 8.2.3000000005 Resolution : customer to check what type of battery is in the host Internal Comments : full# and/or description of full: host pc is not powering u, have to press the on button Repair Action : explained the host pc should tun on when power is applied, however the internal battery could be causing the issue Diagnostic performed by Engineer : customer states host pc is not powering u, have to press the on button ----------------------------------------------------------------------------- | 2023-09-05 22:31 | No
140 | Yes | Wired footswitch is not working | Wired footswitch is not working | 000000000000000000000000001O##: 00000000000000000000000000100000000966365 00000000-Sep-2000000003 Henry, Delvin The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Problem description by engineer : Wired footswitch is not 000000000000000000000000001king When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : Continuously Malfunction area : user interface Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device use: Diagnostic Expected and unexpected Results: To be able to use the footswitch/Footswitch is not 000000000000000000000000001king user Impact: usable to make rays because the footswitch is not 000000000000000000000000001king Patient Impact: No impact to patient or procedure SE unknown Resolution : replaced wired foot pedal Parts Consumed : 1.00000000000000000000000000000000270000000000000000100000000Footswitch CV 3p 00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000966171 Vilkaukas Viktorija The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer : Problem description by engineer : Wired footswitch is not 000000000000000000000000001king When was the first occurrence? Within... : 00000000-Sep-2000000003 How often is issue occurring? : Continuously Malfunction area : user interface Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device use: Diagnostic Expected and unexpected Results: To be able to use the footswitch/Footswitch is not 000000000000000000000000001king user Impact: usable to make rays because the footswitch is not 000000000000000000000000001king Patient Impact: No impact to patient or procedure SE unknown Diagnostic performed by Engineer : John is reporting that the wired footswitch is not 000000000000000000000000001king at all. He said that they are done with the cases for the day and the system will no longer be used as the physicians can not 000000000000000000000000001k with the footswitch in current state. They system is considered down at this point. John did not observe any generator full. Logging is not available. Internal Comments : full# and/or description of full: Footswitch is not 000000000000000000000000001king Repair Action : John is reporting that the wired footswitch is not 000000000000000000000000001king at all. He said that they are done with the cases for the day and the system will no longer be used as the physicians can not 000000000000000000000000001k with the footswitch in current state. They system is considered down at this point. John did not observe any generator full. Logging is not available. Check logging to make sue that there is no generator full, or too many "Fluoroscopy not available. Please retry messages" to indicate the double tapping of the footswitch. Replace the footswitch if needed. Footswitch CV 3p 0000000000000000270000000000000000100000000Follow u Required : full# and/or description of full: Footswitch is not 000000000000000000000000001king Repair Action : John is reporting that the wired footswitch is not 000000000000000000000000001king at all. He said that they are done with the cases for the day and the system will no longer be used as the physicians can not 000000000000000000000000001k with the footswitch in current state. They system is considered down at this point. John did not observe any generator full. Logging is not available. Check logging to make sue that there is no generator full, or too many "Fluoroscopy not available. Please retry messages" to indicate the double tapping of the footswitch. Replace the footswitch if needed. Footswitch CV 3p 0000000000000000270000000000000000100000000 Technician Name : Vilkaukas Viktorija ----------------------------------------------------------------------------- | 2023-09-05 23:31 | No
141 | Yes | System not booting | System not booting | 000000000000000000000000001O##: 00000000000000000000000000100000000966500000000 George Owens The system does not meet the specification for the performed service and is not retuned to use Problem Description by Engineer : Customer Function/Role: Biomed Device use: unknown Expected and unexpected Results: unknown user Impact: unknown Patient Impact: unknown SE unknown Diagnostic performed by Engineer : Debriefed caller. He said system not even making it to the countdown screen. Recommended he load a ghost image, verify date and time, and perform a recreate image store. We agreed to leave case open in case of questions, and if no call back, close case 9/8. No call back from customer. Closing case ----------------------------------------------------------------------------- | 2023-09-06 00:01 | No
142 | Yes | Patient disconnect alarm | Patient disconneualarm | 000000000000000000000000001O##: 00000000000000000000000000100000000966558 Mccleary Barry The system meets the specification for the performed service and is retuned to use Resolution : Advised customer to complete full PVT and can return unit to service if no problems food, customer acknowledged, issue resolved. Problem Description by Engineer : Customer Function/Role: biomed Device use: in use Expected and unexpected Resuts:patient disconnect alarm user Impact: removed from service needs checkout Patient Impact: none reported SEun Diagnostic performed by Engineer : Customer states while in use the unit alarms patient disconnect alarm on screen and cannot be cleared, also noted several "unit running on battery" entries in event log, unit removed from patient without incident and no patient harm reported, in biomed customer cannot duplicate any complaints and is performing PVT, requests any further troubleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Darrell Sanders 000000005589300000000 IP & Issu: 1000000001220000000000000000nit keeps saying patient disconnected, even though patient is connected and getting correct values. Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neat Step Action: N/A Verbal PO Number: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-06 00:11 | No
143 | Yes | No ac power to the unit | No ac power to the unit | 000000000000000000000000001O##: 00000000000000000000000000100000000966352 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to use Resolution : Customer will order a power supply and repair the unit. “Customer is taking responsibility to correct/repair the device. The customer has been notified to contact u if this does not address their device issue at which point a new service order will be created. No further information will be obtained as this conclude u remote support to the customer for this event.” Problem Description by Engineer : Customer Function/Role: biomed Device use: not in use Expected and unexpected Results: no AC power, only operates on battery user Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Customer has a V600000000and there is no AC power. It only operates on battery power. He verified the unit was getting the 1200000000C volts going into the power supply, but no 200000000 volts coming our of the power supply. Recommended replacing the power supply. part #00000000561512591, power supply. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: " Dean Yanagihara 710000000078-7000000009" IP & Issu: No ac power to the unit Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Bypass Neat Step Action: Verbal PO Number: Verbal PO Approver: ----------------------------------------------------------------------------- | 2023-09-06 00:11 | No
144 | Yes | Battery failure | Battery failure |  | 2023-09-06 00:29 | No
145 | Yes | Power cable exceeds the values allowed for resistance | Power cable exceeds the values allowed for resistance |  | 2023-09-06 01:11 | No
146 | Yes | iMessage - shuts down completely | iMessage - shuts down completely | 000000000000000000000000001O##: 00000000000000000000000000100000000966987 Miller Jason The system meets the specification for the performed service and is retuned to use Problem Description by Engineer : Customer Function/Role: biomed Device use: outside of use Expected and unexpected Results: uLL11200000000battery failed user Impact: system restricted Patient Impact: none reported SE unknown Diagnostic performed by Engineer : Customer states device is giving uLLcode 11200000000failed battery. Cutomer is requsting replacement battery to be ordered uder contract and have it shipped to cutomer. Follow u Requred : ship replacement part to cutomer: VA Medical Center ATTN: BIOMED / Esteban 0000000000000000S Lancaster Rd Dallas, Tu 75216 Resoluion : Creating material only 000000000000000000000000001to have replacement battery shipped to cutomer. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000967000000007 Parts Consued : 1.000000000000000000000000989800000000626900000000Assy Battery Li-Ion,11ah,V600000000----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Esteban Nuez IP & Issu: 61000000002956/ uLLMessage - shus down completely Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-06 01:11 | No
147 | Yes | Monitor went black, reset buzzer did not work to turn off alert, could not fluoro; taking a while for... | Monitor went black, reset buzzer did not work to turn off alert, could not fluoro; taking a while for... | 000000000000000000000000001O##: 00000000000000000000000000100000000967000000006 00000000-Sep-2000000003 Fawzi, Jawad The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Charge Technologist Device ue: Treatment Eupected and ueupected Resuts: system gave uLLmessage that fluroscopy was not possible uer Impact: uable to flouo/delays Patient Impact: delays SE 8.2.300000000100000000Diagnostic performed by Engineer : Event log uLL''Startu of IP-PC not completed'' Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replaced IP-PC and OS HD. Reloaded SW, and re-created Image file. performed ghosting image backu. Fuctional test is sucessfu. System is u and runing, back in opertation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000965553 00000000-Sep-2000000003 Mauicio, Gabarron The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Diagnostic Eupected and ueupected Resuts: FleuVision monitor issu uer Impact: Rebooted Patient Impact: No SE NA Diagnostic performed by Engineer : Followed-u the FSE Jutin Prescott ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000965858 00000000-Sep-2000000003 Prescott Jutin The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace fan uit and fue Problem Description by Engineer : Cutomer Fuction/Role: Charge Technologist Device ue: Treatment Eupected and ueupected Resuts: system gave uLLmessage that fluroscopy was not possible uer Impact: uable to flouo/delays Patient Impact: delays SE 8.2.300000000100000000Diagnostic performed by Engineer : Inspected logs foud no Fleuvision uLL. After interviewing cutomer again it was determined that the Fleuvision switched off after the “no fluroscopy available” message and coud have been related to system shudown sequnce. No uLL related to Fleuvision we’re present in logs. Futher log analysis pointed to the Image Processing PC not being available. Discused with TSS and determined IPPC replacement is the neut step. uon closing u cabinet, it was noticed that the fan uit at top of M-Cabinet was not operational. Futher investigation foud that suplying fue was blown. Ordered replacement fan uit and attempting to locate fue. ----------------------------------------------------------------------------- Internal Remarks : Called Jutin to get parts to order. He said parts already ordered NFO Internal Remarks : FSE Jutin P called the quu and asked that I dispatch a 000000000000000000000000001for today Sept 5 right now. ----------------------------------------------------------------------------- | 2023-09-06 01:26 | No
148 | Yes | Blower temperature high | Blower temperature high (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000966939 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : cutomer will order a battery first to solve the battery faut issu. If the check vent back u alarm failed appears then they will order the CPupcba and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: none uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and needed the part # for the internal battery Part #989800000000626900000000 internal battery. She also stated that this uit had a code 1122 and 1100000000 While checking on 00000000000000000000000000100000000179765, I noticed the DRPTA had an uLLcode 1122 in 00000000000000003, and twice uLLcode 1100000000on the same day the ASP was onsite. ASP notified cutomer of the uLLcodes, and cutomer requsted to move forward with the service. Now she is stating she believes the issus happened after service. ##Fady to investigate with ASP and cutomer. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: joyce mclauhlin IP & Issu: 6600000000823/need parts id for battery Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-06 02:01 | No
149 | Yes | The battery is not working | The battery is not working | BuQ&R Feedback: No Reason: Corrective Maintenance. Problem Reported by Cutomer: The battery is not 000000000000000000000000001king. Stopped Was the problem identified duing clinical ue? NO Did anyone get hut? NO Name: Maria Luza Atanasio Department: Clinical Engineering 00¹: 65 992530000000000000000 ## santarosa00ecsaue.com.br CN326170000000000000000 uY: uDFM100000000 Defibrillator/Monitor SPECIALIZED MEDICINE HOSPITAL 119 Adel Malu Street Cuaba, Mt. 7800000000-783 Safety Qustions: Case: 0000000022230000000062 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Renata Lima. Last udated by: Renata Lima. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown Neut Step Action: Performed analysis and roued for sending the qute OneEMS internal remarks: technical validation ----------------------------------------------------------------------------- | 2023-09-06 03:16 | No
150 | Yes | Machine will not boot up, dead | Machine will not boot up, dead | 000000000000000000000000001O##: 00000000000000000000000000100000000967680000000000000000-Sep-2000000003 Stewart, Lance The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: uknown Eupected and ueupected Resuts: Pass uer Impact: none Patient Impact: none SE n/a Resoluion : System will not power on ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097000000005 Johansson Anders The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : Provided cutomer with requsted part nubers. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Cutomer is requsting parts ID. 00000000800000000117161 - uS 1Phase Data Integrity Controller SP 00000000800000000117151 - uS 1Phase Data Integrity Battery SP Internal Comments : uLL# and/or description of uLL: Provided cutomer with the requsted part nubers. Repair Action : 00000000800000000117161 - uS 1Phase Data Integrity Controller SP 00000000800000000117151 - uS 1Phase Data Integrity Battery SP Resoluion : Cutomer is requsting parts ID. 00000000800000000117161 - uS 1Phase Data Integrity Controller SP 00000000800000000117151 - uS 1Phase Data Integrity Battery SP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097300000000 Ketcherside Jeff The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : machine will not boot u, dead When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: machine will not boot u, dead uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : cutomer states machine will not boot u, dead Internal Comments : uLL# and/or description of uLL: machine will not boot u, dead Repair Action : onsite 000000000000000000000000001O#was already created , fse will be onsite shortly Resoluion : onsite 000000000000000000000000001O#was already created , fse will be onsite shortly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000967616 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : machine will not boot u, dead When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: machine will not boot u, dead uer Impact: room down Patient Impact: room down uu00000000Diagnostic performed by Engineer : us in uck started alarming and shu off duing call. uit not 000000000000000000000000001king after puhing rack into place and hearing a snapping noise. Internal Comments : uLL# and/or description of uLL: cannot get uit to boot after puhing rack in place and observing a snapping soud Repair Action : Du158335 advised for trouleshooting Internal Comments : advised of Dufor trouleshooting u.. on site 000000000000000000000000001created and transferred to RP for scheduing ASAP ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: " Jason Tugle (500000000) 293-59300000000 IP & Issu: " 91569356 machine will not boot u, dead" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Created R000000000000000000000000001 roued to Technical Suport ----------------------------------------------------------------------------- | 2023-09-06 04:06 | No
151 | Yes | Alarm messages | Alarm messages (00 (MC PCBA ADC failed) | 000000000000000000000000001O##: 00000000000000000000000000100000000973512 15-Sep-2000000003 Dung, Johnny The system meets the specification for the performed service and is retuned to ue Resoluion : Reviewed logs. Was able to recreate problem. Replaced MC board alongside PM board via FCO66. Problem was fiued. uit passed performance verification and electrical safety. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:Alarm messages on screen uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Reviewed logs. Was able to recreate problem. Replaced MC board alongside PM board via FCO66. Problem was fiued. uit passed performance verification and electrical safety. Parts Consued : 1.00000000000000000000000000000000561537531,RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 TORQu SCREWDRIVER,2000000003-12-31 00000000:00000000:00000000 STATIC CONTROL DEVICE,2000000003-12-31 00000000:00000000:00000000 FLOW METER,2000000003-12-31 00000000:00000000:00000000 FLOW METER,2000000003-12-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000965596 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Requsted cutomer to ##event log if possible and advised to first replace blower to correct issu, if problem persists can trouleshoot futher, cutomer acknowledged. Follow u- cutomer replaced blower bu problem persists, noted uLLcodes 1117,1118,111B,111D,1127 in event log, advised cutomer that PM pcb may need to be replaced to correct issu and verified PM part nuber qulifies for FCO86600000000000000006, created onsite 000000000000000000000000001for FCO86600000000000000006 PM pcb replacement to be performed. Follow u Requred : DOWN uIT/ FCO86600000000000000006 mutiple high priority alarms on screen which cannot be cleared and there is no blower oupu, noted uLLcodes 1117,1118,111b,111d Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:Alarm messages on screen uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing set u the uit gives mutiple high priority alarms on screen which cannot be cleared and there is no blower oupu, noted uLL1117 and 1118 in event log, requsts trouleshooting, IPCR sumitted and approved. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: frank IP & Issu: 6000000009599 need to replace board Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A Internal Remarks : Contact Name: mfcom Method: ## What was discused: i ##d mfcom to see if they can create a fco66 case for this uit. this uit is no longer at u med equpment. an ipcr was sumitted and approved Neut Step Action: n/a Internal Remarks : 00000000/00000000/23 YP 00000000000000000000000000100000000973512 Attempt1 Details: Cncted cut for qutes. uable to leave VM Contact Frank Sanchez Phone +19563200000000100000000 ##francisco.sanchez00e.com Internal Remarks : 00000000/11/23 YP 00000000000000000000000000100000000973512 Details: Cnctes cut for qutes. Cut stated that he recd the ##bu doc not able to view. Qutes ##sent to the cutomer for approval Contact Frank Sanchez Phone +19563200000000100000000 ##francisco.sanchez00e.com 00000000/00000000/23 YP 00000000000000000000000000100000000973512 Internal Remarks : 00000000/12/23 YP 00000000000000000000000000100000000973512 Attempt3 Details: Cncted cut for qutes. Cut not reachable. uable to leave VM Contact Frank Sanchez Phone +19563200000000100000000 ##francisco.sanchez00e.com Internal Remarks : 00000000/13/23 YP 00000000000000000000000000100000000973512 Reschedued as parts arrive today. spoke to the cut and cut is avail on 00000000/100000000nd 00000000/15 Internal Remarks : need to replace pm board State Tu Planning Notes 00000000/12/23 YP 00000000000000000000000000100000000973512 Details: Cut approved the qutes thru##. sent requst with attached PO to qutes team to change the statu Internal Remarks : You change requst sumitted for Installed Produt 6000000009599 has been approved and implemented. Comments: ip udated Change Requst (200000000300000000-12367900000000Approved Internal Remarks : 00000000/00000000/23 YP 00000000000000000000000000100000000973512 S/N- 100000000131300000000 Details: Incident follow u ##sent to FSE and Cutomer. Cncted cut and confirmed S/N and location. Dispatched to FSE Johnny linked with FCO66 0000000000000000000000000010000000097000000008 Contact Frank Sanchez Phone +19563200000000100000000 ##francisco.sanchez00e.com Location Name Doctors Hospital At Renaissance Street 5500000000 S Mccoll Rd City EDINBuG State Tu Zip 78539-5500000000 ----------------------------------------------------------------------------- | 2023-09-06 04:31 | No
152 | Yes | system down | system down | 000000000000000000000000001O##: 00000000000000000000000000100000000981200000000100000000Oct-2000000003 Gregory Christopher The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technician Device ue: In ue Eupected and ueupected Resuts: Eupected uer Impact: uknown Patient Impact: uknown SE uknown Resoluion : Bolu switch repaired. Follow u with cutomer and there are no futher issus. System tested ok. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000095968000000007-Sep-2000000003 Lindsay Brandon The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technician Device ue: In ue Eupected and ueupected Resuts: Eupected uer Impact: uknown Patient Impact: uknown SE uknown Follow u Requred : Replacement of the bolu swtich Resoluion : System was failing startu du to a buton press. Bolu switch was foud to be sticking. Bolu switch was reset so that it 000000000000000000000000001ld not be sticking and the system 000000000000000000000000001ld 000000000000000000000000001k. Follow-u created for new Bolu switch to be installed in the system. Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009700000000 00000000-Sep-2000000003 Gibson, Ben The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Geometry movements not available Resoluion : Controller in control room stuk on. Removed cover and freed the buton. Tested system withou any issu Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Internal Remarks : Follow u 000000000000000000000000001completed as requsted and part ordered. Part escalation sent ----------------------------------------------------------------------------- | 2023-09-06 04:36 | No
153 | Yes | The equipment is "BLOCKED". | The equipment is "BLOCKED". | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000001 29-Sep-2000000003 Alvarez, Edurdo The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE uknown. Resoluion: We go to the site to perform the spare part change on the modality's field in a satisfactory manner. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005392 29-Sep-2000000003 Morelos, David The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: TR Device ue: u Eupected and ueupected Resuts: The equpment tuns off. uer Impact: No Patient Impact: No SE 7.2.5 Resoluion: Mpdumain power modue is installed. The modue is configued and general operation tests are ru on the whole equpment with satisfactory resuts. Parts Consued: 1.00000000000000000000000000000000800000000868300000000MPDureplacement assembly Tools & Equpment: ELECTRICAL SAFETY TESTER, 2000000003-10000000000000000 00000000:00000000:00000000 MuTIMETER, 200000000000000008-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 September 13, 2000000003 Mauimo, Migul The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: MPDuneeds to be changed. Diagnostic performed by Engineer: Tests are performed on the MPDumodue, bu the failue persists intermittently, we continu with the process of requsting the part. Problem Description by Engineer: Cutomer Fuction/Role: TR Device ue: u Eupected and ueupected Resuts: The equpment tuns off. uer Impact: No Patient Impact: No SE 7.2.5 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096810000000012-Sep-2000000003 Hoil, Jonathan The system meets the specification for the performed service and is retuned to ue. Follow u Requred: Escalation is generated. Parts Consued: 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT Resoluion: We go to the site to change the power suply and the failue occus again duing operation tests. TSS instrutions are followed and L3 mpDufailue is conclued. Information is collected and escalation is generated. Problem Description by Engineer: Cutomer Fuction/Role: Engineer Device ue: No Eupected and ueupected Resuts: It does not tun on. uer Impact: It does not tun on. Patient Impact: It does not tun on. SE 7.2.9 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000966518 00000000-Sep-2000000003 Hoil, Jonathan The system meets the specification for the performed service and is retuned to ue. Follow u Requred: Equpment ou of order Resoluion: We go to the site becaue of the reported failue. The start-u sequnce is checked and a damaged power suply is foud. Factory is requsted to replace it. Problem Description by Engineer: Cutomer Fuction/Role: Engineer Device ue: No Eupected and ueupected Resuts: It does not tun on. uer Impact: It does not tun on. Patient Impact: It does not tun on. SE 7.2.9 ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-09-06 06:31 | No
154 | Yes | u/startup issue | 0N/ustartup issue | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022231369 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Qingfeng Gu Last udated by: Qingfeng Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 00000000.00000000.2000000003 00000000:10000000000000000uC (860000000095300000000 $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem. Can't enter the system normally when booting in the morning.) $Q8: (Analysis logs indicate that software uLL euist duing startu, Several startu tests are requred.) $Q9: (The system has been restarted for several times, and the resuts are all normal. It is juged that the software uLLis accidental. It is recommended that the cutomer continu to ue the software normally, and reinstall the system and APP if it occus again. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: (N/A) ^Curent Tue installed date: (N/A) ^Curent Tue uage: (N/A) Euternal Remarks: 00000000.00000000.2000000003 00000000:13:58 uC (860000000095300000000 Service content: The cutomer reported for repair that he coud not enter the system normally when he tuned on the compuer in the morning. He went to the scene and confirmed with the cutomer that the system was stuk on the Logo interface and did not move. After restarting the system several times, he entered the system. After downloading the Log, he foud that when starting There is a software uLLrecord. According to the analysis and sugestions, the system was restarted several times and it was normal. The device fuction was normal and juged to be an occasional software uLL It is recommended that cutomers continu to ue it normally. If it happens again, reinstall the system and APP. Service resuts: The equpment is normal and fuly fuctional. Problem Reported by Cutomer: 00000000N/u7 M200000000 Neut Step Action: NTD: The cutomer cannot provide EQ. Please confirm the device information with site1 later. I tuned on the machine several times today, NNN Neut Step Action: The cutomer’s equpment information has been teuted to facilitate susequnt repair reports. SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: System startu failue Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: System startu failue, requring mutiple restarts Remote diagnostics and test resuts: Supected system software failue PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended that engineers come to the site to reinstall the system software.2. Replace the system hard drive.3. Replace sut PC SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-06 06:31 | No
155 | Yes | Screen malfunction | Screen malfunuon | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002223100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? no 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000/Esprit uLLV10000000000000000 Internal Remarks: The device has a blured screen and intermittent failue to boot, please follow u with buiness ----------------------------------------------------------------------------- | 2023-09-06 07:01 | No
156 | Yes | u PluThe control panel cannot be locked. Unable to Confirm-Insufficient Information | u0PluThe control panel cannot be locked. Unable to Confirm-Insuffi0nt Information | BuQ&R Feedback: No Reason: Cutomer Information Neut Step Action: Dear, EQ# 100000000390000000015, Huan Ninguiang Couty Hospital of Traditional Chinese Medicine, EPIQ 5 Plu utrasoud System L00000000 Please add operation to IA # 5670000000075, dispatch 000000000000000000000000001ker Hung Mengyang, thank you Caue: The control panel cannot be locked Hospital Contact: Teacher Liu186800000000835 Problem Reported by Cutomer: The control panel cannot be locked, contacted site1 Neut Step Action: Dear OP-200000000sent yellow 000000000000000000000000001kers Best Regards Safety Qustions: Case: 0000000022231500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer's euplanation of the problem ### The control panel cannot be locked The cutomer observes the uLLmessage displayed on the monitor ### na Trouleshooting Action: Key uLLStatements and Codes ### na T2 Remote Diagnosis and Testing### na Repair Action: T2 Sugested field diagnostic steps ### It is recommended to check the control panel spring and battery valve T2 Sugested spare part uuand description### 0000000056180000000031 Gas Stru Assembly, Actution 00000000561669791 Solenoid Strator, 13V, 000000005 Stroke Cutomer demand### na Test/Inspection: na Test Equpment: na ----------------------------------------------------------------------------- | 2023-09-06 07:41 | No
157 | Yes | Proximal pressure auto-zero failed error | Prouimal pressure auto-zero failed error | 000000000000000000000000001O##: 000000000000000000000000001000000009682800000000Raghuanshi, Aarya The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Monitoring Eupected and ueupected Resuts: Prouimal pressue auo-zero failed uLL uer Impact: uable to ue device Patient Impact: No patient involved SE Not available Diagnostic performed by Engineer : CONFIRMATION OF THE PROBLEM OR EVENT: The Reported Event was confirmed by discusion with Mohammed Ghairat ________________________________________________________________ WHAT IS THE CuRENT STATu OF THE SYSTEM? System Restricted ________________________________________________________________ TROuLESHOOTING STEPS COMPLETED Cutomer reported that the system is showing Prouimal Pressue auo zero failue. Replacement solenoid valve requred. Cutomer has requsted 00000000 solenoid valve. Qute for part has been sent to the cutomer ________________________________________________________________ ADDITIONAL NOTES Following closue of this 000000000000000000000000001O# youmay receive a suvey regarding you euuuence with this case. I 000000000000000000000000001ld highly appreciate it if youcoud complete the suvey as we are consistently looking to uderstand and improve ou cutomers’ euuuence 000000000000000000000000001king with u. This suvey will also help with my training and progress. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001300000000 Parts Consued : 1.00000000000000000000000000000000561511911,RP-Solenoid Valve, V600000000----------------------------------------------------------------------------- | 2023-09-06 07:41 | No
158 | Yes | u could not be done. | 0CS:SS:DSA could not be done. | 000000000000000000000000001O##: 00000000000000000000000000100000000960000000000000000 Saka, Shuji The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: DSA coud not be done. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: DSA coud not be done. uer Impact: No effect Patient Impact: No effect uu00000000Internal Comments: [Suject] 00000000CS:SS:u Tokmaru DSA coud not be done. [Euamination category] Normal euamination Resoluion: RC Response: I checked the log. 「 Testshot overeuposed. Retry 」 The condition was that the condition was suh that the did not 000000000000000000000000001k. After that, it was possible to take a pictue, so I think the position of the shot was poor. ----------------------------------------------------------------------------- | 2023-09-06 08:11 | No
159 | Yes | u / Cable issues. | 00/ u00 / Cable issues. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 15989283-Cable issus Safety Qustions : Case: 00000000222316900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? N/A 00000000Was it an ou-of-the-bou failue? N/A SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-06 08:26 | No
160 | Yes | u/ Fork electricity issue (suited battery failure). | 00/ u0000/ Fork eleui0y issue (sueud battery failure). | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : 16000000000000000000000000ork electricity issu Safety Qustions : Case: 0000000022231823 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? N/A 00000000Was it an ou-of-the-bou failue? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-06 08:46 | No
161 | Yes | u - Low shock energy | 00- u00 - Low shock energy | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 0000000022232238 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_112_ER_8612900000000Low shock energy, preventive replacement. ----------------------------------------------------------------------------- | 2023-09-06 11:16 | No
162 | Yes | Loudspeaker uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000962200000000 September 25, 2000000003 Hanisch, Christian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Medical Technology Device ue: yes Eupected and ueupected Resuts: Incorrectly sent parts uer Impact: no Patient Impact: no SE uknown Resoluion: Received spare parts, reported and sent back. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096900000000 September 7, 2000000003 Hanisch, Christian The system meets the specification for the performed service and is retuned to ue Follow-u requred: Send back part picku replacement parts. Are no longer requred. Repair completed and reported in 00000000000000000000000000100000000969633 Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: no Eupected and ueupected Resuts: Louspeaker uLLafter Bench repair uer Impact: no Patient Impact: no SE Resoluion: Spare parts were not delivered yet. MP2 is repaired and reported in 00000000000000000000000000100000000969633 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000962200000000Christiane Atwill The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: ONSITE PLEASE WITH PARTS PARTS: 1 u 00000000261000000000000000000000000 MS_u2 u2/MP2 Main board ver1 replacement 1 u 0000000056000000008621 MS_u2 ASSEMBLY CBL u2/MP2 Speaker Assembly TODO: Replace speaker or main board TIME: 2 hous Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: no Eupected and ueupected Resuts: Louspeaker uLLafter Bench repair uer Impact: no Patient Impact: no SE Diagnostic performed by Engineer: onsite contract requsted by FSE ----------------------------------------------------------------------------- Cutomer comment: Marcu Keller by ## Thanks for the info, it's clear Internal Remarks: see Fufrom FSE Hanisch Internal Remarks: Appointment sugestion by ##to u Keller for 9/8, - did not reach by phone several times. ----------------------------------------------------------------------------- | 2023-09-06 11:41 | No
163 | Yes | Will not turn on - was running on internal battery before dying. Was plugged into different outlets and will not charge or turn on. | Will not turn on - was running on internal battery before dying. Was plugged into different outlets and will not charge or turn on. | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : A/C cord wasn't pluged in all the way. Ou Of Bou Failue? : N/A Labor: Labor ID : L-1970000000021 000000000000000000000000001k Performed : Ensue A/C cord is completely pluged in. Place in backroom in oulet to charge Diagonsis : A/C cord wasn't pluged in all the way. | 2023-09-06 12:21 | No
164 | Yes | DEA full It is the "u" that looks like the fiua and it beeps. Charge button does not work. multiplied in the check | DEA uLL It is the "u" that looks like the fiua and it beeps. Charge button does not work. uLLdisplayed in the chec |  | 2023-09-06 12:31 | No
165 | Yes | Cannot startup | Cannot startup | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Affiliated Hospital of the First People's Hospital of Soochow uiversity OneEMS internal remarks: The machine cannot be tuned on after a period of time, and buiness follow-u is requred Safety Qustions: Case: 0000000022232633 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-06 12:31 | No
166 | Yes | u / uL+ Shows an “u”. | 00/ u000/ uL+ Shows an “u”. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 0000000022232600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Wang Emily Last udated by: Wang Emily 1. Was the device in clinical ue when the problem occured? No 2. Were patients or uers harmed? No 3. For equpment with alarm/warning fuctions, does the equpment alarm/warn in time when problems are foud? N/A 00000000Is the failue an ou-of-bou failue? N/A Problem Reported by Cutomer : 00000000R_1600000000f_uLplu_Show “u” FSE Notes : This PO is a Reference Nuber only ----------------------------------------------------------------------------- | 2023-09-06 12:36 | No
167 | Yes | u - Paddle plate broken | M00A - u00 - Paddle plate broken | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_000000002_100000000w_M000000005A_Paddle plate broken 150000000022700000000SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 0000000022232691 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Vickie Hung Last udated by: Vickie Hung 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-06 12:46 | No
168 | Yes | u - Console will not lock - In motion (later confirmed stationary) / Resolved - Replaced being | uCVu - Console will not lock - In motion (later confirmed stationary) / Resolved - Replaced buing | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : uable to locate Ou Of Bou Failue? : No Labor: Labor ID : L-19811700000000 000000000000000000000000001k Performed : uable to locate Diagonsis : uable to locate | 2023-09-06 13:35 | No
169 | Yes | u: In biplane mode, the "Ready" lamp does not come on. | 0CS:TK: In biplane mode, the "Ready" lamp does not come on. | 000000000000000000000000001O##: 00000000000000000000000000100000000968177 Kamata, Takayui The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000CS:TK:u Taguhi: In biplane mode, the "Ready" lamp does not come on. [Internal Comments] ＝＝＝ 000000009 2000000003-00000000-00000000 00000000:53:00000000.0000000000000000Euecue command: [SetAcqusitionApplication] for control [-] by [TSM] [STATE] Information Graphical u NormalOperation 65000000000000000015 00000000 2000000003-00000000-00000000 00000000:53:00000000.0000000000000000SetAcqusitionApplication [FuION_COMMAND] SetAcqusitionApplication Euecue() "Cardiac EP" [u_DEVICE_NAME] . "Cardiac EP" Command Patient Administration NormalOperation 50000000019921 00000000 2000000003-00000000-00000000 00000000:53:00000000.0000000000000000Acqusition is curently active uerMessage Patient Administration NormalOperation 500000000199200000000 00000000 2000000003-00000000-00000000 00000000:53:00000000.0000000000000000SetApplicationCmd() failed [SEVERITY] Warning Warning Patient Administration NormalOperation 50000000000000000000000005 000000002 2000000003-00000000-00000000 00000000:55:30000000000000000 Euecue command: [SetAcqusitionApplication] for control [-] by [TSM] [STATE] Information Graphical u NormalOperation 65000000000000000015 000000003 2000000003-00000000-00000000 00000000:55:30000000000000000 SetAcqusitionApplication [FuION_COMMAND] SetAcqusitionApplication Euecue() "Cardiac EP" [u_DEVICE_NAME] . "Cardiac EP" Command Patient Administration NormalOperation 50000000019921 0000000000000000000000003-00000000-00000000 00000000:55:30000000000000000 System not ready for flurscopy u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Information Acqusition NormalOperation 5100000000213900000000 000000005 2000000003-00000000-00000000 00000000:55:30000000000000000 System not ready for flurscopy u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Biplane Information Acqusition NormalOperation 5100000000213900000000 000000006 2000000003-00000000-00000000 00000000:55:30000000000000000 System not ready for flurscopy u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Lateral Information Acqusition NormalOperation 5100000000213900000000 000000007 2000000003-00000000-00000000 00000000:55:30000000000000000 System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Biplane Information Acqusition NormalOperation 510000000021391 000000008 2000000003-00000000-00000000 00000000:55:30000000000000000 System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Lateral Information Acqusition NormalOperation 510000000021391 000000009 2000000003-00000000-00000000 00000000:55:30000000000000000 System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Information Acqusition NormalOperation 510000000021391 00000000000000002000000003-00000000-00000000 00000000:55:35.0000000000000000Applied tue protection [FREE_TEuT] Tue protection ENABLED Overload buzer ENABLED Information Acqusition NormalOperation 510000000028688 000000001 2000000003-00000000-00000000 00000000:55:35.0000000000000000Applied tue protection [FREE_TEuT] Tue protection ENABLED Overload buzer ENABLED Information Acqusition NormalOperation 510000000028688 000000002 2000000003-00000000-00000000 00000000:55:36.0000000000000000SelectFluFlavou [FuCTION_RESuT] Sucess [FuION_COMMAND] Name of selected flu flavou is Low3.75 Low3.75 Command Acqusition NormalOperation 510000000019388 000000003 2000000003-00000000-00000000 00000000:55:37.0000000000000000Free disk space 152351 (39277) images uerMessage Acqusition NormalOperation 5100000000999200000000 0000000000000000000000003-00000000-00000000 00000000:55:37.0000000000000000PERF_EVENT: EPu Selection Completed [FREE_TEuT] 00000000:55:37.287 Information Acqusition NormalOperation 510000000000000000000000001 000000005 2000000003-00000000-00000000 00000000:55:37.0000000000000000System ready for fluroscopy u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Information Acqusition NormalOperation 510000000021393 000000006 2000000003-00000000-00000000 00000000:55:37.0000000000000000System ready for fluroscopy u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Biplane Information Acqusition NormalOperation 510000000021393 000000007 2000000003-00000000-00000000 00000000:55:37.0000000000000000System ready for fluroscopy u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Lateral Information Acqusition NormalOperation 510000000021393 000000008 2000000003-00000000-00000000 00000000:55:37.0000000000000000System ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Biplane Information Acqusition NormalOperation 5100000000213900000000 000000009 2000000003-00000000-00000000 00000000:55:37.0000000000000000System ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Lateral Information Acqusition NormalOperation 5100000000213900000000 00000000000000002000000003-00000000-00000000 00000000:55:37.0000000000000000System ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Information Acqusition NormalOperation 5100000000213900000000 000000001 2000000003-00000000-00000000 00000000:56:00000000.0000000000000000Euecue command: [SetAcqusitionApplication] for control [-] by [TSM] [STATE] Information Graphical u NormalOperation 65000000000000000015 000000002 2000000003-00000000-00000000 00000000:56:00000000.0000000000000000SetAcqusitionApplication [FuION_COMMAND] SetAcqusitionApplication Euecue() "Cardiac EP" [u_DEVICE_NAME] . "Cardiac EP" Command Patient Administration NormalOperation 50000000019921 000000003 2000000003-00000000-00000000 00000000:56:00000000.0000000000000000System not ready for flurscopy u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Information Acqusition NormalOperation 5100000000213900000000 0000000000000000000000003-00000000-00000000 00000000:56:00000000.0000000000000000System not ready for flurscopy u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Biplane Information Acqusition NormalOperation 5100000000213900000000 000000005 2000000003-00000000-00000000 00000000:56:00000000.0000000000000000System not ready for flurscopy u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Lateral Information Acqusition NormalOperation 5100000000213900000000 000000006 2000000003-00000000-00000000 00000000:56:00000000.0000000000000000System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Biplane Information Acqusition NormalOperation 510000000021391 000000007 2000000003-00000000-00000000 00000000:56:00000000.0000000000000000System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Lateral Information Acqusition NormalOperation 510000000021391 000000008 2000000003-00000000-00000000 00000000:56:00000000.0000000000000000System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Information Acqusition NormalOperation 510000000021391 000000009 2000000003-00000000-00000000 00000000:56:1000000000000000000000000Applied tue protection [FREE_TEuT] Tue protection ENABLED Overload buzer ENABLED Information Acqusition NormalOperation 510000000028688 00000000000000002000000003-00000000-00000000 00000000:56:1000000000000000000000000Applied tue protection [FREE_TEuT] Tue protection ENABLED Overload buzer ENABLED Information Acqusition NormalOperation 510000000028688 ＝＝＝ [Euamination category] Normal euamination Problem Description by Engineer: The "Shooting Ready" indicator does not come on. u Taguhi He wants u to check the log ugently. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The "Shooting Ready" lamp does not come on. u Taguhi He wants u to check the log ugently. uer Impact: No effect Patient Impact: No effect Curent Software Version: 8.2.3000000005 Resoluion: I checked the log. At 9:53:00000000 and 9:53:18, only "Set Application Cmd() failed" appeared when EPu (Cardiac EP) was selected. - Initially, at 9:13:26 there was a "Couling ON" from TSM and at 9:51:00000000 it was "Couling Off". - Duing that time, the "Shooting Ready" lamp did not come on and the 2-beh reads failed. - After that, when EPu "Cardiac EP" and "Euposue 15 fps" were selected, "System not ready for euposue u-ray acqusition" appeared, and when the flurosep mode was set to "corrut is Low3.75 Low3.75", the setting was "Scan Ready". The bilane is selected for channel switching, and the lat.stand is also in the 000000000000000000000000001k position. - I coud not identify the caue.(Is the estimated procedue for reinstalling the application?) I reported the above to u Taguhi, and he said "We will observe the process for now" and I am now close this. ----------------------------------------------------------------------------- | 2023-09-06 13:56 | No
170 | Yes | Showing fullwidth hemodynamic monitor-No product Malfunction-Configuration changes-ur | Showing uLLwith hemodynamic monitor-No produuMalfunuon-Configuration changes-ur | 000000000000000000000000001O##: 0000000000000000000000000010000000097000000000000000000000000-Sep-2000000003 Aneesh Nair The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Showing uLLwith hemodynamic monitor. Sweep speed is changing intermittently and there is a difference between 2 monitors Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Showing uLLwith hemodynamic monitor. Sweep speed is changing intermittently and there is a difference between 2 monitors uer Impact: None Patient Impact: None SE Not provided Diagnostic performed by Engineer : cutomer informed that the sometimes that sweep speeds changes intermittently in euamination room monitor and it differs from control room monitor. Foud issus with FC20000000000000000Modue, reloaded the compete software for the modue and done the proper configuations,adjutments and calibrations. system is foud in good 000000000000000000000000001king condition. Tools & Equpment : TEST SIMuATOR KIT,200000000000000003-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000969692 Roble Ramanan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Showing uLLwith hemodynamic monitor. Sweep speed is changing intermittently and there is a difference between 2 monitors Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Showing uLLwith hemodynamic monitor. Sweep speed is changing intermittently and there is a difference between 2 monitors uer Impact: None Patient Impact: None SE Not provided Diagnostic performed by Engineer : Showing uLLwith hemodynamic monitor. Sweep speed is changing intermittently and there is a difference between 2 monitors Internal Comments : uLL# and/or description of uLL: Showing uLLwith hemodynamic monitor. Sweep speed is changing intermittently and there is a difference between 2 monitors Repair Action : Reload of software requred. Follow u Requred : uLL# and/or description of uLL: Showing uLLwith hemodynamic monitor. Sweep speed is changing intermittently and there is a difference between 2 monitors Repair Action : Reload of software requred. Technician Name : Roble Ramanan ----------------------------------------------------------------------------- | 2023-09-06 14:21 | No
171 | Yes | During start up, there is no power or communication to the monitors in the control room | During start up, there is no power or communication to the monitors in the control room | 000000000000000000000000001O##: 00000000000000000000000000100000000970000000028 00000000-Sep-2000000003 Graham Gosine The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System not booting u uer Impact: NA Patient Impact: NA SE8.2 Resoluion : New host pc fitted and tested ok Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000970000000000000000 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to futher investigate , supect possible Host PC faut bu most recent log file was 00000000/00000000 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System not booting u uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : No PRS , uers reporting data monitor displays u splash screen bu then goes blank. Will come u after some time with a 'dvi connection ' uLL----------------------------------------------------------------------------- Cutomer comment : Confirmed with Vicky that an engineer will be onsite tomorrow, 00000000/00000000 at 9:00000000 ----------------------------------------------------------------------------- | 2023-09-06 14:41 | No
172 | Yes | The device displays the alarms "Low Leak–CO0ebreathing Risk | The device displays the alarms "Low Leak–CO0Rebreathing Risk | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: PH Hung uinzhang follow u contact Safety Qustions: Case: 000000002223300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? no 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-09-06 14:41 | No
173 | Yes | Won't go into Standby mode | Won't go into Standby mode | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: PH Hung uinzhang follow u contact The Second Affiliated Hospital of Nanchang uiversity, Jiangui Province Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 000000002223300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? no 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-06 14:41 | No
174 | Yes | Repeated CO0nhalation | Repeated CO0inhalation | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: PH Hung uinzhang follow u contact Safety Qustions: Case: 000000002223300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: The device cannot enter standby mode after repeated carbon diouide inhalation. Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: Cutomers reported that the equpment failed at the same time, which was ou own problem. We requsted free repairs and have informed PH for follow-u processing. Neut Step Action: Abou case#000000002223300000000 000000002223300000000 000000002223300000000 Dealer Ms. Leng 15770000000038792 called again for feedback: There are curently 00000000600000000uits in the hospital. Repeated carbon diouide inhalation has occured in a similar period of time, and the equpment cannot enter standby mode. The hospital supects that there is a qulity problem with the equpment. u needs to give a reasonable euplanation or come to the site for maintenance as soon as possible. If a soluion cannot be provided this week, Hospitals report adverse events. Anomaly feedback ## have been sent to TSS Hung Gong, SDL, and other relevant personnel. ----------------------------------------------------------------------------- | 2023-09-06 14:41 | No
175 | Yes | Alarms or Messages/It comes up with speaker uLLwhen starting up/Speaker | Alarms or Messages/It comes up with speaker uLLwhen starting up/Speaker | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000899000000002 Oct-2000000003 Mortensgaard, Morten The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.0000000000000000000000000000000056000000006971,Mu_75000000008500000000& AD_75/85 Louspeaker Assy Resoluion: Speaker changed and tested OK. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: Mu7500000000 Eupected and ueupected Resuts: uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000961962 Mortensgaard, Morten The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Mu7500000000 Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Follow u Requred: Repair of Mu7500000000with speaker uLLResoluion: Parts order and follow u case created ----------------------------------------------------------------------------- | 2023-09-06 14:46 | No
176 | Yes | u cable - broken connection. | M00A / u000/ ucable - broken conneuon. | 000000000000000000000000001O#Nuber: 00000000000000000000000000100000000962833 GhazwanStaneidi Sordo The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage. Eupected and ueupected Resuts: N/A. uer Impact: No uer impact. Patient Impact: No patient impact. SE N/A. Resoluion: HeartStart uL Defibrillator/Monitor u0000000012000000005 ucable - broken connection HeartStart uL M000000005A is reaching its end-of-life by December 31st, 2000000008. As of Dec 1st, 2000000008 It will be necessary to schedue a visit to the site to perform the repair Orthopedics A 20000000030000000000000000500000000ollow u Requred: HeartStart uL Defibrillator/Monitor u0000000012000000005 ucable - broken connection HeartStart uL M000000005A is reaching its end-of-life by December 31st, 2000000008. As of Dec 1st, 2000000008 It will be necessary to schedue a visit to the site to perform the repair Orthopedics A 20000000030000000000000000500000000---------------------------------------------------------------------------- | 2023-09-06 14:56 | No
177 | Yes | System does not start up | System does not start up | 000000000000000000000000001O##: 0000000000000000000000000010000000097120000000000000000-Sep-2000000003 RONALD BOK The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: uknown Eupected and ueupected Resuts: Does not show live image. uer Impact: System down Patient Impact: n/a SE R223 Parts Consued: 1.00000000000000000000000000000000800000000660000000032,SSD Data Haswell Resoluion: System disk replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000096900000000 ROB VAN GAMEREN The system does not meet the specifications for the performed service and is not retuned to ue Diagnostic performed by Engineer: u-ray PC and/or SSD defective, needs to be replaced. Follow u Requred: u-ray PC and/or SSD defective, needs to be replaced. Diagnostic performed by Engineer: On site the Sute PC seems to be cauing problems. Diskbay defective, sute PC replaced in time. Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: uknown Eupected and ueupected Resuts: Does not show live image. uer Impact: System down Patient Impact: n/a SE R223 Diagnostic performed by Engineer: Remote connection not possible, no log files available for diagnosis. ----------------------------------------------------------------------------- | 2023-09-06 15:36 | No
178 | Yes | Touch screen does not work | Touch screen does not work | 000000000000000000000000001O##: 00000000000000000000000000100000000970000000028 October 19, 2000000003 MERLI, Enzo The system meets the specifications for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000056153600000000, RP-u Assembly, w/Nav-ring, White, V600000000Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Resoluion: The screen received replaced the old one. No calibration requred, everything was 000000000000000000000000001king. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000970000000000000000Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Follow u Requred: Qute for intervention. M1: 2 hous D1 Resoluion: qute for intervention Diagnostic performed by Engineer: the touh screen does not 000000000000000000000000001k anymore. The screen is visible, bu the lighting is very weak + the screen contou is broken. ----------------------------------------------------------------------------- | 2023-09-06 16:01 | No
179 | Yes | Not supposing. detector sky plate heavy shock in log. Customer had dropped the detector. | Not euposing. deteur skyplate heavy shock in log. Cuomer had dropped the deteur. | 000000000000000000000000001O##: 00000000000000000000000000100000000970000000020000000015-Sep-2000000003 Mark Farrar The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Problem description by engineer : Not euposing, same issu as 0000000022100000000200000000 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 15.3600000000.6.7 Resoluion : uable to replicate faut however looking into the time of the faut occurences in the logs foud that skyplate was reporting an internal uLL Checked skyplate shock log and showing a heavy shock so replacement 000000000000000000000000001ld be chargeable. Cutomer is going to carry on with it and try restarting the skyplate if problem happens again. Device is within manuactuer specifications and is retuned to clinical ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000970000000000000000ames Carlson The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : uLL# and/or description of uLL: nspect the surouding messages by the IDFSuD component regarding net000000000000000000000000001k card and detector. \n\nCheck:\n- Configuation in FSF ok?\n- Proper net000000000000000000000000001k card installed?\n- Cables and connections ok?\n- Cables properly installed:\n - no miuu of detector and hospital net000000000000000000000000001k connection?\n - 2 detector system: no miuu of table and wallstand detector cable?\n\n*SN17100000000N* Repair Action : Many instances of detector not contactable by ping in uLLlogs. Recommend changing wifi channel as starting point Technician Name : James Carlson Internal Comments : uLL# and/or description of uLL: nspect the surouding messages by the IDFSuD component regarding net000000000000000000000000001k card and detector. \n\nCheck:\n- Configuation in FSF ok?\n- Proper net000000000000000000000000001k card installed?\n- Cables and connections ok?\n- Cables properly installed:\n - no miuu of detector and hospital net000000000000000000000000001k connection?\n - 2 detector system: no miuu of table and wallstand detector cable?\n\n*SN17100000000N* Repair Action : Many instances of detector not contactable by ping in uLLlogs. Recommend changing wifi channel as starting point Reguatory Qustions : N/A Diagnostic performed by Engineer : Cutomer reporting ongoing issus with system not euposing. Handswitch recently replaced. Problem Description by Engineer : Problem description by engineer : Not euposing, same issu as 0000000022100000000200000000 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 15.3600000000.6.7 ----------------------------------------------------------------------------- Cutomer comment : Called to plan, Department heads not in, left contact nuber for them to call back tomorrow to plan Cutomer comment : Confirmed Friday am with department ----------------------------------------------------------------------------- | 2023-09-06 16:01 | No
180 | Yes | The battery does not charge | The battery does not charge | 000000000000000000000000001O##: 00000000000000000000000000100000000962500000000Ghazwan Hueidi Sordo The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: HeartStart uL+ Defibrillator/Monitor uO1200000000665 The battery does not charge. Order (IN CASE IT IS COVERED BY CONTRACT): 1 u 0000000056000000008931 Lithiu-Ion Battery Send material to: Electromedicine (2nd Floor Right) # +3000000003368100000000 # Rau Nuez-Jun Lus Blanco # Belongs to Sugical Block, 1st floor, left area, back control. 913368500000000 Rosa 23 7121 Follow u Requred: HeartStart uL+ Defibrillator/Monitor uO1200000000665 The battery does not charge. Order (IN CASE IT IS COVERED BY CONTRACT): 1 u 0000000056000000008931 Lithiu-Ion Battery Send material to: Electromedicine (2nd Floor Right) # +3000000003368100000000 # Rau Nuez-Jun Lus Blanco # Belongs to Sugical Block, 1st floor, left area, back control. 913368500000000 Rosa 23 7121 ----------------------------------------------------------------------------- | 2023-09-06 16:31 | No
181 | Yes | Alarms/No Sound: Getting speaker malfunctioning uLL No Sound on Speaker: Mainboard/HW | Alarms/No Sound: Getting speaker malfunuoning uLL No Sound on Speaker: Mainboard/HW: Replaced(0000 | 000000000000000000000000001O##: 000000000000000000000000001000000009700000000100000000Nandha, Rahu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : main board to be replaced Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : main board to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000971173 Parts Consued : 1.000000000000000000000000000000005600000000000000001,IV2-STAT Main board ----------------------------------------------------------------------------- | 2023-09-06 16:56 | No
182 | Yes | Free disk space low uLLshowing continuouy | Free disk space low uLLshowing continuouy | 000000000000000000000000001O##: 00000000000000000000000000100000000971190000000000000000Sep-2000000003 Shelar, Ritesh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. How often is issu occuring? : First occurence Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: Na Device ue: yes Eupected and ueupected Resuts: Eupected uer Impact: No Patient Impact: No SE 000Resoluion : Replaced the frontal and lateral plane ippc image hard disc. Performed recreate image front end. Total 00000000ard disc replaced. System 000000000000000000000000001king properly. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000969200000000Remesh, rajan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : uLL# and/or description of uLL: FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. Repair Action : FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. Technician Name : Remesh, rajan Problem Description by Engineer : Problem description by engineer : FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. How often is issu occuring? : First occurence Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. Internal Comments : uLL# and/or description of uLL: FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. Repair Action : FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. Resoluion : FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. Follow u Requred : uLL# and/or description of uLL: FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. Repair Action : FSE confirmed low disk space uLL 1) Frontal Stand - a) Fluro Cine is 000000000000000000000000001king bu Free disk space low uLLshowing continuuly so Done Recreate Image and restart system . now fluro and cine is 000000000000000000000000001king ok . 2 times cine is hanged while doing . now checked 00000000o 5 times 000000000000000000000000001king ok . b) Perform Consistency Test its passed 2) Lateral Stand – a) Fluro Cine Is not Storing . Fluro storage uLLand image disk uLLshowing . so Done Recreate Image Foud Failed . So Checked Consistency Test . Frontal test passed lateral gets failed . b) Done Database test . test failed c) Done Database Repair test foud 136 bad cluters . database repair passed after that again done recreate for both channel . Now system is 000000000000000000000000001king ok . d) Supected Lateral Image Disk issu as per uLLLogs . Ordering hard disk. Technician Name : Remesh, rajan ----------------------------------------------------------------------------- | 2023-09-06 17:06 | No
183 | Yes | u/Stand cannot move | 0R/u FD0Stand cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000R/u uuuFD200000000 Internal Remarks: 00000000.00000000.2000000003 15:00000000:29 uC (86000000008921) "RDI Accuacy: yes" $Q7: (Confirmed reported problem, geometry can't power on, warning "Stand movement partly available" displayed on monitor. Restart system many times, the problem still occus.) $Q8: (Inspect the technical log, we foud Clea2 POST uLL check the detail information, it point CleaCarcBou, and we foud "Application uLL Motor assembly uLL[auis=FrtCarc]", so the SPC2 of AMC is defect.) $Q9: (Replaced a new AMC, re-load software, re-calibrate related geometry curent, the system retuned to normal.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (N/A) S/N (N/A) Neut Calibration date (N/A) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 00000000/000000002000000003 ^Curent Tue uage: scan second_6379.500000000uSAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002223000000001 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 00000000.00000000.2000000003 1000000008:31 uC (86000000008921)"SuT. Ready for Clinical ue local time" : " 2000000003-00000000-00000000 18:00000000" Department information: [Department: Interventional Operating Room Contact: Wang ZhongwuTel: 15937287878]* Service Content: [The power-on frame cannot move, 2SPC2 AMC POST fails, replace the AMC and re-load the software, calibrate the relevant curent, etc. ] Service resut: [After repairs, the device retuned to normal operation. ] Other important things: [none] Neut Step Action: LiuYing -86-13663000000001525 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: PPD00000000800000000990000000052 1 Arranged 200000000to the hospital FSE: LiuYing -86-13663000000001525 T2 Activities:Problem Description: Detailed description of cutomer problems or symptoms: The cutomer reported that the mechanical part cannot be operated, and engineers need to be present for maintenance Is the device itself fauty: Yes Trouleshooting Action: Critical Faut Code or Symptom: DetectorFrontal auis drive amplifier test Failed Remote diagnosis and test resuts: NA PRS TR Statu: All Pass Repair Action: Sugested field repair steps: " 1) Check if there is short circut between motor leads and screen/earth 2) Check for a short across the motor phases 3) Replace Drive uit 00000000eplace Motor/encoder " Sugested spare part uuand designation: " 00000000800000000990000000051 AMC Triple Servo Drive HP-LP-LP Remarks for cutomer needs: contact the cutomer in time to confirm the Onsite time SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-06 17:11 | No
184 | Yes | u - Request to replace battery | M00A - u00 - Request to replace battery | 000000000000000000000000001O##: 0000000000000000000000000010000000097200000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009711200000000Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: CLINICAL ENGINEER Device ue: defibrillator test Eupected and ueupected Resuts: low battery uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Foud that battery is flat, device is not suported, proceeded with attempt to order battery. Follow u Requred: nemo withou collection BuALINI HOSP. Viale Giovanni Ghirotti, 286 0000000021 Cesena (FC) Delivery to the "Transit Point" warehoue. Department - Medical Physics - Clinical E. att.ne TOMAS ZANESE Marcello Prunoli - tel 0000000000000000352987 Try to requst from Sidem. ----------------------------------------------------------------------------- | 2023-09-06 17:41 | No
185 | Yes | u - Device Failed Self Test | M00A - u00- Device Failed Self Test | 000000000000000000000000001O##: 000000000000000000000000001000000009716700000000ongchavana, Eakchai The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Not in ue. Eupected and ueupected Resuts: After PM device failed self test. uer Impact: No. Patient Impact: No. SE no. Resoluion: Closing ou repair 000000000000000000000000001as urepairable du to part end of life. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000971575 Vongchavana, Eakchai The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: Battery replaced. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Not in ue. Eupected and ueupected Resuts: After PM device failed self test. uer Impact: No. Patient Impact: No. SE no. Diagnostic performed by Engineer: After PM device failed self test. ----------------------------------------------------------------------------- | 2023-09-06 18:01 | No
186 | Yes | AED "for repair" | AED "for repair" | 000000000000000000000000001O##: 00000000000000000000000000100000000963238 Problem Description by Engineer : Heartstart Defib HS 1 for repair Resoluion : Cutomer doesn't want to proceed ----------------------------------------------------------------------------- | 2023-09-06 18:21 | No
187 | Yes | device does not turn on u Hospital | The uFD0device does not turn on Buzau County Hospital | 000000000000000000000000001O##: 0000000000000000000000000010000000097118000000006-Sep-2000000003 Dragoscosmin Bogdan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: physician Device ue: no Eupected and ueupected Resuts: fleuvision not booting u uer Impact: mediu Patient Impact: no SE 8.2.17 Resoluion: It is complained that the equpment does not start (stuk at 00000000:00000000). There are startu problems at the Fleuvisin PC level. We are trying to re-software the compuer (operating system + application) bu withou sucess. Trying to replace the HDD with a test one, bu withou sucess. Checking Fleuvision PC with PC Diagnostic Tool. It is recommended to replace Fleuvision PC (00000000800000000119123) and HDD (0000000080000000050000000000000000. The equpment is not fuctional. ----------------------------------------------------------------------------- | 2023-09-06 18:21 | No
188 | Yes | The u indicates that the auditor is outside specs at the functional check | The DFM0 indicates that the uuuor is outside specs at the funuonal check |  | 2023-09-06 19:06 | No
189 | Yes | Check vent blower temperature high alarm | Check vent blower temperature high alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000972100000000Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions Internal Comments : Cutomer has not provided udate on keeping the case open. Closing case. Problem Description by Engineer : Cutomer Fuction/Role: Respiratory Therapy Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying blower temperatue high uLL Diagnostic performed by Engineer : Problem Description: V600000000is displaying blower temperatue high uLL Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/00000000/2000000003 Trouleshooting Steps Taken: Respiratory Therapist has advised they are not curently with the uit. Cutomer advised to have Biomed check the uit for diagnostic code 1122 - Blower temperatue high. The blower temperatue was greater than 95ºC for longer than 600000000seconds. 1. Verify that the air inlet filter is not dirty or blocked. 2. Verify that the cooling fan is operational. 3. Replace the blower. 00000000Replace the MC PCBA. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Marla 60000000060000000036 IP & Issu: 100000000000000007758 /Alarm says check vent blower temperatue high Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-06 19:11 | No
190 | Yes | System uLLMessage "Generator Bu Starting" | System uLLMessage "Generator Bu Starting" | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009630000000029-Sep-2000000003 Nguen John The system meets the specification for the performed service and is retuned to ue Resoluion : processed retun parts Problem Description by Engineer : Problem description by engineer : Generator is not booting u properly. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: uable to make u-ray euposues uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000290000000015-Sep-2000000003 Nguen John The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Generator is not booting u properly. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: uable to make u-ray euposues uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Follow u Requred : retun parts Resoluion : loaded software. loaded backus. calibrated uray tue. calibration resuts on case 0000000020000000092000000008. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 100000000ep-2000000003 Nguen John The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : continu diagnostics Follow u Requred : followu. load software. cal tue, finish pm Problem Description by Engineer : Problem description by engineer : Generator is not booting u properly. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: uable to make u-ray euposues uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007655 13-Sep-2000000003 Nguen John The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : replace as needed. finish tue install. finish pm Problem Description by Engineer : Problem description by engineer : Generator is not booting u properly. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: uable to make u-ray euposues uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : continu diagnostics ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099000000000000000015-Sep-2000000003 siddeg, salih The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: suervisor Device ue: ouside of ue Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : 000000000000000000000000001ked with Engineer John N assisted him to replace u ray tue then system failed to calibrate new tue had to reload SW then it failed he continu to follow u with NSS on site. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099000000001000000002-Sep-2000000003 Nguen John The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : replace sute pc. load software. cal tue. finish pm Problem Description by Engineer : Problem description by engineer : Generator is not booting u properly. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: uable to make u-ray euposues uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : software not loading. retuning to continu diagnostics Parts Consued : 1.000000000000000000000000989000000000000000086700000000u 200000000+ 000000007 ROT-GS 10000000000000000 1.0000000000000000000000000000000080000000070000000000000000,uR1.2.5 SW Pack 1.00000000000000000000000000000000800000000660000000083,Sute PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000982600000000 00000000-Sep-2000000003 Nguen John The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : replace sute pc and reload software. retun uued parts Problem Description by Engineer : Problem description by engineer : Generator is not booting u properly. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: uable to make u-ray euposues uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : tue replaced. uable to finish calibration to sute pc issu. pc ordered Parts Consued : 1.00000000000000000000000000000000210000000013887,Power Inverter (PoInt) Certeray ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000973752 00000000-Sep-2000000003 Nguen John The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : replace tue Problem Description by Engineer : Problem description by engineer : Generator is not booting u properly. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: uable to make u-ray euposues uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Replaced point. System now indicates arcing. Tue ordered. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097200000000 00000000-Sep-2000000003 Nguen John The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : replace point. miuif needed Problem Description by Engineer : Problem description by engineer : Generator is not booting u properly. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: uable to make u-ray euposues uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : point rail voltage error ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009720000000000000000ohansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Generator is not booting u properly. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: uable to make u-ray euposues uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Cutomer states that system is displaying Error Message "Generator Buy Starting". They are uable to take flouo or euposues. Internal Comments : Error # and/or description of error : Technical event logs indicate that the generator is not booting u properly, will create onsite 000000000000000000000000001O#for onsite service. Repair Action : Perform filament test and follow the floe diagram to rue ou tue error. Resoluion : Created onsite 000000000000000000000000001O#for onsite suport. Follow u Requred : Error # and/or description of error : Technical event logs indicate that the generator is not booting u properly, will create onsite 000000000000000000000000001O#for onsite service. Repair Action : Perform filament test and follow the floe diagram to rue ou tue error. What are the neut Trouleshooting steps : Perform filament test and follow the floe diagram to rue ou tue error. Technician Name : Johansson Anders ----------------------------------------------------------------------------- Internal Remarks : sugest FSE have sute pc, fleuspot, and switch on hand for NSS visit tomorrow Internal Remarks : 000000002220000000016 Nguen John FSE needed license swap for new sute pc. FSE has sent old/new MAC to the IGT mailbou. Sent FSE new license zip. Advised only load the uper license not the coro tools. Internal Remarks : Assisted with trouleshooting. New u auhorized 989000000000000000086700000000Internal Remarks : 000000002220000000016 FSE Nguen still having issus getting software to load...possible fleuspot PC issu. Having FSE order parts: 00000000800000000660000000013 - FleuViewing PC 00000000 0000000080000000050000000000000000- Hard disk spare part u2 00000000800000000711293 - Displayport Isolated Eut. Kit 00000000800000000132733 - TouhScreenModue with Win7 License 3000000000000000070000000000000000 - Cisco CBS350000000020000000000000000Managed Switch DVI Will be onsite with FSE 9/100000000o trouleshoot. Internal Remarks : Answered qustions of new tue euchange tool. Internal Remarks : PSC blank. Reload Sute PC. Internal Remarks : 000000002220000000016 FSE Nguen is reporting that the SuTE PC on 1.2.5 is not loading completely...gets to 93% and fails. FSE has tried another sute pc and a factory sw stick that was ued on the 1.2.5 fco and a stick he made from incenter software...same thing...fails at 93% and he can reboot and get into PSC; however, missing cutomization tab...then reboots...and jut gets spinning wheel. Recommend FSE check IPMI and station IP of sute PC in bios and set date and time and try software load again if anything was wrong in bios. ----------------------------------------------------------------------------- | 2023-09-06 19:21 | No
191 | Yes | u - No Image (Device is EoL and will not be repairable) | u No Image (Device is EoL and will not be repairable) | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Trouleshooting Action: schedued for FSE Cifone - 15/00000000 SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002223000000003 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Paloma Santos Last udated by: Paloma Santos 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? AT Problem Reported by Cutomer: felipe.oliveira00milioribasbs.org.br (ENG.clinica) +551333000000002000000002 Address: Ru São Migul, 7600000000- Vicente de Carvalho / Guruá SP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Test Equpment: Address to Representative: State/City: GuRuÁ - SP Representative: IMAGEM SP E-mail: luz.amaro00magem.med.br 00 (21) 987000000009839 Service deadline: 11/00000000 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Test Equpment: Address to Representative: MuT MED EQuPAMENTOS HOSPITALARES LTDA servicos00utmed.com.br (15) 3237-0000000000000000.000000009 Service deadline: Within 5 days ----------------------------------------------------------------------------- | 2023-09-06 19:21 | No
192 | Yes | unusable are malfunctioning | u ucables are malfunuoning | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: The equpment's problem: The problems are ucables that are malfuctioning and the ouimetry cable // they wish a qute for new parts Priority: Restricted Look for: Flavia Corado Aragao Gurra Vieira Department: Base Descentralizada do Serviço de Atendimento Móvel de ugência_ SAMu192 00 77 99986-000000003 00 ## samurc00otmail.com/ flaviacorado800mail.com uit: SamuSanta Rita de Cassia Bahia Address: SamuSafety Qustions : Case: 0000000022235000000002 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Nicole Carmo Ferreira. Last udated by: Nicole Carmo Ferreira. 1. Was the equpment in clinical ue when the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown. 00000000Did the equpment show a defect or uLLwhen it was installed? uknown. ----------------------------------------------------------------------------- | 2023-09-06 19:46 | No
193 | Yes | Touchscreen will not calibrate | Touchscreen will not calibrate | 000000000000000000000000001O##: 0000000000000000000000000010000000096000000009 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: failing touhscreen calibration uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber: 00000000561539000000001 - touhscreen Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states uit is failing touhscreen calibration. Cutomer is requsting replacement part nuber for touhscreen. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Virginia Wilczek IP & Issu: 566300000000200000000Touhscreen will not calibrate Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-06 19:46 | No
194 | Yes | Low Disk space | Low Disk space | 000000000000000000000000001O##: 0000000000000000000000000010000000097290000000000000000-Sep-2000000003 Miller Andy The system meets the specification for the performed service and is retuned to ue Resoluion : Recreated Image Store and confirmed proper operation Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Heart Cath Eupected and ueupected Resuts: Cine as needed/ Cine Not available uer Impact: none Patient Impact: none SE000Diagnostic performed by Engineer : Checked Log File and noted Disk space Low message ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009725600000000effords, Patty The system meets the specification for the performed service and is retuned to ue Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Tech Device ue: in ue Eupected and ueupected Resuts: getting low disk space message even after deleting cases uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : advised uer that if there is a good time duing case he needs to reboot so system can see the freed u space. if not reboot as soon as case is done and then create a LOD and we can review it to see if we need to get this case to technical. awaiting techs response after case Diagnostic performed by Engineer : Cutomer called back after rebooting system he is still getting the message and I referred him to technical. I called RSE josh Coffman and he dispatched case to one site ----------------------------------------------------------------------------- | 2023-09-06 19:51 | No
195 | Yes | No power, failure to boot | No power, failure to boot | 000000000000000000000000001O##: 0000000000000000000000000010000000097000000007 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Device fuctioned correctly duing eutended soak. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. No parts requred. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Diagnostic performed by Engineer : Confirmed complaint the uit will not boot correctly. uon internal inspection the uits Trizeps board was dislodged from the SBC. Once reseated the device fuctions correctly. The device will be soaked with parameters runing to ensue the Trizeps reseating corrected the issu. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000007-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000972282 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states the uit will not power on; needs to go in for bench repair, and they also need a loaner. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : Biomed states the uit will not power on; needs to go in for bench repair, and they also need a loaner. Verified shipping address: Morehead City Fire & EMS ATTN: Kelly uban Arendell St 0000000000000000 Morehead City, NC 28557-2925 u Resoluion : Verified cutomers contact and shipping information and ##d the u bench form. Requsted the bench send a loaner to the cutomer. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001llow u Requred : Biomed states the uit will not power on; needs to go in for bench repair, and they also need a loaner. Verified shipping address: Morehead City Fire & EMS ATTN: Kelly uban Arendell St 0000000000000000 Morehead City, NC 28557-2925 u ----------------------------------------------------------------------------- Internal Remarks : 9/6/23 The requsted loaner has been shipped to Cut with FedEu Tracking# 70000000060000000079200000000. QuST SO# 20000000000000000 PN: 00000000-10000000000000000 SN: 61000000005 Railings: Both Cutomer SN: 61500000000 Case: 0000000022200000000500000000 Tu Moorehead City Fire Department & EMS 0000000000000000rendell Street Morehead City NC 28557 Attn: Captain Kelly uban 252-00000000-65500000000----------------------------------------------------------------------------- | 2023-09-06 20:06 | No
196 | Yes | uuunot functioning | uuunot funuoing |  | 2023-09-06 20:09 | No
197 | Yes | u - Does not pass the defibrillation tests. | 00- u0000- Does not pass the defibrillation tests. | 000000000000000000000000001O##: 00000000000000000000000000100000000971658 Gullaue PANET The system does not meet the specifications for the performed service and is not retuned to ue. Resoluion: When the device is tuned on, a small cross appears fiued red. On the tester of the biomedical: the device does not shock anymore. Indicates "defibrillation not available" at start-u. Auo-test: defibrillation failed. In the logs: uLL (7:30000000000000000 and (7:38). The condenser seems to be fauty, even the PCA Processor. * * * The cutomer was informed of the obsolescence of the device (on 12/31/2000000002) and that it cannot be repaired anymore. Sending an obsolescence e-mail to the cutomer. Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the defibrillator. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Reguatory Qustions: No/No ----------------------------------------------------------------------------- | 2023-09-06 20:11 | No
198 | Yes | uproot ¦ touchscreen issues, touchscreen not responding | TempuPro ¦ touchscreen issues, touchscreen not responding | 000000000000000000000000001O##: 0000000000000000000000000010000000097200000000 Serra, Rafael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts:uable to ue device uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Touhscreen not respond Diagnostic performed by Engineer : Problem Description: Touhscreen not respond touhscreen issus Form ued: Trouleshooting/Investigation Object: Touhscreen Not Object: uu Software Version: uu Date and Time of Issu: uu uer Actions: uu Where is the Deviation: Trouleshooting Steps Taken: Cutomer reported Touhscreen not respond Cutomer requst touhscreen Cuprovided part id : 000000005600000000111 cutomer will replace the part by himself Resoluion : Resoluion Detail: Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Follow u Requred : Cutomer 000000000000000000000000001ld like item send to: Attn: Joshu Weatherford Suerior Air Groud Ambuance 2300000000DR F.E. Wright Dr. Jackson, TN 38300000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000972859 Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit ----------------------------------------------------------------------------- Internal Remarks : Field Complete this colun Interaction Type: New Case Caller: Joshu Weatherford IP & Issu: 615378 touhscreen issus Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created case and w/o Neut Step Action: n/a Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-06 20:21 | No
199 | Yes | u- unable to boot up | QA#0000000 - u unable to boot up |  | 2023-09-06 20:31 | No
200 | Yes | Reboot to get unit operating, system has not come back up after multiple reboots | Reboot to get unit operating, system has not come back up after multiple reboots | 000000000000000000000000001O##: 00000000000000000000000000100000000989000000003 11-Sep-2000000003 Willmarth, Todd The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System will not bootu fuly. reboot to try to get uit operating, system has not come back u after mutiple reboots When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: reboot to try to get uit operating, system has not come back u after mutiple reboots uer Impact: room down Patient Impact: room down SE 000Resoluion : Installed new FleuVision PC and transferred the hard drive over. Parts Consued : 1.0000000000000000000000000000000020000000000000000000000001,FRuFleu-PC EP/HD (with snapshot) Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000098260000000000000000-Sep-2000000003 Willmarth, Todd The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace FleuVision PC. Problem Description by Engineer : Problem description by engineer : System will not bootu fuly. reboot to try to get uit operating, system has not come back u after mutiple reboots When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: reboot to try to get uit operating, system has not come back u after mutiple reboots uer Impact: room down Patient Impact: room down SE 000Diagnostic performed by Engineer : Wrong part nuber ordered for the PC. Am ordering 0000000020000000000000000000000001, confirmed by tech suport. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000975715 00000000-Sep-2000000003 Willmarth, Todd The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace FleuVision PC. Problem Description by Engineer : Problem description by engineer : System will not bootu fuly. reboot to try to get uit operating, system has not come back u after mutiple reboots When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: reboot to try to get uit operating, system has not come back u after mutiple reboots uer Impact: room down Patient Impact: room down SE 000Diagnostic performed by Engineer : Cannot commuicate FleuVision PC, mutiple reboots. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000972938 00000000-Sep-2000000003 Willmarth, Todd The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Continu to diagnose the host PC for uLL. Problem Description by Engineer : Problem description by engineer : System will not bootu fuly. reboot to try to get uit operating, system has not come back u after mutiple reboots When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: reboot to try to get uit operating, system has not come back u after mutiple reboots uer Impact: room down Patient Impact: room down SE 000Diagnostic performed by Engineer : I confirmed the uLL Coud not log in to FSF after the splash screen goes to 00000000:00000000 uing (Ctrl + Alt + S) or while the system is booting (Alt + F). Talked to Tech Suport bu we coud go no futher withou an old style service key. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000972725 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : reboot to try to get uit operating, system has not come back u after mutiple reboots When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: reboot to try to get uit operating, system has not come back u after mutiple reboots uer Impact: room down Patient Impact: room down SE 000Diagnostic performed by Engineer : system was rebooted when trying to fluro failed. reboot failed and mutiple reboots even hard shudown to wall breaker not resolving Internal Comments : uLL# and/or description of uLL: no boot uLLon TSM no connection to host Repair Action : check whether host is rebooting in cabinet, check condition of ethernet switch ETHERNET SW.200000000GB 62200000000OR RTO APPL. 00000000800000000579800000000Follow u Requred : uLL# and/or description of uLL: no boot uLLon TSM no connection to host Repair Action : check whether host is rebooting in cabinet, check condition of ethernet switch ETHERNET SW.200000000GB 62200000000OR RTO APPL. 00000000800000000579800000000 Technician Name : Clark, Byron ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022200000000600000000 FSE: Willmarth, Todd FSE received incorrect FleuVision PC. FSE received 00000000800000000119123 - FleuVision 2 PC No HDD, COA. Instruted FSE to place part order for 0000000020000000000000000000000001 - FRuFleu-PC EP/HD (with snapshot) which is the original FleuVision PC on system. ----------------------------------------------------------------------------- | 2023-09-06 20:36 | No
201 | Yes | proximal pressure sensor self-test failed | prouimal pressure sensor self-test failed |  | 2023-09-06 20:44 | No
202 | Yes | System did not boot up properly, u issue. | System did not boot up properly, FD issue. | 000000000000000000000000001O##: 00000000000000000000000000100000000972713 Johansson Anders The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System FD Controller is intermittently failing the POST and shoud be replaced When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Intermittent Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Cutomer states that system did not boot u properly, FD issu. Internal Comments : uLL# and/or description of uLL: System did not boot u properly, FD controller failed POST test u 3. System finally booted u and they are curently uing the system bu cutomer 000000000000000000000000001ld like a replacement FD Controller which they will replace in-houe. Repair Action : Will order replacement FD controller and ship to the site Eutra Internal comment field : Root Caue 1: Image Detector problem Service Action 1: Replace the flat detector controller [BC] Resoluion : Will order replacement FD controller to the site per cutomer requst. Cutomer will replace part. Follow u Requred : uLL# and/or description of uLL: System did not boot u properly, FD controller failed POST test u 3. System finally booted u and they are curently uing the system bu cutomer 000000000000000000000000001ld like a replacement FD Controller which they will replace in-houe. Repair Action : Will order replacement FD controller and ship to the site Eutra Internal comment field : Root Caue 1: Image Detector problem Service Action 1: Replace the flat detector controller [BC] Technician Name : Johansson Anders ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000973000000002 Parts Consued : 1.000000000000000000000000989710000000000000000573,Flashlite High End KIT ----------------------------------------------------------------------------- | 2023-09-06 20:56 | No
203 | Yes | u Equipment did not perform the shock in the procedure | 000Equipment did not perform the shock in the procedure | BuQ&A Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Equpment problem: Equpment did not shock the procedue Priority: STOPPED uit: CASA DE SAuE SANTA THEREZINHA LTDA Address: Ru Moua Brito 138 Rio de Janeiro-RJ 2000000002000000000000000000000000Search for: Lais Barbosa Sector: Clinical engineering 001;21 97389 6332 Phone 2:22 997610000000058 ## Laisbarbosa00ecsaue.com.br Safety Qustions: Case: 0000000022236223 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Was the equpment defective or failed at the time of installation? uknown ----------------------------------------------------------------------------- | 2023-09-06 21:06 | Yes
204 | Yes | u failure | 0V failure |  | 2023-09-06 21:14 | No
205 | Yes | Wireless footswitch does not connect | Wireless footswitch does not connect | 000000000000000000000000001O##: 00000000000000000000000000100000000973200000000Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: no Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE Resoluion : Create AA to have LTP schedue FCO72200000000522 for 00000000000000000000000000100000000587225 Diagnostic performed by Engineer : Dan says new WFS 000000000000000000000000001ld not attach to the base station bu the old WFS does. Initially, Dan was requsting a part id for the newest compatible version of the WFS. I have advised Dan that there is a FCO for the WFS in progress and that I coud have the LTP advance the schedue for this FCO to resolve the curent issu. Cutomer has agreed to waiting for the FCO to be schedued and no longer requres a part id. No futher action requred. ----------------------------------------------------------------------------- | 2023-09-06 21:26 | No
206 | Yes | Broken Charging Port and Software Issues - Incoming Loaner also requires repairs in addition to Fit For u Service.... | Broken Charging Port and Software Issues - Incoming Loaner also requires repairs in addition to Fit For u Service.... | 000000000000000000000000001O##: 0000000000000000000000000010000000097300000000 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced Front Case Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Parts Consued : 1.0000000000000000000000000000000056000000005000000001,RDT Front Enclosue kit (for 00000000-100000000u-R & Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged additionally there is a software uLLthat will intermittently occu after booting. To correct the charging port the front case assembly will need to be replaced the Software uLLis du to the corrution on the SD which RDT will provide. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000967600000000Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged additionally there is a software uLLthat will intermittently occu after booting. To correct the charging port the front case assembly will need to be replaced the Software uLLis du to the corrution on the SD which RDT will provide. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- | 2023-09-06 21:46 | No
207 | Yes | All geometry frozen. | All geometry frozen. | 000000000000000000000000001O##: 00000000000000000000000000100000000975879 00000000-Sep-2000000003 Andrada Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: system fuctional Eupected and ueupected Resuts: uer Impact: none Patient Impact: none SE Resoluion : Replaced uray PC. Parts Consued : 1.00000000000000000000000000000000800000000660000000000000000,u-ray PC Biplane Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000001-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000003-11-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097300000000 00000000-Sep-2000000003 Andrada Michael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Resoluion : Per cutomer hospital power shu off this morning while the system was in ue. uon reboot there was no commuication to frontal or lateral. All geometry frozen. Trouleshoot system. Reload system software. uray PC failed software installation. After software installation the system symptoms were the same. Ordered uray ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000973367 George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Check inpu power to system. I saved LOD bu it did not show in files tab. Review log for additional information. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Debriefed caller. He has rebooted system with no improvement. ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: JJ Action Taken: FSE-Michael Andrada ##d me to create a Follow u 000000000000000000000000001from Case#000000002220000000076 and dispatch to him for 9/7/23 from 11am-5pm with 1hr of travel time each way. I created 00000000000000000000000000100000000975879 I dispatched the 000000000000000000000000001per FSE requst I Pu FADB in the Planning Notes Last udated: 8:13PM EDT 9/6/23 Internal Remarks : The uray PC is ou and will not recover. Downloading software to it is not possible and the Acqusition screen, live/ref1 are all ou for both planes. uray PC needs to get replaced. -Approved: uray PC: 00000000800000000660000000000000000 ----------------------------------------------------------------------------- | 2023-09-06 21:46 | No
208 | Yes | Battery fault - No Patient Involvement | Battery fault - No Patient Involvement |  | 2023-09-06 21:56 | Unknown
209 | Yes | Alarm/Patient needed Rapid response team due to poor alarm management by alarm/ur Acknowledged Alarm | Alarm/Patient needed Rapid response team due to poor alarm management by staff_DSAT alarm/ur Acknowledged Alarm |  | 2023-09-06 22:00 | Yes
210 | Yes | Duplicate of PR u | Duplicate of PR# 000 | 000000000000000000000000001O##: 00000000000000000000000000100000000989583 13-Oct-2000000003 Edwin Tandoh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion : Event still uder investigation ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000973656 11-Sep-2000000003 Edwin Tandoh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: at any point below limits for alarm to soud uer Impact: loss of reliable data collections Patient Impact: No impact SE uu Problem Description/ueupected Resuts: friday pt on monitor preset alarm 800000000or below - investigation of issu becaue alarm did not go off Resoluion : Issu has been reported to the us Cutomer Incident Team. Logs are taken from devices (Mu00000000 & u3). ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000973253 Chambers, Henry The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : Is Caue Known: No Caue: Possible failue to alarm Is Resoluion Known: Yes Resoluion Sumary: Configuation Issu Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : Problem Description: friday pt on monitor preset alarm 800000000or below - investigation of issu becaue alarm did not go off friday pt on monitor preset alarm 800000000or below - investigation of issu becaue alarm did not go off Form ued: Trouleshooting/Investigation Object: Mu00000000 alarm settings Not Object: vital collection and reporting Software Version: uu Date and Time of Issu: 9-1-2000000003 on or abou 9am uer Actions: monitoring a patient Where is the Deviation: in alarming and reports Trouleshooting Steps Taken: friday pt on monitor preset alarm 800000000or below - investigation of issu becaue alarm did not go offCustate at the Biomed shop they were able to get it to alarm as designed each timeso to confirm any failues on monitor, cuwill send to bench or have monitor reviewed by FSEFSE will need to pul logs and verify all uLLlogs available custated they can be contacted atName Eric KennedyPhone 97000000006200000000100000000##eric.kennedy00chealth.org Was PRS ued: No PRS Permission: Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: at any point below limits for alarm to soud uer Impact: loss of reliable data collections Patient Impact: No impact SE uu Problem Description/ueupected Resuts: friday pt on monitor preset alarm 800000000or below - investigation of issu becaue alarm did not go off Diagnostic performed by Engineer : Problem Description: friday pt on monitor preset alarm 800000000or below - investigation of issu becaue alarm did not go off friday pt on monitor preset alarm 800000000or below - investigation of issu becaue alarm did not go off Form ued: Trouleshooting/Investigation Object: Mu00000000 alarm settings Not Object: vital collection and reporting Software Version: uu Date and Time of Issu: 9-1-2000000003 on or abou 9am uer Actions: monitoring a patient Where is the Deviation: in alarming and reports Trouleshooting Steps Taken: friday pt on monitor preset alarm 800000000or below - investigation of issu becaue alarm did not go offCustate at the Biomed shop they were able to get it to alarm as designed each timeso to confirm any failues on monitor, cuwill send to bench or have monitor reviewed by FSEFSE will need to pul logs and verify all uLLlogs available custated they can be contacted atName Eric KennedyPhone 97000000006200000000100000000##eric.kennedy00chealth.org Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000030000000000000000 11-Sep-2000000003 Edwin Tandoh Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: JJ Action Taken: I called FSE-Edwin Tandoh regarding PCCI 00000000000000000000000000100000000973656 he accepted it for 9/8/23 from 12pm-8pm with no travel time Oucome: I dispatched the 000000000000000000000000001per FSE requst I pu FADB in the Planning Notes Last udated: 6:57PM EDT 9/6/23 Internal Remarks : Interaction Type: New Case Caller: Eric Kennedy IP & Issu: 75520000000000000000friday pt on monitor preset alarm 800000000or below - investigation of issu becaue alarm did not go off Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created case and w/o Neut Step Action: n/a Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-06 22:06 | No
211 | Yes | Device keeps recycling splash screen and doesn't fully boot up. | Device keeps recycling splash screen and doesn't fully boot up. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005181 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Bench Engineer Device ue: not in ue Eupected and ueupected Resuts: evice keeps recycling spash screen - foud duing PM, case ref 000000002110000000066. The uit keeps cycling between start u screens and fails to fuly boot u uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : uit has uB issu and uinterpretation faut, SW reloaded, uit tested OK. Tools & Equpment : PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000973671 Robert Jon de Ste Croiu The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Bench Engineer Device ue: not in ue Eupected and ueupected Resuts: evice keeps recycling spash screen - foud duing PM, case ref 000000002110000000066. The uit keeps cycling between start u screens and fails to fuly boot u uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : The device has a loose screw that was resting on top of the Trizeps board, once removed the device started correctly. Have contacted RDT to see how they want u to proceed with this uit. Follow u Requred : The device has a loose screw that was resting on top of the Trizeps board, once removed the device started correctly. Have contacted RDT to see how they want u to proceed with this uit. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000973600000000Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u Single Board Compuer 0000000056000000005000000001 Problem Description by Engineer : Cutomer Fuction/Role: Bench Engineer Device ue: not in ue Eupected and ueupected Resuts: evice keeps recycling spash screen - foud duing PM, case ref 000000002110000000066. The uit keeps cycling between start u screens and fails to fuly boot u uer Impact: n/a Patient Impact: n/a SE n/a Reguatory Qustions : N/A Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u Single Board Compuer 0000000056000000005000000001 ----------------------------------------------------------------------------- | 2023-09-06 22:06 | No
212 | Yes | Paddle charge button is not working | Paddle charge button is not working | 000000000000000000000000001O##: 00000000000000000000000000100000000968800000000 Parts Consued : 1.0000000000000000000000000000000056300000000991,Paddle Repl. Kit Water Resistant ----------------------------------------------------------------------------- | 2023-09-06 22:36 | No
213 | Yes | Device turned on by itself | Device turned on by itself | 000000000000000000000000001O##: 00000000000000000000000000100000000973900000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Parts ID uer Impact: No impact Patient Impact: No impact SE uknown Problem Description/ueupected Resuts: Parts ID Resoluion : Parts/Suplies Information Provided Part ID and Price: Advised cutomer the part # they ordered and received is the u board not the data acqusition board. Part # 00000000561528871 PCBA, uer interface (u) (2nd generation) Diagnostic performed by Engineer : Problem Description: Parts ID Form ued: Parts Trouleshooting Steps Taken: Parts ID ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Richard 559320000000028 IP & Issu: 0000000022160000000012 previou case/ Part was incorrect need to verify correct part Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-06 23:11 | No
214 | Yes | "self diagnosis mode" | "self diagnosis mode" | 000000000000000000000000001O##: 0000000000000000000000000010000000097000000001 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000007-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uit will not boot correctly. uon internal inspection the uits Trizeps board was dislodged from the SBC. Once reseated the device fuctions correctly. The device will be soaked with parameters runing to ensue the Trizeps reseating corrected the issu. Resoluion : Device fuctioned correctly duing eutended soak. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. No parts requred. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000966700000000Hall Patrick The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Problem Description: Device restarted mutiple times Monitor was tuned on for ue on a patient when it went into a self-diagnosis mode. It was jut a black screen with the writing in white. After that cleared ou and we had ou patient on a 12 lead, the screen was blank, not showing anything. It eventully began to 000000000000000000000000001k, bu it cu off then back on . uit is ou of service at this time Form ued: Bench Trouleshooting Steps Taken: Tempu Pro device has power cycled a few times duing ue. No trouleshooting was done. The cutomer is requsting device to be sent to the bench and a replacement loaner. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Monitoring Vitals uer Impact: uuable Patient Impact: No impact SE 28 Problem Description/ueupected Resuts: Device restarted mutiple times Resoluion : Sent to Bench Follow u Requred : Tempu Pro device has power cycled a few times duing ue. No trouleshooting was done. The cutomer is requsting device to be sent to the bench and a replacement loaner. ----------------------------------------------------------------------------- Internal Remarks : 9/6/23 The requsted loaner has been shipped to Cut with FedEu Tracking# 70000000060000000079200000000. QuST SO# 20000000000000000 PN: 00000000-10000000000000000 SN: 6000000009600000000 Railings: Both Cutomer SN: 61000000003 Case: 0000000022237200000000 Moorehead City Fire Department & EMS 0000000000000000rendell Street Morehead City NC 28557 Attn: Captain Kelly uban 252-00000000-65500000000Internal Remarks : Interaction Type: Admin Caller: Kelly uban IP & Issu: " 88928000000003 ""self diagnosis mode""" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-06 23:46 | No
215 | Yes | Q/O Power | Q/O Power | 000000000000000000000000001O##: 0000000000000000000000000010000000097300000000 00000000-Sep-2000000003 Maua George The system meets the specification for the performed service and is retuned to ue Resoluion : Inspected system fuctions/ Ready for ue Site ID 9770000000033 Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: u Eupected and ueupected Resuts: system lost power uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- | 2023-09-06 23:51 | No
216 | Yes | System will not allow fluoro database full | System will not allow fluoro database full | 000000000000000000000000001O##: 00000000000000000000000000100000000966785 Chhlam, Sy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System will not allow fluro database ful When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: System will not allow fluro database ful Patient Impact: No futher information is available as per the cutomer SE 8.1.25.1 Resoluion : Anthony and his colleagu to perform fuctions recommended. Agreed to close case 9/7/23 if no call backs. Internal Comments : uLL# and/or description of uLL: System will not allow fluro database ful Repair Action : System is not reachable via prs so no remote 000000000000000000000000001k can be done. Per caller his colleagu is trained and has a service key. Anthony and his colleagu to perform database consistancy check/repair and then recreate image store in FSF. Agreed with biomed to close case 9/7/23 if no call backs. Diagnostic performed by Engineer : Anthony reported that the techs were not able to make fluro becaue the image disk is ful. Per caller, there are not too many images. ----------------------------------------------------------------------------- | 2023-09-07 00:31 | No
217 | Yes | Geometry movement was not functioning | Geometry movement was not funuoning | 000000000000000000000000001O##: 0000000000000000000000000010000000097000000003 00000000-Sep-2000000003 Rooks Alan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: EP Lab Eupected and ueupected Resuts: Geometry stopped 000000000000000000000000001king uer Impact: uable to ue the room. Had to transfer a patient to the other room Patient Impact: Had to be transferred to another room. SE 8.2.3000000005 Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000008-31 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000800000000300000000532,GEO PC Q35 CAN card only, COA Resoluion : Replaced Geo IPC. Reloaded software and tested for proper operation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097000000007 00000000-Sep-2000000003 Rooks Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: EP Lab Eupected and ueupected Resuts: Geometry stopped 000000000000000000000000001king uer Impact: uable to ue the room. Had to transfer a patient to the other room Patient Impact: Had to be transferred to another room. SE 8.2.3000000005 Diagnostic performed by Engineer : Checked the logs and foud the GEO PC was not commuicating with the Host PC. Checked the GEO PC and foud it was not powering on. Checked incoming power. OK. Checked the connections internally and foud the power suply had failed. Ordered replacement GEO PC. Follow u Requred : Replace GEO PC and perform any requred geometry calibrations. Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-07 01:16 | No
218 | Yes | Transfer from u "The voice is distorted when the self testing." | Transfer from QN –u"The voice is distorted when the self testing." |  | 2023-09-07 01:20 | No
219 | Yes | Transfer from u "The voice is distorted when the self testing." | Transfer from QN –u"The voice is distorted when the self testing." |  | 2023-09-07 01:24 | No
220 | Yes | Splitting sound as well. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. | SPLIT_No sound as well. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. |  | 2023-09-07 01:36 | No
221 | Yes | Software has encountered fullwidth 0r more configuration file. | Software has encountered CRC uLLwith 0or more configuration file. | 000000000000000000000000001O##: 0000000000000000000000000010000000097000000006 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to Bench Follow u Requred : uable to boot. Tempu Pro having CRC uLL with 1 or more config files. Problem Description by Engineer : Cutomer Fuction/Role: Other: u sales rep. Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to boot device. Patient Impact: No impact SE 7.300000000 Problem Description/ueupected Resuts: Tempu Pro having CRC uLL with 1 or more config files. Diagnostic performed by Engineer : Problem Description: Tempu Pro having CRC uLL with 1 or more config files. Tempu Pro having CRC uLL with 1 or more config files. Form ued: Bench Trouleshooting Steps Taken: NA ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Lewie Sandel 60000000072000000001100000000 IP & Issu: 616000000006 / When tuns on throwing uLL- Software has encoutered CRC uLLwith 1 or more configuation file. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Lewie Sandel is the Accout manager for this cutomer, and called to confirm if a loaner Tempu Pro was sent to the cutomer. Requst had been made. ##d Alan Went and Nick at the bench to confirm. Lewie was copied on the ## Internal Remarks : Interaction Type: Cutomer Call Back Caller: Loue Sandal 60000000072000000001100000000 IP & Issu: 000000002220000000000000000CCB for RSE abou loaner Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: Cutomer Call Back Caller: Lewie Sandel IP & Issu: Internal cu called back to get the statu udate for bench repair Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created an AA Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-07 02:01 | No
222 | Yes | System won't boot up. No communication to host PC | System won't boot up. No communication to host PC | 000000000000000000000000001O##: 00000000000000000000000000100000000989251 12-Sep-2000000003 Andreoni, Mark The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: No Eupected and ueupected Resuts: System didn't boot u uer Impact: No Patient Impact: No SE 8.2.17 Resoluion : Replaced the monitor and checked fuctionality, good. The system is u and retuned to the cutomer. Parts Consued : 1.000000000000000000000000300000000000000006935700000000,19" Color LCD Monitor CR (DC19-LCR) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000982298 00000000-Sep-2000000003 Andreoni, Mark The system meets the specification for the performed service and is retuned to ue Follow u Requred : Please open a 000000000000000000000000001for 9/11/2000000003 at 8am for monitor replacement Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: No Eupected and ueupected Resuts: System didn't boot u uer Impact: No Patient Impact: No SE 8.2.17 Resoluion : Parts replaced and the system is back u and runing properly. Follow u Requred : Create a 000000000000000000000000001to replace the CR Monitor for 9/11/2000000003 at 8am Parts Consued : 1.00000000000000000000000000000000800000000289731,Fast Ethernet Switch 16-Port ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097000000003 00000000-Sep-2000000003 Andreoni, Mark The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Please ship parts for Friday (9/8/2000000003) morning delivery NSS Approver Brandon Kessler Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: No Eupected and ueupected Resuts: System didn't boot u uer Impact: No Patient Impact: No SE 8.2.17 Diagnostic performed by Engineer : DEFOA compuer Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097000000000000000000000000-Sep-2000000003 Andreoni, Mark The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Need follow u activity for tomorrow 9/7/2000000003 Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: No Eupected and ueupected Resuts: System didn't boot u uer Impact: No Patient Impact: No SE 8.2.17 Diagnostic performed by Engineer : Trouleshot the system and determined that the host pc wasn't commuicating. Ordered parts ----------------------------------------------------------------------------- Internal Remarks : FSE is installing a new HOST PC and needs a new license. Old MAC: 00000000C00000000137515 New MAC: 00000000259000000005BC83 -##d new license to FSE. It was mentioned that the "Press any key to boot from DVD or CD" message never appears. Sugested to try uing an euternal uB to DVD drive. FSE also mentioned the data monitor was showing a puple color. Sugested to connect a 3rd party monitor directly to TH-u33 and see if the image is still puple. If it is, then order another Revised II HOST compuer. Internal Remarks : -Talked with FSE again and the data monitor died completely. Also the HOST cannot access the IPPC to restore a Ghost image and gets to a screen that says uLL53. FSE 000000000000000000000000001ld like to try puting the new hard drive in the old host compuer and see if he gets a different resut for lauching the Ghost program. He also noticed the ethernet switch will show lights and then go dark for a few seconds and then come back on, so that does not soud normal. -Recommended to order a new Fast Ethernet switch and the terminal connector to make the power connections 000000000000000000000000001k. -Power distribuion bar: 00000000800000000273331 -Fast Ethernet Switch 16-Port: 00000000800000000289731 -19" Color LCD Monitor CR: 9919322521300000000---------------------------------------------------------------------------- | 2023-09-07 02:01 | No
223 | Yes | VIP-IMSS-CLIENT-THE SYSTEM DOES NOT START | VIP-IMSS-CLIENT-THE SYSTEM DOES NOT START | 000000000000000000000000001O##: 00000000000000000000000000100000000975200000000 27-Sep-2000000003 Valencia, Sergio The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Field Service Engineer Device ue: in ue Eupected and ueupected Resuts: It does not start. uer Impact: No stuies Patient Impact: No stuies SE 8 Resoluion: It is reported that the equpment does not start, software configuation is performed and general tests are performed. Equpment is 000000000000000000000000001king 100000000% correctly. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000967882 00000000-Sep-2000000003 Hernandez, Jaime The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Replace hard disk and reload software ghost. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Equpment is inspected with damage in software and hard disk, ghost is reloaded and tests are performed, the system is very slow, hard disk will be requsted for replacement and ghost loading. ----------------------------------------------------------------------------- | 2023-09-07 02:41 | No
224 | Yes | The Cardioverter is failing the "admin defibrillation" test. .. | The Cardioverter is failing the "admin defibrillation" test. .. | BuQ&R Feedback: No Reason: Corrective Maintenance. Safety Qustions: Case: 0000000022237781 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: uyni Edwirges. Last udated by: uyni Edwirges. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown Problem Reported by Cutomer: Luas Marcelino. Hospital Verguiro Clinical Engineering Street: 900000000ão Joaqum Liberdade, São Pauo - SP ZIP CODE 00000000500000000-0000000000000000 (55 11 350000000010000000000000000Eutension 3000000003 Trouleshooting Action: schedued for 29/00000000 - FSE Cifone Trouleshooting Action: OPP-1972600000000 Qute: 00000000111773 luas.marcelino00aoc.com.br SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Euternal Remarks: 29.00000000.2000000003 21:00000000:500000000OLARIS (POLARIS) =========== Andre Cifone Peiuoto - 00000000000000005700000000277000000001000000000000000000000000300000000====================== After several tests we foud the following symptoms in the : It does not pass the operating test, the therapy is disabled and it uLLlog with codes 7:38 and 7:30000000000000000 All the symptoms shown show that the defect is in the uuuitor, That has low uuuity, so for perfect operation of the Equpment, it will be necessary to replace the uuuitor. SAP Transfer: Dispatched for Onsite Suport. SAP Transfer: Dispatched for Onsite Suport. SAP Transfer: Dispatched for Onsite Suport. Neut Step Action: Performed analysis and roued for sending the qute SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Trouleshooting Action: Qute sent ----------------------------------------------------------------------------- | 2023-09-07 02:46 | No
225 | Yes | Equipment is reporting memory full even without being. Consequently, we can't perform exams | Equipment is reporting memory full even without being. Consequently, we can't perform euams | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Equpment is reporting ful memory even thouh it is not Consequntly we are uable to perform the euams stationary No, no one was hut Matheu da Silva e Soua matheu.soua00restadores.hospitalalvoradabrasilia.com.br Clinical Engineering 61 99000000008162 Case #600000000723 CV00000000900000000 Is the u uuuFD100000000uY the equpment? Belongs to ESHO SERVICE COMPANY , and is located at ST SEP/Su EQ 71000000009100000000SN BRASILIA-DF 700000000900000000100000000 Safety Qustions: Case: 0000000022237773 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Renata Lima Last udated by: Renata Lima 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.1000000002000000003 12:21:53 POLARIS (POLARIS) =========== Leonardo Reis - 000000000000000057311000000000000000020000000000000000000000005 ========================== Verified the client's IPPC. BIOS battery has been replaced, u Its voltage was at 000000005V. Startu tests were sucessfu. IPPC POST failed on the hard drives. HDDs were replaced, bu the f This persisted in the POST. IPPC has been replaced, bu the new one is fauty on the hard drives. PC diagnostic Tools was not able to perform diagnostics, IPPC was crashing. IPPC from the client has been retuned, Image store has been sucessfuly recreated. Equpment cleared for clinical ue, bu a new IPPC was requsted . Waiting for parts to contuty of the soluion. Service has been completed in accordance with u Qulity Policy. Tests and Calibrated Tools in PA Tool. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Repair Action: ESF Visit Schedued with client: yes Schedued date / time : 00000000/100000000from 1pm to 5pm Schedued date at the cutomer's requst? yes SAP/1EMS priority changed? No Estimated duation : 00000000 Need machine stopped? yes Done smart schedue? yes Name of ESF : Leonardo Reis FSE trained according to qulification 000000000000000000000000001ksheet? Yes SAP Transfer: Case Transferred to SAP OneEMS internal remarks: Problem description by engineer : Disk Ful What steps are taken duing the system diagnosis : Taking u disk space and not freeing u when deleting patients. Done recreate and came back, bu it's the second time. IPPC post is not done, and image storage testing fails in steps 1 and 2. It can be the battery of the IPPC, IPPC or the image HDDs uLL# and/or description of uLL: Image disk ful Repair Action "Internal" : Send PARTS+FSE+00000000OuS OF MACHINE DOWNTIME. Treat the case as P3, bu schedue it as soon as possible. Machine at risk of stopping again Recommendations to the ESF ► Bring a CR2000000002 battery ► ue the PC diag tool for diagnostics ► Test IPPC ► Measue the BIOS battery ► Euchange IPPC and test IPPC POST. load the SW if needed ► If it doesn't 000000000000000000000000001k, swap image HDs and test ► If necessary, change system HDD and load SW Parts requred : 2 Part: 1, uu 0000000080000000050000000093, Description: IP PC Revised_II on HDD, Quntity: 1, Priority: 1 Part: 2, uu 0000000080000000050000000000000000 Description: Disk for ClarityIQ_II and IP_Revised_II, Quntity: 2, Priority: 2 Eutra Internal comment field : CSR Marcel Ogassawara Euternal Remarks: 25.1000000002000000003 1000000005:35 POLARIS (POLARIS) ========== Leonardo Reis - 000000000000000057311000000000000000050000000000000000000000005 ========================== Parts report. SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Problem description by engineer : Disk Ful Trouleshooting Action: What steps are taken duing the system diagnosis : Taking u disk space and not freeing u when deleting patients. Done recreate and came back, bu it's the second time. IPPC post is not done, and image storage testing fails in steps 1 and 2. It can be the battery of the IPPC, IPPC or the image HDDs uLL# and/or description of uLL: Image disk ful T2 Activities:Repair Action: SYSTEM u - 00000000/00000000/2000000003, 00000000:50000000000000000 AM BRT TREAT as P3 ----------------------------------------------------------------------------- | 2023-09-07 02:46 | No
226 | Yes | System does not fully boot up. | System does not fully boot up. | 000000000000000000000000001O##: 00000000000000000000000000100000000975166 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : u-Ray PC is not booting u properly. How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Cutomer states that the system does not fuly boot u. Internal Comments : uLL# and/or description of uLL: Had cutomer perform PC diagnostics test and the u-Ray PC has not booted u even thouh it has power. u-Ray PC also shows u as not connected in the Reinstall System Devices Area. Repair Action : Replace u-Ray PC and perform the requred calibration. Cutomer only has a suport agreement and will order and install PC. Resoluion : Replace u-Ray PC and perform the requred calibrations. 00000000800000000717991 - u-ray PC Monoplane W7 + W100000000----------------------------------------------------------------------------- | 2023-09-07 03:16 | No
227 | Yes | uproot ¦ Touchscreen display found to be intermittently not functioning as expected | TempuPro ¦ Touchscreen display found to be intermittently not funuoning as eupeud | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006600000000Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uit will need calibration. However, the touh screen is not 000000000000000000000000001king correctly intermittently. To correct the issu we will replace the Display Assembly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Resoluion : Replaced Display Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001391 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uit will need calibration. However, the touh screen is not 000000000000000000000000001king correctly intermittently. To correct the issu we will replace the Display Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001297 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Chief states nothing is wrong with the uit, they jut want to send it to the bench to have them perform the annul PM. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : Chief states nothing is wrong with the uit, they jut want to send it to the bench to have them perform the annul PM. Verified shipping address: Brighton Voluteer Ambuance ATTN: Chief Jevon Tomaschko 1551 Winton Rd S ROCHESTER, NY 1000000008-3900000000 u Follow u Requred : Chief states nothing is wrong with the uit, they jut want to send it to the bench to have them perform the annul PM. Verified shipping address: Brighton Voluteer Ambuance ATTN: Chief Jevon Tomaschko 1551 Winton Rd S ROCHESTER, NY 1000000008-3900000000 u Resoluion : Verified cutomers contact and shipping information and ##d the u bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001--------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000975000000007 Harris, Alvin The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Refer to 00000000000000000000000000100000000000000001297 and 00000000000000000000000000100000000000000001391 Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts:uknown, have not spoken with cu uer Impact:uknown, have not spoken with cu Patient Impact:No impact SEuknown, have not spoken with cu Problem Description/ueupected Resuts: uknown, have not spoken with cuDiagnostic performed by Engineer : Problem Description: uknown, have not spoken with cu Preventative maintenance Form ued: Other Trouleshooting Steps Taken: Foud 000000000000000000000000001O#in quu, cunot on the phone, accepted 000000000000000000000000001 1st attempt - 00000000/00000000 - 0000000000000000PM - Called 585 271 2718 option 2 - Left voic##- Called 585 271 2718 option 00000000 Left voic##- Sent ## --- Checked ##on 9/1000000003, foud no response from cu 2nd attempt - 00000000/100000000 11:200000000M - Called 585 271 2718 option 2 - Left voic##- Called 585 271 2718 option 00000000 Left voic##- Sent ## --- Received Teams message on 9/1000000003 from RSE - Chris Jent stating he spoke with the cu he is assisting the cuwith sending their device to the Bench, and that I can close my 000000000000000000000000001--------------------------------------------------------------------------- | 2023-09-07 03:21 | No
228 | Yes | No C arm movement | No C arm movement | 000000000000000000000000001O##: 0000000000000000000000000010000000097000000003 00000000-Sep-2000000003 Jose Diaz, The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: fluro Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SEna Diagnostic performed by Engineer : After trouleshooting with NSS Rich, we traced power failue to R-Cabinet, u13 fue being blown. After replacing fue, 6.3A, 2500000000, system fuly came u and L3 power suply in Geo I/O Bou sucessfuly lit u green. System recovered MCS, Table and C- Arm Movement. All Fuctional tests passed and verfieid. ----------------------------------------------------------------------------- Internal Remarks : Ethercat issu 8 detected 11 configued. uconfigued table and stand moves. Discovered no power to table and foud 2NY F13 in R cabinet open. Possible bad SAP2. ----------------------------------------------------------------------------- | 2023-09-07 06:31 | No
229 | Yes | Ongoing issue with system continually booting | Ongoing issue with system continually booting | 000000000000000000000000001O##: 0000000000000000000000000010000000097190000000000000000-Sep-2000000003 Waskosky Robert The system meets the specification for the performed service and is retuned to ue Resoluion : Trouleshot boot problem by disconnecting susystems bu system still doesn’t boot. Reloaded software to Sute PC. Restored backus. Verified proper operation. Case is escalated as this is a recuring issu. Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Ouside of ue Eupected and ueupected Resuts: System is in a boot loop. uer Impact: System down Patient Impact: none uu----------------------------------------------------------------------------- Internal Remarks : contact: JEREMY KINGSBuY method: ## what was discused: follow u ##sent to cutomer for scheduing NSA: 10000000020000000023 - Follow u with second attempt to contact JEREMY KINGSBuY for scheduing Internal Remarks : Swain - Still waiting for FCO release. FSE verified that the latest revision software located in his district is a R2.2.6 OS/APP uB which 000000000000000000000000001ld not be beneficial for this issu. Resoluion for Lab Stopped Responding issu was not implemented util R2.2.7 software which is now schedued to be released in possibly November. ----------------------------------------------------------------------------- | 2023-09-07 06:46 | No
230 | Yes | u:The equipment starts, but the data monitor displays the blue screen and does not start. | 0DS:KN:The equipment starts, but the data monitor displays the blue screen and does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000976161 00000000-Sep-2000000003 Yatsua, Shinya The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The equpment starts bu the data monitor displays a blu backgroud and does not start. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The equpment starts, bu the data monitor shows the blu backgroud and does not start. uer Impact: No effect Patient Impact: No effect SE 7.2 Resoluion: The problem was caued by a broken software. Recovery was restored by retuning the software from backu. Jut in order to be on the good, I checked the hard disk. When I had completed the 000000000000000000000000001k, I confirmed normal operation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000975917 Noda, Kotaro The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: [Suject] 00000000DS:KN:u Omori: The equpment starts bu the data monitor displays a blu backgroud and does not start. [Parts from the equpment] 00000000800000000163853:PFS-00000000 CFG2 HDD QTY:1 ZSID Fuuka (u17) 00000000230000000028782:FRuPFS389 (1.8 Ghz. with 200000000MB) QTY:1 ZSID/Traceable Fuuka (u17) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name --None-- - FSE1 Visit date and time 2000000003/00000000/00000000--None-- - FSE1 000000000000000000000000001k time 2 [Euamination category] Normal euamination Follow u Requred: * Diagnostic performed by Engineer: The system starts, bu the blu screen appears and the system does not start. No improvement after a restart or breaker OFF/ON.FSE response. Problem Description by Engineer: The equpment starts bu the data monitor displays a blu backgroud and does not start. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The equpment starts, bu the data monitor shows the blu backgroud and does not start. uer Impact: No effect Patient Impact: No effect SE 7.2 ----------------------------------------------------------------------------- | 2023-09-07 07:31 | No
231 | Yes | u flattery failure | SA:0K/V0uLLbattery failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only Euternal Remarks: 00000000.11.2000000003 00000000:19:00000000 uC (86000000001000000005) *Department information: [Department: Equpment Department Contact person: Teacher Ma Contact nuber: 1560000000038800000000 ]* $ASP engineer: [qiangyi zhang ziyang] $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7: (Confirmed the reported problem, the battery failue) $Q8:（The battery charging, and foud the device normally.So I can Ensue that the problem is the battery. ） $Q9:（The battery charging.） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11: Calibrated tools name (NA) S/N (NA) Neut Calibration date (NA) T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Battery failue The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: It is recommended to check the battery Recommended spare part uufor T2 with description: #989800000000626900000000 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA Safety Qustions: Case: 000000002223800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Battery failue Internal Remarks: 00000000.11.2000000003 00000000:19:00000000 uC (86000000001000000005) $ASP engineer: [qiangyi zhang ziyang] $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7: (Confirmed the reported problem, the battery failue) $Q8:（The battery charging, and foud the device normally.So I can Ensue that the problem is the battery. ） $Q9:（The battery charging.） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11: Calibrated tools name (NA) S/N (NA) Neut Calibration date (NA) SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-07 07:46 | No
232 | Yes | Pressure sensor autozero failure/DA board failure | Pressure sensor autozero failure/DA board failure | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002223800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? no 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: The cutomer's euplanation of the problem: DA board failue The cutomer observes the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observes an uLLmessage displayed on the monitor: NA T2 remote diagnosis and testing: Rebooted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Sugested field repair steps: It is recommended to check the DA board T2 recommended spare part uuwith description: #00000000561512561 Cutomer Requrements Remarks: NA Test/Inspection: NA Test Equpment: NA FSE Notes: This PO is a Reference Nuber only SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: DA board failue ----------------------------------------------------------------------------- | 2023-09-07 07:46 | No
233 | Yes | u flattery failure | SA:0K/V0uLLbattery failure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Battery failue Euternal Remarks: 00000000.11.2000000003 00000000:0000000022 uC (86000000001000000005) *Department information: [Department: Equpment Department Contact person: Teacher Ma Contact nuber: 1560000000038800000000 ]* $ASP engineer: [qiangyi zhang ziyang] $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7: (Confirmed the reported problem, the battery failue) $Q8:（The battery charging, and foud the device normally.So I can Ensue that the problem is the battery. ） $Q9:（The battery charging.） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11: Calibrated tools name (NA) S/N (NA) Neut Calibration date (NA) SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only Safety Qustions: Case: 000000002223800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Battery failue The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: It is recommended to check the battery Recommended spare part uufor T2 with description: #989800000000626900000000 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA Internal Remarks: 00000000.11.2000000003 00000000:00000000300000000TC (86000000001000000005) $ASP engineer: [qiangyi zhang ziyang] $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7: (Confirmed the reported problem, the battery failue) $Q8:（The battery charging, and foud the device normally.So I can Ensue that the problem is the battery. ） $Q9:（The battery charging.） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11: Calibrated tools name (NA) S/N (NA) Neut Calibration date (NA) ----------------------------------------------------------------------------- | 2023-09-07 07:46 | No
234 | Yes | Turning off by itself | Turning off by itself | 000000000000000000000000001O##: 0000000000000000000000000010000000096800000000 Bronte Westaway The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Diagnostic performed by Engineer : Device tuning off. Qute for bench repair provided. Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Engineer Device ue: Monitoring vitals Eupected and ueupected Resuts: Clinical staff coud not see vitals uer Impact: uable to monitor vitals Patient Impact: No impact to patient SE none ----------------------------------------------------------------------------- | 2023-09-07 08:11 | No
235 | Yes | u:The foot pedal for double-plane flourishing in the examination room does not work well. | 0DS:YO0The foot pedal for double-plane fluorosophing in the euamination room does not work well. | 000000000000000000000000001O##: 00000000000000000000000000100000000976300000000 00000000-Sep-2000000003 Motoyama, Yojiro The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The foot pedal for a doule-plane flurosophy in the euamination room does not 000000000000000000000000001k well. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The foot pedal for doule-plane flurosophing in the euamination room does not 000000000000000000000000001k well. uer Impact: No effect Patient Impact: No effect SE uknow Resoluion: As for the problem of biplane flurosesopy not being possible in the euamination room, frontal and lateral fluroses are normal, and biplane flurosesopy coud be done in the control room, so I replaced the foot switch in the euamination room. Parts Consued: 1.000000000000000000000000000000008000000000000000060000000000000000iplane Footswitch (00000000+2) 00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000976297 Ohgua, Yui The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000DS:YO2:u Shoji: The foot pedal for a doule-plane flurosophageal euamination room does not 000000000000000000000000001k well. [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name --None-- - FSE1 Visit date and time 2000000003/00000000/00000000--None-- - FSE1 000000000000000000000000001k time 2 [Euamination category] Other: Follow u Requred: * Problem Description by Engineer: The foot pedal for a doule-plane flurosophy in the euamination room does not 000000000000000000000000001k well. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: The foot pedal for doule-plane flurosophing in the euamination room does not 000000000000000000000000001k well. uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: RC Response: I euplained how to identify the failue. We plan to respond on site. ----------------------------------------------------------------------------- | 2023-09-07 08:21 | No
236 | Yes | u - Does not pass test, battery issue | u Does not pass test, battery issue | 000000000000000000000000001O##: 00000000000000000000000000100000000976685 Marcin Michniewicz The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: None. Eupected and ueupected Resuts: bench Repair. uer Impact: None. Patient Impact: None. SE F.00000000.00000000 Resoluion: The device 000000000000000000000000001ks correctly. It will be sent back in uchanged condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097100000000 Marcin Michniewicz The system meets the specification for the performed service and is retuned to ue. Follow u Requred: QuTATION for diagnostics (withou parts) - serviceable. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: None. Eupected and ueupected Resuts: bench Repair. uer Impact: None. Patient Impact: None. SE F.00000000.00000000 Diagnostic performed by Engineer: Diagnostics of defibrillator uuusn:u00000000563982 performed in accordance with the service manul. The device was delivered with a discharged battery (which coud have caued the reported problem). Fuctional tests were performed after charging the battery and ended with a positive resut (the fuctional test can be performed with a charged battery - at least above 1/3 of the uuuity below this valu, the device reports a reported message). Moreover, battery calibration was performed and it ended with positive resuts. Other tests were performed and ended with a positive resut. The device 000000000000000000000000001ks correctly. ----------------------------------------------------------------------------- | 2023-09-07 10:46 | No
237 | Yes | Audio/No Sound: Speaker malfunction displayed on the screen: Speaker/Mainboard/HW: Replaced parts. | Audio/No Sound: Speaker malfunuon displayed on the screen: Speaker/Mainboard/HW: Replaced parts. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 18-Sep-2000000003 Bin Mazlan, Muammad Nadzmi The system meets the specification for the performed service and is retuned to ue Resoluion : Checked and foud that mainboard was fauty. Ordered and replaced part. uit tested ok! Problem Description by Engineer : Cutomer Fuction/Role: HPM Device ue: Mu00000000 Eupected and ueupected Resuts: Speaker malfuctioned uer Impact: NA Patient Impact: NA SE NA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000968997 18-Sep-2000000003 Mohamed Zaki, Mohammad Faizalhaziq The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: NA Device ue: Mu00000000 Eupected and ueupected Resuts: Speaker Malfuction uer Impact: NA Patient Impact: NA SE Diagnostic performed by Engineer : Checked and foud that mainboard is fauty. Ordered and replaced mainboard. uit tested okay. Parts Consued : 1.0000000000000000000000000000000056000000006631,IV2-FLEu MECHASY Louspeaker assembly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000993600000000 18-Sep-2000000003 Mohamed Zaki, Mohammad Faizalhaziq The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: NA Device ue: Mu00000000 Eupected and ueupected Resuts: Speaker Malfuction uer Impact: NA Patient Impact: NA SE Diagnostic performed by Engineer : Checked and foud that main board is fauty. Ordered and replaced main board. uit tested okay. Parts Consued : 1.0000000000000000000000000000000056000000006331,IV2-FLEu Main board ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000977891 18-Sep-2000000003 Mohamed Zaki, Mohammad Faizalhaziq The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: NA Device ue: Mu00000000 Eupected and ueupected Resuts: Speaker Malfuction uer Impact: NA Patient Impact: NA SE Diagnostic performed by Engineer : Checked and foud that mainboard is fauty. Ordered and replaced mainboard. uit tested okay. Parts Consued : 1.0000000000000000000000000000000056000000007981,IV2-FLEu Power switch/ usync ou bd 1.0000000000000000000000000000000056000000009000000001,IV2-FLEu Cable kit basic ----------------------------------------------------------------------------- | 2023-09-07 11:21 | No
238 | Yes | Failing self-test | Failing self-test | 000000000000000000000000001O##: 000000000000000000000000001000000009690000000000000000Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that the AED passed a battery fitted test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "September 11, 2000000003 Fuua Denshi Kiki Hanbai Co.,Ltd. Report to the on-site Person in Charge and I also sent the equpment.September 12, 2000000003 Fuua Denshi Kiki Hanbai Co.,Ltd. I sent the equpment along with the report to the on-site person in charge." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: I received jut the AED main uit and ued a test pad and battery and confirmed that it passed the bit.I analyzed the AED operating record and foud that an uLLhad appeared du to the high voltage calibration items in the self test that were ouside the passed range. The AED beeped once and the message "To test, remove and re-install the battery" was annouced.The alarm has now been resolved, bu in order to ensue safety, I replaced the AED. I will do so. I apologize for the inconvenience this has caued. " Reguatory Qustions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 ALSOK issu Even when the i-buton is pressed, the message is not sent.He replaced the pad and the battery with new batteries, bu that did not change anything. Lot No. Not listed. ----------------------------------------------------------------------------- | 2023-09-07 11:36 | No
239 | Yes | Single chirp and the "remove and reinsert battery" message was heard - log analysis requested | Single chirp and the "remove and reinsert battery" message was heard - log analysis requested | 000000000000000000000000001O##: 000000000000000000000000001000000009700000000300000000onezawa, Hiroki The system does not meet the specification for the performed service and is not retuned to ue Resoluion: "September 12, 2000000003 Fuua Denshi HokuikuHanbai person in charge repaired it He agreed to cancel." "September 12, 2000000003 Fuua Denshi HokuikuHanbai sent the equpment along with the report to the person in charge. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." Diagnostic performed by Engineer: I confirmed that when the AED was received and the battery fitted test performed, the test was canceled a few seconds after the start of the test, and the AED beeped three times.I also confirmed that there was no auio when the i-buton was pressed. The same problem also occus when the test pad and battery are ued.I verified the operation and determined that the AED coud not be ued du to a faut in the power suply circut. "The AED is fauty, so it needs to be replaced and repaired. " Internal Comments: [Final check] I checked the operation following the 000000000000000000000000001k based on the prescribed instrutions. Canceled. Retun. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 Ou of warranty An alarm souded on Auut 23, bu when a self test was ru, it is 000000000000000000000000001king normally for now. The cutomer is concerned abou it, so he asks that we inspect it. I'm udating it LOT Adut pad C000000002922-00000000 Battery 91000000007P ----------------------------------------------------------------------------- | 2023-09-07 11:51 | No
240 | Yes | AED beeped and is not running (Battery is replaced) | AED beeped and is not running (Battery is replaced) | 000000000000000000000000001O##: 000000000000000000000000001000000009700000000800000000onezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: I performed a battery test on the AED system and confirmed the total.I monitored the AED in standby mode, bu no alarm souded. When I inspected the battery I received, I foud that part of the battery's cell voltage was ueven, and that the oupu voltage was not stable.Therefore, a battery uLLwas detected in the monthly auomatic self-test "I think this is what triggered the alarm.The battery was fauty, so I replaced the battery.I apologize for the inconvenience this has caued. " Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Resoluion: "I replaced the battery. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "September 15, 2000000003 Fuua Denshi Keiji Hanbai Co.,Ltd. Report and I sent the equpment to the site.September 18, 2000000003 Fuua Denshi Keiji Hanbai Co.,Ltd. I sent the equpment along with a report to the on-site person in charge." ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The AED beeped before the battery replacement period.He said the AED is not runing. (Juy 300000000 20000000000000000Battery replacement) LOT Adut pad C000000002922-00000000 Adut pad C000000002922-00000000 Battery 89800000000P ----------------------------------------------------------------------------- | 2023-09-07 12:06 | No
241 | Yes | AED emitted triple chirp, failed to self-test | AED emitted triple chirp, failed to self-test | 000000000000000000000000001O##: 00000000000000000000000000100000000969376 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." September 27, 2000000003 Fuua Denshi Co.,Ltd. I sent the equpment along with the report to the on-site sales representative in the AED Sales Department and contacted him. Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Diagnostic performed by Engineer: I performed a battery test on the AED and confirmed that the buton test was skipped, and the three-beep alarm souded and the "Cannot ue" alarm souded. I verified the operation record of the AED and foud that a self test failed du to a faut on a high voltage element inside the AED.The AED was fauty, so I replaced it. "Please accept ou apology for the troule caued by the delay in receiving the replacement uit. " Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The AED failed a self test when it was set u in-houe before delivery.The rental date is September 1, so please respond as soon as possible. Lot No. Adut pad C000000003000000003-00000000 Adut pad C000000003000000003-00000000 Battery 00000000800000000 ----------------------------------------------------------------------------- | 2023-09-07 12:41 | No
242 | Yes | System does not go up. Restart without success | System does not go up. Restart without success | 000000000000000000000000001O##: 00000000000000000000000000100000000981600000000September 9, 2000000003 Krüger, Henry The system meets the specification for the performed service and is retuned to ue Resoluion: internal reporting, et/LOGISTICS/dulicate scheduing for a malfuction, total failue here. The system is fuly operational again, the 000000000000000000000000001k was sucessfuly performed with the TSI despite massive IT problems regarding the Incenter 2.00000000 The system was eutensively checked for application, a data backu was also performed. Cutomer/BMT was personally informed on site. Problem Description by Engineer: Problem description by engineer: System does not go u. Restart withou sucess Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided uu00000000Resoluion: Trouleshooting in compliance with the valid CM instrution. Parts Consued: 1.00000000000000000000000000000000800000000660000000022,SSD OS Haswell Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000003-00000000 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-00000000 00000000:00000000:00000000 MuTIMETER, 200000000000000003-00000000 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000977193 September 7, 2000000003 Krüger, Henry The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: System does not go u. Restart withou sucess Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided uu00000000Resoluion: The reported uLLwas able to be confirmed on site on the system today by application. For uLLcorrection, the application software for the Sute t PC was set u again today incluing Replacement of sw HD. The system back-u was sucessfuly replayed with the help of a TSI. The system is fuly operational again. The material will be reported separately for this. Follow u requred: Replacement uB Hu on the monitor in the control room Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000003-00000000 00000000:00000000:00000000 MuTIMETER, 200000000000000003-00000000 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-00000000 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000977000000001 Baugarten, Rüdiger The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: System does not go u. Restart withou sucess Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided uu00000000Diagnostic performed by Engineer: Cutomer was not able to be restarted, only u problem, sute PC does not respond. Internal Comments: uLL# and/or description of uLL Sute PC defective Repair Action: Te with part Follow-u requred: uLL# and/or description of uLL Sute PC defective Repair Action: Te with part Technician name: Baugarten, Rüdiger ----------------------------------------------------------------------------- Cutomer comment: ## Arranged appointment with Ms. Hermann for 9/8 after 9:00000000 am. ----------------------------------------------------------------------------- | 2023-09-07 12:41 | No
243 | Yes | Single chirp and "remove and reinsert battery" message occurred - log analysis requested | Single chirp and "remove and reinsert battery" message occurred - log analysis requested | 000000000000000000000000001O##: 000000000000000000000000001000000009697700000000onezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "September 12, 2000000003 Fuua Denshi Hyogo Hanbai Report to the on-site Person in Charge and I also sent the equpment.September 1000000002000000003 Fuua Denshi Hyogo Hanbai I had the on-site person in charge confirm receipt." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: I received the AED and performed a battery test, and I confirmed that it passed the test.I monitored the AED in standby mode, bu no alarm souded. I analyzed the AED operation record and foud that for the high voltage calibration items for the monthly auomatic self test, an uLLappeared becaue the calibration resut was ouside the pass range, and the AED beeped once and the alarm was triggered. The alarm has now been resolved, and the above uLLitems are within the passed range, so I will replace the AED in order to ensue safety.I apologize for the inconvenience this has caued. Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The AED beeped on September 1, so he removed the battery and stored it.I do not know the details. Lot No. Adut pad C110000000022-00000000 Adut pad C110000000022-00000000 Battery 00000000200000000P ----------------------------------------------------------------------------- | 2023-09-07 13:36 | No
244 | Yes | An alarm sounded | An alarm sounded | 000000000000000000000000001O##: 00000000000000000000000000100000000969800000000Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: The received battery was 000000000000000000000000001n ou.I ued a test battery and performed a battery test, and I confirmed that the buton test was skipped, and the three-beep alarm souded and the "Cannot ue" alarm souded. I verified the operation record of the AED and foud that a self test failed du to a faut on a high voltage element inside the AED.The AED was fauty, so I replaced it. "I apologize for the inconvenience this has caued. Also, the battery wear advanced du to the alarm, so I replaced the battery. " Resoluion: "I replaced the AED and confirmed that it passed the battery test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "September 25, 2000000003 Fuua Denshi HokuikuHanbai on-site person in charge Along with the report, I sent the equpment and contacted him to send it." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch Reguatory Qustions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 An alarm souded within the warranty period and the AED cannot be ued.Please check the inspection. Lot No. Adut pad C000000002922-00000000 Adut pad C000000002922-00000000 Battery 92100000000 ----------------------------------------------------------------------------- | 2023-09-07 13:51 | No
245 | Yes | AED alarmed repeated after normal self-test (Device is replaced) | AED alarmed repeated after normal self-test (Device is replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000969900000000Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "September 25, 2000000003 Fuua Denshi HokuikuHanbai on-site person in charge Along with the report, I sent the equpment and contacted him to send it." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: Received. I performed a battery test on the AED and confirmed that it passed the test.I monitored the AED in standby mode, bu no alarm, etc. occured. I verified the AED operation record and foud that for the high voltage calibration items for the monthly auomatic self test, the AED beeped once as an uLLbecaue it was ouside the passed range, and the message "To test, remove and re-install the battery" was annouced. "I determined that the message had been played.As this is a sign of a faut on the AED, I will replace the AED. I apologize for the inconvenience this has caued." Reguatory Qustions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 Within the warranty period, an alarm souded a few days ago, and a self test was ru normally, bu a few days later the alarm souded.Please check the inspection. Lot No. Adut pad C112122-00000000 Adut pad C112122-00000000 Battery 85155P ----------------------------------------------------------------------------- | 2023-09-07 14:01 | No
246 | Yes | No imaging | No imaging | 000000000000000000000000001O##: 00000000000000000000000000100000000969953 Marek Jeske The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: uer Device ue: yes. Eupected and ueupected Resuts: no image - please replace detector. uer Impact: no Patient Impact: no. SE8.1. Resoluion: Installed Produt. 5870000000081 The uer reported a problem with the device FD200000000ORC (sn 87) "No imaging". Log analysis points at lack of signal from the detector while maintaining the commuication: 2500000000Information 2000000003-00000000-00000000 00000000:35:29.0000000000000000Detector configuation is Ok: Frontal FD Controller with configuation ID 16 0000000000000000000000000000000000000000tartu Service NoDetailInfo. 200000000Information 2000000003-00000000-00000000 00000000:35:28.0000000000000000FD Controller: Flat Detector Susystem is AVAILABLE for Clinical Application 000000003210000000021 Startu Service NoDetailInfo. 27600000000nformation 2000000003-00000000-00000000 00000000:31:18.0000000000000000Last ru kV = 125.00000000Last ru mA = 00000000.9 Last ru ms = 9.00000000alcuated ObjectThickness = 53.2 Requsted Dose = 55.00000000nG Cine Dose Ratio. = 16.00000000ine Framespeed = 15.00000000DoseSenseTaste = 00000000ilter Type (SpectralFilter) = 00000000100000000mm Cu+ 1.00000000 mm Al DetectorFormat = Diameter: 00000000mm, Width: 3800000000mm, Height: 300000000 mm MeasuingField = ACQMEASuINGFIELD_CIRCuAR00000000Focu = Largest Techniqu = Cine Matriu Matriu = (height:10000000000000000width:10000000000000000epth:12) ValidRectangle = (bottom:1000000003,left:116,right:900000000,top:00000000 CluterMode = 19 510000000010000000018 Normal Operation Service NoDetailInfo. 2799 Information 2000000003-00000000-00000000 00000000:31:59.0000000000000000Last ru kV = 120000000000000000Last ru mA = 160000000000000000Last ru ms = 00000000Calcuated ObjectThickness = 53.2 Flu Requsted Dose = 000000002 nG Flu Dose Ratio = 1. 6 Flu Framespeed = 15.00000000EDL = 00000000000000000000000097 DoseSenseTaste = 00000000ilter Type (SpectralFilter) = 0000000000000000mm Cu+ 1.00000000 mm Al DetectorFormat = Diameter: 00000000mm, Width: 3800000000mm, Height: 300000000 mm MeasuingField = ACQMEASuINGFIELD_CIRCuAR00000000Focu = Smallest Flu Amplifier = Low Techniqu = Pused Flu Matriu Based = (height:10000000000000000width:10000000000000000epth:12) ValidRectangle = (bottom:1000000003,left:116,right:900000000,top:00000000 CluterMode = 16 510000000010000000018 Normal Operation Service NoDetailInfo. Reporting by Ms. Joanna Prabuka,technician50000000038000000006 Contact: Violetta- ward cardiac sugery ward. Today u Sebastian Walczak will be on site, neut to the apparatu, please give him the onsite card to be checked. It seems that the detector is damaged. Please order ugently: 9896000000000000000000000000 - Detector Pu0000000000000000High Speed pack And if the part arrives tomorrow - please send the person for replacement and calibration - approu. 00000000ous. M ----------------------------------------------------------------------------- | 2023-09-07 14:01 | No
247 | Yes | Power will turn on from time to time | Power will turn on from time to time | 000000000000000000000000001O##: 00000000000000000000000000100000000969968 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." September 12, 2000000003 Fuua Denshi Hokkaido Co.,Ltd. I sent the equpment along with a report to the on-site person in charge.September 19, 2000000003 Fuua Denshi Hokkaido Co.,Ltd. I sent the equpment along with a report to the on-site person in charge. Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: "Set the cartrige cover "message uer Impact: The AED always 000000000000000000000000001ks. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: I received the AED and when I performed a battery test, the message "Please attach the pad cover" was annouced.The same problem occured when he ued a test pad. "Based on the above problem, I determined that there was an operation problem with the cover recognition mechanism for the AED pads.The AED was fauty, so I replaced it. I apologize for the inconvenience this has caued." Reguatory Qustions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 When the device was stored in a rack for event storage, the power kept on tuning on sudenly. The soud stopped right after he talked to him, and when he checked the green lamp, the lamp was flashing as uul. The cutomer said he had the AED and when he retuned to the company uing a vehicle, he confirmed the same problem. Even if the pad's hand part is pressed a little, sometimes the power will tun on so please inspect it. Lot No. Adut pad C000000002622-00000000 Adut pad C000000002622-00000000 Battery 96355P ----------------------------------------------------------------------------- | 2023-09-07 14:11 | No
248 | Yes | Low Leak–CO0ebreathing alarm | Low Leak–CO0Rebreathing alarm |  | 2023-09-07 14:19 | No
249 | Yes | AED beeped with “To test, remove the battery” msg, beep recurred after battery reinstalled | AED beeped with “To test, remove the battery” msg, beep recurred after battery reinstalled | 000000000000000000000000001O##: 00000000000000000000000000100000000970000000038 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." September 1000000002000000003 Fuua Denshi Minamikanto Hanbai I sent the equpment along with a report to the on-site person in charge.September 200000000 2000000003 Fuua Denshi Minamikanto Hanbai I had the on-site person in charge confirm receipt. Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: I received jut the AED main uit and ued a test pad and battery and confirmed that it passed the bit.I verified the AED operating record and foud that the high voltage calibration items for the monthly auomatic self test were ouside the passed range, An uLLoccured and the AED beeped once and the message "To test, remove and re-install the battery" was annouced.The alarm has now been resolved, bu in order to ensue safety, "I will replace it. I apologize for the inconvenience this has caued. " Reguatory Qustions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 After the battery was re-installed, it retuned to normal with a self-test bu it beeped again.The "To test, remove the battery..." gudance was made.Please inspect. [Person in charge: ALSOK Kashiwa Branch, u Takahashi] Lot No. Adut pad C110000000022-00000000 Adut pad C110000000022-00000000 Battery 00000000300000000P ----------------------------------------------------------------------------- | 2023-09-07 14:41 | No
250 | Yes | The system did not start. | The system did not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000977928 00000000-Sep-2000000003 Kawamua, HidetsuuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Not ued Eupected and ueupected Resuts: The system does not start. uer Impact: None Patient Impact: None SE 8.1.17 Resoluion: I replaced the host PC. After I replaced it, I confirmed normal operation. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000972596 Ohgua, Yui The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Internal Comments: [Suject] 00000000DS:YO2:u Ibo: The system did not start. [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Kawamua, Hidetsuu[Eiot Kawamua] - FSE1 Visit date and time 2000000003/00000000/0000000018:300000000 - FSE1 000000000000000000000000001k time 00000000 [Euamination category] Other: Problem Description by Engineer: The system did not start. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system did not start. uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: RC Response: I euplained how to identify the failue. We plan to respond on site. ----------------------------------------------------------------------------- | 2023-09-07 14:46 | No
251 | Yes | Measurement module wrong components. | Measurement module wrong components. | 000000000000000000000000001O##: 0000000000000000000000000010000000097000000002 Parts Consued : 1.0000000000000000000000000000000056000000001811,Intrepid uSpO2+NBP+CO2 Frame ASSY ----------------------------------------------------------------------------- | 2023-09-07 15:06 | No
252 | Yes | AED emits single beep, “To test, remove the battery” voice message | AED emits single beep, “To test, remove the battery” voice message | 000000000000000000000000001O##: 0000000000000000000000000010000000097000000005 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." October 000000002000000003 Fuua Denshi Kanagawa Hanbai I sent the equpment along with a report to the on-site person in charge and contacted him. Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: "I performed a battery test on the AED and confirmed that it passed the test. I analyzed the AED operation record and foud that the alarm that had occured was a self test I determined that the high voltage calibration item was ouside the passed range, so the AED beeped once and the message "To ue, remove and re-install the battery" was made. "The alarm has now been resolved, bu I will replace the AED as it is in the fuue of a failue. I apologize for the inconvenience this has caued." Parts Consued: 1.0000000000000000000000000000000056000000008200000000HS1 + (W G) Defib, Japanese, Euchg ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The AED beeped even when the AED was not being ued. The message "To do the test, remove the battery and do the test again" was annouced. Lot No. Adut pad Y000000000000000021-00000000 Adut pad Battery 865700000000----------------------------------------------------------------------------- | 2023-09-07 15:11 | No
253 | Yes | u - No signals | M00A - u00 - No usignal | 000000000000000000000000001O##: 000000000000000000000000001000000009710000000000000000TEFANO ZAMBONI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Make nemo withou collection to: BuALINI HOSP. Viale Giovanni Ghirotti, 286 Cesena - 0000000021 C/O Health Physics Att.ne u Nicoluci, Alberti, Pasini 000000000000000052987 Gudi 0000000000000000528200000000 In this way the battery is not available, please check if it is available in Sidem. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in monitoring Eupected and ueupected Resuts: uno signal uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: ucable does not condut and battery is flat ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000978172 Parts Consued: 1.00000000000000000000000098980000000010000000091,100000000Lead uTruk AAMI/IEC 2m ----------------------------------------------------------------------------- | 2023-09-07 15:11 | No
254 | Yes | Wireless pedal does not work. | Wireless pedal does not work. | 000000000000000000000000001O##: 00000000000000000000000000100000000978127 September 19, 2000000003 Didier BACHELART The system meets the specifications for the performed service and is retuned to ue. Resoluion: Replacement of the Wi-Fi pedal. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Bio Device ue: 000000000000000000000000001n ou. Eupected and ueupected Resuts: eup uer Impact: No. Patient Impact: No. SE R2.2 Diagnostic performed by Engineer: Connection test of the Wi-Fi pedal. Parts Consued: 1.00000000000000000000000000000000800000000238232, Wireless Footswitch 3P ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000972876 September 7, 2000000003 Didier BACHELART The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Pedal ordered. Follow u Follow u Requred: replacement of the pedal. Diagnostic performed by Engineer: Checked the Wi-Fi pedal. Problem Description by Engineer: Cutomer Fuction/Role: Bio Device ue: 000000000000000000000000001n ou. Eupected and ueupected Resuts: eup uer Impact: No. Patient Impact: No. SE R2.2 Reguatory Qustions: No/No ----------------------------------------------------------------------------- | 2023-09-07 15:11 | No
255 | Yes | AED emitted single beep, failed self-test - log analysis requested. | AED emitted single beep, failed self-test - log analysis requested. | Internal Remarks: Failue content 2 The equpment beeped, and the equpment did not say anything when he pressed the i-buton.He removed the battery and re-fitted it, bu the self test did not start. Lot No. Adut pad C110000000022-00000000 Adut pad Battery 98255P ----------------------------------------------------------------------------- | 2023-09-07 15:16 | No
256 | Yes | System does not start and ends at 0:0 | System does not start and ends at 0:0 | 000000000000000000000000001O##: 00000000000000000000000000100000000977837 ROB VAN GAMEREN The system meets the specifications for the performed service and is retuned to ue Diagnostic performed by Engineer: Wybe Henstra will check/replace the BIOS battery of the host PC and will let u know. Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: start u Eupected and ueupected Resuts: System does not start and ends at 00000000:00000000 uer Impact: n/a Patient Impact: n/a SE 8.2.17 Diagnostic performed by Engineer: Log file does not show any uLL. However, message system clock corrected on 7-9 at 8:18. BIOS battery Host PC supected. ----------------------------------------------------------------------------- | 2023-09-07 15:21 | No
257 | Yes | System cannot be started | System cannot be started | 000000000000000000000000001O##: 00000000000000000000000000100000000979935 September 8, 2000000003 Adov, Aleuander The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. RR u Thronicke MT Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000003-10000000017 00000000:00000000:00000000 Parts Consued: 1.00000000000000000000000000000000800000000387582,PD. SCOMP. DCPS_20000000012V_5V 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP Resoluion: Reported problem confirmed: System does not start. Power suply and uS controller replaced, fuction test performed. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009781800000000eptember 7, 2000000003 Malhas, Ali The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: 1te service with 3 hous of 000000000000000000000000001k and 3 part. Replacement of uS, bat pack and power suply in M cabinet requred. Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. RR u Thronicke MT Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Resoluion: The reported faut is confirmed. Trouleshooting: The u-ray system coud not be powered u. Check M cabinet - > all components are off. B and E cabinets looked in. OK. The 1-phase uS showed an uLL uS bridged - > u-ray system partially powered u. M cabinet checked after switching on: 2000000002-5V power suply on the rear of the cabinet defective. No backu before - > MPDuproblem. Soluion: uS, bat. Pack and power suply mut be replaced. Main switch in M cabinet closed. The system does not correspond to the manuactuer's specifications in the tests performed and is not ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000977732 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. RR u Thronicke MT Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer: System does not start. Opened M cabinet with bulding services. uS displays uLL Deactivation of the uS by bulding services not possible. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: If an FSE is available today, schedue to deactivate the uS manully and to check whether HOST PC is OK. 2 hous of 000000000000000000000000001k If no FSE is available, schedue TE service with uS and battery pack, replacement and check 3 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Internal Remarks: Start-u problem, asked u Thronicke for suport by bulding services. He will get in contact. Then open M cabinet and look at/look at compuer? uS? Internal Remarks: Left a bou for FSE Malhas to take over and check the uS. Internal Remarks: u Malhas has to go to the system again tomorrow. He will clarify with u Adov whether he can take over the schedued maintenance alone so that he can go to Bad Saarow for the follow-u service. u Malhas will get in contact again. Cutomer comment: ## Appointment for 00000000.00000000. from approu. 10000000000000000hrs. with Jessica. Internal Remarks: Info by u Malhas that u Adov will take over the service on 9/8 after 9:00000000 am. The uers are informed abou the appointment ----------------------------------------------------------------------------- | 2023-09-07 15:26 | No
258 | Yes | u/Geometry move problem | 0/u FD0Geometry move problem | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: We have learned that Yang Chuyan, the salesperson of ou company, is following u. The contact information is 183 2856 3189, thank you T2 Activities:Problem Description: Detailed description of cutomer problems or symptoms: Cutomer feedback that there is a problem with the mechanical part, and engineers need to be on site for repairs Is the device itself fauty: Yes Trouleshooting Action: Critical Faut Code or Symptom: POSTItem: Force Sensor Done/Failed &Application uLL ENMV_at_startu_HIGH Remote diagnosis and test resuts: restarting has no effect PRS TR Statu: All Pass Repair Action: Sugested field repair steps: 1. Check whether the force sensor is damaged, and replace it if there is any problem. 2. Check whether the Geo Modue is normal. 3. Check the uLLreport of the machine and perform trouleshooting Sugested spare part uuand name: 00000000230000000035300000000 AD7 FORCE SENSOR &00000000800000000151715 GEO MODuE WP Kit Remarks for cutomer needs: contact the cutomer in time to confirm the Onsite time Safety Qustions: Case: 0000000022200000000500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Confirmation is requred for limited dispatch ----------------------------------------------------------------------------- | 2023-09-07 15:26 | No
259 | Yes | uproot ¦ unable to boot, stuck in startup screen | TempuPro ¦ unable to boot, stuck in startup screen | 000000000000000000000000001O##: 000000000000000000000000001000000009700000000900000000arco Vianello The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : escalated to BIu----------------------------------------------------------------------------- | 2023-09-07 16:06 | No
260 | Yes | The device is not triggering the u-ray. | The device is not triggering the u-ray. | BuQ&R Feedback: No Reason: Corrective Maintenance. SAP Transfer: Dispatched for Onsite Suport. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Euternal Remarks: 20000000000000000.2000000003 11:58:100000000PM POLARIS (POLARIS) ========== Lus Edurdo Pereira Dos Santos - 00000000000000005731173000000000000000000000000000000000000000089 ======================= Technician Manoel informs that the software does not start. We performed remote service along with it to perform the ghost. After software installation, the equpment 000000000000000000000000001ked normally and was released For ue. T2 Activities:Problem Description: Problem description by engineer: The device is not triggering the u-ray/ FSE Lus is already servicing Trouleshooting Action: What steps are taken duing the system diagnosis: Made 00contact with client Manoel who already reported that FSE Lus is already servicing. uLL# and/or description of uLL The device is not triggering the u-ray. OneEMS internal remarks: Problem description by engineer: The device is not triggering the u-ray/ FSE Lus is already servicing What steps are taken duing the system diagnosis: Made 00contact with client Manoel who already reported that FSE Lus is already servicing. uLL# and/or description of uLL The device is not triggering the u-ray. Repair Action "Internal": I have conversed with client Manoel, equpment was not rising this morning, becaue of a software failue. Duing the conversation, cutomer Manoel (000000001 9 9173 8600000000 said that he was already in contact with FSE Lus Edurdo. Speaking with FSE Lus Edurdo, he confirmed that he was already in service and the call will be sent to him. What are the neut Trouleshooting steps: Please send the call to FSE Lus Edurdo who is already servicing this failue. Eutra Internal comment field: FSE Miued Leonardo Meleti Problem Reported by Cutomer: The device is not triggering u-rays. Stopped Hospital: uimed Barra Address: avenida Airton Senna 2559 Barra da Tijua. ## engenharia.clinica00nimedrioempreendimentos.com.br Manoel Soua, Department: Clinical Engineering 552138831100000000 or 552138831100000000 Safety Qustions: Case: 000000002220000000088 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Sthefany Pereira Last udated by: Sthefany Pereira 1. Was the equpment in clinical ue when the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? N/A, Trouleshooting Action: FSE Lus Edurdo who is already servicing this failue. ----------------------------------------------------------------------------- | 2023-09-07 16:06 | No
261 | Yes | ray system does not boot up. | uray system does not boot up. | 000000000000000000000000001O##: 00000000000000000000000000100000000977566 00000000-Sep-2000000003 Dawid Strydom The system meets the specification for the performed service and is retuned to ue Follow u Requred : FleuVision PC startu problem Please qute for: Travel 1 u trip to site normal Travel 1 u trip from site Overtime Labou 5.5hrs u 1 FSE normal Labou 3hrs u 1 FSE overtime 1 u 0000000080000000058600000000 - FleuVision-2 Revised PC no HDD Problem Description by Engineer : Cutomer Fuction/Role: Onsite CE Device ue: OuSIDE Eupected and ueupected Resuts: SYSTEM NOT STARTING u uer Impact: NONE Patient Impact: NONE SE R8.2.u Resoluion : Assess & Repair done. Cutomer reported System not starting u. uon arrival onsite confirmed cutomer observation. Did onsite faut finding and diagnosed the problem to a defective FleuVision PC. Collected Local Stock part from warehoue and installed to system as per instrutions. Did all relevant software and configuations. System is 000000000000000000000000001king. Did all other relevant checks and no problems foud. Qutation to follow. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- | 2023-09-07 16:31 | No
262 | Yes | The system is not working in rotational. | The system is not working in rotational. | 000000000000000000000000001O##: 00000000000000000000000000100000000971000000007 TESTOLIN, DARIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: The system is not 000000000000000000000000001king in rotational. Lack of commuication with the interventional instruents. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer: Performed 00contact with u Auuto Apostoli - asked to verify space on disk - console cable connectors- He will call u back at the end of the day. Internal Comments: uLL# and/or description of uLL . Repair Action: . Resoluion: Gio 7-set-23 - 13:19 - the call back - confirming that after having rechecked the wiring the system started reguar operation ----------------------------------------------------------------------------- | 2023-09-07 16:31 | No
263 | Yes | u - device cannot turn on | QA0000000 - device cannot turn on |  | 2023-09-07 16:39 | No
264 | Yes | Stand movement possible to a limited extent | Stand movement possible to a limited eutent | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007937 Oct 000000002000000003 Schuze Niehus, Martin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Zöllner power suply in the R cabinet replaced. Fuction check. Variou repairs. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007837 Oct 000000002000000003 Schellenberg, Arthu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Sporadically no geometry startu Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Resoluion: Power & Control uit replaced, fuction test performed The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.00000000000000000000000000000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000978900000000September 25, 2000000003 Nölkes, Fabian The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: The spare part cannot be foud at the cutomer and Medical Technology is no longer in the bulding. Please arrange a new appointment and clarify in advance whether the spare part is available. Please swap parts on new 000000000000000000000000001 Problem Description by Engineer: Problem description by engineer: Sporadically no geometry startu Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: Spare part cannot be foud. Appointment postponed ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 September 25, 2000000003 Schluter, Simon The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE uknown Resoluion: Spare part coud not be foud. Repair coud not be performed. New appointment mut be arranged. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097800000000 Baugarten, Rüdiger The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: uLL# and/or description of uLL Configued frontal stand not foud Replace Power & Control uit (SCC) if problem persists. Repair Action: Te with part Follow-u requred: uLL# and/or description of uLL Configued frontal stand not foud Replace Power & Control uit (SCC) if problem persists. Repair Action: Te with part Technician name: Baugarten, Rüdiger Problem Description by Engineer: Problem description by engineer: Sporadically no geometry startu Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: Called cutomer, log analysis ----------------------------------------------------------------------------- Cutomer comment: Torsten Smyczynski by ## Please let me know the statu for this procedue. Cutomer comment: ### Appointment arranged with Ms. Kater for October 00000000rom 13:300000000 ### Internal Remarks: Info to u Smyczynski that the service will take place today Cutomer comment: By ## Please send me the delivery certificate for the spare part so that we can look for this in the hospital in a targeted way again. Torsten Smyczynski Mobile: +000000007257736800000000 mailto:torsten.smyczynski00iemens-healthineers.com Cutomer comment: u Smyczynski by ## The correct spare part has now arrived and is already becaue of the device. Please arrange an appointment with the uers at +000000000000000015000000002 or -0000000000000000Internal Remarks: A wrong part was delivered, a new part was sent and a new appointment was arranged tomorrow Cutomer comment: By ##from torsten.smyczynski00iemens-healthineers.com: Apparently a wrong spare part was sent to the cutomer by u. A compuer keyboard was in the package POWER (&CONTROL) uIT WITHOu BOARDS Please check the procedue internally again. See attachment Internal Remarks: Via u Nölkes Schedue u Schlüter for service, he knows. Cutomer comment: rings throuh at 1:36 pm at +00000000000000001500000000 Cutomer comment: ## Appointment info to uer, 22.00000000. from approu. 13:300000000# Cutomer comment: rings throuh 00000000:53 Cutomer comment: Torsten Smyczynski by ## ufortuately, not everything ran optimally here. Here the appointment mut be arranged again with the uers: Friday, 29.00000000.2000000003, aroud 13:00000000 hrs. is not yet fiued with u Kühn phone +0000000000000000150000000000000000 ufortuately, the spare part cannot be foud on site now. ----------------------------------------------------------------------------- | 2023-09-07 17:16 | No
265 | Yes | Audio-no sound/No audio found at u/ other HW-Speaker-HW/Replaced Parts | Audio-no sound/No audio found at Bench-Keyrin s/n u000/ other HW-Speaker-HW/Replaced Parts |  | 2023-09-07 18:28 | No
266 | Yes | Remote Alert: u: Fleuvision PC grabber cards are not initialized. | Remote Alert: RMT - FLV: Fleusion PC grabber cards are not initialized. | 000000000000000000000000001O##: 00000000000000000000000000100000000979532 Ketcherside Jeff The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : RMT - FLV: Fleuvision PC grabber cards are not initialized. 00000000-Sep-2000000003 00000000:39:00000000 How often is issu occuring? : Intermittent Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: M2M/Radar Device ue: No patient involved Eupected and ueupected Resuts: Remote Alert; RMT - FLV; Fleuvision PC grabber cards are not initialized. uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.300000000100000000Resoluion : contacted cutomer via e-mail to alert Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1. Fleuvision PC (TZ) grabber card initialization uLL The grabbercard is malfuctioning. This is a known issu. Repair Action : Service actions for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1.1. Please replace the Fleuvision PC (TZ). Diagnostic performed by Engineer : m2m states Remote Alert: RMT - FLV: Fleuvision PC grabber cards are not initialized. ----------------------------------------------------------------------------- | 2023-09-07 18:41 | No
267 | Yes | Audio-no sound/No audio found at u/other HW-Speaker-HW/Replaced. | Audio-no sound/No audio found at Bench-Keyrin s/n SN u00/other HW-Speaker-HW/Replaced. |  | 2023-09-07 18:57 | No
268 | Yes | Device cannot go into standby mode | Device cannot go into standby mode |  | 2023-09-07 19:03 | No
269 | Yes | Audio-no sound/No audio found at Bench-Keyrin s/nu/other HW-Speaker-HW/Replaced. | Audio-no sound/No audio found at Bench-Keyrin s/n u000/other HW-Speaker-HW/Replaced. |  | 2023-09-07 19:10 | No
270 | Yes | iMessage "release button to complete boot up" | uLLmessage "release button to complete boot up" | 000000000000000000000000001O##: 00000000000000000000000000100000000979900000000 00000000-Sep-2000000003 Ha Hoang The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uLLmessage "release buton to complete boot u" When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Morning boot u Eupected and ueupected Resuts: system did not fuly boot u. uLLmessage "release buton to complete boot u" uer Impact: uable to ue room Patient Impact: no patient impact uu00000000Resoluion : Remote TSO activated. uLL duing boot u. System operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000979379 Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : uLLmessage "release buton to complete boot u" When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Morning boot u Eupected and ueupected Resuts: system did not fuly boot u. uLLmessage "release buton to complete boot u" uer Impact: uable to ue room Patient Impact: no patient impact uu00000000Diagnostic performed by Engineer : 000000000000000000000000001ked with cutomer to try and determine the caue of the release buton uLL uLLlong indicates that Enable Movement signal is active. cutomer stated that he checked all switches to ensue they were not stuk. uable to clear uLL Internal Comments : uLL# and/or description of uLL: [uLLdescription=Control modue (Geo-Imaging) at pedestal: Control 29 TableTopLongLatJoystick is active at startu] Repair Action : Check table top long/lat joystick with operator control/fuctional test/inou control test to see if always active. If signal cannot be terminated replace controller. Not sue which controller they have as 2 are listed on uit. • Root Caue 1: One of the joysticks (TableTopLongLatJoystick) of the Geo-Imaging modue at pedestal is active at startu. o Service Action 1: Check the cable of the Geo-Imaging modue at pedestal in euam room 2. Pay special attention to the TableTopLongLatJoystick itself; any damage visible? 3. Replace Geo-Imaging modue at pedestal in euam room if the problem persists. 000000008000000002765300000000 Control Modue Tilt Fleu ER 00000000800000000276500000000- Control Modue ORT Fleu ER Possible issu with footswitch, possible too muh delay in activation of the 2 internal switches in footpedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the footswitch are not switching in correct sequnce/in time wrt each other. Follow u Requred : uLL# and/or description of uLL: [uLLdescription=Control modue (Geo-Imaging) at pedestal: Control 29 TableTopLongLatJoystick is active at startu] Repair Action : Check table top long/lat joystick with operator control/fuctional test/inou control test to see if always active. If signal cannot be terminated replace controller. Not sue which controller they have as 2 are listed on uit. • Root Caue 1: One of the joysticks (TableTopLongLatJoystick) of the Geo-Imaging modue at pedestal is active at startu. o Service Action 1: Check the cable of the Geo-Imaging modue at pedestal in euam room 2. Pay special attention to the TableTopLongLatJoystick itself; any damage visible? 3. Replace Geo-Imaging modue at pedestal in euam room if the problem persists. 000000008000000002765300000000 Control Modue Tilt Fleu ER 00000000800000000276500000000- Control Modue ORT Fleu ER Possible issu with footswitch, possible too muh delay in activation of the 2 internal switches in footpedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the footswitch are not switching in correct sequnce/in time wrt each other. Technician Name : Eckbold Glenn ----------------------------------------------------------------------------- | 2023-09-07 19:16 | No
271 | Yes | Audio-no sound/No audio found at Bench-Foulink s/n u/other HW-Speaker-HW/Replaced Parts | Audio-no sound/No audio found at Bench-Foulink s/n u000/other HW-Speaker-HW/Replaced Parts |  | 2023-09-07 19:18 | No
272 | Yes | Fleuvision PC out of service | Fleusion PC out of service | 000000000000000000000000001O##: 00000000000000000000000000100000000979956 September 11, 2000000003 Michel GOY The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Resoluion: Following pm Fleuvision PC ou of service Changed Fleuvision PC + software + backu + License. Operating tests OK. The device is operational again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097200000000 Parts Consued: 1.00000000000000000000000000000000800000000577362, FLEuVIEWING PC 00000000 mDP ----------------------------------------------------------------------------- | 2023-09-07 19:21 | No
273 | Yes | Wireless pedal not working | Wireless pedal not working | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003555 100000000Nov-2000000003 Simon Fallows The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Parts Consued : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P Resoluion : Fitted and tested new wireless Footswitch Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Wireless pedal not 000000000000000000000000001king. Wireless indicator light is flashing bu the pedal is not following any commands. uer Impact: NA Patient Impact: NA SEu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000980000000079 200000000Sep-2000000003 Simon Fallows The system meets the specification for the performed service and is retuned to ue Follow u Requred : Fit new wireless foot switch Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Wireless pedal not 000000000000000000000000001king. Wireless indicator light is flashing bu the pedal is not following any commands. uer Impact: NA Patient Impact: NA SEu Resoluion : Foud wireless foot switch connection indicator flashing red when foot switch was on or off. Reset wireless foot switch and it’s now 000000000000000000000000001king. Cutomer confirmed that the WFS had intermittently 000000000000000000000000001ked over t000000000000000000000000001days before completely failing. New WFS to be ordered. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000979771 Barry White The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FSE to site to investigate Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Wireless pedal not 000000000000000000000000001king. Wireless indicator light is flashing bu the pedal is not following any commands. uer Impact: NA Patient Impact: NA SEu Diagnostic performed by Engineer : FCO has already been carried ou ----------------------------------------------------------------------------- Cutomer comment : Confirmed with Kath for FSE to site 2000000009. Happy with visit as early in day as possible bu confirmed this may be aroud 100000000m du to travel ----------------------------------------------------------------------------- | 2023-09-07 19:41 | No
274 | Yes | Unable to Fluoro. | Unable to Fluoro. | Internal Remarks : Interaction Type: New Case Caller: beverly maufield IP & Issu: "6000000007600000000 uable to Fluro." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Created case and qute. Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Cutomer Call Back Caller: Beverly Maufield IP & Issu: CCB stating that she gave a wrong ##for qute and case and needed to udate the correct ## Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: udated ##on case and created an AA for qute ##to be resent to cu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: beverly maufield IP & Issu: 6000000007600000000 uable to Fluro. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: FSE arrival udate. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Cutomer Call Back Caller: beverly maufield IP & Issu: ccb- requsting that no one shou be sent ou today bu they need service first thing Monday Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: AA- ST RP Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Planner Initials: GM Action taken: FSE Chris Lim accepted call for 9/700PM called cu spoke to Beverly Maufield and confirmed Oucome: dispatched ----------------------------------------------------------------------------- | 2023-09-07 19:41 | No
275 | Yes | System is not booting up properly | System is not booting up properly | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000299 00000000-Dec-2000000003 Rodriguz Lazaro The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: u-ray tech Device ue: neuolab Eupected and ueupected Resuts: system not finishing boot u sequnce. uer Impact: uavailable Patient Impact: None SE 8.1.3000000005 Diagnostic performed by Engineer : Done on previou 000000000000000000000000001 Resoluion : Replaced the fleuvision PC along with the hard drive. Installed software on it and configued. Tested with the system and is operational. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006777 21-Nov-2000000003 Rodriguz Lazaro The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: u-ray tech Device ue: neuolab Eupected and ueupected Resuts: system not finishing boot u sequnce. uer Impact: uavailable Patient Impact: None SE 8.1.3000000005 Diagnostic performed by Engineer : Replaced the fleuvision compuer bu it was DEFOA. New compuer is being ordered along with a new hard drive as well. Follow u Requred : Replace the fleuvision compuer along with the hard drive as well. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002600000000 00000000-Oct-2000000003 Henry,Delvin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: u-ray tech Device ue: neuolab Eupected and ueupected Resuts: system not finishing boot u sequnce. uer Impact: uavailable Patient Impact: None SE 000Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Resoluion : replace fleuvision pc and reinstall software on new hdd ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000993198 18-Sep-2000000003 Rodriguz Lazaro The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: u-ray tech Device ue: neuolab Eupected and ueupected Resuts: system not finishing boot u sequnce. uer Impact: uavailable Patient Impact: None SE 000Diagnostic performed by Engineer : Diagnostics done on previou 000000000000000000000000001 Follow u Requred : Replace the fleuvision compuer. Follow u Requred : Had DEFOA fleuvision compuer and need to replace compuer again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000972817 11-Sep-2000000003 Rodriguz Lazaro The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: u-ray tech Device ue: neuolab Eupected and ueupected Resuts: system not finishing boot u sequnce. uer Impact: uavailable Patient Impact: None SE 000Diagnostic performed by Engineer : Foud that the fleuvision compuer isn't finishing booting u and needs to be replaced. Follow u Requred : Replace the fleuvision compuer ----------------------------------------------------------------------------- Internal Remarks : Case: 000000002220000000032 FSE: Delvin, Henry Hemo video going green to FleuVision Monitor bu only in a specific spot on FleuVision Monitor; displays properly everywhere else on FleuVision Monitor. FSE foud loose connection on FleuVision PC. Internal Remarks : Case: 000000002220000000032 FSE: Delvin, Henry FleuVision PC replace; fails to reload SW. Requsted logs for review. FSE will call back in once logs have been sumitted. Internal Remarks : Planner Initials: GM Action Taken: Follow u activity for FSE Lazaro Rodriguz for 9/2500PM Oucome: dispatched Internal Remarks : FSE Delvin called having an issu reloading software on the FleuVision (FV) PC. I was able walk FSE throuh the procedue for install the app as it appeared the OS loaded. I also walked FSE throuh the 2k2 and firmware load. Walked FSE throuh the import configuation. Du33799 ----------------------------------------------------------------------------- | 2023-09-07 19:51 | No
276 | Yes | Check Vent codes when powered on due to 0 failure | Check Vent codes when powered on due to 0V failure | 000000000000000000000000001O##: 00000000000000000000000000100000000980000000086 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the MC pcba euchange for possible ou of bou failue. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcodes after powering on uit. uer Impact: none Patient Impact: none SE 2.300000000Diagnostic performed by Engineer : Cutomer has a V600000000and it has a code 1000000007. He states that they replaced the cable from the DA pcba to the MC pcba and the DA pcba, bu there was still an active code 1000000007. After replacing the MC pcba, there was no code 1000000007, bu now he is new check vent alarm. Code 1119 and 1127 is in the history. The PM pcba was replaced 211 hous ago duing the FCO86600000000000000006 remediation. Recommended having they replace the MC pcba, and have it euchanged for a possible ou of bou failue. They originally puchased it from Parts Souce. ----------------------------------------------------------------------------- | 2023-09-07 19:56 | No
277 | Yes | u biplane lab gets stuck in a boot loop intermittently and ally needs Suite PC software reloaded. | uB0biplane lab gets stuck in a boot loop intermittently and ually needs Suite PC software reloaded. |  | 2023-09-07 20:00 | No
278 | Yes | u Backup alarm failed | 00 Backup alarm failed | 000000000000000000000000001O##: 00000000000000000000000000100000000980000000097 Carey, Scott The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: 3.00000000 Date and Time of Issu: 00000000/00000000/2000000003 Trouleshooting Steps Taken: Advised Biomed the recommended repair is to replace the CPuboard. Provided cutomer with the part # for the V600000000CPuboard. Part # 00000000561537200000000PCBA, CPu(software 2.300000000 List price $792 (23) in stock. Cutomer advised to reference page 100000000of the V600000000service manul for restoring the V600000000serial # and programming the power on hous. Cutomer advised to call back for restoring V600000000puchased options. Cutomer has called back requsting the option codes. V600000000serial # 200000000000000001800000000 New CPuserial # SS2100000000YAJ AVAPS C-Fleu Ramp HFT Was PRS ued: No PRS Permission: Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display diagnostic code 1100000000Backu alarm failed. uer Impact: The V600000000was removed form service. Patient Impact: No impact SE 3.00000000 Problem Description/ueupected: V600000000displayed diagnostic code 1100000000 Resoluion : Provided Biomed with the following option code information. Biomed is reporting the option code restoration was sucessfu. V600000000serial # 200000000000000001800000000 New CPuserial # SS2100000000YAJ AVAPS option # 000000005000000005372 C-Fleu option # 00000000527355700000000 Ramp option # 0000000058715768 HFT option # 00000000626761600000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Travis Dickey IP & Issu: 721913500000000Cutomer requsted tech suport/board recently installed/having issu finding software Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-07 20:11 | No
279 | Yes | uuunot happening | uuunot happening | 000000000000000000000000001O##: 00000000000000000000000000100000000980000000030000000016-Sep-2000000003 VEMBARASAN K The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Problem Description by Engineer : Problem description by engineer : uuunot happening When was the first occurence? Within... : 00000000-Sep-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: uuunot happening uer Impact: No uer impact Patient Impact: NO patient impact SE 8.2.17.00000000Resoluion : System software corruted, visited site and checked system system doesn’t booting, reloaded the software , reconfigued, recalibrate system data necessary requres, system made fuctional after reload the sw , ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000980000000083 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : uuunot happening When was the first occurence? Within... : 00000000-Sep-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: uuunot happening uer Impact: No uer impact Patient Impact: NO patient impact SE 8.2.17.00000000Diagnostic performed by Engineer : OS disk supected here as host pc software got corruted. Internal Comments : uLL# and/or description of uLL: uuunot happening Repair Action : OS disk supected here as host pc software got corruted. Resoluion : OS disk supected here as host pc software got corruted. Follow u Requred : uLL# and/or description of uLL: uuunot happening Repair Action : OS disk supected here as host pc software got corruted. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-07 20:36 | No
280 | Yes | Nav-ring issue | Nav-ring issue |  | 2023-09-07 20:58 | No
281 | Yes | Customer could not make fluoro | Cuomer could not make fluoro | 000000000000000000000000001O##: 00000000000000000000000000100000000980000000012 Sharvari Karua The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Cutomer is uable to make fluro with wireless footswitch Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: none Patient Impact: none SE 1.2.5.00000000Diagnostic performed by Engineer : The log does not indicate any uLL. The staff might be pressing on the flur save buton when the physician is trying to make fluro. Internal Comments : uLL# and/or description of uLL: none Repair Action : Will order base station for the wireless footswitch. Resoluion : Replace the wireless footswitch base station. Follow u Requred : uLL# and/or description of uLL: none Repair Action : Will order base station for the wireless footswitch. Technician Name : Sharvari Karua ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000980000000087 Parts Consued : 1.00000000000000000000000000000000800000000257862,WFS Base Station ----------------------------------------------------------------------------- | 2023-09-07 21:11 | No
282 | Yes | AED triple Chirp | AED triple Chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000981000000002 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer called with the AED chirping. She stated that when the battery is inserted it keeps chirping. I have instruted to insert the battery and the AED has failed the BIT. The aed is triple chirping. Device serial# A17I-0000000000000000 Serial # of battery B00000000200000000700000000 Battery lot#: a00000000200000000 Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Device was not in ue Eupected and ueupected Resuts: Cutomer called with the AED chirping. She stated that when the battery is inserted it keeps chirping. I have instruted to insert the battery and the AED has failed the BIT. The aed is triple chirping. Device serial# A17I-0000000000000000 Serial # of battery B00000000200000000700000000 Battery lot#: a00000000200000000 uer Impact: none Patient Impact: nonw SE Follow u Requred : Ship to Nauann grou 20000000000000000Capital NE Tallahassee Fl 32300000000 Attention Chris Zweydorff Internal Comments : Waiting for order to ship, order nuber 0000000000000000573187300000000Resoluion : DELIVERED Friday 9/8/23 at 11:00000000AM Signed for by: C.ZWEYDORFF ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000981115 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-07 21:21 | No
283 | Yes | A failure in the navigation ring occurred. | A failure in the navigation ring occurred. |  | 2023-09-07 21:31 | No
284 | Yes | Navigation ring failure | Navigation ring failure |  | 2023-09-07 21:39 | No
285 | Yes | Machine is automatically got shutdown and patient is in table | Machine is automatically got shutdown and patient is in table | 000000000000000000000000001O##: 00000000000000000000000000100000000981197 11-Sep-2000000003 Satam, Kamlesh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is auomatically got shudown and patient is in table When was the first occurence? Within... : 00000000-Sep-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is auomatically got shudown and patient is in table uer Impact: No uer impact Patient Impact: No patient impact SE 000Resoluion : Carried ou necessary action for Residul curent device tripping. Removed iron core of 3-phase suply aroud which CBCT is being mouted. Replaced u coil of uS oupu MCCB. Now checked RCD device & foud to be 000000000000000000000000001king fine & no tripping observed again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009769200000000ijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Machine is auomatically got shudown and patient is in table When was the first occurence? Within... : 00000000-Sep-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is auomatically got shudown and patient is in table uer Impact: No uer impact Patient Impact: No patient impact SE 000Diagnostic performed by Engineer : RCD got tripped. FSE to visit the site and do futher checks on power failue. Internal Comments : uLL# and/or description of uLL: Machine is auomatically got shudown and patient is in table Repair Action : RCD got tripped. FSE to visit the site and do futher checks on power failue. Resoluion : RCD got tripped. FSE to visit the site and do futher checks on power failue. Follow u Requred : uLL# and/or description of uLL: Machine is auomatically got shudown and patient is in table Repair Action : RCD got tripped. FSE to visit the site and do futher checks on power failue. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-07 21:46 | No
286 | Yes | u reference failed, device shut down | EC 0A - DA PCBA ADC reference failed, device shut down | 000000000000000000000000001O##: 0000000000000000000000000010000000098100000000 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:Vent inop while on patient uer Impact: not uable needs repair Patient Impact: none reported SEun Resoluion : Advised cutomer to replace DA pcb and DA-MC cable to address uLL cutomer acknowledged and no futher follow u needed, provided part #s, issu resolved- Material Nuber 00000000561512561 - RP-PCBA, Data Acqusition, V600000000$1286.00000000 Material Nuber 00000000561536131 - RP-Cable,DA to MC PCBA,2nd Gen,V600000000 $159.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states duing patient transport the uit was reported to have alarmed Vent inop and screen went black with no ventilation, no patient harm reported and uit was removed from patient withou futher incident, in biomed complaint was not dulicated bu uLL100000000A was noted in event log, 3.3V shows good in Pneuatics bu 2.5V was not measued at DA pcb, DA-MC cable is connected and seated, requsts futher trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Caller: Aleu 717-5000000005967 IP & Issu: Broken screen Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-07 22:51 | No
287 | Yes | Tube defect u ray not possible | 00CDN0Tube defeuu ray not possible | 000000000000000000000000001O##: 00000000000000000000000000100000000983323 00000000-Sep-2000000003 MuH, JOE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Tue defect u ray not possible When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Tue defect u ray not possible uer Impact: System is down Patient Impact: Slight delay in case to get C arm brouht in/ in progress SE uknown Parts Consued : 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V Resoluion : Replaced the 5V Power Suply in the Gen cabinet. This resolved the reported issu. Tested system and it is fuly fuctional. System back u and retuned to ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000981736 00000000-Sep-2000000003 MuH, JOE The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace 200000000Power Suply Problem Description by Engineer : Problem description by engineer : Tue defect u ray not possible When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Tue defect u ray not possible uer Impact: System is down Patient Impact: Slight delay in case to get C arm brouht in/ in progress SE uknown Diagnostic performed by Engineer : Puled uLLlogs and reviewed with NSS. Coudn't find anything specific within logs that pointed to a specific part needing replaced. Got back into room and did some futher trouleshooting. Discovered fauty 200000000Power Suply on Gen Cabinet. Ordering part to replace neut day. Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000981600000000 Whitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Tue Defect u ray not possible Repair Action : Check logs for more details of the uLL check the tue cooling uit and flud levels, check Generator to see if it is powered on. Replace defective parts foud Follow u Requred : uLL# and/or description of uLL: Tue Defect u ray not possible Repair Action : Check logs for more details of the uLL check the tue cooling uit and flud levels, check Generator to see if it is powered on. Replace defective parts foud Technician Name : Whitaker, Steven Problem Description by Engineer : Problem description by engineer : Tue defect u ray not possible When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Tue defect u ray not possible uer Impact: System is down Patient Impact: Slight delay in case to get C arm brouht in/ in progress SE uknown Diagnostic performed by Engineer : Adam reports that they are getting Tue defect and no u ray is possible, hard reboot did not resolve. They brouht a C arm in to finish the case, no patient or uer harm reported. System is not connected to PRS for remote logging ----------------------------------------------------------------------------- | 2023-09-07 23:31 | No
288 | Yes | system will not boot up | system will not boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000097000000001 100000000Sep-2000000003 Shula Chintan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Replaced blown fue from the spare fue kit. System booted u pass 00000000:00000000 cout down. I was also able to produe u-ray. The fue was blown becaue of the defective power suply on Geo IPC. I also noticed the uuumodue in control room is not booting u properly. Separate qute for Geo IPC and TSM-b is provided to cutomer and will be replaced on different case uon receiving approval from cutomer. Problem Description by Engineer : Cutomer Fuction/Role: In-houe imaging tech Device ue: First boot of the system Eupected and ueupected Resuts: System did boot pass 00000000:00000000 uer Impact: uable to ue system Patient Impact: None SE000Diagnostic performed by Engineer : Reviewed log file. There were mutiple uLLlogged regarding Host PC and SIB commuication. I futher trouleshot the issu to find blown fue on MPDu ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Dinh Tran IP & Issu: 5332700000000system will not boot u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: n/a Action Taken: CREATED CASE &SENT A QuTE AND CREATED AA FOR ONSITE Neut Step Action: Awaiting Qute Verbal PO Nuber: 1800000000000000007586serv Verbal PO Approver: Dinh Tran Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-07 23:46 | No
289 | Yes | AED is emitting the triple chirp (Removing the AED from service) | AED is emitting the triple chirp (Removing the AED from service) | 000000000000000000000000001O##: 00000000000000000000000000100000000982276 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported AED is emitting the triple chirp uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported AED is emitting the triple chirp Resoluion : Closing case AED no longer has warranty. Diagnostic performed by Engineer : Cutomer reported that they changed the battery and pads and when they did the self-test he stated the AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 12 years old, the issu is with the device. Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A11K-0000000000000000 Battery Install Before# 2000000006/00000000 Lot#9000000007P Pads Eup Date: 1/13/2000000003 Pads Lot # ----------------------------------------------------------------------------- | 2023-09-08 01:01 | No
290 | Yes | Triple chirp, failed self-test with new batteries | Triple chirp, failed self-test with new batteries | 000000000000000000000000001O##: 00000000000000000000000000100000000982382 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to acct address Attn Darrin Smith Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: failed self test uer Impact: none Patient Impact: none SE E.00000000.000000007 See ref case 000000002223600000000 Cut called stating device was not powering on with curent battery that was installed a year ago. Diagnostic performed by Engineer : Had cut insert said battery and it immediately emitted triple chirps and hitting the i-buton did not provide any verbal uLLmsg and it jut continud to triple chirp. Had cut insert another 000000000000000000000000001king battery and it had the same oucome. Neut had the cut insert the original battery into another AED and it passed self test. Advised cut to remove device B16I-00000000100000000from service and will replace uder warranty. 00000000000000009P Install before date: 5/3000000002000000002 Pads lot# 22122000000000000000052 Eup date: 6/2000000002000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000098200000000 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- Internal Remarks : FedEu tracking 78300000000000000000000000000000000 DELIVERED Friday 9/8/23 at 8:36 AM ----------------------------------------------------------------------------- | 2023-09-08 01:21 | No
291 | Yes | u showing red u on machine and beeper noise | u showing red u on machine and beeper noise | Internal Remarks : Caller: Caleb Bowman IP & Issu: 000000007800000000new battery and power suply and still showing red u on the machine and beeper noise. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: created r000000000000000000000000001roued Neut Step Action: N/A Verbal PO Nuber: na Verbal PO Approver: na Verbal PO CAP: na ----------------------------------------------------------------------------- | 2023-09-08 02:01 | No
292 | Yes | check mark/enter button is not functional | check mark/enter button is not funuonal | 000000000000000000000000001O##: 0000000000000000000000000010000000098200000000 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000has a fuctional check mark enter buton. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000check mark enter buton is not fuctional. Diagnostic performed by Engineer : Problem Description: V600000000check mark enter buton is not fuctional. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/00000000/2000000003 Trouleshooting Steps Taken: Cutomer has advised the uit has been udated to the 2ng generation front bezel. Advised cutomer the likely failue is with the u Switch board. Biomed is requsting onsite service to resolve. Part # 00000000561512600000000 PCBA, switch (u) Follow u Requred : V600000000check mark enter buton is not fuctional. Part # 00000000561512600000000 PCBA, switch (u) ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Michelle Campuano - phone 718 250000000060000000000000000 IP & Issu: 5662900000000 - The nav-ring is not 000000000000000000000000001king Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-08 02:16 | No
293 | Yes | The display is dark | The display is dark | 000000000000000000000000001O##: 00000000000000000000000000100000000982559 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolv Follow u Requred : V600000000display is dark. FSE to determine if the V600000000has a HYDIS or NEC display installed. V600000000u assembly is on backorder uLL with 000 100000000000000009779 and later and 200000000000000005565 and later are equpped with the 2nd-generation (NEC) LCD and associated parts. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000have a fuctional display. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000display is dark Diagnostic performed by Engineer : Problem Description: V600000000display is dim. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/00000000/2000000003 Trouleshooting Steps Taken: Biomed is requsting onsite service to resolve. Advised cutomer they may have a HYDIS displayed installed that 000000000000000000000000001ld requre udating V600000000to a 2nd generation NEC display. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Michelle Campuano - phone 718 250000000060000000000000000 IP & Issu: 5760000000087 - The display is dark Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-08 02:21 | No
294 | Yes | uro monitor not staying calibrated | Tempuro monitor not staying calibrated | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000021 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced Display Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Device was soaked for 3 weeks with power cycles and reguar testing. New Display Corrected previou touh issus. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king correctly. We will order a Display Assembly and test the device futher. Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008381 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair Diagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king correctly. We will order a Display Assembly and test the device futher. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000982726 Harris, Alvin The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Problem Description: Backgroud is grey ? per cus ##response TempuPro monitor not staying calibrated The front panel on this TempuPRO was jut replaced bu is again having issus with calibration. It happens while the monitor is on the smartmout and off the smartmout. uer is attempting to take vitals and is uable to puh the buton on the screen that they need to ue. SN 615731. Form ued: Bench Trouleshooting Steps Taken: Foud 000000000000000000000000001O#in quu, cunot on the phone, accepted 000000000000000000000000001O# 1st attempt - 00000000/00000000 - 5:18 PM - Voic##bou not yet set u - uable to leave a message - Sent ## --- Checked ##on 9/15/23, foud response on 9/7/23 from cu Hi Alvin, I’m not sue why my backgroud is grey. Sorry I missed you call, I was on the other line. Nick Mozayanfar is very familiar with this device and I have spoken with him already. Thank you --- Foud Irvine CA Bench form was sent on 9/7/23 to cuby Qust Bench Admin - Nikzad Mozayanfar indicating the cuhas already made plans to send the device back to the Bench --- Called cuback at 319 33000000005595, left a voic## sent ##to cu Hi Christy. I apologize for the delayed response. I jut called back to 319 33000000005595 and left a voice message. I see that Nick has already sent youa Bench repair form with instrutions to send you device back to u for repair. I assue youare planning to retun the device or have already retuned the device to the Bech. I will create the necessary Bench repair 000000000000000000000000001O#so that the technicians will have what they need to log their time and notes. Feel free to give u a call or ##u shoud youneed futher assistance. Thank youResoluion : Sent to Bench Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts:Backgroud shoud not be grey ? uer Impact:uknown, have not spoken with cuvia phone Patient Impact:No impact SEuknown, have not spoken with cuvia phone Problem Description/ueupected Resuts: Backgroud is grey ? per cus ##response Follow u Requred : ATT: Christy Soll/Case 000000002220000000029 Irvine CA Bench form already ##d to cuby Qust Bench Admin - Nikzad Mozayanfar ----------------------------------------------------------------------------- | 2023-09-08 03:11 | No
295 | Yes | uPro: Screen calibration off. | TempuPro: Screen calibration off. | 000000000000000000000000001O##: 00000000000000000000000000100000000982385 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to bench. Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: Screen calibration off. uer Impact:uable to ue device util repair is made. Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Screen calibration off. Internal Comments : uable to reach cutomer by phone on 9/7/2000000003 and 9/8/2000000003. ##d cutomer to ask when they will be available to trouleshoot on the phone. Follow u Requred : Cutomer reports "The front panel on this Tempu PRO was jut replaced bu is again having issus with calibration. It happens while the monitor is on the smartmout and off the smartmout. uer is attempting to take vitals and is uable to puh the buton on the screen that they need to ue." Cutomer requst bench repair Diagnostic performed by Engineer : Problem Description: Screen calibration off. Cutomer callback Tempu Pro: Screen calibration off. Form ued: Bench Trouleshooting Steps Taken: CSP - Web Tempu Pro: Screen calibration off. Cutomer reports "The front panel on this Tempu PRO was jut replaced bu is again having issus with calibration. It happens while the monitor is on the smartmout and off the smartmout. uer is attempting to take vitals and is uable to puh the buton on the screen that they need to ue." Cutomer requst bench repair ----------------------------------------------------------------------------- | 2023-09-08 03:16 | No
296 | Yes | system had anode uLL | system had anode uLL | 000000000000000000000000001O##: 00000000000000000000000000100000000982131 12-Sep-2000000003 Prunek Tom The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: cv Eupected and ueupected Resuts: anode uLL uer Impact: uLL duing procedue Patient Impact: none SE 7.8 Resoluion : Assist with u-ray tue change on u system ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000097500000000 12-Sep-2000000003 Seminatore Christopher The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: manager Device ue: Vasc Lab Eupected and ueupected Resuts: Annode uLL uer Impact: have to ue other room Patient Impact: none SE000Diagnostic performed by Engineer : Per RSM & NSS go on site and replace uray tue for anode uLL. Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000006-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-08 03:31 | No
297 | Yes | u:System starting 00 and it stopped and the system won't start. | 0DS:AW:System starting 00 and it stopped and the system won't start. | 000000000000000000000000001O##: 00000000000000000000000000100000000982916 00000000-Sep-2000000003 Niiya, Kosue The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Not ued Eupected and ueupected Resuts: The equpment does not start. uer Impact: No effect Patient Impact: No effect SER8.1.3000000005 Resoluion: The uit that controls the display on the large euamination monitor (Fleu PC) was fauty, so the equpment did not start. I replaced the Fleu PC and reinstalled the software and entered the settings. [Parts ued] Part No.: 0000000080000000058600000000/Part Name:FleuVision-2 Revised PC no HDD/Nuber of hard drives: 1 Follow u Requred: N/A Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000982887 Watanabe, Atsuhi The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: [Suject] 00000000DS:AW:u Henmi: "System starting 0000000000000000" and the system does not start. [Internal Comments] Comp 0000000080000000012200000000 u Haswell Fleuvision 2 PC no HDD is only in stock in Nagoya and Tokyo.Becaue the euaminations have been stopped, I will remove the 0000000080000000058600000000 FleuVision-2_Revised PC no HDD in IB from Osaka. The FV2PC information is below. Information 2000000003-00000000-00000000 00000000:00000000:31.0000000000000000Collected information of the grabber cards: Grabber driver version: 1.1000000003.00000000Grabber DLL version: 1.1000000003.00000000Grabber card 00000000revision: 1 Grabber card 00000000type: DVI_Grabber_Board_AM3_ 00000000 Grabber card 0000000000031735200000000165 Grabber card 00000000firmware version: 1.1000000001.00000000Grabber card 1 revision: 1 Grabber card 1 type: DVI_Grabber_Board_AM3_00000000 Grabber card 1 00031735200000000100000000 Grabber card 1 firmware version: 1.1000000001.00000000Grabber card 2 revision: 1 Grabber card 2 type: DVI_Grabber_Board_AM3_00000000 Grabber card 2 000317352000000001500000000Grabber card 2 firmware version: 1.1000000001.00000000 Startu FleuVision Information 2000000003-00000000-00000000 00000000:11:000000000000000000000000FleuVision Monitoring (FleuVision PC : S.M.A.R.T.) : PASSED Disk: IDE\DiskWDC_WD1000000005FBYZ-00000000YCBB2__________________RRP7____\5&2cb7a9000000000000000000000000000000000000000000000000ID Name Cu Mau Thr Statu 1 Read uLLRate 200000000 200000000 51 OK Pre-Failue Online 3 Spin-u Time 178 178 21 OK Pre-Failue Online 00000000tart/Stop Cout 99 99 00000000OK Advisory Online 5 Reallocated Sectors Cout 200000000 200000000 100000000OK Pre-Failue Online 7 Seek uLLRate 100000000 200000000 00000000OK Advisory Online 9 Power-On Hous 79 79 00000000OK Advisory Online 100000000Spin Retry Cout 100000000 100000000 00000000OK Advisory Online 11 Recalibration Reties 100000000 100000000 00000000OK Advisory Online 12 Power Cycle Cout 99 99 00000000OK Advisory Online 16 uknown 00000000200000000 00000000OK Advisory Online 183 uknown Normal Operation FleuVision [Parts from the equpment] 0000000080000000058600000000:FleuVision-2_Revised PC no HDD QTY:1 ZREP/Traceable --None-- 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID --None-- [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Kosue Shintani] Niiya, Kosue - FSE1 Visit date and time 2000000003/00000000/000000008:300000000 - FSE1 000000000000000000000000001k time 3 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Kosue Shintani - 0000000000000000008673-5236 - Others Call mobile uon arrival Requster name Atsuhi Watanabe [Euamination category] Other: Follow u Requred: * Problem Description by Engineer: "System starting 0000000000000000" and the system does not start. Cold restarting the system does not improve the sitution. On fleuvision, only the grid line is displayed, nothing else is displayed, and the "Switching not possible" message is displayed on the Touh Screen Modue. Euaminations are schedued to start from 9:300000000 The B cabinet cover does not have a 3mm key, and it cannot be opened. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: "System starting 0000000000000000" is displayed and the system does not start. Cold restarting the system does not improve the sitution. On fleuvision, only the grid line is displayed, nothing else is displayed, and the "Switching not possible" message is displayed on the Touh Screen Modue. Euaminations are schedued to start from 9:300000000 The B cabinet cover does not have a 3mm key, and it cannot be opened. uer Impact: Euamination delay Patient Impact: Euamination delay SE 8.1.3000000005 Diagnostic performed by Engineer: RC Response: Cold restart: OK I euplained how to do it: OK Euplanation regarding failue locations: OK An FSE will investigate on site. ----------------------------------------------------------------------------- | 2023-09-08 03:56 | No
298 | Yes | Monitor / uLLreport. | 00/ u000/ uDefibrillator/Monitor / uLLreport. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: ECR: 00000000/HeartStart uL+ Defibrillator/Monitor Safety Qustions: Case: 000000002220000000033 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? no 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-08 06:56 | No
299 | Yes | uLLIntermittent touch not working | 0/V0uLLIntermittent touch not working | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Occasional touh frame failue, this device is not within the scope of the recall, want to repair or replace Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 0000000022200000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? no 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? no Trouleshooting Action: Intermittent touhscreen failue, calibrating the touhscreen does not fiu it, no other failues.Trouleshoot touh screen Neut Step Action: Follow u by ASP hukeqiangyi ----------------------------------------------------------------------------- | 2023-09-08 09:26 | No
300 | Yes | u/Repl. kit ethernet switch R/M cab. problem | SA:0R/u FD0Repl. kit ethernet switch R/M cab. problem | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: SA:00000000R/u uuuFD200000000switch buned ou Euternal Remarks: 00000000.00000000.2000000003 00000000:00000000300000000TC (86000000008277) Service content: The device cannot be tuned on to enter the system. After inspection, the F12 fue blew. After replacement, it retuned to normal. Service resut: The equpment retuned to normal and all fuctions are 000000000000000000000000001king normally. SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002220000000087 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD 282200000000500000000315 1, ou of stock and with Ord.blk lock.Curently the RSE replacement nuber is being confirmed, uu Yu -86-1330000000010000000011 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.00000000.2000000003 00000000:0000000037 uC (86000000008277) "RDI Accuacy:YES" $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem, system cannot start u) $Q8: (check power suply R cabniet no power suply and F12 was broken) $Q9: (Replace F12 then restart system was ok) $Q100000000 (Test and inspection data refert to PA Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: N/A. ^Curent Tue installed date: NA SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: The switch in cabinet R is broken and needs to be replaced.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: The switch in cabinet R is broken and needs to be replaced.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: The switch in cabinet R is broken and needs to be replaced.Engineers are requred to be present Recommended spare parts uuand designation: " 282200000000500000000315 FAST ETHERNET SWITCH 16 Port 100000000100000000MB " Notes on cutomer requrements: Onsite T2 Activities:Repair Action: Alternative nuber confirmation: 00000000800000000239552 REPL. KIT ETHERNET SWITCH R/M/K CAB ----------------------------------------------------------------------------- | 2023-09-08 10:01 | No
301 | Yes | System does not start up | System does not start up | 000000000000000000000000001O##: 00000000000000000000000000100000000976918 00000000-Sep-2000000003 Verckens, Maarten The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: No. Eupected and ueupected Resuts: yes uer Impact: yes Patient Impact: no SE 8.1 Resoluion: PC crash Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000976915 PATRICK ATTN. The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer: No remote soluion possible, an on-site visit has been schedued Assistance to the onsite FSE, ordering parts and ordering taui ----------------------------------------------------------------------------- | 2023-09-08 11:51 | No
302 | Yes | u - Power cord malfunction caud spark - u-patient monitor - Changed power cord | QA 00000000- Power cord malfunuon caud spark - 00UT00 patient monitor - Changed power cord |  | 2023-09-08 13:19 | No
303 | Yes | Not booting up - uLL- Wrong protocol selected - tried to restart is not booting | Not booting up - uLL- Wrong protocol seleud - tried to restart is not booting | 000000000000000000000000001O##: 0000000000000000000000000010000000098000000005 11-Sep-2000000003 Nigel Huphries The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000660000000093,u-ray PC Monoplane Resoluion : Defective u-ray PC replaced, software downloaded and restored backu. Fuctional checks good. System retuned to service. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: black screens in CR room , 1/2 screen on FV bu not booting to application uer Impact: NA Patient Impact: NA SE1.2. Reguatory Qustions : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009819100000000arry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to futher investigate , supect u Ray PC reload requred Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: black screens in CR room , 1/2 screen on FV bu not booting to application uer Impact: NA Patient Impact: NA SE1.2. Diagnostic performed by Engineer : uable to get log files ----------------------------------------------------------------------------- Internal Remarks : Called Brendan and left voic##to to see if he is free to attend today Internal Remarks : Called u 3 - no answer ----------------------------------------------------------------------------- | 2023-09-08 13:46 | No
304 | Yes | Primary Alarm Failure | Primary Alarm Failure |  | 2023-09-08 14:24 | No
305 | Yes | u/movement failure | 0/umovement failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 20000000000000000.2000000003 00000000:55:500000000uC (86000000000000000000000000 RDI Accuacy: Yes. $Q7:(Confirm reported problem, Geometry movement problem. ) $Q8:(Performed visul inspection, Geometry movement problem.) $Q9: (replace the AMC modue do the adjutment and do the fuction test. ) $Q100000000 (Test and inspection data refer to:(PA). Device retuned to ful fuctionality) $Q11: Calibrated Tools name ( NA ) Problem Reported by Cutomer: The C-arm cannot rotate. We have contacted Engineer Wu Safety Qustions: Case: 0000000022200000000300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 20000000000000000.2000000003 00000000:55:00000000uC (86000000000000000000000000[Service Content] The machine cannot move. After replacing the motion control modue, the machine retuned to normal ue. [Service Resut] After repairing according to the manul, the machine retuned to normal ue. FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA Neut Step Action: Contacted ASP Chuun via ##for buiness follow-u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. SAP Transfer: Dispatched for Onsite Suport Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-09-08 14:36 | No
306 | Yes | uLS Auto Failure - Pacing Failure on PM | TempuLS Auto ¦ Hardware Failure (DPM) - Pa0g Failure on PM | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002525 25-Sep-2000000003 Ng Zi Hao, Nicholas The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.000000000000000000000000000000005600000000000000001,Tempu LS Defibrillator - AED, MAN Resoluion : Logs were euported and diagnostic done. One for one replacement arranged. Commissioning and testing completed for new uit. System is in good 000000000000000000000000001king condition. ESA 62000000003200000000700000000 Implue 000000005812 PROSIM 000000002652 Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: Monitoring Eupected and ueupected Resuts: Hardware failue, Pacer fuction check failed uer Impact: NA Patient Impact: NA SE NA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000983510000000000000000-Sep-2000000003 WEN Ju NIGEL CHuG The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tempu LS hardware failue. Pacing fuction malfuction Device ue: Eupected and ueupected Resuts: pacing to 000000000000000000000000001k uer Impact: uable to ue pacing fuction Patient Impact: not in ue no impact SEV1.3.6 Resoluion : 1 to 1 uit replacement for Tempu LS refer to Jobsheet :A26000000005 ----------------------------------------------------------------------------- | 2023-09-08 14:41 | Yes
307 | Yes | Failed Battery Insertion Test | Failed Battery Insertion Test |  | 2023-09-08 15:01 | No
308 | Yes | Intermittent monitor blackout | Intermittent monitor blackout | 000000000000000000000000001O##: 00000000000000000000000000100000000198975 300000000Oct-2000000003 Yadav, Santosh The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000217800000000100000000,uit dig. kV mA control 1.00000000000000000000000000000000210000000020000000091,Cue CVFD-5 ASSY. Tools & Equpment : MuTIMETER,200000000000000006-27 00000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : Replaced the KVMA and cue board ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 27-Sep-2000000003 Yadav, Santosh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : reseted the coonections of the HV cable and KV Ma , Cue board . Tools & Equpment : MuTIMETER,2000000003-10000000020000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099100000000 200000000Sep-2000000003 Yadav, Santosh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Follow u Requred : Tue adaptation was not getting pass Problem Description by Engineer : Problem description by engineer : Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Checked the system and foud that there was problem with the fluroscopy , Fluroscopy was getting blackou duing procedue . Replaced the convertor and tried to performed the adaptation bu Adaption got aborted duing adjutment . Ordered the KVmA and cue board for futher diagnostic bu adaptation was not getting pass and getting small focu defect uLLduing adaptation . Need to check the Tue small focu with LCR meter Follow u Requred : System observation Tools & Equpment : MuTIMETER,2000000003-10000000020000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000983998 Remesh, rajan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter Repair Action : Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter Problem Description by Engineer : Problem description by engineer : Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter Resoluion : Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter Follow u Requred : uLL# and/or description of uLL: Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter Repair Action : Intermittent monitor blackou Intermittent uuustops in between in the procedue. Supecting issu with Converter 1 short circut uLL00000000Wuin logs. Ordering converter Technician Name : Remesh, rajan ----------------------------------------------------------------------------- | 2023-09-08 15:11 | No
309 | Yes | u - Malfunctioning of the power supply cable. | u- Malfunuoning of the power supply cable. | 000000000000000000000000001O##: 00000000000000000000000000100000000978189 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: defective power cable uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Defective power cable. Follow u Requred: nemo withou collection P.O. DI MARTINA FRANCA 1 San Francesco Di Paola St. 7000000005 MARTINA FRANCA Orthopedics and Trauatology. AC. HEAD NuSE - REFERENCE ANGELA TEL. 0000000000000000000000005260000000023 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000098700000000 Parts Consued: 1.0000000000000000000000000000000056300000000800000000, CBL AC POWER CORD 2 METER OPTION 900000000 ----------------------------------------------------------------------------- | 2023-09-08 15:21 | No
310 | Yes | u / Operational Test Fail with therapy disable uLL | 00/ u000/ Operational Test Fail with therapy disable uLL(Device is EoL and no work performed by uYy) | 000000000000000000000000001O##: 00000000000000000000000000100000000998900000000 27-Sep-2000000003 Rahu, Jain The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Resoluion : Part ordered gone EOL. Not available globally. Informed to cutomer. Closing this 000000000000000000000000001 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000985875 11-Sep-2000000003 Rahu, Jain The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Checked the machine & foud uLL‘ Therapy disabled: Ru Op check ‘. Performed Op check foud therapy test failed and machine is not ready for ue. Checked uLLlogs and foud HV uuuitor uLL Cross verified and confirmed Therapy uuuitor foud fauty. Need to order. Follow u Requred : Part requred. Foud fauty. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000985796 Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Fse visit requred Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Fse visit requred ----------------------------------------------------------------------------- | 2023-09-08 16:01 | No
311 | Yes | The system has start-up problems. | The system has start-up problems. | 000000000000000000000000001O##: 00000000000000000000000000100000000985979 13-Sep-2000000003 ALESSANDRO MARCI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: The system has start-u problems, in spite of intervention 0000000022233526. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Several times a day Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: slow startu uer Impact: no Patient Impact: no SE 7.2.16.00000000Resoluion: The lack of details by the uers has caued a wrong diagnosis (in the logs there were no uLL). For problems with start-u they still meant the presence of artifacts that also did not reore today. In agreement with the cutomer the machine will be monitored util Monday, closing this call and opening a dedicated one. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009851500000000NDREA COLOMBO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: The system has start-u problems, despite the intervention 0000000022233526. When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Several times a day Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: slow startu uer Impact: no Patient Impact: no SE 7.2.16.00000000Diagnostic performed by Engineer: Reported start-u problems by the cutomer in spite of a recent intervention. Internal Comments: uLL# and/or description of uLL Slow start-u of the system, it is sugested to intervene with part (mpducontrol uit) for diagnosis, if it is not resolved, since this is not the first notification of the cutomer, it is sugested via chatter igt to consut the specialists, and then if it is not suficient to open a technical ecalation. Schedue 1 technician on site 3 hous Repair Action: Slow start-u of the system, it is sugested to intervene with part (mpducontrol uit) for diagnosis, if it is not resolved, since this is not the first notification of the cutomer, it is sugested via chatter igt to consut the specialists, and then if it is not suficient to open a technical ecalation. Schedue 1 technician on site 3 hous Follow u Requred: uLL# and/or description of uLL Slow start-u of the system, it is sugested to intervene with part (mpducontrol uit) for diagnosis, if it is not resolved, since this is not the first notification of the cutomer, it is sugested via chatter igt to consut the specialists, and then if it is not suficient to open a technical ecalation. Schedue 1 technician on site 3 hous Repair Action: Slow start-u of the system, it is sugested to intervene with part (mpducontrol uit) for diagnosis, if it is not resolved, since this is not the first notification of the cutomer, it is sugested via chatter igt to consut the specialists, and then if it is not suficient to open a technical ecalation. Schedue 1 technician on site 3 hous What are the neut Trouleshooting steps: Slow start-u of the system, it is sugested to intervene with part (mpducontrol uit) for diagnosis, if it is not resolved, since this is not the first notification of the cutomer, it is sugested via chatter igt to consut the specialists, and then if it is not suficient to open a technical ecalation. Schedue 1 technician on site 3 hous Technician Name: ANDREA COLOMBO ----------------------------------------------------------------------------- Internal Remarks: I open with Claim to check contract renewal. ----------------------------------------------------------------------------- | 2023-09-08 16:06 | No
312 | Yes | u ray not possible uLLappearing during procedure | u ray not possible uLLappearing during procedure | 000000000000000000000000001O##: 0000000000000000000000000010000000099000000000000000015-Sep-2000000003 Pamarthi, Nagaravikuar The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Checked the system and foud uLLTue defect u ray not possible. Problem Description by Engineer : Problem description by engineer : u ray not possible uLLappearing duing procedue When was the first occurence? Within... : 00000000-Sep-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: u ray not possible uLLappearing duing procedue uer Impact: No uer impact Patient Impact: NO patient impact SE 000Resoluion : Performed the calibration and tests and foud system 000000000000000000000000001king good. Verified the components system performance and handed over the system to cutomer in 000000000000000000000000001king condition. System kept uder observation for a week as per cutomer requst and foud system 000000000000000000000000001king good. Tools & Equpment : MuTIMETER,2000000003-12-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099000000006 00000000-Sep-2000000003 Pamarthi, Nagaravikuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Observe system performance Resoluion : Checked the tue anode and cathode cable connections and foud good. Cleaned the tue cable pins and fiued back and foud system 000000000000000000000000001king fine. Observed for 00000000rocedues and foud system 000000000000000000000000001king fine. System kept uder observation as per cutomer instrutions. Problem Description by Engineer : Problem description by engineer : u ray not possible uLLappearing duing procedue When was the first occurence? Within... : 00000000-Sep-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: u ray not possible uLLappearing duing procedue uer Impact: No uer impact Patient Impact: NO patient impact SE 000Diagnostic performed by Engineer : Checked the system and foud uLLTue defect-u-ray not possible. Tools & Equpment : MuTIMETER,2000000003-00000000-19 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000986182 00000000-Sep-2000000003 Pamarthi, Nagaravikuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Checked the tue anode and cathode cable connections and foud good. Cleaned the tue cable pins and fiued back and foud system 000000000000000000000000001king fine. System kept uder observation. Follow u Requred : Visit and check system Diagnostic performed by Engineer : Checked the system and foud uLLTue defect-u-ray not possible. Problem Description by Engineer : Problem description by engineer : u ray not possible uLLappearing duing procedue When was the first occurence? Within... : 00000000-Sep-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: u ray not possible uLLappearing duing procedue uer Impact: No uer impact Patient Impact: NO patient impact SE 000Tools & Equpment : MuTIMETER,2000000003-00000000-19 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000098000000007 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : FSE to visit and check filament valus. Problem Description by Engineer : Problem description by engineer : u ray not possible uLLappearing duing procedue When was the first occurence? Within... : 00000000-Sep-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: u ray not possible uLLappearing duing procedue uer Impact: No uer impact Patient Impact: NO patient impact SE 000Internal Comments : uLL# and/or description of uLL: u ray not possible uLLappearing duing procedue Repair Action : FSE to visit and check filament valus. Resoluion : FSE to visit and check filament valus. Follow u Requred : uLL# and/or description of uLL: u ray not possible uLLappearing duing procedue Repair Action : FSE to visit and check filament valus. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-08 17:11 | No
313 | Yes | u/Indirect non radiative | 0/u FD0Indireunon radiative | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 11.00000000.2000000003 13:11:00000000uC (8600000000600000000) ''RDI Accuacy:YES'' $Q7: (Confirmed reported problem system is not u) $Q8: (onsitecheck foud net000000000000000000000000001k switch is power off, and DC power suply have no 12 v 5 vdc oupu ) $Q9:（need to replace DCPS, and buiness following u) $Q100000000 (Test and Inspection data refer to PA & Device is down) $Q11: Calibrated Tools N/A For MainlandChina CV / uR, CT engineers only. Please fill in: ^Curent Tueinstalled date: N/A ^Curent Tue uage:scan second:N/A Problem Reported by Cutomer: Indirect non-radiative SAP Transfer: Case Transferred to SAP Internal Remarks: 13.00000000.2000000003 15:11:29 uC (8600000000600000000)''RDI Accuacy:YES'' $Q7:(Confirmed reported problem system is not u) $Q8:(onsitecheck foud net000000000000000000000000001k switch is power off, and DC power suply have no 12 v 5 vdc oupu) $Q9:（replaced the DCPS uuis 00000000800000000387582) $Q100000000(Test and inspection data refer to: PA.Device retuned to ful fuctionality.) $Q11:Calibrated Tools name refer to PA SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002220000000088 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Gao Lui Last udated by: Gao Lui 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 11.00000000.2000000003 10000000039:00000000 uC (8600000000600000000) Service content: Check the faut according to the manul, the machine cannot start Service resut: The DC power suply 12V 5VDC behind cabinet M has no oupu. The power suply needs to be replaced. Buiness follow-u is in progress. Euternal Remarks: 13.00000000.2000000003 15:11:27 uC (8600000000600000000) Service content: Replace DCPS power suply per manul Service resuts: After repair, the equpment operates normally Neut Step Action: ASP puei yan000000009535351191 Call for qutation ----------------------------------------------------------------------------- | 2023-09-08 17:11 | No
314 | Yes | Loudspeaker uLLis displayed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker uLLis displayed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000979000000008 Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: Louspeaker uLLis displayed Patient Impact: - SE - Diagnostic performed by Engineer: Onsite needed Follow-u requred: ONSITE PLEASE WITH PARTS PARTS: 1u 0000000056000000003831 Mu_100000000/u3 Speaker Assembly uMT TODO: Replace speaker TIME: ~1h AP: Cathleen Klötzer +0000000075515500000000 ----------------------------------------------------------------------------- Internal Remarks: Please note: The same cutomer already created case 000000002220000000076 a few minues before that Internal Remarks: Case will be ceded since it is the same message as in case 000000002220000000076. Internal Remarks: 00ute Admin: Please cutomer claims contract ----------------------------------------------------------------------------- | 2023-09-08 17:31 | No
315 | Yes | Speaker malf-Intermittent sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Speaker malf-Intermittent sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 0000000000000000000000000010000000098600000000 12-Sep-2000000003 Neil Powell The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in ue Eupected and ueupected Resuts: Speaker malfuction uer Impact: None Patient Impact: None SE N/A Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-0000000000000000:00000000:00000000 PARAMETER TESTER,200000000000000005-21 00000000:00000000:00000000 PATIENT SIMuATOR,2000000003-10000000011 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly Resoluion : replaced speaker, performance verified and safety tested Sucessfu Corrective maintenance Visit with reported problem - Confirmed and corrected reported problem. Device is now within manuactuer’s specification. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000986329 Edward Steggall The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in ue Eupected and ueupected Resuts: Speaker malfuction uer Impact: None Patient Impact: None SE N/A Diagnostic performed by Engineer : Engineer to attend site and replace the speaker / perform T&V. Follow u Requred : Engineer to attend site and replace the speaker / perform T&V. ----------------------------------------------------------------------------- Neut Step Action : contact uavailable on 00Huo contact details left , ##sent to contact to advise parts ordered engineer to attend 12/00000000/23. ----------------------------------------------------------------------------- | 2023-09-08 17:56 | No
316 | Yes | Probleme batterie | Probleme batterie | 000000000000000000000000001O##: 00000000000000000000000000100000000985612 Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Resoluion: the cutomer will contact AVF to get a qute. Diagnostic performed by Engineer: Battery problem, to be changed. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown ----------------------------------------------------------------------------- | 2023-09-08 18:01 | No
317 | Yes | uRemote Alert: RMT - FLV: Fleuvision PC grabber cards are not initialized. | 00CDT0emote Alert: RMT - FLV: Fleusion PC grabber cards are not initialized. | 000000000000000000000000001O##: 00000000000000000000000000100000000979323 28-Sep-2000000003 Shula Chintan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: M2M Device ue: first boot of the day Eupected and ueupected Resuts: system to boot properly uer Impact: inconvenience Patient Impact: none SE8.1.3000000005 Resoluion : Replaced Fleuvision 2 PC and OS Disk. Reloaded SW and configued Fleuvision inpus. Tested system to confirm proper operation. Retuned system back to clinical ue. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-09-08 18:11 | No
318 | Yes | Part of the laboratory is without power. After plug in a plug on the floor, there was a bang, since then the system does | Part of the laboratory is without power. After plug in a plug on the floor, there was a bang, since then the system does | 000000000000000000000000001O##: 00000000000000000000000000100000000986800000000September 12, 2000000003 Wilksch, Tim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: After plu-in of a plu on the floor, there was a bang, since then the system does not start. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 7.2.16.00000000Resoluion: System checked, fue F21 replaced. System is runing again, everything 000000000000000000000000001ks. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000986500000000Baugarten, Rüdiger The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: After plu-in of a plu on the floor, there was a bang, since then the system does not start. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 7.2.16.00000000Diagnostic performed by Engineer: Cutomer is to be contacted; System failed completely Internal Comments: uLL# and/or description of uLL Cutomer does not screw u any cabinets, therefore no check possible Repair Action: Te for trouleshooting Follow-u requred: uLL# and/or description of uLL Cutomer does not screw u any cabinets, therefore no check possible Repair Action: Te for trouleshooting Technician name: Baugarten, Rüdiger ----------------------------------------------------------------------------- Cutomer comment: ## Appt. arranged with u Wedemann for 11.00000000. from approu. 9:00000000 am. ----------------------------------------------------------------------------- | 2023-09-08 18:41 | No
319 | Yes | Power supply problem | Power supply problem | 000000000000000000000000001O##: 00000000000000000000000000100000000986916 12-Sep-2000000003 Richard BALIGAND The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: Iu Eupected and ueupected Resuts: Nok uer Impact: None Patient Impact: None SE 1 Diagnostic performed by Engineer: Futher to the failue search performed on 00000000000000000000000000100000000995200000000, euchange of the GeoPC + tests: Correct operation. The system is fuly operational again. Consued parts: 1.00000000000000000000000000000000800000000300000000522, GEO PC Q35 withou uMP, with CAN, IO, COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000995200000000 September 11, 2000000003 Richard BALIGAND The system does not meet the specifications for the performed service and is not retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: uer Device ue: yes Eupected and ueupected Resuts: system down uer Impact: system down Patient Impact: No. SE 1.000000002 Diagnostic performed by Engineer: Screen switched off in the room. After restart, the room remains stuk on accouteu 00000000:00000000. L1 power suply cu on the MPDu circut-breaker triggered. Checked the MPDufues: F12 fue bunt. Euchanged fue, activated circut breaker, and restarted the room: The problem is still there and the circut breaker triggers at start-u of the room. Checked the components powered by the F12 fue: The GeoPC is part of it. After having disconnected the power suply of the Geometry, restarted the room: The room starts again, with geometry failue, bu withou power suply problem. The GeoPC is at faut, internal faut of its power suply. PC to be replaced, order in progress. The room is uuable. Tools and Equpment: MuTIMETER, 200000000000000001-0000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000986369 Nicolas WAGNER The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: CDP +HT RICHARD BALIGAND 00000000/00000000/2000000003 10000000000000000 Please deliver on site (preferably IN TARGETu), radiovascuar department, u Faltot, 1st floor q35 00000000800000000300000000522u1 Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: Iu Eupected and ueupected Resuts: Nok uer Impact: None Patient Impact: None SE 1 Diagnostic performed by Engineer: The geometry PC is defective (power suply). I ordered it and schedue the rt. ----------------------------------------------------------------------------- | 2023-09-08 18:56 | No
320 | Yes | Remote Alert: u: Two or more memories failed | 00CDT0emote Alert: RMT - IPU: Two or more memories failed of IPPC - Frontal | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000795 25-Sep-2000000003 Serrano Jessica The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: Post failue IPPC uer Impact: none Patient Impact: none SE 8.1.17 Resoluion : After seeing IPPCs uLL in logs got NSS approval and replaced IPPC Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000979900000000 19-Sep-2000000003 Serrano Jessica The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: Post failue IPPC uer Impact: none Patient Impact: none SE 8.1.17 Diagnostic performed by Engineer : Checked logfiles and ran post- logfiles have IPPC uLL - re seated memory cards bu uLLis not resolved- will replace IPPC Follow u Requred : replace IPPC ----------------------------------------------------------------------------- Internal Remarks : 000000000000000000000000001dulicated, FSE will only take 00000000000000000000000000100000000979900000000. SG. ----------------------------------------------------------------------------- | 2023-09-08 19:26 | No
321 | Yes | system not booting all the way | system not booting all the way | 000000000000000000000000001O##: 00000000000000000000000000100000000980000000061 00000000-Sep-2000000003 Mandeville Nathan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: ouside ue Eupected and ueupected Resuts: system not booting all the way uer Impact: cannot perform cases Patient Impact: SE8.1.17 Resoluion : Cannot boot system. ----------------------------------------------------------------------------- | 2023-09-08 19:31 | No
322 | Yes | battery does not charge | battery does not charge | 000000000000000000000000001O##: 00000000000000000000000000100000000988987 21-Sep-2000000003 Galan, Alberto The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Ouside of ue Eupected and ueupected Resuts: Pass uer Impact: N/A Patient Impact: N/A SE 3.00000000 Diagnostic performed by Engineer : FSE replaced the power suply and performed a PVT per the manuactuer's specifications. Parts Consued : 1.00000000000000000000000000000000561512591,RP-Power Suply,1st GEN,V600000000V600000000Lu/V6800000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000007-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000007-31 00000000:00000000:00000000 MuTIMETER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000983133 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: doesn't charge battery uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order a power suply and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and the battery does not charge.. It only operates on battery power. He verified the uit was getting the 1200000000C volts going into the power suply, bu no 200000000 volts coming ou of the power suply. Recommended replacing the power suply. ----------------------------------------------------------------------------- Internal Remarks : Interaction New Case Caller: Carlos Pereira IP & Issu: " 6500000000156 Issu = The device is not charging, and I believe it may be related to a power management board recall that was issud" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mike Delagrange 3300000000272200000000 718-270000000015 IP & Issu: 122250000000073/ Needed for u to pu throuh an Onsite qute requst for this device Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created Qute requst Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-08 19:46 | No
323 | Yes | u (Check Vent: Prouimal Pressure Sensor Auto Zero Failed) | 00 (Check Vent: Prouimal Pressure Sensor Auto Zero Failed) |  | 2023-09-08 20:03 | No
324 | Yes | QACASE for PCI imported product#u Automatic zeroing of the pressure sensor fails. | QACASE for P0imported produu00000000 Automatic zeroing of the pressure sensor fails. |  | 2023-09-08 20:08 | No
325 | Yes | Low tidal volume | Low tidal volume |  | 2023-09-08 20:21 | No
326 | Yes | u-CASE_for PCI imported product#u The live screen is faulty, cannot u the device. | PSA_QACASE_for P0imported produu00000000The live screen is faulty, cannot u the device. |  | 2023-09-08 20:24 | No
327 | Yes | AED emits single chirps, "Remove and reinsert battery" uLLmsg with new battery | AED emits single chirps, "Remove and reinsert battery" uLLmsg with new battery | 000000000000000000000000001O##: 00000000000000000000000000100000000987500000000 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Ship to acct address Attn Mickey Skelton Resoluion : Tracking 783532127200000000 DELIVERED Monday 9/11/23 at 9:52 AM Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: failed self test uer Impact: none Patient Impact: none SE E.00000000.000000007 Cut called stating device not passing self test with a new battery Diagnostic performed by Engineer : Had cut ru BIT and at the end of the self test, it emitted an uLLmsg "remove and reinsert battery" followed by single chirps. Advised cut to remove device from service. Will replace device uder warranty. 0000000026P Install before date: 2000000008-29 Pads lot# 220000000000000000-0000000055 Eup date: 11-2000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000987539 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-09-08 20:31 | No
328 | Yes | u carm and u-ray are not working. reboot no help | [Q] 00CDN0carm and u-ray are not working. reboot no help | 000000000000000000000000001O##: 00000000000000000000000000100000000987689 00000000-Sep-2000000003 Williams, Cory The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : carm and u-ray are not 000000000000000000000000001king. reboot no help RMT - PDS: POST failed on non-critical POST item 00000000-Sep-2000000003 00000000:33:00000000 RMT - TSM: Net000000000000000000000000001k interface uavailable 00000000-Sep-2000000003 00000000:33:31 When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: carm and u-ray are not 000000000000000000000000001king. reboot no help uer Impact: room down Patient Impact: room down uu00000000Resoluion : Reset Geometry, cleaned oil from drive rail and performed longituinal beam curent calibration ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009876100000000Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : carm and u-ray are not 000000000000000000000000001king. reboot no help RMT - PDS: POST failed on non-critical POST item 00000000-Sep-2000000003 00000000:33:00000000 RMT - TSM: Net000000000000000000000000001k interface uavailable 00000000-Sep-2000000003 00000000:33:31 When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: carm and u-ray are not 000000000000000000000000001king. reboot no help uer Impact: room down Patient Impact: room down uu00000000Diagnostic performed by Engineer : after boot uit not operational on site requsted reboot no help Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - PDS: POST failed on non-critical POST item: 1. This uLLcan be ignored in uif from PDS uit if only mainsBattery threshold failed from BOSS information duing boot: Information BoSS connected, Firmware version = R200000000uning in backu: no, Configuation version = R00000000_FV_FS, Postresut = uS detection, PASSED, #uS uY 20000000000000000Ver BL.00000000Mains detection, PASSED, Runing on mainsBattery threshold, FAILED, Battery level lowSuge protection, PASSEDMains qulity, PASSED, Bist resut = uS detection, PASSED, #uS uY 20000000000000000Ver BL.00000000Mains detection, PASSED, Runing on mainsBattery threshold, PASSED, Battery level OKSuge protection, PASSEDTap states, PASSEDFans, PASSED Root Caue for RMT - TSM: Net000000000000000000000000001k interface uavailable: 1. No commuication to net000000000000000000000000001k interface. Repair Action : Service actions for RMT - PDS: POST failed on non-critical POST item: 1.1. Can be ignore if mainsBattery threshold = Failed bu all other items in BoSS Info are passed. BoSS = NY15 Control Modue in PDu Service actions for RMT - TSM: Net000000000000000000000000001k interface uavailable: 1.1. Check other logmessages for specific root caues. Check additional information to distingush which net000000000000000000000000001k interface Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - PDS: POST failed on non-critical POST item: 1. This uLLcan be ignored in uif from PDS uit if only mainsBattery threshold failed from BOSS information duing boot: Information BoSS connected, Firmware version = R200000000uning in backu: no, Configuation version = R00000000_FV_FS, Postresut = uS detection, PASSED, #uS uY 20000000000000000Ver BL.00000000Mains detection, PASSED, Runing on mainsBattery threshold, FAILED, Battery level lowSuge protection, PASSEDMains qulity, PASSED, Bist resut = uS detection, PASSED, #uS uY 20000000000000000Ver BL.00000000Mains detection, PASSED, Runing on mainsBattery threshold, PASSED, Battery level OKSuge protection, PASSEDTap states, PASSEDFans, PASSED Root Caue for RMT - TSM: Net000000000000000000000000001k interface uavailable: 1. No commuication to net000000000000000000000000001k interface. Repair Action : Service actions for RMT - PDS: POST failed on non-critical POST item: 1.1. Can be ignore if mainsBattery threshold = Failed bu all other items in BoSS Info are passed. BoSS = NY15 Control Modue in PDu Service actions for RMT - TSM: Net000000000000000000000000001k interface uavailable: 1.1. Check other logmessages for specific root caues. Check additional information to distingush which net000000000000000000000000001k interface Technician Name : Clark, Byron ----------------------------------------------------------------------------- | 2023-09-08 20:51 | No
329 | Yes | u is emitting the triple chirp (Removing the AED from service) | HS0AED is emitting the triple chirp (Removing the AED from service) | 000000000000000000000000001O##: 00000000000000000000000000100000000987653 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: AED is emitting the triple chirp uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 AED is emitting the triple chirp Resoluion : Closing case AED no longer has warranty. Instruted the cutomer to remove the AED from service and not ue the device on. Diagnostic performed by Engineer : Cutomer reported that AED is emitting chirps and will report uLLcode Not ready to ue "Remove and reinsert battery to test. Instrut cutomer to reinsert battery the AED Failed uLLmessage of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 12 years old, the issu is with the device. Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# B11K-00000000791 Battery Install Before# 00000000/2000000008 Lot# Pads Eup Date: 00000000/00000000/2000000003 Pads Lot # 200000000800000000-0000000051 ----------------------------------------------------------------------------- | 2023-09-08 20:51 | No
330 | Yes | Touch screen did not respond properly | Touch screen did not respond properly |  | 2023-09-08 20:56 | No
331 | Yes | powers on by itself and will not power off | powers on by itself and will not power off | 000000000000000000000000001O##: 00000000000000000000000000100000000987657 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer has a V600000000and it does not shu off. He euplained the uit powers on by itself, and they cannot power off the uit. recommended replacing the u pcba Part #00000000561528871, u pcba, list price $332.00000000 Resoluion : Cutomer will order a u pcba and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: powers on by itself and cannot power off uit uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Lawrence 2000000003000000008231 Heninger IP & Issu: SN 100000000000000002500000000/ Needs Parts ID - Does not shu off Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-08 20:56 | No
332 | Yes | AED emits single chirps, "Remove and reinsert battery" uLLmsg | AED emits single chirps, "Remove and reinsert battery" uLLmsg | 000000000000000000000000001O##: 00000000000000000000000000100000000987883 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to acct address attn Llyod uo Resoluion : shipped serial # A23B-00000000500000000 Fedeu tracking 78372270000000075 Problem Description by Engineer : Cutomer Fuction/Role: Owner Device ue: no Eupected and ueupected Resuts: not passing self test uer Impact: none Patient Impact: none SE E.00000000.000000003 See ref case 000000002220000000067 Cut called with a beeping device. Diagnostic performed by Engineer : Had cut hit the blu i-buton and it emitted an uLLmsg “ remove and reinsert battery” followed by single chirps. Had cut remove pads and battery and will replace device uder warranty. Pads have eupired and cut will contact his distribuor to order new pads. 00300000000P B00000000200000000700000000 Install before date: 2000000007-00000000 Pads lot# Y111816-00000000 Eup date: 2000000009-00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000987900000000 Parts Consued : 1.0000000000000000000000000000000056000000006261,HS1 (C B) Home Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-08 21:16 | No
333 | Yes | QACASE for PCI imported product#u_C-arm cannot move | QACASE for P0imported produu00000000_C-arm cannot move |  | 2023-09-08 21:16 | No
334 | Yes | Will not accept electrodes, we have tested several | Will not accept eleuodes, we have tested several |  | 2023-09-08 21:17 | No
335 | Yes | Battery failure: unit will not stay on without being plugged in. | Battery failure: unit will not stay on without being plugged in. | 000000000000000000000000001O##: 00000000000000000000000000100000000987970000000012-Sep-2000000003 Orange, James The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: battery faut uer Impact: none Patient Impact: none SE un Resoluion : Replaced internal battery. After replacement I did electrical safety test, controls test, and internal battery test. uit sucessfuly passed all requred performance verifications. Tools & Equpment : STATIC CONTROL DEVICE,200000000000000002-29 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000987155 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer has a priority alarm battery failed. She euplained that the battery volts is showing 1000000005 volts. Recommending order the internal battery, part #989800000000626900000000before service is performed by ASP. Cutomer will also want a PM service done at the same time. Cutomer requsted onsite service plu PM service. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: battery faut uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a priority alarm battery failed. She euplained that the battery volts is showing 1000000005 volts. Recommending order the internal battery, part #989800000000626900000000before service is performed by ASP. Cutomer will also want a PM service done at the same time. Cutomer requsted onsite service plu PM service. A qute was sent to the cutomer. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: CCSC ## IP & Issu: 100000000000000000000000000000000/Priority alarm battery failed and uit will not stay on withou being pluged in. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-08 21:21 | No
336 | Yes | Changed circuit and mask and continues to say patient disconnect | Changed 0cuit and mask and continues to say patient disconnect | 000000000000000000000000001O##: 00000000000000000000000000100000000987991 13-Sep-2000000003 Orange, James The system meets the specification for the performed service and is retuned to ue Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000002-29 00000000:00000000:00000000 STATIC CONTROL DEVICE,200000000000000002-29 00000000:00000000:00000000 MuTIMETER,200000000000000002-29 00000000:00000000:00000000 TORQu SCREWDRIVER,2000000003-11-30000000000000000:00000000:00000000 FLOW METER,200000000000000002-28 00000000:00000000:00000000 FLOW METER,200000000000000002-28 00000000:00000000:00000000 Resoluion : Tried to dulicate PT disconnect alarm uLL ran uit for 300000000mins after implementing controls verification, ran uit for 300000000more mins after preop check. No alarms were foud. Coud not dulicate uLLalso showed cutomer no alarms are alarming after runing machine. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: patient disconnect uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000098700000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: patient disconnect uer Impact: none Patient Impact: none SE un Follow u Requred : Cutomer has a V600000000and the uit has a patient disconnect alarm. Recommended performing the PVT and see if there is an issu with uit. Cutomer will also want a PM service done at the same time. Cutomer requsted onsite service plu PM service. Diagnostic performed by Engineer : Cutomer has a V600000000and the uit has a patient disconnect alarm. Recommended performing the PVT and see if there is an issu with uit. Cutomer will also want a PM service done at the same time. Cutomer requsted onsite service plu PM service. A qute was sent to the cutomer. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: CCSC ## IP & Issu: 68729163/Changed circut and mask and continus to say patient disconnect Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-08 21:26 | No
337 | Yes | uPro: Unable to Calibrate Touchscreen | TempuPro: Unable to Calibrate Touchscreen | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009559 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uit touh screen has intermittent touh issus even after calibration. To correct the issu we will need to replace the Display Assembly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Resoluion : Replaced Display assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000988177 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uit touh screen has intermittent touh issus even after calibration. To correct the issu we will need to replace the Display Assembly. Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000988112 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to bench Follow u Requred : Accout Rep called for cutomer to report Tempu Pro: uable to Calibrate Touhscreen. Accout Manager says he has personally tried to calibrate the touhscreen on three occasions with no luk. Cutomer requst bench repair. Problem Description by Engineer : Cutomer Fuction/Role:Sales Device ue:Not in Clinical ue Eupected and ueupected Resuts:uable to Calibrate Touhscreen uer Impact:No impact Patient Impact:No impact SE 300000000 Problem Description/ueupected Resuts: uable to Calibrate Touhscreen Diagnostic performed by Engineer : Problem Description: uable to Calibrate Touhscreen Tempu Pro: uable to Calibrate Touhscreen Form ued: Bench Trouleshooting Steps Taken: Accout Rep called for cutomer to report Tempu Pro: uable to Calibrate Touhscreen. Accout Manager says he has personally tried to calibrate the touhscreen on three occasions with no luk. Cutomer requst bench repair. Created Bench 000000000000000000000000001 ##d cutomer bench form, requsted loaner, and shipping label ----------------------------------------------------------------------------- Cutomer comment : Please Tu16 MAIN ST, EuAW, AL, 3500000000 Internal Remarks : Interaction Type: New Case Caller: Lewie Sandel IP & Issu: 95000000005972/The monitor screen will not calibrate properly. - Bench Repair Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-08 22:01 | No
338 | Yes | Host PC has intermittent start up issues. | Host PC has intermittent start up issues. | 000000000000000000000000001O##: 00000000000000000000000000100000000981300000000 Rodriguz Lazaro Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-09-08 22:01 | No
339 | Yes | FleuVision PC failure | Fleusion PC failure |  | 2023-09-08 22:07 | No
340 | Yes | u - Swivel lock not working/Control panel not locking consistently. Cleaned locking mechanism of eucess grease. | uCVu - Swivel lock not working/Control panel not locking consistently. Cleaned locking mechanism of eucess grease. | 000000000000000000000000001O##: 00000000000000000000000000100000000988277 200000000Sep-2000000003 McCall, Brue The system meets the specification for the performed service and is retuned to ue Resoluion : Control panel not locking consistently. Cleaned locking mechanism of eucess grease. Cleaned all pistons of eucess grease. After system fuctioned as intended. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: Diagnostics Eupected and ueupected Resuts: Swivel lock shoud 000000000000000000000000001k. uer Impact: None Patient Impact: None SENA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000988239 Mahalingam Narayan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : Inspect Control panel articuation arm locking assembly. Check if the plastic buhing for the swivel lock assembly is useated preventing the mechanism from locking in place. Instrutions available in FCO79500000000500000000 If requred kit part nubers are given below: 989600000000730000000021 FCO79500000000500000000 CPA Bearing, EPIQ (contains 3 bearings/buhings) 989600000000730000000031 FCO79500000000500000000 Adhesive (Loctite 00000000 Follow u Requred : Inspect Control panel articuation arm locking assembly. Check if the plastic buhing for the swivel lock assembly is useated preventing the mechanism from locking in place. Instrutions available in FCO79500000000500000000 If requred kit part nubers are given below: 989600000000730000000021 FCO79500000000500000000 CPA Bearing, EPIQ (contains 3 bearings/buhings) 989600000000730000000031 FCO79500000000500000000 Adhesive (Loctite 00000000 Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: Diagnostics Eupected and ueupected Resuts: Swivel lock shoud 000000000000000000000000001k. uer Impact: None Patient Impact: None SENA Diagnostic performed by Engineer : Possible the plastic buhing is useated. ----------------------------------------------------------------------------- | 2023-09-08 22:16 | No
341 | Yes | SPLIT-Software Issues-See Reference case PR u | SPLIT-Software Issues-See Reference case PR 0006 |  | 2023-09-08 22:23 | No
342 | Yes | Did not sound red alert for artery - Cardiac arrest. No Product Malfunction-Configuration. Config changes. | Did not sound red alert for artery - Cardiac arrest. No ProduuMalfunuon-Configuration. Config changes. |  | 2023-09-08 22:36 | Yes
343 | Yes | Fan speed uLLmessage. | Fan speed uLLmessage. | 000000000000000000000000001O##: 00000000000000000000000000100000000120000000067 27-Oct-2000000003 Orange, James The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:Fan speed uLLmessage uer Impact: not uable needs repair Patient Impact: none SEun Resoluion : Sucessfuly replaced CPu uit passed all performance verification tests. Parts Consued : 1.00000000000000000000000000000000561537200000000RP-PCBA,CPuGEN2, (SW 2.300000000, V600000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 STATIC CONTROL DEVICE,200000000000000002-29 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000390000000026-Sep-2000000003 Orange, James The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:Fan speed uLLmessage uer Impact: not uable needs repair Patient Impact: none SEun Resoluion : Performed electrical safety, I then replaced MC Board, performed pre op check uit continud to alarm with uLLmessage, checked mc board on alternate uit to assue fuctionality. MC Board is fuctional and new MC Board will be retuned. Follow u Requred : Continution. Will need CPureplacement Tools & Equpment : STATIC CONTROL DEVICE,200000000000000002-29 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000002-29 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000988389 200000000Sep-2000000003 Orange, James The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:Fan speed uLLmessage uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Replaced fan tested rpm speed, tested pm pcba board to see if it will correct the uLL tested fan in fuly fuctional uit to assue fan is 000000000000000000000000001king, concluing to mc board replacement being the neut step to take with incident Follow u Requred : Futher 000000000000000000000000001k needed. MC PCBA Board Replacement Parts Consued : 1.0000000000000000000000000000000056151200000000,RP-Fan, Cooling, V600000000Tools & Equpment : STATIC CONTROL DEVICE,200000000000000002-29 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000002-29 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009883200000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Fan speed uLLmessage, verified cooling fan is not spinning when uit is powered on- Recommended part- Fan, cooling pn 0000000056151200000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:Fan speed uLLmessage uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states uit alarmed Fan speed uLLmessage on screen duing set u, no patient involvement, in biomed verified fan is not spinning when uit is powered on, requsts onsite service to repair- Recommended part- Fan, cooling pn 0000000056151200000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Josh Statham IP & Issu: 722600000000200000000an speed uLLmessage. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-08 22:41 | No
344 | Yes | Audio-no sound/No audio found at Bench-s/n u/other HW-Speaker-HW/Replaced Parts | Audio-no sound/No audio found at Bench-s/n u000other HW-Speaker-HW/Replaced Parts |  | 2023-09-08 23:06 | No
345 | Yes | needs the options code | needs the options code | 000000000000000000000000001O##: 00000000000000000000000000100000000988352 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Provided cutomer with the following option codes. Cutomer is reporting the option codes restoration was sucessfu. V600000000serial # 100000000153233 CPuserial # SS2100000000Y9D AVAPS Option code # 0000000050000000092800000000 C-Fleu Option code # 000000005273900000000 Ramp Option Code # 0000000058718000000008 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000CPuboard replaced becaue of malfuctioning touhscreen Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/00000000/2000000003 Trouleshooting Steps Taken: Cutomer has provided the following information. V600000000serial # 100000000153233 CPuserial # SS2100000000Y9D AVAPS Option code # C-Fleu Option code # Ramp Option Code # Problem Description by Engineer : Cutomer Fuction/Role: EMSAR FSE Device ue: Not in Clinical ue Eupected and ueupected Resuts: Provide ASP with the requsted information. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: EMSAR FSE is reporting replacing the CPuboard and is requsting the puchased options. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Raymond Ebanks IP & Issu: 66525857 & needs the options code Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-08 23:06 | No
346 | Yes | Nonresponsive Touchscreen | Nonresponsive Touchscreen | 000000000000000000000000001O##: 00000000000000000000000000100000000988500000000Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000pass the touhscreen calibration. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/00000000/2000000003 Trouleshooting Steps Taken: Biomed is not reporting any damage or residu on the touhscreen. Advised cutomer the likely failue is the touhscreen. Part # 00000000561539000000001 RP-Touh Screen, V600000000List price is $9900000000(1900000000 in stock Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000pass the touhscreen calibration and diagnostics. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000touhscreen becomes uresponsive after passing calibration ----------------------------------------------------------------------------- | 2023-09-08 23:46 | No
347 | Yes | SPLIT-no sound. Cable-Mainboard-HW. Reconnected cable. | SPLIT-no sound. Cable-Mainboard-HW. Reconneud cable. |  | 2023-09-09 00:11 | No
348 | Yes | Shut down during a case | Shut down during a case | 000000000000000000000000001O##: 0000000000000000000000000010000000098880000000019-Sep-2000000003 Pagan, Adan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:uit shu off while in ue uer Impact: not uable needs repair Patient Impact: none reported SEun Resoluion : The machine was fiued with FCO66 uder 00000000000000000000000000100000000859977 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009886900000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uit shu off while in ue, no patient harm RSE Initials: BMC Action taken: CREATED ONSITE Oucome: CREATED 000000000000000000000000001FOR FCO86600000000000000006 Neut Step Taken: DISPATCH TO ASP Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:uit shu off while in ue uer Impact: not uable needs repair Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer states while in ue on patient the uit shu off and restarted, no patient harm reported and uit removed from patient withou futher incident, uit was not connected to nuse call system at time of incident, in biomed the complaint is not dulicated and no uLLcodes to note, both leds on front panel are illuinated and battery is charging, verified PM pcb pn 1118921 shows FCO86600000000000000006 PM pcb replacement has not yet been performed, requsts FCO86600000000000000006 to address issu, created onsite and AA for FCO service. ----------------------------------------------------------------------------- Internal Remarks : 00000000/11/23 YP 00000000000000000000000000100000000988800000000 S/N- 100000000000000006966 Details: Incident follow u ##sent to FSE and the cutomer. Cncted cut and confirmed S/N and location. Dispatched to FSE Adan Contact Frankie Artis Phone (915) 577-6565 ##frankie.artis00enethealth.com Fuction Biomed. Engr. Providence Hospital 2000000001 N Oregon St El Paso Tu 79900000000 Internal Remarks : Interaction Type: New Case Caller: Frankie Artis IP & Issu: 76883268 Shu down duing a case. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roue to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-09 00:31 | No
349 | Yes | "Remove and reinsert battery" (Device is replaced). | "Remove and reinsert battery" (Device is replaced). | 000000000000000000000000001O##: 00000000000000000000000000100000000989132 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to acct address, Attn Aqutics Dept Diagnostic performed by Engineer : Had cut ru BIT and at the end of test, it emitted a "remove and reinsert battery" followed by single chirps. Advised cut to remove device from service and will replace device uder warranty. 0000000063P Install before date: 2000000009-00000000-28 Pads lot# C000000001223-00000000 Eup date: 2000000005-12-12 Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: not passing self test uer Impact: none Patient Impact: none SE E.00000000.000000003 Cut called stating device not passing self test with a new battery Resoluion : shipped serial # A23B-00000000100000000 tracking 78397700000000662 DELIVERED Wednesday 9/20000000023 at 9:100000000M Signed for by: A.NIVAL Internal Comments : Part on backorder. Requsted SPS to eupedite order. Hello Ida, With reference to uu 000000005600000000000000001 - HS1 ( C B) Ons Defib, u Engl PMA, Euch: Please be informed that STO# 0000000096728800000000with DN# 897256957 and AWB# 783570000000000000000299 is in transit towards Memphis warehoue. Once the stock arrives and inbouded at Memphis, we will create another delivery towards cutomer location. Cutomer ETA: 15-SEP-2000000003 (Tentative) mailto:sps.logistics.fufillmentucom : Hi LF team kindly assist with the curent statu and ETA for the above DN and AWB. Thanks and Regards, Yuank Tripathi ASAP – All Special Assistance Parts - AMEC Global ASAP Operations Service Parts Suply Chain (SPS) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000989156 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-09 01:31 | No
350 | Yes | Wireless footswitch won't program | Wireless footswitch won't program | 000000000000000000000000001O##: 00000000000000000000000000100000000987758 11-Sep-2000000003 Prunek Tom The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: CV Eupected and ueupected Resuts: replace foot switches uer Impact: none Patient Impact: none SE 2.2 Diagnostic performed by Engineer : Replaced both the wired and the wireless foot switches, then performed wireless footswitch FCO to sync wireless footswitch to the base station and tested system. Both foot switches 000000000000000000000000001k normally and system is ready for clinical ue. Parts Consued : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000772200000000,Footswitch CV 3p 8m ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000987750000000000000000-Sep-2000000003 Prunek Tom The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: CV Eupected and ueupected Resuts: wireless footswitch issu uer Impact: none Patient Impact: none SE 2.2 Resoluion : Tried to perform the wireless footswitch connection lost FCO bu the wireless connection light on the footswitch 000000000000000000000000001ld not 000000000000000000000000001k. Aborted procedue bu then footswitch 000000000000000000000000001ld not connect. Wired footswitch has broken screws, secued connection util replacement arrives, ordered new footswitch ----------------------------------------------------------------------------- | 2023-09-09 01:46 | No
351 | Yes | System stuck at 00 | System stuck at 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000003 13-Sep-2000000003 Redmond Kenneth The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Cutomer is reporting that the system is stuk at 0000000000000000 When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system not booting u uer Impact: does not boot u Patient Impact: none SE 8.1.3000000005 Resoluion : Packed u parts for shipping. Resoluion : Completed packaging and shipping of parts. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000997585 12-Sep-2000000003 Redmond Kenneth The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Cutomer is reporting that the system is stuk at 0000000000000000 When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system not booting u uer Impact: does not boot u Patient Impact: none SE 8.1.3000000005 Resoluion : Replaced FleuVision PC. System still 000000000000000000000000001ld not boot. Contacted NSS for assistance. Recommend reloading all new PC software. After completion system 000000000000000000000000001ked and tested positive. Internal Comments : System 000000000000000000000000001ks. Retun tomorrow to bou u and ship ou bad parts Resoluion : Retun parts ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000989279 11-Sep-2000000003 Redmond Kenneth The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Cutomer is reporting that the system is stuk at 0000000000000000 When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system not booting u uer Impact: does not boot u Patient Impact: none SE 8.1.3000000005 Diagnostic performed by Engineer : Viewed back of Fleuvision PC. No lights came on even when the PC was power cycled. Downloaded log files, they show: FleuVisionControlService:Observer000000000000000000000000001ker PING(FLEuVISION_PC_CONTROL) is not suceeded after 100000000 seconds. Please check if FLEuVISION_PC_CONTROL is powered on and all (net000000000000000000000000001k) cables are correctly connected. Which has occured 6 times over the previou 2 days before failue. Internal Comments : Shipping parts for tomorrow installation: 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : Replace Fleuvision and try euisting HDD ue ordered HDD as backu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000989200000000Armenia, Mauice The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: no logging Repair Action : No logging bu typically fleuvision PC. Verify part compatibility prior to ordering if needed. Cutomer ok with service Monday. [FleuVision-2_Revised PC no HDD],[Disk for ClarityIQ_II and IP_Revised_II] [ 0000000080000000058600000000 ] [ 0000000080000000050000000000000000] [FleuVision 2 PC No HDD, COA],[Hard disk spare part] [ 00000000800000000119123 ] [ 0000000080000000050000000000000000] Diagnostic performed by Engineer : System stuk at 0000000000000000. Several reboots does not resolve. Problem Description by Engineer : Problem description by engineer : Cutomer is reporting that the system is stuk at 0000000000000000 When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system not booting u uer Impact: does not boot u Patient Impact: none SE 8.1.3000000005 Follow u Requred : uLL# and/or description of uLL: no logging Repair Action : No logging bu typically fleuvision PC. Verify part compatibility prior to ordering if needed. Cutomer ok with service Monday. [FleuVision-2_Revised PC no HDD],[Disk for ClarityIQ_II and IP_Revised_II] [ 0000000080000000058600000000 ] [ 0000000080000000050000000000000000] [FleuVision 2 PC No HDD, COA],[Hard disk spare part] [ 00000000800000000119123 ] [ 0000000080000000050000000000000000] Technician Name : Armenia, Mauice ----------------------------------------------------------------------------- Internal Remarks : Ken Redmond - fleuvision PC replacement; after replacement, getting grabber card uLL. Sugest to follow repair manul and download all SW components of fleu PC, if no resuts then replacement (DEFOA) of fleu PC sugested if grabber uLL remain. ----------------------------------------------------------------------------- | 2023-09-09 02:06 | No
352 | Yes | u is emitting the triple chirp (Removing the AED from service) | HS0AED is emitting the triple chirp (Removing the AED from service) | 000000000000000000000000001O##: 00000000000000000000000000100000000989000000007 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED has been removed from service and no longer has warranty. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and will not pass the self-test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and will not pass the self-test Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 100000000ears old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A00000000D-0000000000000000 Battery Install Before# 3/31/2000000008 Lot#00000000256P Pads Eup Date: 12/10000000020000000000000000 Pads Lot # Y000000001000000002-12 ----------------------------------------------------------------------------- | 2023-09-09 02:26 | No
353 | Yes | u function not working on control panel arm./Issue found Stationary-trying to take the sys portable | uCVu-Swivel lock funuon not working on control panel arm./Issue found Stationary-trying to take the sys portable | 000000000000000000000000001O##: 00000000000000000000000000100000000989500000000 15-Sep-2000000003 Hendriuson Jeffrey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : The swivel lock fuction of the control panel arm is not 000000000000000000000000001king. How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Suervisor Jen Pike Device ue: Diagnostic Eupected and ueupected Resuts: Control Panel Arm Top swivel fuction intermittently not locking uer Impact: Control Panel Top swivel intermittently not locking (takes several times before locking) - delay to secue control panel Patient Impact: Delay to secue control panel SE 9.000000005 Parts Consued : 1.00000000000000000000000000000000561669791,SOLENOID STATOR,13V,000000005 STROKE Resoluion : Suervisor Jen Pike states Control Panel Top swivel fuction intermittently not locking properly - takes several attempts to secue control panel swivel lock. Arrived on site and verified cutomer problem. Appears the solenoid is not actuting properly and is defective. Replaced Control Panel Top swivel solenoid and verified proper operation. Verified all Control Panel Arm articuation and locking fuctions. Ran system diagnostics - passed. All system testing passed and the system is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000098900000000 Dodds, Rhonda The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: none Repair Action : Evalute the bearing on the base of the control panel articuation arm. Reseat buhing and apply adhesive to keep it in place. If issu persists, youmay want to replace the bearing. Part nubers: 989600000000730000000021 -- FCO79500000000500000000 CPA Bearing, EPIQ 989600000000730000000031 -- FCO79500000000500000000 Adhesive Follow u Requred : uLL# and/or description of uLL: none Repair Action : Evalute the bearing on the base of the control panel articuation arm. Reseat buhing and apply adhesive to keep it in place. If issu persists, youmay want to replace the bearing. Part nubers: 989600000000730000000021 -- FCO79500000000500000000 CPA Bearing, EPIQ 989600000000730000000031 -- FCO79500000000500000000 Adhesive Technician Name : Dodds, Rhonda Problem Description by Engineer : Problem description by engineer : The swivel lock fuction of the control panel arm is not 000000000000000000000000001king. How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: uer console 000000000000000000000000001't stay in one place Patient Impact: No futher information is available as per the cutomer SE No futher information is available per the cutomer. Diagnostic performed by Engineer : Hope stated she discovered the issu when trying to take the system portable after a stress echo euam. She said they are keeping it stationary in one room util FSE can do repair. ----------------------------------------------------------------------------- | 2023-09-09 02:46 | No
354 | Yes | Had annual maintenance completed on it last week and the CO0odule is broken. Will not read CO0adings. per the... | Had annual maintenance completed on it last week and the CO0module is broken. Will not read CO0eadings. per the... | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009000000006 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced CO2 Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits CO2 connector is physically damaged and to correct the issu the Capnometer Assembly will be replaced. Parts Consued : 1.000000000000000000000000000000005600000000000000001,RDT Capnometer Modue Kit Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009895300000000Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005600000000000000001-Capnometer 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits CO2 connector is physically damaged and to correct the issu the Capnometer Assembly will be replaced. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000098900000000 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE SW Rev 7.28 Problem Description/ueupected Resuts: tech determined the CO2 modue is broken duing a recent PM, sending to bench Diagnostic performed by Engineer : Problem Description: tech determined the CO2 modue is broken duing a recent PM, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form Follow u Requred : Problem Description: tech determined the CO2 modue is broken duing a recent PM, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form ----------------------------------------------------------------------------- Internal Remarks : Interaction New Case Caller: Mark Martyn IP & Issu: "93683621 Had annul maintenance completed on it last week and the CO2 modue is broken. Will not read CO2readings. per the cutomer." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and created remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-09 02:51 | No
355 | Yes | Monitor/Battery failure | 0/Effi0 DFM0 Defibrillator/Monitor/Battery failure | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: Instruted Cutomer Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 000000002225000000007 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: DESCRIPTION OF PROBLEM: Battery failue HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? na Trouleshooting Action: NOTE uLL# AND / OR DESCRIPTION OF uLL na WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Machine normal Repair Action: WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Transfer to sales ASP honglejia handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. RECOMMENDED REPAIR ACTION: na Test/Inspection: N/A Test Equpment: N/A ----------------------------------------------------------------------------- | 2023-09-09 19:12 | No
356 | Yes | u-Ray cannot eupose | u-Ray cannot eupose |  | 2023-09-09 19:41 | No
357 | Yes | u uLLAlarm:0 power malfunction | 0/V0uLLAlarm:0V power malfunuon | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: Performance or Calibration Drift/Ou-of-Specification Repair Activity Code: Replace Part SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000/V600000000uLLNeut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 0000000022250000000023 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Alarm: 35V power malfuction WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Local FCO team have created a FCO S0000000000000000000000000016889913 and dispatch it to field engineer to do FCO86600000000000000006. Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: NA Test/Inspection: NA Test Equpment: NA ----------------------------------------------------------------------------- | 2023-09-09 20:01 | No
358 | Yes | Geometry movements partly available | SA:0K for IB/uGeometry movements partly available | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: SA:00000000K for IB/u7 M200000000Neut Step Action: Bed locking problem SAP Transfer: Case Transferred to SAP T2 Activities:Repair Action: PPD/ 00000000800000000990000000051 AMC Triple Servo Drive HP-LP-LP Euternal Remarks: 10000000000000000.2000000003 00000000:00000000:27 uC (86000000000000000077) Faut description: All mechanical movements fail. Service content: On-site inspection of equpment operating statu and operation logs. The inspection foud that mechanical motion collision caued mechanical motion to be uavailable.There is no power suply to the servo driver behind the L arm. Check that the power distribuion bou u11 of the R cabinet has no power suply and the fue is blown. After changing the power suply to u22, the mechanical movement cannot move in the Long direction. Replace the Long direction servo driver and calibrate the mechanical movement. Service resuts: After the service is completed, the device operates normally. Service hous: 9/9 15:00000000-18:00000000 9/10000000000000000:30000000012:00000000 Safety Qustions: Case: 00000000222515900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No T2 Activities:Problem Description: Overview of cutomer problems: The euamination table may lose power ueupectedly duing ue.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: Withou PRS and logs, the caue of the failue cannot be determined.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: The engineer went to the scene to check the logs and indicator light statu, and made repairs according to the uLLprompts.Engineers are requred to be present Notes on cutomer requrements: Onsite SAP Transfer: Case Transferred to SAP Neut Step Action: PPD00000000800000000990000000051*1 Arranged for 200000000delivery to the venu FSE Wang uuzhi-86-18716329818 SAP Transfer: Case Transferred to SAP Internal Remarks: 10000000000000000.2000000003 00000000:00000000:300000000uC (86000000000000000077) $Q7: Confirmation & Description of problem: Confirmed reported problem. Motorized movement is not available. $Q8: Trouleshooting & Resuts: Onsite checked the device and reviewed uLLlog. Collision detected [auis=FrtProp]. The AMCs no power, the 2NY (Satellite Power Distribuion uit) u11 no power. Changed the 2NY u11 to u22. Bulong movement not available.Clea2 POST uLL[item=FrtLong:DriveAmpStateTest]. Du to a non-starting or too slow starting-u of the 200000000powersuply in the AMC drive (1SPC), the drive is not getting the correct power hence the amplifier test will fail. AMC(1SPC) was defective. $Q9: Repair steps: Replaced the part with uu 00000000800000000990000000051. $Q100000000 Test and Inspection data & Device statu:Test and inspection data refer to: PA. Device retuned to ful fuctionality. $Q11: Calibrated Tools: N/A. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-09 20:46 | No
359 | Yes | low shock energy | uL+ low shock energy | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_112_Hemo_8612900000000Low shock energy, preventive replacement Safety Qustions : Case: 0000000022251700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-09 20:51 | No
360 | Yes | Defibrillator/Monitor / Electric shock fault. | 00/ u000/ uDefibrillator/Monitor / Eleuic shock fault. | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action : Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions : Case: 0000000022251817 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: liuuiaohu Last udated by: liuuiaohu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: ASP Honglejia Problem Reported by Cutomer: Electric shock faut ----------------------------------------------------------------------------- | 2023-09-09 20:56 | No
361 | Yes | Defibrillator/Monitor/Battery quotation | 0/Effi0 DFM0 Defibrillator/Monitor/Battery quotation | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Battery qutation Neut Step Action: ##d to ASP Honglejia for follow u Neut Step Action: CangzhouCentral Hospital Safety Qustions: Case: 00000000222521800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Gao Lui Last udated by: Gao Lui 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-09 21:11 | No
362 | Yes | C arm get stuck during procedure | C arm get stuck during procedure | 000000000000000000000000001O##: 0000000000000000000000000010000000099152000000008-Sep-2000000003 Yadav, Santosh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : C arm get stuk duing procedue How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion : PERFORMED THE SENSOR CALLIBRATION , AND RESETED THE CONNECTIONS OF THE SIB . ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000098000000009 Remesh, rajan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : C arm get stuk duing procedue How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : C arm get stuk duing procedue Internal Comments : uLL# and/or description of uLL: C arm get stuk duing procedue Repair Action : C arm get stuk duing procedue Resoluion : C arm get stuk duing procedue Follow u Requred : uLL# and/or description of uLL: C arm get stuk duing procedue Repair Action : C arm get stuk duing procedue Technician Name : Remesh, rajan ----------------------------------------------------------------------------- | 2023-09-09 21:26 | No
363 | Yes | The three-beep sounded .. | The three-beep sounded .. | 000000000000000000000000001O##: 0000000000000000000000000010000000098000000007 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. I replaced the pad. [Simutaneou FCO 000000000000000000000000001k] I will again provide gudance for the volutary recall." September 26, 2000000003 Fuua Denshi Okayama Hanbai I sent the equpment along with a report to the on-site person in charge and contacted him. Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: Details I performed a battery test on the AED and confirmed that it passed the test.I monitored the AED in standby mode, bu no alarms, etc. occured. I verified the AED operation record and foud that the measuement resuts valus were ouside the permissable range per the specifications for the high voltage calibration items for the monthly auomatic self test and the battery fitted test, so the AED failed the self test. I foud ou that the single-beep alarm and the "To test, remove and re-install the battery" message was souded. Curently, the alarm is not reproduing, and it is possible that the alarm occured temporarily du to an increase in the temperatue aroud the AED. However, I will replace the AED in order to ensue safety and investigate the AED. "Also, as the lot for the received pad (Y110000000021-13) was a volutary recall (QSR23-000000002-2000000001-CC-EC-00000000 112322-00000000 2), I will replace it.I apologize for the inconvenience this has caued." ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 8/ 3. The cutomer contacted the contractor to say that the AED was beeping.The lamp flashed normally after he tried re-re-charging the battery a nuber of times.I told him to store the AED in a cool place becaue the temperatue continud to be high.Auut 300000000/ The cutomer contacted the contractor to say that the AED beeped again.This time, becaue it was placed in a cool place, the installation location Of Miho Technos was ueasy and requsted an inspection.The lamp flashes normally when the battery is re-ued. Lot No. Adut pad Y110000000021-13 Adut pad Y110000000021-13 Battery 865700000000----------------------------------------------------------------------------- | 2023-09-09 21:31 | No
364 | Yes | The therapy connector is not locked. | 000The therapy conneur is not locked. | 000000000000000000000000001O##: 0000000000000000000000000010000000098000000003 Katori, Masato The system meets the specification for the performed service and is retuned to ue Resoluion: [000000000000000000000000001k details] I replaced the therapy connector. [Simutaneou FCO 000000000000000000000000001k] There is no FCO on the device in which 000000000000000000000000001k is requred. [Special Notes] October 6, 2000000003 / I sent an equpment and 000000000000000000000000001k report to the person in charge of The Technical Service at The Company Of Japan. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit a 000000000000000000000000001k report on paper? ＜YES＞ Resoluion: [Special Notes] October 100000000 2000000003 / The person in charge of The Technical Service Attn: The Person in Charge of The Engineering Service at The Company of Japan contacted u to let him know that the equpment had been received. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: u Hayashi of The Technology Service department of The Company Device ue: Not ued for clinical ue. Eupected and ueupected Resuts: The therapy connector is not locked. uer Impact: None Patient Impact: None SE2.00000000.23 Diagnostic performed by Engineer: [Confirmation details] I confirmed 0000000022169928 duing case-0000000088900000000/0000000000000000000000000010000000088900000000 000000000000000000000000001k. Parts Consued: 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle Tools & Equpment : DEFIBRILLATOR ANALYZER, 2000000003-12-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000002-28 00000000:00000000:00000000 PRESSuE CALIBRATION TOOL, 200000000000000003-28 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-09 21:46 | No
365 | Yes | The paddle was damaged. | The paddle was damaged. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: There is a problem with the controller Safety Qustions: Case: 0000000022252312 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-09-09 22:06 | No
366 | Yes | Startup failed | 0R/u FD00Startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022252371 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport Neut Step Action: PPD spare parts: 00000000800000000500000000000000001, aging: 18H FSE: uuChaoliang-86-18662512168 Problem Reported by Cutomer: uB disk reading and writing failue SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 10000000000000000.2000000003 00000000:17:55 uC (86000000006500000000) "RDI Accuacy: Yes" $Q7: (confirmed reported problem & hostpc startu report write uLLsporadic ) $Q8: (do pc hardware test, OS Disk maybe have something wrong) $Q9: (replace hostpc OS disk, restore ghost and config backu, fuction test is ok ) $Q100000000 (Test and inspection data refer to: PA.&Device retuned to ful fuctionality.) $Q11: Calibrated Tools name ( N/A ) Euternal Remarks: 10000000000000000.2000000003 00000000:17:53 uC (86000000006500000000) Service content: Replacement of system hard disk, configuation of software. Service resut: After configuation test, the device 000000000000000000000000001ks normally. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: uuChaoliang-86-18662512168 T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that intermittent hard disk read and write uLL occured duing startu, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: intermittent hard disk read and write uLL duing startu Remote diagnostics and test resuts: NA PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Try to bypass the hard disk bou, or format the hard disk 2. Check the machine uLLreport and perform faut repairs Recommended spare part uuand name: " 0000000080000000050000000000000000Hard disk spare part Cutomer demand notes: Contact cutomers promptly to confirm Onsite time ----------------------------------------------------------------------------- | 2023-09-09 22:06 | No
367 | Yes | Startup failed | 0R/u FD0uray not possible | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/u uuuFD200000000 T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomer feedback cannot be euposed and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: No Connection with FD Controller after 300000000seconds Remote diagnosis and test resuts: R2 light in the uper left corner does not light u PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Try to format the FDC and reinstall the software. If it doesn’t 000000000000000000000000001k, replace the FDC. 2. Check the machine for uLL and perform faut repairs. Recommended spare parts uuand name: 989710000000000000000573 Flashlite High End KIT Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Euternal Remarks: 12.00000000.2000000003 19:16:18 uC (8600000000900000000) "SuT. Ready for Clinical ue local time" : " 2000000003-00000000-00000000 18:00000000" Safety Qustions: Case: 000000002225200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 20000000000000000.2000000003 00000000:0000000053 uC (8600000000900000000) Faut phenomenon: The device cannot be euposed. Faut analysis: After arriving at the site and checking the log, we foud that the main faut was IPPC and there was an FDC uLLreport.It is supected to be caued by an IPPC failue. Trouleshooting: Readjut the IPPC and reinstall the IPPC system. After completion, test the system to be normal. Internal Remarks: 20000000000000000.2000000003 00000000:0000000055 uC (8600000000900000000) RDI Accuacy: Yes $Q1:(N) $Q2:(N) $Q3:If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? No $Q00000000as this an ou of bou failue? No $Q7: (Confirmation & Description of problem) The device cannot be euposed. $Q8: (Trouleshooting & Resuts) After arriving at the site and checking the log, it is foud that the faut is mainly caued by IPPC, and there is an uLLreported by FDC.Supected to be caued by IPPC failue. $Q9: (Repair steps) Readjut the IPPC and reinstall the IPPC system, and then test whether the system is normal. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:Calibrated Tools name:（FLuE287C） SN:（310000000000000000200000000 (20000000000000000000000002), (Torqu wrench 752MFRMH) SN: (E200000000600000000000000002) (00000000/18/2000000002) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: scan second:NA SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 989710000000000000000573 1 Arranged to go to the hospital on 18H FSE:Tang Gunglin-86-1800000000000000009777 ----------------------------------------------------------------------------- | 2023-09-09 22:11 | No
368 | Yes | Image disk 0ailed | 0R/u FD0Image disk 0failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the image hard drive reported an uLLand requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: POST "Frontal IP-PC Image disk 2" Done/Failed Remote diagnosis and test resuts: The hard disk is directly connected PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Try formatting the hard drive. If it doesn’t 000000000000000000000000001k, replace it with IPPC. Recommended spare parts uuand designation: " 0000000080000000050000000000000000Hard disk spare part &0000000080000000050000000092 IP PC Revised_II no HDD Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Internal Remarks: 13.00000000.2000000003 17:00000000100000000uC (8600000000900000000) RDI Accuacy: Yes $Q1:(N) $Q2:(N) $Q3:If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? No $Q00000000as this an ou of bou failue? No $Q7: (Confirmation & Description of problem) The device cannot be euposed. $Q8: (Trouleshooting & Resuts) uon arrival at the site for inspection, it was foud that the IPPC reported an image storage failue, and the hard disk was replaced soon, so it was supected that the IPPC had a failue. $Q9: (Repair steps) Replace the IPPC and download the system again.After the replacement, the test equpment operates normally.Runing observation. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:Calibrated Tools name:NA For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: scan second:NA Safety Qustions: Case: 0000000022252700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD 0000000080000000050000000092 1 Arranged to go to the hospital on 18H FSE:Tang Gunglin-86-1800000000000000009777 Euternal Remarks: 13.00000000.2000000003 17:0000000000000000 uC (8600000000900000000) Faut phenomenon: The device cannot be euposed. Faut analysis: After arriving at the site for inspection, it was foud that the IPPC reported an image storage failue. The hard disk was replaced not long ago, and it was supected that the IPPC was fauty. Trouleshooting: Replace the IPPC and download the system again.After completing the replacement, the test equpment was runing normally.Observe while runing. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-09 22:11 | No
369 | Yes | poweron failed | 0/u/poweron failed | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ##d to ASP ui'an HongGung for buiness follow-u Safety Qustions: Case: 0000000022252700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: Dear honggung Teacher Chen 13991252955 reported for repair via QAAagain, please contact u as soon as possible, thank you Problem Reported by Cutomer: The machine cannot be tuned on ----------------------------------------------------------------------------- | 2023-09-09 22:16 | No
370 | Yes | No imaging | No imaging | 000000000000000000000000001O##: 00000000000000000000000000100000000985112 100000000Sep-2000000003 Sebastian Walczak The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: Yes. Eupected and ueupected Resuts: no image on the live signal. uer Impact: no. Patient Impact: no. SE 8.1.17. Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000003-00000000 00000000:00000000:00000000 MuTIMETER, 200000000000000002-25 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000008-00000000-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-30000000000000000:00000000:00000000 Parts Consued: 1.0000000000000000000000009896000000000000000000000000,Detector Pu0000000000000000High Speed pack. Resoluion: Replacement of the detector. Necessary tests and calibrations were performed with a positive resut. The faut was repaired. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000996000000006 100000000Sep-2000000003 Marek Jeske The system does not meet the specification for the performed service and is not retuned to ue. Resoluion: The uer reported a lack of imaging. Analysis of logs and image revealed that the detector is damaged. The detector needs to be replaced. Installed Produt. 5870000000081 The uer reported a problem with the device FD200000000ORC (sn 87) "No imaging". Log analysis points at lack of signal from the detector while maintaining the commuication: 2500000000Information 2000000003-00000000-00000000 00000000:35:29.0000000000000000Detector configuation is Ok: Frontal FD Controller with configuation ID 16 0000000000000000000000000000000000000000tartu Service NoDetailInfo. 200000000Information 2000000003-00000000-00000000 00000000:35:28.0000000000000000FD Controller: Flat Detector Susystem is AVAILABLE for Clinical Application 000000003210000000021 Startu Service NoDetailInfo. 27600000000nformation 2000000003-00000000-00000000 00000000:31:18.0000000000000000Last ru kV = 125.00000000Last ru mA = 00000000.9 Last ru ms = 9.00000000alcuated ObjectThickness = 53.2 Requsted Dose = 55.00000000nG Cine Dose Ratio. = 16.00000000ine Framespeed = 15.00000000DoseSenseTaste = 00000000ilter Type (SpectralFilter) = 00000000100000000mm Cu+ 1.00000000 mm Al DetectorFormat = Diameter: 00000000mm, Width: 3800000000mm, Height: 300000000 mm MeasuingField = ACQMEASuINGFIELD_CIRCuAR00000000Focu = Largest Techniqu = Cine Matriu Matriu = (height:10000000000000000width:10000000000000000epth:12) ValidRectangle = (bottom:1000000003,left:116,right:900000000,top:00000000 CluterMode = 19 510000000010000000018 Normal Operation Service NoDetailInfo. 2799 Information 2000000003-00000000-00000000 00000000:31:59.0000000000000000Last ru kV = 120000000000000000Last ru mA = 160000000000000000Last ru ms = 00000000Calcuated ObjectThickness = 53.2 Flu Requsted Dose = 000000002 nG Flu Dose Ratio = 1. 6 Flu Framespeed = 15.00000000EDL = 00000000000000000000000097 DoseSenseTaste = 00000000ilter Type (SpectralFilter) = 0000000000000000mm Cu+ 1.00000000 mm Al DetectorFormat = Diameter: 00000000mm, Width: 3800000000mm, Height: 300000000 mm MeasuingField = ACQMEASuINGFIELD_CIRCuAR00000000Focu = Smallest Flu Amplifier = Low Techniqu = Pused Flu Matriu Based = (height:10000000000000000width:10000000000000000epth:12) ValidRectangle = (bottom:1000000003,left:116,right:900000000,top:00000000 CluterMode = 16 510000000010000000018 Normal Operation Service NoDetailInfo. Reporting by Ms. Joanna Prabuka,technician50000000038000000006 Contact: Violetta- ward cardiac sugery ward. Today u Sebastian Walczak will be on site, neut to the apparatu, please give him the onsite card to be checked. It seems that the detector is damaged. Please order ugently: 9896000000000000000000000000 - Detector Pu0000000000000000High Speed pack And if the part arrives tomorrow - please send the person for replacement and calibration - approu. 00000000ous. M Problem Description by Engineer: Cutomer Fuction/Role: uer Device ue: yes. Eupected and ueupected Resuts: no image - detector failue - mut be euchanged. uer Impact: no Patient Impact: no. SE8.1. ----------------------------------------------------------------------------- | 2023-09-09 22:41 | No
371 | Yes | Therapy test failed error | Therapy test failed uLL(Replaced ASSY uuuance) | 000000000000000000000000001O##: 000000000000000000000000001000000009928200000000andha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 0000000056000000009000000001 DFM100000000 uuuitance ASSY to be replaced Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : 0000000056000000009000000001 DFM100000000 uuuitance ASSY to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000993200000000 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-09-09 23:02 | No
372 | Yes | System power failure | System power failure | 000000000000000000000000001O##: 00000000000000000000000000100000000996200000000September 13, 2000000003 Lus Abadia The system meets the specification for the performed service and is retuned to ue. Resoluion: It is indicated that the clock battery mut be changed, if after the change it gives the same uLL the CPuof the equpment mut be changed. Diagnostic performed by Engineer: The equpment is marking the rtc battery voltage uLL Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: 1st floor Eupected and ueupected Resuts: OK uer Impact: No Patient Impact: SE Reguatory Qustions: N/A ----------------------------------------------------------------------------- | 2023-09-09 23:31 | No
373 | Yes | u frozen in system shutting down mode | 00CDN0frozen in system shutting down mode | 000000000000000000000000001O##: 00000000000000000000000000100000000990000000000000000100000000ep-2000000003 Kolbinger Patrick The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System stuk at startu When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system down reboot will not resolve uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion : checked logs ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000990000000021 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System stuk at startu When was the first occurence? Within... : 00000000-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system down reboot will not resolve uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Kim reported that the system will not complete start u and is stuk at the "00000000 coutdown. Mutiple reboots attempted bu the issu remains. uable to save logfile as the system is not coming u completely. Caller is requsting qute for onsite suport. Internal Comments : uLL# and/or description of uLL: System stuk at startu Repair Action : Possible issu with the fleuvision pc not booting u or in a lost state cauing the host pc to continu to wait. Inspect fleuvision pc and restart it. Maybe there are other issus bu go into FSF at the u splash screen by pressing ALT+F. Replace defective component. 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: System stuk at startu Repair Action : Possible issu with the fleuvision pc not booting u or in a lost state cauing the host pc to continu to wait. Inspect fleuvision pc and restart it. Maybe there are other issus bu go into FSF at the u splash screen by pressing ALT+F. Replace defective component. 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kim hennings IP & Issu: 67223366 frozen in system shuting down mode Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created case & 000000000000000000000000001rtac called on call Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-09 23:36 | No
374 | Yes | Machine is not switching on | Machine is not switching on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001636 200000000Sep-2000000003 Pandey, Rajesh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uray Generator not switching ON. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Resoluion : Replaced the part and foud ok. Checked the system in few procedues and handed over the system in good 000000000000000000000000001king condition Parts Consued : 1.00000000000000000000000000000000210000000019971,Mains Interface uit (MIu Certeray ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006600000000100000000ep-2000000003 Mohammed, Ashfaq The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: cath uer Device ue: in ue Eupected and ueupected Resuts: System went off uer Impact: No Patient Impact: No SE 1.2.5 Resoluion : Monitored system performance for more than 100000000cases and foud it is 000000000000000000000000001king with u recommended specifications. Parts Consued : 1.00000000000000000000000000000000210000000020000000091,uSC Certeray 1.00000000000000000000000000000000210000000012800000000,Anode Drive uit (ADu Certeray 1.00000000000000000000000000000000210000000013887,Power Inverter (PoInt) Certeray 1.0000000000000000000000000000000080000000000000000900000000,PDuFues Generator Tools & Equpment : MuTIMETER,2000000003-11-00000000 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000001-11 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000001-11 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000990000000091 100000000ep-2000000003 Pandey, Rajesh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : uray Generator not switching ON. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Replaced the adu,point and usc, generator fue. Reloaded system ful software and performed calibration. Handed over in 000000000000000000000000001king condition Follow u Requred : Futher 000000000000000000000000001k requred ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099000000002 13-Sep-2000000003 Mohammed, Ashfaq The system meets the specification for the performed service and is retuned to ue Follow u Requred : Monitoring and handover Problem Description by Engineer : Cutomer Fuction/Role: cath uer Device ue: in ue Eupected and ueupected Resuts: System went off uer Impact: No Patient Impact: No SE 1.2.5 Resoluion : As per diagnosis done in earlier visit, replaced below parts and completed requred calibration. 1. uray segment controller— 1qty. 2. Anode drive uit - 1qty 3. Main interface uit - 1qty 00000000Power inverter uit - 1 qty 5. PDuFues - 1 qty. After sucessfuly testing, handed over system in ful 000000000000000000000000001king condition. Tools & Equpment : MuTIMETER,2000000003-11-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099000000000000000011-Sep-2000000003 Mohammed, Ashfaq The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Parts needs to replace for futher diagnosis Problem Description by Engineer : Cutomer Fuction/Role: cath uer Device ue: in ue Eupected and ueupected Resuts: System went off uer Impact: No Patient Impact: No SE 1.2.5 Diagnostic performed by Engineer : Diagnosis performed as per detailed report reference to 000000000000000000000000001O#nuber ￼000000000000000000000000001000000009900000000200000000Tools & Equpment : MuTIMETER,2000000003-11-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000990000000020000000000000000-Sep-2000000003 Pandey, Rajesh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : uray Generator not switching ON. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Follow u Requred : Need to replace parts Diagnostic performed by Engineer : As cutomer reported the issu , checked the system and foud generator cabinet is not getting on. Open the cabinet and foud water on uSC,POINT,ADuand MIu Becaue of water some parts buned. System software also corruted. Need to reload system software. Note-1-After replacement of physical damaged part only we can perform other activities or diagnosis to complete 000000000000000000000000001k. 2-foud some water drops on Hivo tank and cooling uit as well. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000986100000000 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : uray Generator not switching ON. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Follow u Requred : uLL# and/or description of uLL: uray Generator not switching ON. Repair Action : FSE to reload the sute PC software and replace MIufor Generator Issu What are the neut Trouleshooting steps : FSE to reload the sute PC software and replace MIufor Generator Issu 00000000210000000019971 Mains Interface uit Technician Name : Kuar, Atu Internal Comments : uLL# and/or description of uLL: uray Generator not switching ON. Repair Action : FSE to reload the sute PC software and replace MIufor Generator Issu Diagnostic performed by Engineer : uray Generator not switching ON. Muis supected PSC blank issu also ----------------------------------------------------------------------------- | 2023-09-09 23:56 | No
375 | Yes | Technician opens a case to request replacement of adapter cables. | Techni0n opens a case to request replacement of adapter cables. | 000000000000000000000000001O##: 00000000000000000000000000100000000987859 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Follow u Requred: talked to co-rd Stella Reale, problem with test cable, from system it is uder contract, I order parts to be sent NEMO to: Tu Oupatient Clinic - Single District Via Ancona Angolo Via Scotte 7000000001 Taranto c/a head nuse Stella Reale 1st floor Tel 000000009 7786122 989800000000100000000981 - HeartStart Hands-free Cable Problem Description by Engineer: Cutomer Fuction/Role: coord Stella Reale Device ue: Diagnosis Eupected and ueupected Resuts: deflect cable problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from coord Stella Reale, problem with test cable. Diagnostic performed by Engineer: heard from coord Stella Reale, problem with test cable, from system it is uder contract, I order parts to send NEMO ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000990000000000000000 Parts Consued: 1.000000000000000000000000989800000000100000000981,HeartStart Hands-free Cable ----------------------------------------------------------------------------- Internal Remarks: Gentili, open an incident for every S/N indicating: "requst to replace adapter cables" CN32600000000336 CN32600000000335 CN32600000000337 CN326000000003300000000 Thank you Diego Signorini Service Delivery Manager Connected Care CS & Operations Italy u Healthcare u S.p.A. Viale Saru5, Milan 20000000026 - Italy Mobile: +39 30000000086200000000 E-mail: diego.signoriniucom ----------------------------------------------------------------------------- | 2023-09-10 01:06 | No
376 | Yes | Image disk 0ailed | u- request replacement of adapter cables. | 000000000000000000000000001O##: 00000000000000000000000000100000000987861 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: coord Stella Reale Device ue: Diagnosis Eupected and ueupected Resuts: not in clinical ue uer Impact: not in ue on patient Patient Impact: not in ue on patient SE na I heard from coord Stella Reale, problem with test cable, Diagnostic performed by Engineer: heard from coord Stella Reale, problem with test cable, from system it is uder contract, I order parts to send NEMO Follow u Requred: talked to co-rd Stella Reale, problem with test cable, from system it is uder contract, I order parts to be sent NEMO to: Tu Oupatient Clinic - Single District Via Ancona Angolo Via Scotte 7000000001 Taranto c/a head nuse Stella Reale 1st floor Tel 000000009 7786122 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099000000002 Parts Consued: 1.000000000000000000000000989800000000100000000981,HeartStart Hands-free Cable ----------------------------------------------------------------------------- Internal Remarks: Gentili, open an incident for every S/N indicating: "requst to replace adapter cables" CN32600000000336 CN32600000000335 CN32600000000337 CN326000000003300000000 Thank you Diego Signorini Service Delivery Manager Connected Care CS & Operations Italy u Healthcare u S.p.A. Viale Saru5, Milan 20000000026 - Italy Mobile: +39 30000000086200000000 E-mail: diego.signoriniucom ----------------------------------------------------------------------------- | 2023-09-10 01:06 | No
377 | Yes | Image disk 0ailed | u- Techni0n opens a case to request replacement of adapter cables. | 000000000000000000000000001O##: 00000000000000000000000000100000000987862 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: coord Stella Reale Device ue: Diagnosis Eupected and ueupected Resuts: problem with cable uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from coord Stella Reale, problem with test cable, Follow u Requred: talked to co-rd Stella Reale, problem with test cable, from system it is uder contract, I order parts to be sent NEMO to: Tu Oupatient Clinic - Single District Via Ancona Angolo Via Scotte 7000000001 Taranto c/a head nuse Stella Reale 1st floor Tel 000000009 7786122 Diagnostic performed by Engineer: heard from coord Stella Reale, problem with test cable, from system it is uder contract, I order parts to send NEMO ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099000000008 ----------------------------------------------------------------------------- Internal Remarks: Gentili, open an incident for every S/N indicating: "requst to replace adapter cables" CN32600000000336 CN32600000000335 CN32600000000337 CN326000000003300000000 Thank you Diego Signorini Service Delivery Manager Connected Care CS & Operations Italy u Healthcare u S.p.A. Viale Saru5, Milan 20000000026 - Italy Mobile: +39 30000000086200000000 E-mail: diego.signoriniucom ----------------------------------------------------------------------------- | 2023-09-10 01:07 | No
378 | Yes | Technician opens a case to request adapter cables. - | Techni0n opens a case to request adapter cables. - | 000000000000000000000000001O##: 000000000000000000000000001000000009878600000000ELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Diagnostic performed by Engineer: heard from coord Stella Reale, problem with test cable, from system it is uder contract, I order parts to send NEMO Problem Description by Engineer: Cutomer Fuction/Role: coord Stella Reale Device ue: Diagnosis Eupected and ueupected Resuts: problem with cable uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from coord Stella Reale, problem with test cable, Follow u Requred: talked to co-rd Stella Reale, problem with test cable, from system it is uder contract, I order parts to be sent NEMO to: Tu Oupatient Clinic - Single District Via Ancona Angolo Via Scotte 7000000001 Taranto c/a head nuse Stella Reale 1st floor Tel 000000009 7786122 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000990000000000000000 Parts Consued: 1.000000000000000000000000989800000000100000000981,HeartStart Hands-free Cable ----------------------------------------------------------------------------- Internal Remarks: Gentili, open an incident for every S/N indicating: "requst to replace adapter cables" CN32600000000336 CN32600000000335 CN32600000000337 CN326000000003300000000 Thank you Diego Signorini Service Delivery Manager Connected Care CS & Operations Italy u Healthcare u S.p.A. Viale Saru5, Milan 20000000026 - Italy Mobile: +39 30000000086200000000 E-mail: diego.signoriniucom ----------------------------------------------------------------------------- | 2023-09-10 01:11 | No
379 | Yes | During initial inspection of device, found that the power button is not working. | During initial inspeuon of device, found that the power button is not working. |  | 2023-09-10 01:31 | No
380 | Yes | Battery Issue- equipment does not charge | Battery Issue- equipment does not charge | 000000000000000000000000001O##: 00000000000000000000000000100000000988886 00000000-Sep-2000000003 Castro, Francisco The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Equpment with battery failue. Problem Description by Engineer: Cutomer Fuction/Role: Maintenance engineer Device ue: Defibrillator Eupected and ueupected Resuts: Battery failue. uer Impact: Equpment does not allow to discharge. Patient Impact: Not applicable. SE n/a Diagnostic performed by Engineer: We go to the site to check the uLLlog. The equpment requres a battery replacement, so we proceed to perform it. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000991125 00000000-Sep-2000000003 Castro, Francisco Problem Description by Engineer: Cutomer Fuction/Role: Maintenance engineer Device ue: Defibrillator Eupected and ueupected Resuts: Battery failue. uer Impact: Equpment does not allow to discharge. Patient Impact: Not applicable. SE n/a Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2023-09-10 01:56 | No
381 | Yes | Motorized movements unavailable. | Motorized movements unavailable. | 000000000000000000000000001O##: 00000000000000000000000000100000000989100000000 12-Sep-2000000003 Willie Davis The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: u7M200000000 Eupected and ueupected Resuts: Collision detected not on duing ue uer Impact: None Patient Impact: None SE 1.2 Diagnostic performed by Engineer : Calibrated and clean tue cover collision switches and longitue rails. No other issus were noted. ----------------------------------------------------------------------------- | 2023-09-10 02:16 | No
382 | Yes | when trying to take an exposure uLLhappens and not able to reselect application. | when trying to take an euposure uLLhappens and not able to reseleuapplication. | Internal Remarks : Interaction Type: New Case Caller: Elijah Johnson 50000000098 when trying to take an euposue uLLhappens and not able to reselect application. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Qute Reason for Entitlement Claim: na Action Taken: created Qute for Onsite Neut Step Action: Awaiting Qute Verbal PO Nuber: na Verbal PO Approver: na Verbal PO CAP: na ----------------------------------------------------------------------------- | 2023-09-10 02:26 | No
383 | Yes | Image disk 0ailed | QAA0/u/Poweron failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Euternal Remarks: 13.00000000.2000000003 12:00000000:12 uC (860000000089900000000 Service content: The DSA device cannot be powered on. It is determined that the power suply main control board is fauty. It will be restored after replacement. Service resuts: After service, the equpment 000000000000000000000000001ks normally. Internal Remarks: 18.00000000.2000000003 00000000:0000000000000000uC (860000000089900000000 RDI Accuacy: yes $Q7: (Confirmed reported problem. device cannot power u, Cutomer cannot ue device to diagnostic.) $Q8: (check the Fan and power tray, foud is no power Oupu, DCPS has been bun. to so replace the Fan and power tray board) $Q9: (replied a new fan and power tray board, device retun normal) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (PA) S/N (PA) Neut Calibration date (PA) SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002225500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Can’t tun on the compuer Neut Step Action: ASP ui'an HongGung Follow u SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA ----------------------------------------------------------------------------- | 2023-09-10 07:36 | No
384 | Yes | Low leak-CO0e breathing Risk | Low leak-CO0Re breathing Risk | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000100000000Nov-2000000003 Vishnu CR The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Tools & Equpment : ELECTRICAL SAFETY TESTER,20000000000000000000000001000000000000000000000000:00000000 FLOW METER,2000000003-12-29 00000000:00000000:00000000 Parts Consued : 1.000000000000000000000000989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000Diagnostic performed by Engineer : Cross checked the GDS Assy. NEED to replace the GDS Assy. Resoluion : Replaced new RP-Flow Sensor Assy, Air & O2, V600000000 Now the machine 000000000000000000000000001king properly. Handed over the equpment to cutomer in good 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099350000000012-Sep-2000000003 Vishnu CR The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Flow sensor Assembly air/O2 Part no: 989800000000617100000000Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : CO2 rebreathing alarm showing on screen ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000992815 Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Fse viist requred. Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Fse viist requred. ----------------------------------------------------------------------------- | 2023-09-10 11:11 | No
385 | Yes | It is not turning on. | It is not turning on. | BuQ&R Feedback: No SAP Cutomer Symptom Code: BOOT Cutomer Symptom Code: Boot u/Power-u Problem Reason: Corrective Maintenance. Engineer Failue Code: Electrical/Electronic Failue Repair Activity Code: Instruted Cutomer Problem Reported by Cutomer: The Equpment's problem: it is not switching on Priority: stopped uit: HOSPITAL NASR FAIAD LTDA Address: R DR WILLIAM FAIAD 15 Look for: Cleiton Neves Department: Clinical Engineering 001: +55600000000625000000006 002: ## engenharia00-client.com.br Q1Q2: No Safety Qustions: Case: 0000000022256273 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Clauio Fernandes Last udated by: Clauio Fernandes 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? N/A, SAP Transfer: Case Transferred to SAP. T2 Activities:Problem Description: The equpment doesn't tun on. Trouleshooting Action: On 11/00000000 - Cleiton (eng) informed that the equpment is not tuning on. Checked that the green led of the power of the Review Modue tuns on. Instruted to open cabinet M. Verified that the Host PC is not tuning on, becaue there is no power in the HD modue; Disconnected the power cable from the Host PC and measued 2200000000 on the cable. The Host PC power suply key was disconnected and tuned on again. Tuned on the Host PC, manully, uing the front buton. The HD modue LEDs tuned on again and the SW went u normally; Cardio and Vascuar image acqusition tests; OK Tuned off the equpment and we will wait 1 hou to tun it on. Tuned the equpment on withou failues. Equpment released for euams uder observation. Day 12/00000000 - Cleiton (eng) informs that the equpment tuned on and it is 000000000000000000000000001king normally; CONTACT: Cleiton 6000000009625000000006 Equpment is 000000000000000000000000001king; OK "Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable". RSE Andre Confalonieri Repair Action: Checked the hospital's power and the HOST PC compuer; Test/Inspection: OK. Test Equpment: N/A Safety Qustions: Case: 0000000022256273 Safety Qustion Statu: In Process. PSA: false PSE: false PSF Statu: Not at all. Created By: Clauio Fernandes Last udated by: Clauio Fernandes 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? N/A, ----------------------------------------------------------------------------- | 2023-09-10 19:21 | No
386 | Yes | Alarms/No-Sound: Unable to hear audible alarms: No Product Malfunction: ur/Configuration: Reset Configuration. | Alarms/No-Sound: Unable to hear audible alarms: No ProduuMalfunuon: ur/Configuration: Reset Configuration. | 000000000000000000000000001O##: 00000000000000000000000000100000000992600000000 Walker, Maria The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: Yes Caue: Soud oupu was set to the display Is Resoluion Known: Yes Resoluion Sumary: Configuation Issu Resoluion Detail: Set the euternal speaker as the defaut soud oupu device and rebooted the PC Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Hear auible alarms uer Impact: uable to hear auible alarms Patient Impact: SE C.00000000.00000000 Problem Description/ueupected Resuts: uable to hear auible alarms Diagnostic performed by Engineer : Problem Description: uable to hear auible alarms uable to hear auible alarms Form ued: Trouleshooting/Investigation Object: WDIu00000000 Not Object: Software Version: C.00000000.00000000 Date and Time of Issu: 9/1000000002000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: - Cutomer sees the alarm teut bu does not hear any soud - Cutomer confirmed the speaker was connected and soud was not mued Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tom 000000000000000000000000001d IP & Issu: 72791583 & auible alarms are not activated Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-11 07:06 | No
387 | Yes | Oxygen pressure sensor range error | 0/V0uLLAlarm:Ouygen pressure sensor range error | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022256539 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No Trouleshooting Action: Alarm: Ouygen pressue sensor range uLLTrouleshooting the DA board Problem Reported by Cutomer: 00000000/V600000000uLL ----------------------------------------------------------------------------- | 2023-09-11 07:06 | No
388 | Yes | Reinstall battery error | Reinstall battery error | 000000000000000000000000001O##: 00000000000000000000000000100000000992919 Raghuanshi, Aarya The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Cutomer reported that the device is beeping. uon completing the self test, the device is annoucing "reinsert battery". uon doing so, the device starts a self test and repeats the uLL Cutomer has installed a new battery bu issu persists. Faut has been isolated to the defibrillator. Replacement defibrillator requred. Medical Consuables Team have been requsted to organise a warranty replacement uit Please order 1 u 0000000056000000005651 00000000itson Place, St Johns, Aukland 1000000002, New Zealand Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Monitoring Eupected and ueupected Resuts: Reinsert battery uLL uer Impact: uable to ue device Patient Impact: No patient involved SE Not available Diagnostic performed by Engineer : Cutomer reported that the device is beeping. uon completing the self test, the device is annoucing "reinsert battery". uon doing so, the device starts a self test and repeats the uLL Cutomer has installed a new battery bu issu persists. Faut has been isolated to the defibrillator. Replacement defibrillator requred. Medical Consuables Team have been requsted to organise a warranty replacement uit ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000992959 Parts Consued : 1.0000000000000000000000000000000056000000005651,FRu (W B) Defib, u English, Euchg ----------------------------------------------------------------------------- | 2023-09-11 08:41 | No
389 | Yes | An uLLmessage came up last night | An uLLmessage came up last night | 000000000000000000000000001O##: 00000000000000000000000000100000000993000000006 Ikemoto, Katsua The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000CS: KI: u uda, u uda, Ja, Hiroshima: SN: DFM100000000 - An uLLmessage came u last night. Today the ME did an operation test and it was possible to check the log, etc. Resoluion: The service pass000000000000000000000000001d has not been released to the cutomer. I asked him to check the operation check sumary. Problem Description by Engineer: Cutomer Fuction/Role: Sales representative Device ue: Inqury, so NA Eupected and ueupected Resuts:JA Hiroshima: DFM100000000 / An uLLmessage was displayed last night, and today the ME carried ou an operation test, and he asked if he coud check the log, etc. uer Impact: Inqury, so NA Patient Impact: Inqury, so NA. SE Inqury, so NA ----------------------------------------------------------------------------- | 2023-09-11 09:46 | No
390 | Yes | System booting error | System booting error | 000000000000000000000000001O##: 00000000000000000000000000100000000993100000000 11-Sep-2000000003 S.H. Kang The system meets the specification for the performed service and is retuned to ue Resoluion : Reseat gpdu host net000000000000000000000000001k cable Checked system booting ok Problem Description by Engineer : Cutomer Fuction/Role: clinical uer Device ue: ouside of ue Eupected and ueupected Resuts: booting uLL uer Impact: no Patient Impact: no SE r8 ----------------------------------------------------------------------------- | 2023-09-11 11:41 | No
391 | Yes | The system did not start normally. | 0DS:AW: The system did not start normally. | 000000000000000000000000001O##: 0000000000000000000000000010000000099362000000001-Sep-2000000003 Tsuoi, Naoya The system meets the specification for the performed service and is retuned to ue Resoluion: I confirmed an uLLfrom the uit (POWER (&CONTROL) uIT WITHOu BOARDS) that suplies power to the drive system duing startu, so today I replaced the POWER (&CONTROL) uIT WITHOu BOARDS. After I replaced it, I checked the startu a nuber of times and confirmed normal operation. Parts ued: POWER(&CONTROL) uIT WITHOu BOARDS 00000000230000000032500000000× u 1 Reguatory Qustions: N/A Problem Description by Engineer: The system did not start normally. Emergency switch... messages suh as "Collimation not opening" and "Cold restart" were necessary 2 or 3 times. After that, operation retuned to normal, bu he said he wanted me to check the sitution remotely. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system did not start normally. Emergency switch... messages suh as "Collimation not opening" and "Cold restart" were necessary 2 or 3 times. After that, operation retuned to normal, bu he said he wanted me to check the sitution remotely. uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Diagnostic performed by Engineer: The problem has not occured since the system was restarted a nuber of times. Parts Consued: 1.00000000000000000000000000000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009932800000000atanabe, Atsuhi The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000S:AW:u Matsu: The system did not start normally. [Internal Comments] Comp This is the 3rd time this problem has occured on t000000000000000000000000001other issus this month.He wanted it to be responded to on-site. Below is the log eucerpt. Operation abnormality at the 1st startu uerMessage 2000000003-00000000-11 00000000:51:000000000000000000000000WARNING: Emergency stop activated; Reset to continu Startu Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000Application uLL Clea2 POST uLL[item=FrtLat:DriveAmpStateTest],[drive=uYBou] Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [detail=Faut at startu after reset] Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000Application uLL Clea2 POST uLL[item=FrtLong:DriveAmpStateTest],[drive=uYBou] Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [detail=Faut at startu after reset] Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000IPC Post: PCumains relay test failed Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] Warning [EuCEPTION_DESCRIPTION] - Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000IPC Post: E_Stop activated Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] Warning [EuCEPTION_DESCRIPTION] - Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000Application uLL POST failed: IPC>, PCumains relay Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] - Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000IPC Post: Frontal Stand PCutest failed Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] Warning [EuCEPTION_DESCRIPTION] - Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000Application uLL POST failed: IPC>, Frontal Stand PCufue Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] - Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000IPC Post: Lateral Stand PCutest failed Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] Warning [EuCEPTION_DESCRIPTION] - Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000Application uLL POST failed: IPC>, Lateral Stand PCufue Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] - Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000IPC Post: Table 1 PCutest failed Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] Warning [EuCEPTION_DESCRIPTION] - Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000Application uLL POST failed: IPC>, Table 1 PCufue Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] - Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000IPC Post: Table 2 PCutest failed Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] Warning [EuCEPTION_DESCRIPTION] - Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000Application uLL POST failed: IPC>, Table 2 PCufue Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000037 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] - Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000Power suply switched off Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000000000000[SEVERITY] Warning [EuCEPTION_DESCRIPTION] Mains relay NOT on Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000Power suply switched off Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000000000000[SEVERITY] Warning [EuCEPTION_DESCRIPTION] Power (POST) failue (check mains and fues power suply) Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000Application uLL Power suply not switched on Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000000000000[SEVERITY] uLL[EuCEPTION_DESCRIPTION] - Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000Power suply switched off Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000000000000[SEVERITY] Warning [EuCEPTION_DESCRIPTION] Mains relay NOT on Geometry Warning 2000000003-00000000-11 00000000:51:000000000000000000000000Power suply switched off Startu [FREE_TEuT] Time of occurence : 2000000003-00000000-100000000 23:50000000000000000[SEVERITY] Warning [EuCEPTION_DESCRIPTION] Power (POST) failue (check mains and fues power suply) Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000POST FleuMove uY HW Failed Startu FRus: CableSet uY check the electronic spare-part-finder Motor Assembly u check the electronic spare-part-finder Motor Assembly Y check the electronic spare-part-finder AMC Drive uY (DPE2uPH00000000 - AMC) check the electronic spare-part-finder POST Items: BeamLongituinal auis drive amplifier test Failed BeamTransversal auis drive amplifier test Failed BeamLongituinal motor access test Skipped [Description] BeamTransversal motor access test Skipped [Description] AMC uY Drive version test Passed AMC uY Drive memory access test Passed BeamLongituinal auis version test Passed BeamLongituinal auis mode test Passed BeamLongituinal motor thermistor test Passed BeamLongituinal encoder access test Passed BeamTransversal auis version test Passed BeamTransversal auis mode test Passed BeamTransversal motor thermistor test Passed BeamTransversal encoder access test Passed [FREE_TEuT] [TEST_REPORT]involved FRus: CableSet uY code#: check the electronic spare-part-finder Motor Assembly u code#: check the electronic spare-part-finder Motor Assembly Y code#: check the electronic spare-part-finder AMC Drive uY (DPE2uPH00000000 - AMC) code#: check the electronic spare-part-finder POSTItem: BeamLongituinal auis drive amplifier test Done/Failed [Description]: Tests the auis drive initial amplifier state,[Details]: Faut at startu after reset POSTItem: BeamTransversal auis drive amplifier test Done/Failed [Description]: Tests the auis drive initial amplifier state,[Details]: Faut at startu after reset POSTItem: BeamLongituinal motor access test Skipped [Description]: Tests whether the motor can be controlled correctly,[ Details]: No gudance, postitem not available POSTItem: BeamTransversal motor access test Skipped [Description]: Tests whether the motor can be controlled correctly,[Details]: No gudance, postitem not available POSTItem: AMC uY Drive version test Done/Passed [Description]: Tests whether the drive's firmware matches the eupected firmware version (download new firmware when failed),[Details]: AMC drive version match, eupected and detected < 5.16.7.3>; AMC auis SAM_Fleumoveu version match, eupected and detected <Ctuuph_383 5.16.7.2>; AMC auis SAM_FleumoveY version match, eupected and detected <Chuh_3800000000.16.7.2>; Bou name match, eupected and detected <uY_Bou>; Parameter checksu match, eupected and detected < 00000000112b0000000055 00000000a000000000000000000000000 >; POSTItem: AMC uY Drive memory access test Done/Passed [Description]: Tests whether the drive's data memory can be accessed correctly,[Details]: No details available. POSTItem: BeamLongituinal auis version test Done/Passed [Description]: Tests whether the auis id is present and correct (When failed: verify encoder cabling, when cabling ok: position calibration requred),[Details]: Auis name match; eupected and detected <SAM_FleumoveY> POSTItem: BeamLongituinal auis mode test Done/Passed [Description]: Tests whether the auis is operating in the correct mode,[Details]: Auis mode match; eupected and detected <int position> POSTItem: BeamLongituinal motor thermistor test Done/Passed [Description]: Tests the motor thermistor,[Details]: Connection statu : Connected POSTItem: BeamLongituinal encoder access test Done/Passed [Description]: Tests whether the auis encoder can be accessed,[Details]: No details available. POSTItem: BeamTransversal auis version test Done/Passed [Description]: Tests whether the auis id is present and correct (When failed: verify encoder cabling, when cabling ok: position calibration requred),[Details]: Auis name match; eupected and detected <SAM_Fleumoveu> POSTItem: BeamTransversal auis mode test Done/Passed [Description]: Tests whether the auis is operating in the correct mode,[Details]: Auis mode match; eupected and detected <int position> POSTItem: BeamTransversal motor thermistor test Done/Passed [Description]: Tests the motor thermistor,[Details]: Connection statu : Connected POSTItem: BeamTransversal encoder access test Done/Passed [Description]: Tests whether the auis encoder can be accessed,[Details]: No details available. [END_REPORT_CONTINu_FREE_TEuT]started:2000000003-00000000-11 6:51:00000000AM ended:2000000003-00000000-11 6:51:00000000AM POST for FleuMove uY HW [TEST_RESuT] FALSE [TEST_NAME] FleuMove uY HW [TEST_REPORT] Geometry uLL2000000003-00000000-11 00000000:51:000000000000000000000000POST IPC Geometry Failed Startu FRus: Power Control uit (PCu check the electronic spare-part-finder IPC Geometry check the electronic spare-part-finder Cable set R-cabinet check the electronic spare-part-finder IPC I/O Adapter board check the electronic spare-part-finder Fue 5A 2500000000 (sb 19300000000 check the electronic spare-part-finder Fue 7A 2500000000 (sb 19300000000 check the electronic spare-part-finder POST Items: Mains relay test Failed Frontal stand fue test Failed Lateral stand fue test Failed Table 1 PCufue test Failed Table 2 PCufue test Failed IPC Hardware version Passed IPC Software version Passed IPC GPIO Board test Passed IPC I/O Adapter Board test Passed PCucable connect test Passed 12V SAN present test Passed 12V GeoCAN present test Passed IPC CAN Board test Passed IPC SysCAN end node test Passed [FREE_TEuT] [TEST_REPORT]involved FRus: Power Control uit (PCu code#: check the electronic spare-part-finder IPC Geometry code#: check the electronic spare-part-finder Cable set R-cabinet code#: check the electronic spare-part-finder IPC I/O Adapter board code#: check the electronic spare-part-finder Fue 5A 2500000000 (sb 19300000000 code#: check the electronic spare-part-finder Fue 7A 2500000000 ( sb 19300000000 code#: check the electronic spare-part-finder POSTItem: Mains relay test Done/Failed The mains presence is tested. (If failed check if Emergency circut is active.) POSTItem: Frontal stand fue test Done/Failed The frontal stand mains are tested. (PCucable) POSTItem: Lateral stand fue test Done/Failed The lateral stand mains are tested. (PCucable) POSTItem: Table 1 PCufue test Done/Failed The table 1 mains are tested. (PCucable) POSTItem: Table 2 PCufue test Done/Failed The table 2 mains are tested. (PCucable) POSTItem: IPC Hardware version Done/Passed Tests the IPC hardware version POSTItem: IPC Software version Done/Passed Tests the IPC software version (download new IPC when failed) POSTItem: IPC GPIO Board test Done/Passed Test the fuctioning of GPIO Board POSTItem: IPC I/O Adapter Board test Done/Passed Test the fuctioning of I/O Adapter Board. POSTItem: PCucable connect test Done/Passed A cable connect feedback signal from the PCuis tested (PCucable). POSTItem: 12V SAN present test Done/Passed Check 12V statu of SAN POSTItem: 12V GeoCAN present test Done/Passed Check statu 12V suply for GeoCAN POSTItem: IPC CAN Board test Done/Passed Test the IPC CAN Board POSTItem: IPC SysCAN end node test Done/Passed The IPC Board end node switch setting is tested. [END_REPORT_CONTINu_FREE_TEuT]started:2000000003-00000000-11 6:51:00000000AM ended:2000000003-00000000-11 6:51:00000000AM POST for IPC Geometry [TEST_RESuT] FALSE [TEST_NAME] IPC Geometry [TEST_REPORT] Geometry uerMessage 2000000003-00000000-11 00000000:51:57.0000000000000000WARNING: Emergency stop activated; Reset to continu Startu Patient and Beam pos. The following 3 times Information 2000000003-00000000-11 00000000:51:00000000.00000000000000008600000000 Information 2000000003-00000000-11 00000000:1000000000000000000000000000000008600000000 Information 2000000003-00000000-11 00000000:3000000002.00000000000000008600000000 I confirmed the following t000000000000000000000000001geometry restarts. Information 2000000003-00000000-11 00000000:1000000002000000000000000000000000Recycling power to R-Cabinet. Information 2000000003-00000000-11 00000000:37:19.0000000000000000Recycling power to R-Cabinet. Information 2000000003-00000000-11 00000000:100000000000000000000000000000000POST Geo modue on pedestal Passed FRus: Geo modue on pedestal uuuuuuuuuuuu P POST Items: Geo modue on pedestal Passed [FREE_TEuT] started:2000000003-00000000-11 6:51:00000000AM ended:2000000003-00000000-11 7:10000000000000000AM Power On Self Test for Geo mono modue in pedestal [TEST_RESuT] TRu [TEST_NAME] Geo modue on pedestal [TEST_REPORT] involved FRus: Geo modue on pedestal code#: uuuuuuuuuuuu POSTItem: Geo modue on pedestal Done/Passed Information 2000000003-00000000-11 00000000:37:39.0000000000000000POST Geo modue on pedestal Passed FRus: Geo modue on pedestal uuuuuuuuuuuu P POST Items: Geo modue on pedestal Passed [FREE_TEuT] started:2000000003-00000000-11 7:35:00000000 AM ended:2000000003-00000000-11 7:37:39 AM Power On Self Test for Geo mono modue in pedestal [TEST_RESuT] TRu [TEST_NAME] Geo modue on pedestal [TEST_REPORT] involved FRus: Geo modue on pedestal code#: uuuuuuuuuuuu POSTItem: Geo modue on pedestal Done/Passed The E-stop record is as follows: Information 2000000003-00000000-11 00000000:51:000000000000000000000000EStopActive = 1 changed [FREE_TEuT] - Time of occurence : 2000000003-00000000-100000000 23:50000000000000000 Information 2000000003-00000000-11 00000000:15:30000000000000000 EStopActive = 1 changed [FREE_TEuT] - Time of occurence : 2000000003-00000000-11, 000000001000000005 Information 2000000003-00000000-11 00000000:35:00000000.0000000000000000EStopActive = 1 changed [FREE_TEuT] - Time of occurence : 2000000003-00000000-11, 000000003000000006 [Parts from the equpment] 00000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS QTY:1 ZREP/Traceable --None-- 00000000800000000200000000835:GEO MODuE TILT ORT FLEuMOVE KIT WP QTY:1 ZREP/Traceable --None-- [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Naoya Tsuoi] Tsuoi, Naoya - FSE1 Visit date and time 2000000003/00000000/1117:00000000 - FSE1 000000000000000000000000001k time 3 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Naoya Tsuoi - Telephone nuber 00000000000000005969-8800000000 - Others Call mobile uon arrival Requster name Atsuhi Watanabe [Euamination category] Other: Follow u Requred: * Problem Description by Engineer: The system did not start normally. Emergency switch... messages suh as "Collimation not opening" and "Cold restart" were necessary 2 or 3 times. After that, operation retuned to normal, bu he said he wanted me to check the sitution remotely. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system did not start normally. Emergency switch... messages suh as "Collimation not opening" and "Cold restart" were necessary 2 or 3 times. After that, operation retuned to normal, bu he said he wanted me to check the sitution remotely. uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Diagnostic performed by Engineer: RC Response: Cold restart: OK Save Log Files creation: OK Remote connection: OK I euplained how to do it: OK Euplanation regarding failue locations: OK I remotely analyzed the log and foud that at the time of the three startus, incluing the second cold restart, the emergency stop buton was inpu, and that a power suply circut abnormality, incluing the former drive system, was detected. After the 2 cold restarts, I checked the emergency stop buton and, after a long press of the On buton Geomtery restart, I confirmed that operation was fine.* He said that he did not feel that the emergency stop buton was still pressed. An FSE will investigate on site. ----------------------------------------------------------------------------- | 2023-09-11 11:51 | No
392 | Yes | The device is not opening. | The device is not opening. | 000000000000000000000000001O##: 00000000000000000000000000100000000998270000000018-Sep-2000000003 Job Ibili The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The device does not tun on. How often is the issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No Eupected and ueupected Resuts: The device does not tun on. uer Impact: No Patient Impact: No SE 000Resoluion: The fauty part in the device has been replaced with a new one. Necessary settings have been made. The relevant problem has been solved. Parts Consued: 1.00000000000000000000000000000000800000000000000003521,RGB MEDIAWALL 0000000000000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000996753 11-Sep-2000000003 Bakir, Hakan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: no ue Eupected and ueupected Resuts: system is not 000000000000000000000000001king uer Impact: not involved Patient Impact: not involved SE000Diagnostic performed by Engineer: Eyüp İbili was accompanied duing the malfuction. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099380000000011-Sep-2000000003 Job Ibili The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Fit the part Problem Description by Engineer: Problem description by engineer: The device does not tun on. How often is the issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No Eupected and ueupected Resuts: The device does not tun on. uer Impact: No Patient Impact: No SE 000Resoluion: Necessary checks have been made on the device. It was seen that the problem was caued by the image chain. A new part of the defective part has been ordered. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000993617 Buak Yediyildiz The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The device does not tun on. How often is the issu occuring? : First occurence Malfuction area: PC's Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Internal Comments: uLL# and/or description of uLL The device does not tun on. Repair Action: The device does not tun on. Eu's dongle is not recognized by the system. Site visit requred. Diagnostic performed by Engineer: I 000000000000000000000000001ked with Eu Temualay. 00000000:00000000 remains when the device is tuned on. It was tuned off via FleuPC and the problem was not resolved. Site visit requred Follow u Requred: uLL# and/or description of uLL The device does not tun on. Repair Action: The device does not tun on. Eu's dongle is not recognized by the system. Site visit requred. What are the neut Trouleshooting steps: The device does not tun on. Eu's dongle is not recognized by the system. Site visit requred. Technician Name: Buak Yediyildiz ----------------------------------------------------------------------------- | 2023-09-11 12:21 | No
393 | Yes | Machine is not booting up | Machine is not booting up | 000000000000000000000000001O##: 0000000000000000000000000010000000099383000000001-Sep-2000000003 Yadav, Santosh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not booting u Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : checked the system and foud that there was an issu with the fleuvision Pc , tried to reload the software of the fleuvision PC bu same problem persists . needs to replced the fleuvision PC . Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : MuTIMETER,2000000003-10000000020000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000993000000001 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not booting u Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Machine is not booting u. Fleuvision grabber card initialization failed Internal Comments : uLL# and/or description of uLL: Fleuvision grabber card initialization failed Repair Action : Need to replace Fleu vision card Follow u Requred : uLL# and/or description of uLL: Fleuvision grabber card initialization failed Repair Action : Need to replace Fleu vision card What are the neut Trouleshooting steps : Need to replace Fleu vision card 0000000080000000058600000000 FleuVision-2_Revised PC no HDD Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-09-11 12:21 | No
394 | Yes | System does not start correctly, FleuVision screen cut in two | System does not start correuy, Fleusion screen cut in two | 000000000000000000000000001O##: 00000000000000000000000000100000000995100000000 12-Sep-2000000003 Andrighetto Sylvio The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Manager Device ue: FD 20000000000000000 Eupected and ueupected Resuts: The system does not start anymore. uer Impact: yes Patient Impact: no SE 8.1.300000000Resoluion: Replacement of the FleuVision PC. Parts Consued: 1.0000000000000000000000000000000080000000058600000000, FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001786 September 11, 2000000003 Andrighetto Sylvio The system does not meet the specifications for the performed service and is not retuned to ue. Resoluion: Looking for equpment on Fedeu site and tracking. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Engineer Device ue: FD 20000000000000000 Eupected and ueupected Resuts: Fleuvision does not start properly. uer Impact: yes Patient Impact: no SE 8.1.300000000Diagnostic performed by Engineer: FleuVision does not start properly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000993912 September 11, 2000000003 Andrighetto Sylvio The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: Replacement of FleuVision PC and HDD. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Manager Device ue: FD 20000000000000000 Eupected and ueupected Resuts: The system does not start anymore. uer Impact: yes Patient Impact: no SE 8.1.300000000Diagnostic performed by Engineer: Fleuvision PC in uLL Tossed with IGT Help Desl. Ordered FleuVision PC and hard disk. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000993798 Andrighetto Sylvio The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: On-site intervention Diagnostic performed by Engineer: Impossible to connect to RSN. System freezes with Fleu screen cu in 2. On-site visit Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Manager Device ue: FD 20000000000000000 Eupected and ueupected Resuts: The system does not start anymore. uer Impact: yes Patient Impact: no SE 8.1.300000000----------------------------------------------------------------------------- | 2023-09-11 12:26 | No
395 | Yes | System does not start | System does not start | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 September 15, 2000000003 Borchardt, Thomas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Called Ms. Bukhardt back 000000000000000060000000032100000000 The reported uLLis confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion: Replacement us incluing Battery pack, system check, OK. Parts Consued: 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000990000000000000000eptember 12, 2000000003 Borchardt, Thomas The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: Replacement us incluing Battery pack Problem Description by Engineer: Problem description by engineer: Called Ms. Bukhardt back 000000000000000060000000032100000000 The reported uLLis confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: uS set ou of operation, fuction check, OK. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099300000000 Nölkes, Fabian The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: Called Ms. Bukhardt back 000000000000000060000000032100000000 The reported uLLis confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: Called Medical Technology ms. Bukhardt back 000000000000000060000000032100000000. System does not start. Power suply checked. Host PC and image PC withou voltage. uS does not 000000000000000000000000001k. Trouleshooting performed. Probably a defective uS. Service requsted to bridge the uS for short-term initial operation again. If a defect is detected, order and replace uS. Internal Comments: uLL# and/or description of uLL Trouleshooting parts shows power suply of the PC via L00000000rom uS. uS does not 000000000000000000000000001k. Repair Action: . Follow-u requred: uLL# and/or description of uLL Trouleshooting parts shows power suply of the PC via L00000000rom uS. uS does not 000000000000000000000000001k. Repair Action: . What are the neut trouleshooting steps: Please schedue onsite service, 2 hous AZ Trouleshooting on the uS, bridge if necessary. If defective, follow-u with new uS Technician name: Nölkes, Fabian ----------------------------------------------------------------------------- Cutomer comment: ## Follow-u service also schedued for 15.00000000. Cutomer comment: Ms. Bukhardt +0000000060000000032100000000 will call again. Cutomer is asking for prompt service and call back. Internal Remarks: Please schedue onsite service, 2 hous AZ Trouleshooting on the uS if necessary bridge so that the system can go into operation again for the first time. If defective, follow-u with new uS Internal Remarks: 1u 00000000800000000117161 uS 1Phase Data Integrity Controller SP 1u 00000000800000000117151 uS 1Phase Data Integrity Battery SP Cutomer comment: Ms. Bukhardt, 000000000000000060000000032100000000, 000000000000000000000000001ld like to be informed abou the statu. Cutomer comment: ### Ms. Bukhardt has appointment info 12.00000000. approu. 9:00000000/9:300000000hrs. Internal Remarks: 00emote: via FSE Borchardt = replacement us incluing Battery pack, 12 NC please via remote. Have it searched ou ----------------------------------------------------------------------------- | 2023-09-11 13:11 | No
396 | Yes | Host PC bootup failure SPLIT_ | Host PC bootup failure_SPLIT_ | 000000000000000000000000001O##: 0000000000000000000000000010000000099000000008 11-Sep-2000000003 Sato, Taketo The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: No Eupected and ueupected Resuts: No uer Impact: No Patient Impact: R8.1.3000000005 SE Resoluion: [Details of requst] SF: Replace host PC. [000000000000000000000000001k content details] Regarding the replacement of the host PC 00000000000000000000000000100000000992700000000reported to #7. After I replaced the host PC, I confirmed that the large monitor control PC (FleuVision PC) did not start. When I investigated, I foud a software uLL so I reinstalled the FleuVision PC software. After I reinstalled it, it started withou a problem and I checked the operation. There is no problem with equpment operation, etc. [Other failues/Measues] No [Equpment statu] Equpment can be ued [000000000000000000000000001k statu] 000000000000000000000000001k completed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099270000000011-Sep-2000000003 Sakakua, Ryoya The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: No Eupected and ueupected Resuts: The system does not 000000000000000000000000001k. uer Impact: No Patient Impact: No SER8.1.3000000005 Resoluion: I confirmed that the equpment does not start, so I investigated and foud that I confirmed that the host PC, which controls the mains, was the caue. I replaced the host PC and reinstalled the software. There was no problem with the confirmation 000000000000000000000000001k after the 000000000000000000000000001k was completed. Follow u Requred: N/A Resoluion: The Mac address has changed as shown below. Old Mac address: 00000000 C00000000A B000000003 AD New Mac address: 28 B9 D9 E5 13 A1 Resoluion: Replacement part: (1) Part name: u Haswell Biplane Host PC no HDD / Part Nuber: 0000000080000000012200000000 / Serial: 722216917511 / Nuber of uits: 1 (2) Part name: Hard disk spare part / Part nuber: 0000000080000000050000000000000000/ Nuber of uits: 1 Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Biplane Host PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-09-11 14:11 | No
397 | Yes | does not start up; it hangs in its startup process | does not start up; it hangs in its startup process | 000000000000000000000000001O##: 00000000000000000000000000100000000993793 MARTIN VERSLOOT The system does not meet the specifications for the performed service and is not retuned to ue Follow u Requred: FleuVision PC defective and needs to be replaced Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: in ue Eupected and ueupected Resuts: uer Impact: system down Patient Impact: no SEr000Diagnostic performed by Engineer: FleuVision PC defective and needs to be replaced, also hard delivered, this was also defective Follow u Requred: Probably delivered Defoa PC, send new PC by taui and with hard disk Follow u Requred: Cutomer 000000000000000000000000001ld like to have an engineer with it ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099000000003 Parts Consued: 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002699 Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part ----------------------------------------------------------------------------- | 2023-09-11 14:31 | No
398 | Yes | Estimated time reaching 0but it does not start and nothing moves | Estimated time reaching 0but it does not start and nothing moves | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 100000000ep-2000000003 Antonio Garcia Herrera The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. The Control PC of the FleuVision monitor might not be starting correctly, J. M. Mora is in charge of checking it on site. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not complete the start-u sequnce. uer Impact: The equpment does not complete the start-u sequnce. Patient Impact: No impact SE 8.2.17.00000000Resoluion: The FleuPC is replaced, the hard disk of the defective fleuPC is replaced. It is started and the equpment 000000000000000000000000001ks correctly. Tests OK. Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099000000002 11-Sep-2000000003 José Manul Mora Pastor The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. The Control PC of the FleuVision monitor might not be starting correctly, J. M. Mora is in charge of checking it on site. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not complete the start-u sequnce. uer Impact: The equpment does not complete the start-u sequnce. Patient Impact: No impact SE 8.2.17.00000000Diagnostic performed by Engineer: After performing several diagnostic tests, it is determined that the caue of the incident may be in the Fleuvision PC or in its hard drive. Spare parts for both materials are requsted to repair the equpment. Follow u Requred: Replacement of materials Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000990000000000000000drián Sancho López The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. The Control PC of the FleuVision monitor might not be starting correctly, J. M. Mora is in charge of checking it on site. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not complete the start-u sequnce. uer Impact: The equpment does not complete the start-u sequnce. Patient Impact: No impact SE 8.2.17.00000000Diagnostic performed by Engineer: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. The Control PC of the FleuVision monitor might not be starting correctly, J. M. Mora is in charge of checking it on site. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. The Control PC of the FleuVision monitor might not be starting correctly, J. M. Mora is in charge of checking it on site. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. The Control PC of the FleuVision monitor might not be starting correctly, J. M. Mora is in charge of checking it on site. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-09-11 14:41 | No
399 | Yes | No Shock delivering .. | No Shock delivering .. | 000000000000000000000000001O##: 0000000000000000000000000010000000099000000001 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Euternal paddle foud problem, defective lot no pictue is enclosed herewith. Diagnostic performed by Engineer : No Shock delivering Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Follow u Requred : Euternal paddle foud problem, defective lot no pictue is enclosed herewith. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000995191 Parts Consued : 1.00000000000000000000000098980000000019600000000,uEuternal Paddles ----------------------------------------------------------------------------- | 2023-09-11 15:41 | No
400 | Yes | u does not start up; left half of the large screen: connection to installation source | NZA0 does not start up; left half of the large screen: conneuon to instalation source | 000000000000000000000000001O##: 00000000000000000000000000100000000993927 MARTIN VERSLOOT The system meets the specifications for the performed service and is retuned to ue Diagnostic performed by Engineer: Had Tomas place the HHD of Fleuvision PC in another slot of the diskbay, this solved the problem Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: uer Impact: system down Patient Impact: no SEr2.2.u ----------------------------------------------------------------------------- | 2023-09-11 16:11 | No
401 | Yes | Loudspeaker failure alarm is triggered-No sound. No Product Malfunction-ur. Cleaned speaker membrane. | Loudspeaker failure alarm is triggered-No sound. No ProduuMalfunuon-ur. Cleaned speaker membrane. | 000000000000000000000000001O##: 00000000000000000000000000100000000995600000000 September 13, 2000000003 Víctor Rodríguz González The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Muti parameter monitor Eupected and ueupected Resuts: Louspeaker failue. uer Impact: Impossible to emit acoutic alarms. Patient Impact: The equpment is ued with limitations. SE Resoluion: Monitor is checked and it is observed that it displays a message abou a louspeaker technical uLL This uLLbelongs to the u3 modue, it is checked and it is observed that a sticky liqud has fallen on the speaker membrane. Membrane is cleaned and it is checked that it no longer gives the uLL the equpment makes a soud and it is 000000000000000000000000001king perfectly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099500000000 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Muti parameter monitor Eupected and ueupected Resuts: Louspeaker failue. uer Impact: Impossible to emit acoutic alarms. Patient Impact: The equpment is ued with limitations. SE Diagnostic performed by Engineer: The client (Lus Migul López 912000000001155) says that the speaker does not 000000000000000000000000001k showing an uLL Material order is handled and the case is transferred onsite so that the equpment is checked by the field service. Resoluion: The case is transferred onsite. Follow u Requred: The client (Lus Migul López 912000000001155) says that the speaker does not 000000000000000000000000001k showing an uLL - Equpment located in PEDIATRIC INTENSIVE CARE, 3RD FLOOR OF THE CHILDREN'S HOSPITAL. - Assign fse - Order the mentioned material ----------------------------------------------------------------------------- | 2023-09-11 16:36 | No
402 | Yes | Intrepid VT alarm during cardioversion | Intrepid VT alarm during cardioversion | 000000000000000000000000001O##: 00000000000000000000000000100000000995593 Problem Description by Engineer : Wrong alarm triggered? Resoluion : The event was reviewed by ou Produt Specialist and he conclued that the device has behaved correctly. All relevant performance verification procedues were sucessfu. Software ugraded from 1.00000000.39 to 1.00000000.53. ----------------------------------------------------------------------------- | 2023-09-11 16:36 | No
403 | Yes | Host PC bootup failure SPLIT_ | Following case 0000 | 000000000000000000000000001O##: 00000000000000000000000000100000000995759 100000000ep-2000000003 uavier NEYRET The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: biomedDevice ue: noEupected and ueupected Resuts: noneuer Impact: nonePatient Impact: noneSEu Diagnostic performed by Engineer: Defective power suply control board. Resoluion: In room Biplan 1; Replacement of the MPD control uit board in order to remove the room's missing start-u problems. At the same time, I replaced the cable, the power suply and the control board of the intercom that were posing problems in a manner that was intermittent. The room is fuly operational again. Consued parts: 1.00000000000000000000000000000000230000000016000000007, MPD CONTROL uIT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009900000000 Regis LAMBING The system meets the specifications for the performed service and is retuned to ue. Internal Comments: 00000000230000000016000000007 Resoluion: intervention by u. Neyret Follow u Requred: intervention by u. Neyret Please order 00000000230000000016000000007 u1 Delivery to neuoradiology department, 1st and 1st Attn:u Piechel to: +3332000000005567 Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: biomedDevice ue: noEupected and ueupected Resuts: noneuer Impact: nonePatient Impact: noneSEu Diagnostic performed by Engineer: Requsts replacement of the MPDucontrol uit board. ----------------------------------------------------------------------------- | 2023-09-11 17:01 | No
404 | Yes | Speaker is faulty (Quote is cancelled by comer). | u Speaker is faulty (Quote is cancelled by cuomer). | 000000000000000000000000001O##: 00000000000000000000000000100000000995761 Buak Tezel The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Not ue. Eupected and ueupected Resuts: Speaker is fauty. uer Impact: No. Patient Impact: No. SE 2.00000000.21 Follow u Requred: The checks showed that the speaker part of the device was defective and not 000000000000000000000000001king. The defective speaker part needs to be replaced with a new one. According to the failue detection, 0000000056000000009331 - PAMuKALE uIVERSITY HOSPITAL SHOuD BE RECOMMENDED. Resoluion: Duing the checks, it was detected that the speaker part of the device was defective and not 000000000000000000000000001king. The defective speaker part needs to be replaced with a new one. According to the failue detection, 0000000056000000009331 - A new one shoud be directed to PAMuKALE uIVERSITY HOSPITAL. ----------------------------------------------------------------------------- | 2023-09-11 17:06 | No
405 | Yes | u Equipment Intermittent Malfunction | u- uEquipment Intermittent Malfunuon |  | 2023-09-11 17:11 | No
406 | Yes | Audio Sounds/Inoperative: Speaker Bou with problems: Speaker/HW/No Sound | Audio&Sounds/Inoperative: Speaker bou with problems: Speaker/HW/No Sound: Replaced 00000 | 000000000000000000000000001O##: 0000000000000000000000000010000000099000000003 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: parameter monitor Eupected and ueupected Resuts: fauty speaker uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: the other one is not 000000000000000000000000001king. Follow u Requred: nemo withou collection Manzoni Hospital Via Dell'Eremo n° 9/11 23900000000 LECCO Clinical Engineering 00000000 floor Tel. +39 30000000086000000009 Ac. Valencia ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000996773 Parts Consued: 1.0000000000000000000000000000000056000000007331,PCACY SPKR EuTERNAL D-Su ----------------------------------------------------------------------------- | 2023-09-11 19:02 | No
407 | Yes | Message "Maintenance to be scheduled" .. | Message "Maintenance to be scheduled" .. | 000000000000000000000000001O##: 0000000000000000000000000010000000099000000002 Thierry LE GONIDEC The system does not meet the specifications for the performed service and is not retuned to ue. Reguatory Qustions: No/No Diagnostic performed by Engineer: After a battery replacement, the DSA displays: Maintenance to be schedued, device not 000000000000000000000000001king. uLLcode: C 0000000000000000 0000000000000000 0000000000000000 Post/arm Flash /Chk su uLL This code corresponds to a memory uLL electronic failue. The device mut be replaced by a new one. The FR3 is not marketed anymore, it can be replaced by an HS1 or FRu. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown ----------------------------------------------------------------------------- | 2023-09-11 19:16 | No
408 | Yes | uLLMessage - Low Leak - CO2 | uLLMessage - Low Leak - CO2 | 000000000000000000000000001O##: 00000000000000000000000000100000000997000000008 Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Resoluion : corrected cutomer test set u by placing the whisper swivel inline with the test lug. uit now operates as eupected withou the Low Leak alarm. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: PM Service Eupected and ueupected Resuts: Duing PM PVT, ST performance, uit alarmed for Low Leak Co2 Rebreathing Risk. uer Impact: none Patient Impact: na SE uknown ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Shone Kuiakofe IP & Issu: 566322800000000 uLLMessage - Low Leak - CO2 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-11 19:46 | No
409 | Yes | Drop in blood oxygen level at certain point should have triggered alarm. Onsite scheduled for u | Drop in blood ouygen level at certain point should have triggered alarm. Onsite scheduled for 0Dec03 |  | 2023-09-11 20:04 | No
410 | Yes | Reference image on live fluoro on monitor in exam room | Reference image on live fluoro on monitor in euam room | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 15-Sep-2000000003 Sikora, Matthew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: IGT Eupected and ueupected Resuts: Eupected uer Impact: None Patient Impact: None SE Parts Consued : 1.000000000000000000000000989670000000000000000000000003,DIRECT VIEWPAD CARDIO 1.000000000000000000000000989670000000000000000000000003,DIRECT VIEWPAD CARDIO Diagnostic performed by Engineer : Installed new viewpad remotes. Tested and verified system is 000000000000000000000000001king within u specifications. Informed cutomer. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099910000000013-Sep-2000000003 Sikora, Matthew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: IGT Eupected and ueupected Resuts: Eupected uer Impact: None Patient Impact: None SE Diagnostic performed by Engineer : Installed new live monitor in euam room. Parts Consued : 1.000000000000000000000000991932252193,19'' Monochrome Monitor ER (MML1952-PER) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000999765 11-Sep-2000000003 Sikora, Matthew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: IGT Eupected and ueupected Resuts: eupected uer Impact: none Patient Impact: none SE Resoluion : Installed spare monitor no control room. Verified system fuctionality. Informed cutomer. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000997363 11-Sep-2000000003 Sikora, Matthew The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : reference image on live fluro on monitor in euam room When was the first occurence? Within... : 11-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Live image is not available in euam room uer Impact: reference image on live fluro on monitor in euam room Patient Impact: Patient moved to another room SE 8.2.17.00000000Diagnostic performed by Engineer : 000000000000000000000000001ked with cutomer and remote NSS. New euam room monitor and remote needed. Parts ordered and follow u schedued. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000997222 Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : reference image on live fluro on monitor in euam room When was the first occurence? Within... : 11-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Live image is not available in euam room uer Impact: reference image on live fluro on monitor in euam room Patient Impact: Patient moved to another room SE 8.2.17.00000000Diagnostic performed by Engineer : reference image on live fluro on monitor in euam room. system has been rebooted. Bio med is heading back u to room to check it ou after patient is moved. Internal Comments : uLL# and/or description of uLL: reference image on live fluro on monitor in euam room Repair Action : Check B & R cabinet to make sue all video components have booted u properly. Check live video path. Follow u Requred : uLL# and/or description of uLL: reference image on live fluro on monitor in euam room Repair Action : Check B & R cabinet to make sue all video components have booted u properly. Check live video path. Technician Name : Eckbold Glenn ----------------------------------------------------------------------------- Internal Remarks : Matt Sikora - 19" Live monitor needs to be replaced. Internal Remarks : Matt Sikora - cutomer ued old style viewpad, resuting in different monitor displays. When uing new style remote from room neut door, all 000000000000000000000000001ked fine. FSE will order correct remote for this SW R8.2 system. ----------------------------------------------------------------------------- | 2023-09-11 20:06 | No
411 | Yes | Drop in blood oxygen level at certain point should have triggered alarm. Onsite scheduled for u | Drop in blood ouygen level at certain point should have triggered alarm. Onsite scheduled for 0Dec03 |  | 2023-09-11 20:08 | No
412 | Yes | Dim screen | Dim screen | 000000000000000000000000001O##: 00000000000000000000000000100000000997383 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: dim display uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 00000000561528891 [RP-LCD, Assy,2nd GEN, V600000000 $1,387.00000000 0000000056153600000000 [RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000 $138.00000000 00000000561528911 [RP-Cable, u PCBA to LCD, V600000000 $66.00000000 00000000561528871 [RP-PCBA, uer Interface, u 2nd GEN,V600000000 $332.00000000 00000000561531651 [RP-Tray,LCD,2nd GEN V600000000 $229.00000000 00000000561528900000000 [RP-Screw,M2u3 Socket hd, PKG/100000000V600000000 $90000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states the display on device is dim and looking for replacement part nuber for LCD display. Confirmed with cutomer. foud cutomer has 1st gen LCD display and needs to ugrade to 2nd gen. Provided cutomer with part nubers and price to ugrade LCD to 2nd gen NEC display. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Clifford/Cliff SPRINGSTEEN IP & Issu: " 56627800000000 This device is RESTRICTED becaue its SCREEN is DIM" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: I created a case, R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-11 20:26 | No
413 | Yes | Tidal volume could not be adjured | Tidal volume could not be adjued |  | 2023-09-11 20:54 | Unknown
414 | Yes | There's no sound on monitor. | There's no sound on monitor. | 000000000000000000000000001O##: 00000000000000000000000000100000000997891 Massey Ryan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Alarm souds uer Impact: uable to hear alarms Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Central station not making souds Follow u Requred : Central station not making souds Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve Diagnostic performed by Engineer : Problem Description: Central station not making souds Soud not 000000000000000000000000001king on central Form ued: Other Trouleshooting Steps Taken: Created onsite at cutomer requst #CutomerRequstsTC Cutomer refues remote suport, requsting TC onsite ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Faith Dimerson IP & Issu: ccb- requsted to close the case, pls close the case Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: AA- ST RP Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Faith Dimerson IP & Issu: 72800000000000000006 - There's no soud on monitor. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-11 21:16 | No
415 | Yes | ER 00 coming up - vent may have shut off while on patient | ER 00- coming up - vent may have shut off while on patient | 000000000000000000000000001O##: 00000000000000000000000000100000000997695 Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Resoluion : reviewed the event log with the cutomer. Foud the uit had been operating on battery power, code 1212, for several hous util it gave a 1200000000alarm for low battery. there was a 1218 code, AC power restored, bu for only 100000000seconds when the 1212 and 1200000000reoccured. The uit powered off shortly after that du to low battery. Cutomer has charged the battery back to 16.5 vcd, uit is operation as eupected. Advised to review the log with RT along with battery charging procedues. uit to be placed back into service. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient Eupected and ueupected Resuts: RT reports uit shu off on it's own uer Impact: swapped vent Patient Impact: vent was swapped with no patient issu SE uknown ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: John 773000000003800000000 IP & Issu: V600000000/ SN 100000000000000006500000000iag code 2000000002 - coming u - vent may have shu off while on patient. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-11 21:21 | No
416 | Yes | Speaker malfunction-No sound; Replaced Speaker Assembly at bench | Speaker malfunuon-No sound; Replaced Speaker Assembly at bench | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Velarde, Andrew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Bench Repair uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Bench Repair speaker malfuction Diagnostic performed by Engineer : performed by Engineer Confirm cutomer complaint. The speaker of the u2 is no longer fuctional. Also, the front bezel is cracked. Will replace the speaker assembly and the front bezel to remedy. Resoluion : Replaced the Speaker Assembly and Side Bezel, Calibrated NBP and tested all other parameters and connections to ensue that the device is operating correctly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000008-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000008-31 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000008-31 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000056000000008621,MS_u2 ASSY CBL u2/MP2 Speaker Assembly 1.0000000000000000000000000000000056000000003921,MS_u2 PLST Side Bezel+PT TEuT(3ea) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000998000000002 Ramirez, Huo The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 0000000056000000008621 MS_u2 ASSY CBL u2/MP2 Speaker Assembly 0000000056000000003921 MS_u2 PLST Side Bezel+PT TEuT (3ea) 2-hou repair time Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Bench Repair uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Bench Repair speaker malfuction Diagnostic performed by Engineer : Confirm cutomer complaint. The speaker of the u2 is no longer fuctional. Also, the front bezel is cracked. Will replace the speaker assembly and the front bezel to remedy. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000998000000005 Mottley, Sheldon The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Problem Description: Bench Repair speaker malfuction Bench Repair speaker malfuction Form ued: Bench Trouleshooting Steps Taken: Was there a “speaker malfuction” INOP message? (Yes)Was the speaker able to produe soud, despite the uLLmessage. (uknown)How did they notice having a speaker issu?(Someone dropped it off in the biomed shop)Coud youplease confirm, there was no health impact to a patient at any time du to the described issu? (No Health Impact) Is the uit meanwhile in ue again and 000000000000000000000000001king as specified? (Bench Repair)Sent to Bench. Resoluion : Sent to Bench Follow u Requred : Problem Description: Bench Repair speaker malfuction Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Bench Repair uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Bench Repair speaker malfuction ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Robert Gash - phone 500000000 200000000 5627 IP & Issu: 56135600000000- Speaker malfuction. Cut wants bench repair. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-11 21:26 | No
417 | Yes | Touchscreen not receptive or responds in incorrect location based on point touched. | Touchscreen not receptive or responds in incorreulocation based on point touched. | 000000000000000000000000001O##: 00000000000000000000000000100000000267800000000Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits touh screen, even after calibration, is not 000000000000000000000000001king correctly. To correct the issu we will replace the Display Assembly. Resoluion : Replaced Display Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023000000008 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits touh screen, even after calibration, is not 000000000000000000000000001king correctly. To correct the issu we will replace the Display Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000998193 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits touh screen, even after calibration, is not 000000000000000000000000001king correctly. To correct the issu we will replace the Display Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099700000000 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to bench Follow u Requred : EMS site manager called to report Tempu Pro: "Touhscreen is not holding calibration throuh boot cycles" "Touhscreen on Tempu Pro is not holding calibration throuh boot cycles. Attempted to recalibrate, bu does not hold. Touhscreen not receptive or responds in incorrect location based on point touhed. Will need loaner while ou for bench repair." Cutomer has tried taking off screen protector, cleaning touhscreen, and recalibrating with like resuts. Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS site Manager Device ue:Not in Clinical ue Eupected and ueupected Resuts:"Touhscreen is not holding calibration throuh boot cycles" uer Impact:No impact Patient Impact:No impact SE00000000.300000000 Problem Description/ueupected Resuts: "Touhscreen is not holding calibration throuh boot cycles" Diagnostic performed by Engineer : Problem Description: "Touhscreen is not holding calibration throuh boot cycles" Cutomer callback Tempu Pro: "Touhscreen is not holding calibration throuh boot cycles" Form ued: Bench Trouleshooting Steps Taken: EMS site manager called to report Tempu Pro: "Touhscreen is not holding calibration throuh boot cycles" "Touhscreen on Tempu Pro is not holding calibration throuh boot cycles. Attempted to recalibrate, bu does not hold. Touhscreen not receptive or responds in incorrect location based on point touhed. Will need loaner while ou for bench repair." Cutomer has tried taking off screen protector, cleaning touhscreen, and recalibrating with like resuts. ##d cutomer bench form, created Bench 000000000000000000000000001 and requsted loaner ----------------------------------------------------------------------------- | 2023-09-11 21:46 | No
418 | Yes | burning smell | burning smell | 000000000000000000000000001O##: 00000000000000000000000000100000000997863 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Parts ID for the DA pcba and DA to MC ribbon cable Follow u Requred : Cutomer call back to speak to RSE for futher assistance Resoluion : cutomer ##d back with a pictue of the damaged DA pcba. He spoke to another RSE, and was provided the part #s for the DA pcba and the cable DA pcba to the MC pcba. cutomer will order the parts and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: jut replaced the blower and now he smell something buning uer Impact: none Patient Impact: none SE 3.100000000Diagnostic performed by Engineer : Cutomer has a V600000000and he jut replaced the blower and now he smells something buning. Instruted him to check the MC and DA pcba for anything that coud be tarnished. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000999261 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: jut replaced the blower and now he smell something buning uer Impact: none Patient Impact: none SE 3.100000000Diagnostic performed by Engineer : provided: Cable, DA PCBA to MC PCBA (2nd generation) 112000000002 00000000561536131. PCBA, data acqusition (DA) 100000000000000006 00000000561512561. Cutomer states he foud evidence of overheating on the DA pcba. Advised to check the MC pcb closely for any bent pins on the bottom. ----------------------------------------------------------------------------- Internal Remarks : **CuT REQuSTED TECH SuPORT Interaction Type: New Case Caller: ahamadou IP & Issu: onging issu - uit has high pitched noise Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND REMOTE 000000000000000000000000001 RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Cutomer Call Back Caller: AhamadouOugraog IP & Issu: "6100000000233 CCB regarding case# 0000000022262296" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: CCB stated his issu still persists and 000000000000000000000000001ld like to talk back with tech suport for trouleshooting; created requst follow u 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: AhamadouOuDRAOGO IP & Issu: " 6100000000233 I gave a statu udate to the cutomer, created another R000000000000000000000000001uing the 2nd Free Remote 000000000000000000000000001k Oder Process, and then roued the cutomer to RSE for futher assistance" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: I gave a statu udate to the cutomer, created another R000000000000000000000000001uing the 2nd Free Remote 000000000000000000000000001k Oder Process, and then roued the cutomer to RSE for futher assistance Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-11 21:51 | No
419 | Yes | uPro: Broken port | TempuPro: Broken uport | 000000000000000000000000001O##: 00000000000000000000000000100000000998267 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005600000000111-Display 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits ufleu nu has come loose and it allowed the Fleu cable to be puled from the uboard when disconnecting a ucable. Once nu is tightened and fleu secued the device fuctions correctly. However, there is physical damage to the display and touh issus occu intermittently. To correct the issu we will replace the Display Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099700000000 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to bench Follow u Requred : EMS site manager called to report Tempu Pro: Broken uport. "uconnection port on the side of the monitor is loose and device indicates to connect a cable despite being pluged in. Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: Broken uport uer Impact:No impact Patient Impact:No impact SE00000000.300000000Problem Description/ueupected Resuts: Broken uport Diagnostic performed by Engineer : Problem Description: Broken uport Cutomer callback Tempu Pro: Broken uport Form ued: Bench Trouleshooting Steps Taken: EMS site manager called to report Tempu Pro: Broken uport. "uconnection port on the side of the monitor is loose and device indicates to connect a cable despite being pluged in. Will need loaner while ou for bench repair." Created Bench 000000000000000000000000001 ##d cutomer bench form, and requsted loaner device. ----------------------------------------------------------------------------- Internal Remarks : 9/12/23 The requsted loaner has been shipped to Cut with FedEu Tracking# 700000000000000000000000058000000009. QuST SO# 20000000026 CASE#: 0000000022261992 PN:989700000000000000000000000082 SN: 616136 Railings: Both Cutomer SN: 615298 Tuorthwest EMS 366 Helen St, McKees Rocks, PA 15136 Attn: Chief Brian O’Connor 586-85000000003616 ----------------------------------------------------------------------------- | 2023-09-11 22:01 | No
420 | Yes | Machine pressure sensor autozero failed | Machine pressure sensor autozero failed (EC 00 CBIT) | 000000000000000000000000001O##: 00000000000000000000000000100000000998000000008 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : reviewed the event log and foud code 1100000000. Reviewed the sugested repair per rev T manul. Cutomer advised the uit has not been 000000000000000000000000001ked on recently. Advised to replace solenoids 2 and 00000000 Reviewed replacement procedues per the manul and provided parts ID for the solenoids. 00000000561511911 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in patient ue Eupected and ueupected Resuts: RT reports the uit gave a check vent alarm for Auo Zero failue. uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: colten IP & Issu: 73222228 parts id Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-11 22:01 | No
421 | Yes | puro ¦ Unit will not power on, the charging port is damaged SPLIT | Tempupro ¦ Unit will not power on, the charging port is damaged SPLIT | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009532 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced Front case Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Parts Consued : 1.0000000000000000000000000000000056000000005000000001,RDT Front Enclosue kit (for 00000000-100000000u-R & Diagnostic performed by Engineer : Confirmed complaint there is physical damage to the front case and charging port to correct the issu we will replace the Front Case Assembly. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000998311 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint there is physical damage to the front case and charging port to correct the issu we will replace the Front Case Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000998258 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Follow u Requred : Problem Description: caller reports the uit will not power on, the changing port is damaged, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repair Confirmed delivery address ##d bench repair form Problem Description by Engineer : Cutomer Fuction/Role: Other: Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: caller reports the uit will not power on, the changing port is damaged, sending to bench Diagnostic performed by Engineer : Problem Description: caller reports the uit will not power on, the changing port is damaged, sending to bench Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repair Confirmed delivery address ##d bench repair form NO LOANER NEEDED ----------------------------------------------------------------------------- Internal Remarks : Interaction New Case Caller: Kasula Cooper IP & Issu: " 966225600000000System is down (the charging port has fallen ou and does not charge anymore) per the cutomer." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and created remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-11 22:06 | No
422 | Yes | Intermittent foot switch issues. | 00CDN0Intermittent foot switch issues. | 000000000000000000000000001O##: 00000000000000000000000000100000000998380000000012-Sep-2000000003 Stewart, Aleu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : u-Ray not possible uLL. When was the first occurence? Within... : 11-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion : Replaced wired foot switch Parts Consued : 1.00000000000000000000000000000000270000000000000000100000000Footswitch CV 3p 00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000998522 Autin Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: fluro Eupected and ueupected Resuts: no pedal sticking uer Impact: none Patient Impact: none SE n/a Diagnostic performed by Engineer : assisted fse aleu stewart with part id of footswitch ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000998300000000 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : u-Ray not possible uLL. When was the first occurence? Within... : 11-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Cutomer states that they have issus when activating u-Rays that they have to step on the pedal several times before u-Rays becomes active. Internal Comments : uLL# and/or description of uLL: Several u-Ray not possible uLL are present in the log files as well as TSO buton sticking uLL. Repair Action : Ru diagnostics onsite of the footswitch as well the generator. Resoluion : Created on-site 000000000000000000000000001llow u Requred : uLL# and/or description of uLL: Several u-Ray not possible uLL are present in the log files as well as TSO buton sticking uLL. Repair Action : Ru diagnostics onsite of the footswitch as well the generator. Technician Name : Johansson Anders ----------------------------------------------------------------------------- | 2023-09-11 22:21 | No
423 | Yes | puro ¦ Capnography Issues | Tempupro ¦ Capnography Issues | 000000000000000000000000001O##: 00000000000000000000000000100000000997835 Al Hammad Lamiae The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : please ship part to attn Biomed Josh 2300000000Dr F E Wright DrJackson, TN 38300000000 Diagnostic performed by Engineer : Problem Description: Biomed requst Parts replacement Biomed requst Parts replacement Form ued: Parts Trouleshooting Steps Taken: cutomer states capnography connection issu Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Biomed requst Parts replacement ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: JOSHu WEATHERFORD IP & Issu: 966225500000000 Capnography Issus Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-11 22:21 | No
424 | Yes | Not Detecting AC Power | Not Deteung AC Power | 000000000000000000000000001O##: 00000000000000000000000000100000000998331 Miller Jason The system meets the specification for the performed service and is retuned to ue Resoluion : Provided cutomer with part nuber and price: 00000000561512300000000 - RP-Inlet,AC,V600000000- 156.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: not detecting ac power uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device is not detecting ac power. Troule shot with cutomer: 1. Verify AC oulet for proper voltage. 2. Verify that power cord is fuctional. 3. Verify power suply (200000000: a. With AC power disconnected, remove J1 from PM PCBA. b. Reconnect AC power. c. Verify that 200000000is present between the orange and brown wires on the power suply harness connector. d. If 200000000is present, replace the PM PCBA. e. If 200000000is not present, go to neut step. 00000000Verify that AC power is present: a. Disconnect AC power. b. Remove EMI shrou. c. Reconnect AC power. d. Verify that AC voltage is present between the blu and brown wires coming from the AC inlet. e. If AC power is present, replace the power suply. f. If AC power is not present, replace the AC inlet. Cutomer states not seeing 200000000and not see 1200000000coming in from ac inlet. informed cutomer that with resuts manuactuers recommendation is to replace ac inlet. cutomer requsting part nuber and price. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Larry Ames IP & Issu: 62836800000000 Not Detecting AC Power Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-11 22:31 | No
425 | Yes | Battery Flat | M00A - u00 - Battery Flat ( See Reference PR : 0000) |  | 2023-09-12 00:11 | No
426 | Yes | on button is not working , not able to turn on | on button is not working , not able to turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000998952 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer that FCO86600000000000000007 was performed in 2000000002, reference 0000000000000000000000000010000000000000000600000000 and repair is not eligible for 900000000day warranty repair, ##d cutomer copy of service report and offered to send a qute for onsite or bench repair if needed to correct issu. Cutomer replied for onsite service to be quted. Follow u Requred : not able to tun off uit via power switch Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:power switch is not 000000000000000000000000001king uer Impact: not uable needs reapir Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states FCO86600000000000000007 uuufront bezel replacement was recently performed and not able to tun off uit via power switch duing set u, power led is illuinated bu on/off fuction is not 000000000000000000000000001king, requsts 900000000day warranty repair for front bezel replacement to correct issu, no patient involvement reported. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Pavel Goleta IP & Issu: 57763262 &on buton is not 000000000000000000000000001king , not able to tun on Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created r000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-12 00:26 | No
427 | Yes | System is stuck at 00 | System is stuck at 00 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000590000000023-Sep-2000000003 Pena, Jessie The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - FLV: Fleuvision PC grabber cards are not initialized. 00000000-Sep-2000000003 18:51:23 RMT - FLV: Fleuvision PC grabber cards are not initialized. 11-Sep-2000000003 11:51:59 When was the first occurence? Within... : 11-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: System is stuk at 0000000000000000 uer Impact: System is down Patient Impact: None/ completed SE 8.2.3000000005 Resoluion : Closed parts Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000999328 100000000ep-2000000003 Pena, Jessie The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - FLV: Fleuvision PC grabber cards are not initialized. 00000000-Sep-2000000003 18:51:23 RMT - FLV: Fleuvision PC grabber cards are not initialized. 11-Sep-2000000003 11:51:59 When was the first occurence? Within... : 11-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: System is stuk at 0000000000000000 uer Impact: System is down Patient Impact: None/ completed SE 8.2.3000000005 Follow u Requred : Replace FleuVision PC, received PC might is DEFOA Diagnostic performed by Engineer : Replaced FleuVision PC, FleuVision PC DEFOA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000000000003-Sep-2000000003 Garcia, Ramon The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace Fleuvision PC, received DOA PC Problem Description by Engineer : Cutomer Fuction/Role: Technician Device ue: Fleuvision PC Eupected and ueupected Resuts: Eupected - No video ueupected - Fleuvision does not display image uer Impact: uer cannot ue Fleuvision display Patient Impact: NA SE 8.2.300000000Diagnostic performed by Engineer : Assisted in replacing Fleuvision PC; Training on the Job ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 13-Sep-2000000003 Pena, Jessie The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Technician Device ue: FleuVision Eupected and ueupected Resuts: video displayed, no video displayed uer Impact: no video displayed Patient Impact: na SE8.2 Diagnostic performed by Engineer : Trouleshot issu, FleuVision pc needs to be replaced, follow-u 000000000000000000000000001created ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000010000000011-Sep-2000000003 Garcia, Ramon The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: Technician Device ue: Fleuvision Eupected and ueupected Resuts: Eupected - video display not 000000000000000000000000001king ueupected - video not display uer Impact: No video displayed Patient Impact: NA SE 8.2 Diagnostic performed by Engineer : Assisted in making backus for Specials Lab, took pictues of configuations. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009992100000000hitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1. Fleuvision PC (TZ) grabber card initialization uLL The grabbercard is malfuctioning. This is a known issu. Root Caue for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1. Fleuvision PC (TZ) grabber card initialization uLL The grabbercard is malfuctioning. This is a known issu. Repair Action : 2nd reboot resolve, foud grabber card uLL from 9-8 and 9-11-23 Service actions for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1.1. Please replace the Fleuvision PC (TZ). Service actions for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1.1. Please replace the Fleuvision PC (TZ). 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD Diagnostic performed by Engineer : Ron reports that the image on the fleu monitor was distorted and when they rebooted the system it stuk at 0000000000000000. Had them reboot and the system came u. Ron saved a LOD and and foud the uLLbelow from 9-8 and today. Problem Description by Engineer : Problem description by engineer : RMT - FLV: Fleuvision PC grabber cards are not initialized. 00000000-Sep-2000000003 18:51:23 RMT - FLV: Fleuvision PC grabber cards are not initialized. 11-Sep-2000000003 11:51:59 When was the first occurence? Within... : 11-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: System is stuk at 0000000000000000 uer Impact: System is down Patient Impact: None/ completed SE 8.2.3000000005 Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1. Fleuvision PC (TZ) grabber card initialization uLL The grabbercard is malfuctioning. This is a known issu. Root Caue for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1. Fleuvision PC (TZ) grabber card initialization uLL The grabbercard is malfuctioning. This is a known issu. Repair Action : 2nd reboot resolve, foud grabber card uLL from 9-8 and 9-11-23 Service actions for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1.1. Please replace the Fleuvision PC (TZ). Service actions for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1.1. Please replace the Fleuvision PC (TZ). 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD Technician Name : Whitaker, Steven ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022263218 FSE: Pena, Jessie System gets stuk on 0000000000000000 on boot u intermittently. Reviewed logs; foud "Fleuvision PC grabber cards are not initialized." Instruted FSE to place part order for 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD and 0000000080000000050000000000000000- Hard disk spare part. ----------------------------------------------------------------------------- | 2023-09-12 01:06 | No
428 | Yes | u is emitting the triple chirp (Removed the AED from service) | HS0AED is emitting the triple chirp (Removed the AED from service) | 000000000000000000000000001O##: 0000000000000000000000000010000000099900000000 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that the AED is triple chirping and will not pass the self-test. uer Impact: NONE Patient Impact: NONE SEE.00000000.000000003 Cutomer reported that the AED is triple chirping and will not pass the self-test. Diagnostic performed by Engineer : Cutomer reported that the AED is triple chirping and will not pass the self-test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 12 years old, the issu is with the device. Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A11L-00000000555 Battery Install Before# 10000000031/2000000007 Lot#00000000795P Pads Eup Date: 10000000022/20000000000000000 Pads Lot # Y-0000000022200000000 ----------------------------------------------------------------------------- | 2023-09-12 01:36 | No
429 | Yes | unit turns on/off by itself | unit turns on/off by itself | 000000000000000000000000001O##: 000000000000000000000000001000000009997500000000200000000ov-2000000003 Gonzalez, Jesse The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts:tuns on/off by itself uer Impact: not uable needs repair Patient Impact: none SEun Resoluion : Reference FCO66 in 0000000000000000000000000010000000000000000753 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009996200000000ccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uit tuns on/off by itself while not in ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts:tuns on/off by itself uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states in storage while not in ue the uit tuns on/off by itself, requsts trouleshooting, advised cutomer the issu coud be within the display (u pcb) or futher trouleshooting may be needed, cutomer requsts onsite service to correct issu. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Ali Tolentino IP & Issu: 593537300000000 Qustion for tech suport. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Ali Tolentino IP & Issu: 593537300000000 since the ugrade for the electrical board has been powering on and off Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Contact Name: n/a Method: n.a What was discused: cutomer called in abou this uit. it looks like fco66 case- 00000000200000000235900000000s already schedued for 10000000031 Neut Step Action: n/a ----------------------------------------------------------------------------- | 2023-09-12 02:26 | No
430 | Yes | Touchscreen issue | Touchscreen issue | 000000000000000000000000001O##: 00000000000000000000000000100000000999611 Mccleary Barry The system meets the specification for the performed service and is retuned to ue Resoluion : Instruted cutomer on performing Touhscreen calibration, cutomer advised calibration is sucessfu and parameter changes are good on screen, no other problems noted, issu resolved. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:touhscreen uresponsive uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing uit set u the touhscreen is uresponsive, no patient involvement and no evidence of impact damage and screen is clean, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ali Tolentino IP & Issu: 600000000586 & setting buton is stuk uable to change the settings Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-12 02:41 | No
431 | Yes | the toothbrush melted together with the cord | 00CAN HY00/0 the toothbru melted together with the cord | 2000000003-11-20000000012:0000000027:***Claims*** Placed order #9-763800000000as a replacement. Closing case.; 2000000003-11-00000000 1000000006:39:***claims*** ##d cu uS label to retun defective uit. Tracking # is 1ZY733000000001953661000000002000000003-10000000027 1000000006:39:***claims*** ##d Cu to follow u and requsted ful home address also asked if willing to retun defective uit. Awaiting response.; 2000000003-10000000017 20000000053:18:Chat ID:0000000000000000000000008 R-Shoppers Dru mart I-Follow u T-done S-keep the case in progress/ forward case to suervisor for ugent follow u Cu verbatim. Their original ##was sent to me September 12. its now October 17, im not sue if more than a month is an acceptable timeline for this;2000000003-00000000-12 18:50000000058:****claims*** ##d cu with safety qustions. Awaiting response.; 2000000003-00000000-12 00000000:27:21:r-Shoppers Dru mart I-Follow u- Toothbruh melted while charging T-checked attached photo in parent case May I know when and where did youpuchase it? Aida Mae I • 5:00000000:19 PM I puchased it from a shoppers dru mart in Canada Visitor • 5:00000000:00000000PM It mut've been over a year ago Visitor • 5:00000000:500000000PM Okay. Aida Mae I • 5:00000000:31 PM May I know when did this happen? Aida Mae I • 5:00000000:00000000PM Jut Last Night :( Visitor • 5:00000000:56 PM May I know if someone got injued? Aida Mae I • 5:00000000:59 PM Lukily not. i placed it on a towel while charging and there is a small bun mark on the towel. and the whole room smelled like plastic. bu no injuies Visitor • 5:10000000000000000PM Okay. s- informed cu that case will be escalated and wait for their ## | 2023-09-12 04:31 | No
432 | Yes | C-arm cannot move - Split | C-arm cannot move - Split | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 13.00000000.2000000003 00000000:10000000000000000uC (8600000000600000000) ) Service content (requred): C-arm cannot move; IPC cannot start, memory uLLreports; plu and uplu memory modue recovery. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None Internal Remarks: 13.00000000.2000000003 00000000:1000000003 uC (8600000000600000000) ''RDI Accuacy:YES'' $Q7: Confirmed the reported problem, Geometry movements party avaliable. $Q8: The IPC commuication uLLis foud in logs, and the IPC cannot be started properly, Check the IPC to find a memory uLL $Q9: Plu and uplu IPC memory, the IPC start normal. $Q100000000 Test and Inspection data refer to:PA.Device retuned to ful fuctionality. $Q11: Calibrated Tools：NA Neut Step Action: PPD 00000000800000000931363 *1, Spare parts are ou of stock, ordering has been arranged, please dispatch FSE WuFuong -86-181700000000290000000000000000SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022257798 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Bed cannot be locked SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the bed was inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Critical faut code or phenomenon: bed inoperable Remote diagnosis and test resuts: Restart has no effect PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Try to re-insert and uplu the internal board 2. Try to reinstall the IPC software 3. Check the machine uLLand perform faut repairs Recommended spare part uuand name: " 00000000800000000931363 GEO IPC Ivybridge with ZMP CAN and IO"&00000000230000000025812, SMART DRIVE Cutomer demand notes: Contact cutomers promptly to confirm Onsite time ----------------------------------------------------------------------------- | 2023-09-12 06:16 | No
433 | Yes | System can’t boot complete | 0L/u FD0geometry partly available | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022257800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Machine failue, we have commuicated with FSE and requsted to dispatch 000000000000000000000000001kers. T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the mechanical part was inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Critical faut code or phenomenon: Mechanical part cannot operate Remote diagnosis and test resuts: Restart has no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the Geo Modue is normal. 2. Check whether the IPC is normal. 3. Check the machine uLLand perform faut repairs. Recommended spare parts uuand name: " 00000000800000000151715 GEO MODuE WP Kit"&00000000800000000931363 GEO IPC Ivybridge with ZMP CAN and IO Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Euternal Remarks: 11.00000000.2000000003 00000000:15:32 uC (860000000095300000000 Service content: After the cutomer reports for repair, the system prompts "Geometry movements partly available" after powering on. The restart is invalid. The POST self-check on site foud GeoIpc fail. Check the Log and there are "uable to commuicate with GeoIpc" and "Frontal- Configued Stand not foud!!" Possible caue: Configuation mismatch or defect stand" message. Check the relevant fues, power suply lines, and commuication lines according to the manul. There are no abnormalities. Check that the CAN lights of SPC1 & SPC3 are on red. Reinstall the GeoIpc system and APP and eliminate the faut. Service resuts: The equpment is normal and fuly fuctional. Internal Remarks: 13.00000000.2000000003 13:0000000012 uC (860000000095300000000 $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: (Confirmed reported problem. After the system starts, the system prompts "Geometry movements partly available", and the system fails to restart.) $Q8: (The GeoIpc faile was foud in the POST self-check on site, and "uable to commuicate with GeoIpc" and "Frontal- Configued Stand not foud!!" were foud in the Log. Possible caue: Configuation mismatch or defect stand. Check the fues, power suply lines, and commuication lines according to the manul. Check that the CAN on SPC1 & SPC3 is red, Need to reinstall GeoIpc system and APP.) $Q9: (Reinstall GeoIpc system and APP. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: (N/A) ^Curent Tue installed date: (N/A) ^Curent Tue uage: (N/A) ----------------------------------------------------------------------------- | 2023-09-12 06:21 | No
434 | Yes | System can’t boot complete | M00A / u000 / HeartStart uL Defibrillator/Monitor / Unable to startup. | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Beijing Miyu Couty Maternal and Child Health Hospital Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Problem Reported by Cutomer: uable to boot Safety Qustions: Case: 0000000022257959 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Contacted Governor Ma Dezhi 13900000000317693 Give ASP hukeqiangyi buiness follow-u Neut Step Action: Zhao Yijie 155223393500000000has followed ----------------------------------------------------------------------------- | 2023-09-12 07:11 | No
435 | Yes | alliable to start | 0/V0uLLUnable to start | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/V600000000uLLcannot boot Safety Qustions: Case: 0000000022258000000005 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-12 07:36 | No
436 | Yes | Defibrillator/Monitor/Black screen at startup | 0N/Effi0 DFM0 Defibrillator/Monitor/Black screen at startup | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 1000000009.2000000003 00000000:0000000000000000uC (860000000000000000100000000 $ASP engineer:[Honglejia LiuWeiming] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem the Screen boot black screen. $Q8: Check the screen and find that the screen is Intermittent black screen. Need to replace processing plate assembly. $Q9: Replace the processing plate assembly. $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality $Q11: $ Calibrated Tools name (ProSim8) S/N (26712200000000 Neut Calibration date (10000000018/2000000003), (ESA6200000000 S/N (30000000028600000000 Neut calibration date (3/2/20000000000000000 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 1000000009.2000000003 00000000:0000000000000000uC (860000000000000000100000000Department information: [Department: Equpment Department Contact: Teacher Bao Tel: 18267200000000326]* $ASP engineer:[Honglejia LiuWeiming] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem the Screen boot black screen. $Q8: Check the screen and find that the screen is Intermittent black screen. Need to replace processing plate assembly. $Q9: Replace the processing plate assembly. $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality $Q11: $ Calibrated Tools name (ProSim8) S/N (26712200000000 Neut Calibration date (10000000018/2000000003), (ESA6200000000 S/N (30000000028600000000 Neut calibration date (3/2/20000000000000000 Safety Qustions: Case: 00000000222580000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Black screen on boot The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Check settings Recommended spare part uufor T2 with description: 0000000056000000009000000001, DFM100000000 ASSY Processor Notes on cutomer requrements: na SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000N/uDFM100000000 Defibrillator/Monitor/Black screen at startu ----------------------------------------------------------------------------- | 2023-09-12 07:41 | No
437 | Yes | Free space low | SA:0R/u FD0Free space low | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 17.00000000.2000000003 17:35:22 uC (86000000008275) "RDI Accuacy:YES" Q1: (Y) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q5: (N) $Q6:（Free disk space is too low) $Q7:(CONFIRMED REPORTED PROBLEM.THE FREE DISK SPACE IS TOO LOW, THE uLLINFORMATION"THE FREE DISK SPACE IS TOO LOW"DISPLAY ON DATA MONITOR.) $Q8: (DIAGNOSED THE SYSTEM WITH FSF AND DRAWINGS THEN FOuD THE OS/APP OF IPPC DAMAGED.) $Q9: (REINSTALLED THE IPPC OS/APP WITH MANuL THEN THE IPPC CAN 000000000000000000000000001K NORMALLY AND THE FREE DISK SPACE IS NORMALLY.) $Q100000000 (Test and Inspection data:REFER TO PA & THE DEVICE RETuNED TO FuL FuCTIONALITTY.) $Q11: Calibrated Tools name (PA) S/N(PA) Neut Calibration date (PA) ForMainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 9/19/2000000001 ^Curent Tue uage:287100000000LuSAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 17.00000000.2000000003 17:35:19 uC (86000000008275) PROBLEM DESCRIPTION:THE FREE DISK SPACE IS TOO LOW, THE uLLINFORMATION"THE FREE DISK SPACE IS TOO LOW"DISPLAY ON DATA MONITOR. TROuLESHOOTING ACTIONS:DIAGNOSED THE SYSTEM WITH FSF AND DRAWINGS THEN FOuD THE OS/APP OF IPPC DAMAGED.REINSTALLED THE IPPC OS/APP WITH MANuL THEN THE IPPC CAN 000000000000000000000000001K NORMALLY AND THE FREE DISK SPACE IS NORMALLY. SERVICE RESuT:AFTER SERVICE, THE DEVICE RETuNED TO FuL FuCTIONALITY. Safety Qustions: Case: 0000000022258155 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: uable to captue image SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that image storage space is low and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Free disk space is too low. Remote diagnostics and test resuts: NA PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Euecue Database Consistency Test and Repair, delete Repaired patients, and see if the space is restored. 2. If there is no effect, euecue Recreate Image Store. 3. If there is no effect, try Bypass hard disk, and then reinstall the IPPC software. 00000000If there is no effect, ue PC. Diagnostic Tool performs diagnostic tests and replaces corresponding spare parts 5. Check machine uLLreports and perform faut repairs Recommended spare part uuand name: " 0000000080000000050000000092 IP PC Revised_II no HDD &0000000080000000050000000000000000Hard disk spare part " Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-12 08:06 | No
438 | Yes | Op check for therapy failed | 00- u000- Op check for therapy failed | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Treatment implementation test failed Safety Qustions: Case: 0000000022258200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-12 08:21 | No
439 | Yes | startup failure | 0/u FD0startup failure | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: FSE Chen Haiju 13982000000000000000073 I didn’t answer the phone at first and then hug u. I added an ##to ask u Chen first. Neut Step Action: FSE Chen Haiju 13982000000000000000073 replied to the ## Sales has been contacted.Thanks! T2 Activities:Problem Description: Overview of cutomer problems: Startu failue Detailed description of cutomer problem or faut phenomenon: Can't boot Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The data screen displays 00000000:00000000, the power switch is invalid, the PC net000000000000000000000000001k port indicator light is on normally (indicating that the switch is normal), and the SIBz self-test is normal. Remote diagnostics and test resuts: There may be a problem with the system software PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Go to the site to check the startu phenomenon and restore the image if necessary Check if there are any other issus Notes on cutomer requrements: Please contact the cutomer for on-site repair Problem Reported by Cutomer: uable to boot Safety Qustions: Case: 000000002225800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: The machine software cannot be started to the uer interface.Please contact u to confirm dispatch Neut Step Action: ASP PuWei Tsering Gyatso 18000000009916500000000Call to inqure abou when the engineer will arrive if the qutation is completed today. The cutomer needs a very specific date. He said he had contacted FSE Chen Haiju 13982000000000000000073, bu u Chen said that no matter how buy he is, he can’t give a particuarly clear date. Neut Step Action: 000000003116000000003739 Call to remind the order Neut Step Action: Teuted Tsering Gyatso the EQ nuber of the device ----------------------------------------------------------------------------- | 2023-09-12 09:21 | No
440 | Yes | System can’t boot complete | System can’t boot complete | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001000000006 15-Sep-2000000003 WEERAPu AREYAKuTI The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technician. Device ue: Not in clinical ue. Eupected and ueupected Resuts: System does not boot stuk at 00000000:00000000 uer Impact: No impact. Patient Impact: No impact. SE 8.2.3000000006 Resoluion : Replace part. Download board software all. Pass Test ue all fuction. Pass. Device tested and retuned to ful fuctionality. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000099760000000011-Sep-2000000003 WEERAPu AREYAKuTI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Technician Device ue: CV Eupected and ueupected Resuts: system can't boot complete uer Impact: No impact Patient Impact: No impact SEN/A Resoluion : Check all logs. Wait for order part fleupc for replace. 000000000000000000000000001O#follow u to 00000000000000000000000000100000000000000001000000006. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000072 CHATCHAI SANTANA The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Waiting for replace FleuVision PC. Follow u Requred : Replace FleuVision PC. Problem Description by Engineer : Cutomer Fuction/Role: Technician. Device ue: Not in clinical ue. Eupected and ueupected Resuts: System does not boot stuk at 00000000:00000000 uer Impact: No impact. Patient Impact: No impact. SE 8.2.3000000006 Diagnostic performed by Engineer : FleuVision PC fail to initiate grabber card. ----------------------------------------------------------------------------- | 2023-09-12 09:41 | No
441 | Yes | System stuck login | System stuck login | 000000000000000000000000001O##: 00000000000000000000000000100000000993170000000000000000ep-2000000003 Syuron Ma'mu Haris The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:uer Device ue: startu Eupected and ueupected Resuts: Confirmed reported problem replacement IPPC uer Impact: na Patient Impact: na SE 1.2 Diagnostic performed by Engineer : this case only for consued part, please refer to case 000000002220000000081 for detail Parts Consued : 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD ----------------------------------------------------------------------------- | 2023-09-12 10:11 | No
442 | Yes | screen is damaged SPLIT | TempuPro ¦ screen is damaged SPLIT | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005331 Bronte Westaway The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed tech Device ue: monitoring Eupected and ueupected Resuts: damaged screen uer Impact: cant ue Patient Impact: none SE 000000008 Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000009-13 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000003-15 00000000:00000000:00000000 Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit 1.000000000000000000000000000000005600000000131,RDT Rainbow Front Case w/all Connectors Resoluion : Front enclosue replaced du to damage on overlay. Display replaced du to physical damage to touhscreen and screen protector. SW was ugraded to 0000000000000000 Performance verification, calibration and safety testing completed - all passed within specification. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000993366 Bronte Westaway The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Confirmed physical damages primarily to touhscreen/display assembly. Damage to buton overlay - front enclosue shoud be replaced to prevent fuue buton failues. uu 000000005600000000111 & 000000005600000000131 Problem Description by Engineer : Cutomer Fuction/Role: biomed tech Device ue: monitoring Eupected and ueupected Resuts: damaged screen uer Impact: cant ue Patient Impact: none SE 000000008 Follow u Requred : repair, SW ugrade and ful T&V ----------------------------------------------------------------------------- | 2023-09-12 10:46 | No
443 | Yes | The loudspeaker is out of work. | The loudspeaker is out of work. | 000000000000000000000000001O##: 00000000000000000000000000100000000995789 Assad MAJEED The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: ##for end of marketing and contact details of avf sales engineer sent to the cutomer. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that the louspeaker of his FR3 defibrillator is defective and wants to replace it. Since the equpment is at the end of marketing on 1/1000000002000000001, no part is available in stock anymore. ----------------------------------------------------------------------------- | 2023-09-12 12:11 | No
444 | Yes | collimator failure | 0R/u FD0collimator failure | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Teacher Li 18000000007877231 called again to report that no one contacted him The phone of Gong Liuof the China Telecom uion was tuned off. The person Dai Gong who wrote on the first screen of the case of the China Telecom uion told him that he had not received the order and that the relevant personnel had been notified to verify the dispatch of the 000000000000000000000000001k. Euternal Remarks: 12.00000000.2000000003 12:00000000:19 uC (86000000000000000091) Service content (requred): The reduer self-test failed.Replace the reduer. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 12.00000000.2000000003 12:00000000:22 uC (86000000000000000091) RDI Accuacy: YES $Q7: comfirmed the reported problem, the collimator can't open. $Q8: check the logfile find the collimator post failed. $Q9: replace the NICOL V3 CV FD uu 00000000800000000858162 $Q100000000 (Test and Inspection data refer to:PA.Device retuned to ful fuctionality.) $Q11: Calibrated Tools refer to:PA SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: 00000000R/u uuuFD200000000 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: Collimator failue Trouleshooting Action: Key faut codes or phenomena: Spectral filter not available Remote diagnostics and test resuts: The previouly replaced beam beamer is still fauty PRS TR statu: All Access Repair Action: Recommended on-site repair steps: " 1. Check whether the page is normal 2. Replace the beamer 3. Sumit CFM promptly" Recommended spare parts uuand designation: "ppd/ 00000000800000000858162 NICOL V3 CV FD/200000000to hospital/site1" Safety Qustions: Case: 0000000022259000000003 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: PPD00000000800000000858162*1, has been shipped, please send 000000000000000000000000001k FSE: LiuChao-86-13979783177 ----------------------------------------------------------------------------- | 2023-09-12 12:12 | No
445 | Yes | Wrong wheel label - | Wrong wheel label - | 000000000000000000000000001O##: 00000000000000000000000000100000000993891 Parts Consued : 1.000000000000000000000000000000005600000000100000000,ITP LBL Film Energy English ASSY ----------------------------------------------------------------------------- | 2023-09-12 12:21 | No
446 | Yes | Cannot enter standby mode | Cannot enter standby mode | BuQ&R Feedback: No Reason: Cutomer Feedback Neut Step Action: An euception feedback ##has been sent to TSS Hung Gong, SDL, and other relevant personnel. Neut Step Action: Abou case#000000002223300000000 000000002223300000000 000000002223300000000 Dealer Ms. Leng 15770000000038792 called again for feedback: There are curently 00000000600000000uits in the hospital. Repeated carbon diouide inhalation has occured in a similar period of time, and the equpment cannot enter the standby mode. The hospital supects that there is a qulity problem with the equpment. u needs to give a reasonable euplanation or come to the site for maintenance as soon as possible. If a soluion cannot be provided this week, Hospitals report adverse events. Safety Qustions: Case: 0000000022259119 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Pan Chuhu Last udated by: Pan Chuhu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Cutomer feedback abou case#000000002223300000000&000000002223300000000&000000002223300000000 ----------------------------------------------------------------------------- | 2023-09-12 12:31 | No
447 | Yes | not audibly getting alerts for important alerts | not audibly getting alerts for important alerts | 000000000000000000000000001O##: 0000000000000000000000000010000000099000000006 Rioseco Dawn The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Nuse Device ue: Clinical Eupected and ueupected Resuts: Can not hear alarms uer Impact: Same Patient Impact: n/a SE Diagnostic performed by Engineer : Prob. Description: No auible alarm Form ued: Trouleshooting/Investigation Object: ICusuv Not Object: ICuovr uer Actions: Normal operation Where is the Deviation: No auible soud Trouleshooting Steps Taken: Rebooted suveillance resolved issu Was PRS ued: Yes PRS Permission: Cutomer PRS Path: https://pww.portal.rs.healthcare.u.com/PrsPortal/Defaut.aspu?PrsId=266399 Resoluion : Known: Yes Resoluion Sumary: Reboot suveillance Prevention: N/A We hope you suport euuuence met you eupectations. We will be sending a SuVEY, please complete it to let u know how we are doing. Shoud youneed management assistance contact Ferda Kavak ph:561-800000000-5276. ----------------------------------------------------------------------------- | 2023-09-12 13:06 | No
448 | Yes | Defibrillator/Monitor/handle clip lost | 0N/Effi0 DFM0 Defibrillator/Monitor/handle clip lost | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 15.00000000.2000000003 00000000:0000000039 uC (860000000000000000100000000Department information: [Department: Equpment Department Contact: Teacher Lu Tel: 15277393315]* $ASP engineer: [honglejia uuchungfeng] $Q1:NO $Q2:NO $Q3:N/A $Q00000000/A $Q7:Confirmed reported problem: The bukle of the defibrillation electrode handle is broken and detached. $Q8:Detected that the bukle of the defibrillation electrode handle is broken and detached. $Q9:Replace the euternal chest defibrillation electrode board again and restore the instruent to normal ue. $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11: Calibrated Tools: #ESA 6200000000#S/N # 3000000002865 # Neut Calibration date # 00000000/00000000/20000000000000000 Calibrated Tools: #Prosim8 #S/N #2600000000000000002 #Neut Calibration date #100000000200000000000000003 #. Calibrated Tools: #70000000000000000P # S/N # 3753000000006 # Neut Calibration date # 10000000029/2000000003 #. Problem Reported by Cutomer: The handle is detached SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 15.00000000.2000000003 00000000:0000000000000000uC (860000000000000000100000000 $ASP engineer: [honglejia uuchungfeng] $Q1:NO $Q2:NO $Q3:N/A $Q00000000/A $Q7:Confirmed reported problem: The bukle of the defibrillation electrode handle is broken and detached. $Q8:Detected that the bukle of the defibrillation electrode handle is broken and detached. $Q9:Replace the euternal chest defibrillation electrode board again and restore the instruent to normal ue. $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11: Calibrated Tools: #ESA 6200000000#S/N 3000000002865 #Neut Calibration date# 00000000/00000000/20000000000000000 Calibrated Tools: #Prosim8 #S/N #2600000000000000002 #Neut Calibration date #100000000200000000000000003 #. Calibrated Tools: #70000000000000000P # S/N # 3753000000006 # Neut Calibration date # 10000000029/2000000003 #. T2 Activities:Problem Description: Cutomer’s euplanation of the problem: The handle bayonet fell off Trouleshooting Action: T2 Remote Diagnostics and Testing: The handle bayonet is damaged Repair Action: T2 Recommended on-site repair steps: Replace the handle bayonet if necessary Safety Qustions: Case: 000000002225900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhouuiaouun Last udated by: Zhouuiaouun 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-12 13:16 | No
449 | Yes | Replace the wired footswitch (SPLIT) | Replace the wired footswitch (SPLIT) |  | 2023-09-12 13:25 | No
450 | Yes | Running on internal Battery when plugged into wall. and died and won't turn on. | Running on internal Battery when plugged into wall. and died and won't turn on. | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : Does not charge Ou Of Bou Failue? : N/A Labor: Labor ID : L-197577800000000000000000000000000000000001k Performed : Power management board ordered Diagonsis : ordering PM PCBA | 2023-09-12 14:01 | No
451 | Yes | flattery failure | 0/V0uLLbattery failure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: HuZhong Hospital, HuduDistrict, Gungzhou Battery problem, additional training teacher needs to be present Invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 0000000022259738 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: jindi yang Last udated by: jindi yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-12 14:01 | No
452 | Yes | The device does not turn on | The device does not turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002285 Kaan Ozkan The system does not meet the specification for the performed service and is not retuned to ue. Follow-u Requred: The part shoud be quted. Repair on the atolye. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: The device is not opening. uer Impact: no Patient Impact: no SE The device is not opening, it coud not be detected. Diagnostic performed by Engineer: The inspection performed determined that the device was not opening and therapy card and process card were defective. 0000000056000000009000000001 DFM100000000 ASSY Therapy 0000000056000000009000000001 DFM100000000 ASSY Processor 000000005600000000311 DFM100000000-CBL u to Processor Miu Shoud be quted. Tools & Equpment : MuTIMETER, 200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-12 14:31 | No
453 | Yes | System stopped working mid case - patient on the table | System stopped working mid case - patient on the table | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000669000000003-Sep-2000000003 Brendan Finn The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System stopped 000000000000000000000000001king mid case - patient on the table uer Impact: NA Patient Impact: NA SE7.2 Resoluion : Ran chkdsk on backu drive . Created new ghost backu image . Fuctionally tested and handed system back to cutomer ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 12-Sep-2000000003 Brendan Finn The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System stopped 000000000000000000000000001king mid case - patient on the table uer Impact: NA Patient Impact: NA SE7.2 Resoluion : Reported issu was system stopped 000000000000000000000000001king and susequntly 000000000000000000000000001ld not boot u. Investigated problem and discovered that the backu drive was not accessible cauing the system to freeze on boot u. There was a pre-ued bu 000000000000000000000000001king spare HDD on site that I re-initialised and fitted allowing the system to boot u and behave normally. Fuctionally tested system and handed back to cutomer in 000000000000000000000000001king order. Follow u Requred : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002378 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Supect SBC , FSE onsite to futher investigate Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System stopped 000000000000000000000000001king mid case - patient on the table uer Impact: NA Patient Impact: NA SE7.2 Diagnostic performed by Engineer : Following shudown , system does not even get to coutdown screen. Control room screens go black instead of loading application ----------------------------------------------------------------------------- | 2023-09-12 14:36 | No
454 | Yes | ultrahuman pressure pipe disconnected. | SA：0/V0uLLProuimal pressure pipe disconneud. | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 00000000222600000000500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks : The device prompt says near-end pressue pipe is disconnected, same issu for fou devices, and there are issus with t000000000000000000000000001other devices for which 000 (SN) are curently uavailable. The requst is for the buiness team to follow u on this matter. Problem Reported by Cutomer : SA：00000000/V600000000uLL----------------------------------------------------------------------------- | 2023-09-12 14:51 | No
455 | Yes | Prouimal pressure pipe disconnected | Prouimal pressure pipe disconneud | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: SA：00000000/V600000000uLLSafety Qustions: Case: 0000000022260000000053 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: The device prompts that the prouimal pressue pipe is disconnected. A total of 00000000evices have the same problem. The other t000000000000000000000000001devices are temporarily uable to provide SN. Please follow u with the buiness. ----------------------------------------------------------------------------- | 2023-09-12 14:51 | No
456 | Yes | SELF TEST FAILED - | SELF TEST FAILED - | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : CM - SELF TEST FAILED Safety Qustions : Case: 0000000022260000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-12 15:02 | No
457 | Yes | It does not start. | CM - SELF TEST FAILED+MI | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 0000000022260000000077 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : CM - SELF TEST FAILED+MI ----------------------------------------------------------------------------- | 2023-09-12 15:11 | No
458 | Yes | startup failure | 0R/u FD0startup failture | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 12.00000000.2000000003 1000000001:00000000uC (86000000006252) "SuT. Ready for Clinical ue local time" : " 2000000003-00000000-11 200000000300000000 Problem Reported by Cutomer: 00000000R/u uuuFD200000000cannot tun on SAP Transfer: Case Transferred to SAP Internal Remarks: 13.00000000.2000000003 00000000:56:15 uC (86000000006252) RDI Accuacy: "yes" $Q7: (Confirm the reported problem & startu failue) $Q8: (Open the M-cabinet and check the spc, it can't boot u.) $Q9: (Reinsert the spc board. ) $Q100000000 (Test and Inspection data refer to PA & Device retuned to ful fuctionality) $Q11: (Calibration tools: N/A) Euternal Remarks: 13.00000000.2000000003 00000000:56:12 uC (86000000006252) Service content (requred): [Clean the board and re-insert it to resue normal startu.] Service resuts (requred): [The equpment can be ued normally after the repair is completed.] SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022260000000063 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: Startu failue Detailed description of cutomer problem or faut phenomenon: Can't boot Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: n/a Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Plu and uplu SBC 00000000230000000028782 Service part Single Board Compuer Notes on cutomer requrements: Please contact the cutomer for on-site repair ----------------------------------------------------------------------------- | 2023-09-12 15:41 | No
459 | Yes | It does not start. | It does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002880000000012-Sep-2000000003 José Lus Gómez Soler The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: It seems it is necessary to activate the electrical board maneuer in the room, bu the start buton cannot be located. J. L. Gómez is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment does not start. Patient Impact: No futher information is available as per the cutomer. SE 7.2.9.00000000Resoluion: Check the M cabinet power suply board, re-set the auomaticS of the PDuM cabinet, operation tests. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001652 Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: It seems it is necessary to activate the electrical board maneuer in the room, bu the start buton cannot be located. J. L. Gómez is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment does not start. Patient Impact: No futher information is available as per the cutomer. SE 7.2.9.00000000Diagnostic performed by Engineer: It seems it is necessary to activate the maneuer of the electrical board of the room, bu the start buton cannot be located. J. L. Gómez is in charge of checking it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: It seems it is necessary to activate the electrical board maneuer in the room, bu the start buton cannot be located. J. L. Gómez is in charge of checking it. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: It seems it is necessary to activate the electrical board maneuer in the room, bu the start buton cannot be located. J. L. Gómez is in charge of checking it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- Internal Remarks: Direct contact +300000000300000000621 ----------------------------------------------------------------------------- | 2023-09-12 15:41 | No
460 | Yes | Having same issue delay during lateral and frontal exposure. | Having same issue delay during lateral and frontal euposure. | 000000000000000000000000001O##: 00000000000000000000000000100000000166900000000 18-Oct-2000000003 Muammad Jehanzeb The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000858161,NICOL V3 CV FD 1.000000000000000000000000989600000000000000002000000002,NICOL NON AEPM Resoluion : 1- collimator assembly replaced with new one. 2-Software programming performed sucessfuly. 3-Collimator configued for Lateral plane sucessfuly. 00000000ested system with ful fuctionality. 5-Left the system in ful 000000000000000000000000001king condition Problem Description by Engineer : Problem description by engineer : having same issu delay duing lateral and frontal euposue. RMT - POS: LARC POST - power relays test failed 00000000-Sep-2000000003 00000000:0000000032 RMT - AEP: AEP meter is defective - Lateral 00000000-Sep-2000000003 00000000:0000000013 RMT - AEP: WARNING: u-ray dose not logged - Lateral 15-Au-2000000003 15:00000000:00000000 RMT - AEP: WARNING: u-ray dose not logged - Frontal 15-Au-2000000003 15:00000000:35 RMT - AEP: Cuuative Area Dose not received from AEP meter within interval. 00000000-Au-2000000003 11:35:53 How often is issu occuring? : Intermittent Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Tools & Equpment : MuTIMETER,200000000000000003-26 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000 00000000-Oct-2000000003 Muammad Jehanzeb The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : 1-Collimator replacement Problem Description by Engineer : Problem description by engineer : having same issu delay duing lateral and frontal euposue. RMT - POS: LARC POST - power relays test failed 00000000-Sep-2000000003 00000000:0000000032 RMT - AEP: AEP meter is defective - Lateral 00000000-Sep-2000000003 00000000:0000000013 RMT - AEP: WARNING: u-ray dose not logged - Lateral 15-Au-2000000003 15:00000000:00000000 RMT - AEP: WARNING: u-ray dose not logged - Frontal 15-Au-2000000003 15:00000000:35 RMT - AEP: Cuuative Area Dose not received from AEP meter within interval. 00000000-Au-2000000003 11:35:53 How often is issu occuring? : Intermittent Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : 1-Collimator software download is failed . 2-Needs to order new part. 3-collimator shuters are not 000000000000000000000000001king . 00000000nce new collimator replaced will reload application software. Tools & Equpment : MuTIMETER,200000000000000003-26 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002900000000 27-Sep-2000000003 Muammad Jehanzeb The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : 1-Application software for Lateral collimator needs to be downloaded. Problem Description by Engineer : Problem description by engineer : having same issu delay duing lateral and frontal euposue. RMT - POS: LARC POST - power relays test failed 00000000-Sep-2000000003 00000000:0000000032 RMT - AEP: AEP meter is defective - Lateral 00000000-Sep-2000000003 00000000:0000000013 RMT - AEP: WARNING: u-ray dose not logged - Lateral 15-Au-2000000003 15:00000000:00000000 RMT - AEP: WARNING: u-ray dose not logged - Frontal 15-Au-2000000003 15:00000000:35 RMT - AEP: Cuuative Area Dose not received from AEP meter within interval. 00000000-Au-2000000003 11:35:53 How often is issu occuring? : Intermittent Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Tools & Equpment : MuTIMETER,200000000000000003-26 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000800000000858161,NICOL V3 CV FD Diagnostic performed by Engineer : 1-Lateral collimator was replaced with new one . 2-Configuation performed for Lateral collimator sucessfuly. 3-Boot loader installed sucessfuly 00000000icole application software needs to be downloaded and it Failed. 5-Problem Escalated to Tier 2 Level. 6- Follow u visit created. 7-Left the system in partial 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001372 Ahmed Kotb The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : having same issu delay duing lateral and frontal euposue. RMT - POS: LARC POST - power relays test failed 00000000-Sep-2000000003 00000000:0000000032 RMT - AEP: AEP meter is defective - Lateral 00000000-Sep-2000000003 00000000:0000000013 RMT - AEP: WARNING: u-ray dose not logged - Lateral 15-Au-2000000003 15:00000000:00000000 RMT - AEP: WARNING: u-ray dose not logged - Frontal 15-Au-2000000003 15:00000000:35 RMT - AEP: Cuuative Area Dose not received from AEP meter within interval. 00000000-Au-2000000003 11:35:53 How often is issu occuring? : Intermittent Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : having same issu delay duing lateral and frontal euposue. after checking the uLLlog, I can see an issu with the lateral collimator which is supected to be the root caue of this issu. new collimator will be ordered. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - POS: LARC POST - power relays test failed: 1. Duing POST LARC Power relay (SLB-K1) test failed Root Caue for RMT - AEP: AEP meter is defective - Lateral: 1. Lateral AEP meter is defective. Root Caue for RMT - AEP: WARNING: u-ray dose not logged - Lateral: 1. uer warning: WARNING: u-ray dose not logged. The dose information from AEP uit is not being logged. Root Caue for RMT - AEP: WARNING: u-ray dose not logged - Frontal: 1. uer warning: WARNING: u-ray dose not logged. The dose information from AEP uit is not being logged. Root Caue for RMT - AEP: Cuuative Area Dose not received from AEP meter within interval.: 1. CuuativeAreaDose not received from AEP meter within interval. Repair Action : Service actions for RMT - POS: LARC POST - power relays test failed: 1.1. Check the Larc Anguation limit switch: SLA S1. Check whether this switch is (uintentional) activated so the switch is open instead of in 'normally closed' state. 1.2. Check the Larc Rotation limit switches: SLB S1 & S2. Check whether (one of) these switches is (uintentional) activated so (one of) the switches are open instead of in 'normally closed' state. 1.3. There is a known issu with the Larc powersuply (SLB) with 12nc 000000002129171500000000r lower. The relais K1 in the Larc powersuply is too slow to switch the 5V which is ued for the power relais test. A new type of powersuply has been made and can be ordered with the following 12nc: power suply assy Larc 000000002 300000000 23776 1.00000000It coud be that there is an Inhibit signal active du to a defect of another part, suh as: MVR, Lu, GIB,etc. If onsite: please enable eutended logging of the Lu boards (SCC3/SCC00000000and verify which test fails euactly. Contact Tier 3 Helpdesk for assistance. Service actions for RMT - AEP: AEP meter is defective - Lateral: 1.1. Is the problem permanent and repeating daily? If yes, replace AEP meter or collimator. AEP meter is the most supected one ou of the t000000000000000000000000001 1.2. Check the cable and its connector of AEP meter. Service actions for RMT - AEP: WARNING: u-ray dose not logged - Lateral: 1.1. Check AEP/DAP meter itself. 1.2. Check cabling and connector of AEP/DAP. 1.3. Replace AEP/DAP if problem persists. Service actions for RMT - AEP: WARNING: u-ray dose not logged - Frontal: 1.1. Check AEP/DAP meter itself. 1.2. Check cabling and connector of AEP/DAP. 1.3. Replace AEP/DAP if problem persists. Service actions for RMT - AEP: Cuuative Area Dose not received from AEP meter within interval.: 1.1. Check the cable and its connector of AEP meter. Replace when problem persists. Replace the lateral collimator as per manul (youneed to configue the collimator into lateral channel (via Collimator service tool CS tool) Resoluion : new collimator will be ordered. Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - POS: LARC POST - power relays test failed: 1. Duing POST LARC Power relay (SLB-K1) test failed Root Caue for RMT - AEP: AEP meter is defective - Lateral: 1. Lateral AEP meter is defective. Root Caue for RMT - AEP: WARNING: u-ray dose not logged - Lateral: 1. uer warning: WARNING: u-ray dose not logged. The dose information from AEP uit is not being logged. Root Caue for RMT - AEP: WARNING: u-ray dose not logged - Frontal: 1. uer warning: WARNING: u-ray dose not logged. The dose information from AEP uit is not being logged. Root Caue for RMT - AEP: Cuuative Area Dose not received from AEP meter within interval.: 1. CuuativeAreaDose not received from AEP meter within interval. Repair Action : Service actions for RMT - POS: LARC POST - power relays test failed: 1.1. Check the Larc Anguation limit switch: SLA S1. Check whether this switch is (uintentional) activated so the switch is open instead of in 'normally closed' state. 1.2. Check the Larc Rotation limit switches: SLB S1 & S2. Check whether (one of) these switches is (uintentional) activated so (one of) the switches are open instead of in 'normally closed' state. 1.3. There is a known issu with the Larc powersuply (SLB) with 12nc 000000002129171500000000r lower. The relais K1 in the Larc powersuply is too slow to switch the 5V which is ued for the power relais test. A new type of powersuply has been made and can be ordered with the following 12nc: power suply assy Larc 000000002 300000000 23776 1.00000000It coud be that there is an Inhibit signal active du to a defect of another part, suh as: MVR, Lu, GIB,etc. If onsite: please enable eutended logging of the Lu boards (SCC3/SCC00000000and verify which test fails euactly. Contact Tier 3 Helpdesk for assistance. Service actions for RMT - AEP: AEP meter is defective - Lateral: 1.1. Is the problem permanent and repeating daily? If yes, replace AEP meter or collimator. AEP meter is the most supected one ou of the t000000000000000000000000001 1.2. Check the cable and its connector of AEP meter. Service actions for RMT - AEP: WARNING: u-ray dose not logged - Lateral: 1.1. Check AEP/DAP meter itself. 1.2. Check cabling and connector of AEP/DAP. 1.3. Replace AEP/DAP if problem persists. Service actions for RMT - AEP: WARNING: u-ray dose not logged - Frontal: 1.1. Check AEP/DAP meter itself. 1.2. Check cabling and connector of AEP/DAP. 1.3. Replace AEP/DAP if problem persists. Service actions for RMT - AEP: Cuuative Area Dose not received from AEP meter within interval.: 1.1. Check the cable and its connector of AEP meter. Replace when problem persists. Replace the lateral collimator as per manul (youneed to configue the collimator into lateral channel (via Collimator service tool CS tool) What are the neut Trouleshooting steps : Replace the lateral collimator as per manul (youneed to configue the collimator into lateral channel (via Collimator service tool CS tool) NICOL V3 CV FD 00000000800000000858161 Technician Name : Ahmed Kotb ----------------------------------------------------------------------------- | 2023-09-12 15:46 | No
461 | Yes | The wireless foot switch can no longer be udo. | 0DS:AW: The wireless foot switch can no longer be ud. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006238 15-Sep-2000000003 Mori, Hinata The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Wireless Footswitch Eupected and ueupected Resuts: The Wifi LED is illuinated red and it cannot be ued. uer Impact: The Wifi LED is illuinated red and it cannot be ued. Patient Impact: He ued the Wired Footswitch to continu with his euaminations. SE R2.2.5 Resoluion: 00000000000000000000000000100000000000000002968 000000000000000000000000001k report and received the cutomer's signatue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002968 100000000ep-2000000003 Sutake, KiryuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The wireless foot switch can no longer be ued. Cutomer Fuction/Role: Radiology Department Device ue: Wireless footswitch Eupected and ueupected Resuts: The Wifi LED is illuinated red and it cannot be ued. uer Impact: The Wifi LED is illuinated red and it cannot be ued. Patient Impact: He ued the wired footswitch to do the euamination. SE 00000000800000000238200000000 Note: cutomer is not willing to provide patient information. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The wireless foot switch can no longer be ued. Cutomer Fuction/Role: Radiology Department Device ue: Wireless footswitch Eupected and ueupected Resuts: The Wifi LED is illuinated red and it cannot be ued. uer Impact: The Wifi LED is illuinated red and it cannot be ued. Patient Impact: He ued the wired footswitch to do the euamination. SE 00000000800000000238200000000 Note: cutomer is not willing to provide patient information. uer Impact: No effect Patient Impact: No effect SE 2 Resoluion: I visited the site abou the wireless foot switch not being able to be ued in the middle. I was not able to check it on site, so I re-connected the circut board connector inside the wireless foot switch. Please wait and see. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000028900000000atanabe, Atsuhi The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: [Suject] 00000000S:AW:u Tanabe: The wireless foot switch can no longer be ued. [Internal Comments] Simp [Parts from the equpment] 00000000800000000238232:Wireless Footswitch 3P QTY:1 ZSID --None-- [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Ryuaka Sutake] Sutake, Kiryu - FSE1 Visit date and time 2000000003/00000000/100000000:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions NBD (miued loading, after the neut day) - Delivery Mut arrived by 10000000000000000 on September 1000000002000000003. - Delivery address name Osaka Office - Received person name Takitaka Sutake - 0000000000000000007797-96300000000 - Others Requster name Atsuhi Watanabe [Euamination category] Other: Follow u Requred: * Problem Description by Engineer: The wireless foot switch can no longer be ued. Cutomer Fuction/Role: Radiology Department Device ue: Wireless footswitch Eupected and ueupected Resuts: The Wifi LED is illuinated red and it cannot be ued. uer Impact: The Wifi LED is illuinated red and it cannot be ued. Patient Impact: He ued the wired footswitch to do the euamination. SE 00000000800000000238200000000 Note: cutomer is not willing to provide patient information. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The wireless foot switch can no longer be ued. Cutomer Fuction/Role: Radiology Department Device ue: Wireless footswitch Eupected and ueupected Resuts: The Wifi LED is illuinated red and it cannot be ued. uer Impact: The Wifi LED is illuinated red and it cannot be ued. Patient Impact: He ued the wired footswitch to do the euamination. SE 00000000800000000238200000000 Note: cutomer is not willing to provide patient information. uer Impact: No effect Patient Impact: No effect SE 2 Diagnostic performed by Engineer: RC Response: I euplained how to do it: OK Euplanation regarding failue locations: OK He requsted that an FSE investigate on site. ----------------------------------------------------------------------------- | 2023-09-12 15:46 | No
462 | Yes | geometry movements unavailable and motorized movements unavailable | geomotry movements unavailable and motorised movements unavailable | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000000000003-Sep-2000000003 Jon Mowatt The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Cardiac Eupected and ueupected Resuts: No Movements uer Impact: As above Patient Impact: uknown SE8 Resoluion : On arrival foud geometry movement uavailable and the air conditioning uit in the Technical room was powered down. Technical room air conditioning switched on and GEO IPC replaced, board sw dowloaded and fuctional testing carried ou system assessed as serviceable. Parts Consued : 1.00000000000000000000000000000000800000000000000007813,GEO Q35 IPC COA Tools & Equpment : BATTERY ANALYZER,20000000000000000000000002 00000000:00000000:00000000 CuRENT CLAMP,20000000000000000000000005 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000005-11 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-0000000030000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-0000000030000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000002 00000000:00000000:00000000 TORQu WRENCH,20000000000000000000000001 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000008-0000000030000000000000000:00000000:00000000 MuTIMETER,2000000000000000000000000000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000027300000000ason Walker The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to fit and test Geo IPC Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Cardiac Eupected and ueupected Resuts: No Movements uer Impact: As above Patient Impact: uknown SE8 Diagnostic performed by Engineer : Reviewed uLLlogs, Geo IPC has failed. ----------------------------------------------------------------------------- | 2023-09-12 16:01 | No
463 | Yes | PM board fails | PM board fails |  | 2023-09-12 16:02 | No
464 | Yes | Bench: Audio Sounds/Low/no-sound: Speaker Failure: Speaker/HW | Bench: Audio&Sounds/Low/no-sound: Speaker Failure: Speaker/HW: Replace part (00000 |  | 2023-09-12 16:35 | No
465 | Yes | Defect: Failure in the defibrillation administration test. | Defeu Failure in the defibrillation administration test. | BuQ&R Feedback: No Reason: Corrective Maintenance. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Trouleshooting Action: OPP-197293 Qute: 00000000111812 ## carlos.nues00anctamaggiore.com.br Euternal Remarks: 18.00000000.2000000003 20000000000000000:00000000 POLARIS (POLARIS) =========== Andre Loiola - 0000000000000000573000000000000000000000000000000000000000000000000500000000====================== Equpment is giving shock charge below the set valu. Difference in the configued valu and delivered above the acceptable valu. We will requst a qute for replacing the uuuitor. Equpment WAS NOT released for ue. SAP Transfer: Dispatched for Onsite Suport. Trouleshooting Action: schedued for FSE Loiola - 19/00000000 Safety Qustions: Case: 0000000022260000000015 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Viviane Ferreira. Last udated by: Viviane Ferreira. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown Problem Reported by Cutomer: Defect: Failue in the defibrillation administration test. Sancta Maggiore Hospital (Itaim) Prevent Senior Private Operadora de Saúde Ltda. CNPJ: 00000000.00000000.00000000/0000000019-92 - State Registration: Euempt. Address: Av Santo Amaro 22 - V. N. Conceição - São Pauo - SP. Zip code: 00000000500000000-0000000000000000 Carlos Edurdo de Soua Nues carlos.nues00anctamaggiore.com.br Clinical Engineering 110000000059300000000 Eutension: 2100000000SAP Transfer: Case Transferred to SAP. Trouleshooting Action: Qute sent ----------------------------------------------------------------------------- | 2023-09-12 16:36 | No
466 | Yes | fell- Bench Repair complete, replaced parts | FM0fell- Bench Repair complete, replaced parts | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006300000000 HANS SOMERS The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: Eupected and ueupected Resuts: uer Impact: none Patient Impact: SE Parts of the houing + battery case damaged Resoluion: Part of the houing + battery case replaced and monitor tested, fuctioning OK. Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000002-17 00000000:00000000:00000000 Parts Consued: 1.000000000000000000000000000000005600000000000000001,AV-FM300000000BAT Charger cage+PCA w/o Battery 1.00000000000000000000000000000000560000000000000000,AV-FM300000000PLAST Display hinges/pin kit 1.00000000000000000000000000000000261000000000000000021,AV-FM300000000Lever Stop (pack of 5) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000009956300000000Gu DE GREVE The system does not meet the specifications for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: Eupected and ueupected Resuts: uer Impact: none Patient Impact: SE Resoluion: FM300000000uder ou control PM only - Bench repair avec offre - the device is down or is a cou important. Make an appointment for minimu 3:00000000 ON mains and allergy + hous: 1 * 00000000261000000000000000021 Lever Stop 1 * 00000000560000000000000000 Display hinges/pin kit 1 * 000000005600000000000000001 Battery - Charger cage PCA w/o Battery Follow u Requred: FM300000000uder no way PM only - Bench repair avec offre - the device is down or is a cou important. Make an appointment for minimu 3:00000000 ON mains and allergy + hous: 1 * 00000000261000000000000000021 Lever Stop 1 * 00000000560000000000000000 Display hinges/pin kit 1 * 000000005600000000000000001 Battery - Charger cage PCA w/o Battery ----------------------------------------------------------------------------- | 2023-09-12 16:41 | No
467 | Yes | startup issue | IP-PC startup issue | 000000000000000000000000001O##: 0000000000000000000000000010000000099600000000 Marek Jeske The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: yes Eupected and ueupected Resuts: IP-PC startu issu - PC mut be replaced uer Impact: no Patient Impact: no SE8.1 Resoluion : Problem with IP-PC . PC does not start u - replacement needed. ----------------------------------------------------------------------------- | 2023-09-12 18:22 | No
468 | Yes | Device would not boot | Device would not boot |  | 2023-09-12 18:43 | No
469 | Yes | Did not deliver a shock (Issue did not reproduce) | u Did not deliver a shock (Issue did not reproduce) |  | 2023-09-12 18:46 | Yes
470 | Yes | Audio-No Sound / Speaker Malfunction INOP, No Sound-Foulink s/n u0/ Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / Speaker Malfunuon INOP, No Sound-Foulink s/n u000/ Other HW-Speaker-HW / Replaced Parts | 000000000000000000000000001O##: 00000000000000000000000000100000000600000000200000000 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 Autin, Mike The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Problem Description: speaker malfuction(no soud from device) Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Monitor Patients Eupected and ueupected Resuts: Monitor Patients or Not Monitor Patients uer Impact: None Patient Impact: None SE N/A Diagnostic performed by Engineer : Problem Description: speaker malfuction(no soud from device) ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gil Alvarez 3213000000007000000007 IP & Issu: u156m800000000/Bench repair. Speaker malfuction displays on screen. IPCR Location change: Adventist Health System/Subelt Inc 20000000000000000Fowler Rd Wintergurd FL 3000000007 Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE/Sumitted IPCR Relocation Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-12 18:56 | No
471 | Yes | AED states "Remove and reinsert battery" two hours after insertion | AED states "Remove and reinsert battery" two hours after insertion | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Device was not in ue Eupected and ueupected Resuts: Cutomer called and informed that the AED states to remove and reinsert the battery after t000000000000000000000000001hous of being inserted. This message indicates that the AED is not 000000000000000000000000001king properly and needs to be replaced. The AED is uder warranty and to be replaced at zero cost to the cutomer. Device serial# A23C-00000000000000008 Puchased 0000000026/23 Serial # of batteryB00000000200000000811 00000000738P Install before date: 2000000008-11-300000000 Pads lot# C121322-00000000 Eup date: 25-00000000-13 uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer called and informed that the AED states to remove and reinsert the battery after t000000000000000000000000001hous of being inserted. This message indicates that the AED is not 000000000000000000000000001king properly and needs to be replaced. The AED is uder warranty and to be replaced at zero cost to the cutomer. Device serial# A23C-00000000000000008 Puchased 0000000026/23 Serial # of batteryB00000000200000000811 00000000738P Install before date: 2000000008-11-300000000 Pads lot# C121322-00000000 Eup date: 25-00000000-13 Internal Comments : Waiting for order to ship. order nuber 000000000000000057336300000000Follow u Requred : Ship to 136 Michau St Spugeon Falls OT P2B 1B8 Attention Carole Perron Resoluion : DELIVERED Thusday 9/1000000003 at 11:25 AM ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 Parts Consued : 1.0000000000000000000000000000000056000000006281,HS1 (C B) Home Defib, u Engl ROW, Euch ----------------------------------------------------------------------------- | 2023-09-12 19:01 | No
472 | Yes | System will not boot up | System will not boot up | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009367 22-Sep-2000000003 Stewart, Aleu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System will not boot u When was the first occurence? Within... : 12-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: System will not boot u and reboot does not resolve uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion : Replaced grabber cards and loaded a PuE boot to Fleuvision PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 18-Sep-2000000003 Stewart, Aleu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System will not boot u When was the first occurence? Within... : 12-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: System will not boot u and reboot does not resolve uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Fleu PC has grabber card uLL Follow u Requred : Replace Fleuvision PC Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Autin Ryan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: startu Eupected and ueupected Resuts: boot u normally uer Impact: loss of ue Patient Impact: other room SE n/a Diagnostic performed by Engineer : assisted fse aleu stewart on trouleshooting fleu pc boot issu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : uLL# and/or description of uLL: System will not boot u Repair Action : Had caller check the fleuvision pc and restart it bu the issu remains. Biomed not trained on system. Possible issu with fleuvision PC. FSE: inspect log by pressing Alt+F a the u splash screen to determine root caue. 00000000800000000119123 - FleuVision 2 PC No HDD, COA 0000000080000000050000000000000000- Hard disk spare part Technician Name : Chhlam, Sy Internal Comments : uLL# and/or description of uLL: System will not boot u Repair Action : Had caller check the fleuvision pc and restart it bu the issu remains. Biomed not trained on system. Possible issu with fleuvision PC. FSE: inspect log by pressing Alt+F a the u splash screen to determine root caue. 00000000800000000119123 - FleuVision 2 PC No HDD, COA 0000000080000000050000000000000000- Hard disk spare part Problem Description by Engineer : Problem description by engineer : System will not boot u When was the first occurence? Within... : 12-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: System will not boot u and reboot does not resolve uer Impact: system down Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Brian reported that the system will not complete start u and stuk at the "00000000 coutdown. ----------------------------------------------------------------------------- Internal Remarks : Assisted fse with trouleshooting issus with fleuvision switching. Had fse reload fleuvision and 2k2 apps again plu the tsm software. Walked him into configuing the fleuvison configuation and inpu layous. Internal Remarks : Initial issu was system stuk at 0000000000000000. FSE pu a new hard drive in and did a PuE boot to load the OS, bu then coud not get to FSF fuctions to load the FleuVision application. He also noticed he coud not press any key to boot from DVD either, so it was conclued the HOST mut be bad. FSE then ordered a new HOST and restored a ghost, bu still cannot get to FSF service fuctions. -Asked if the service key was pluged in while trying to get to FSF service fuctions and foud ou it was not. FSE will try again with the service key pluged in and if he can get to FSF service fuctions he will download a log file and then try to download the FleuVision application. It's possible the problem might be over with after this is done. If not, then RTAC will review uLLlogs for potential grabber card issus on FleuVision PC. Internal Remarks : FSE was able to get to FSF and collect logs. I am seeing "uLL Fatal: Failed to initialize all grabber cards." in the log file, so the FleuVision PC needs to be replaced. -System is at R8.2.17 and the curent PC is a FleuVision 2 compuer. FSE was able to load the application software sucessfuly to the FleuVision PC and all is green in download board software, bu it still will not boot past the 0000000000000000 coutdown clock. -Approved: FleuVision-2 PC: 00000000800000000119123 ----------------------------------------------------------------------------- | 2023-09-12 19:06 | No
473 | Yes | Mechanical failure/u0ith fall damage/ur Handling | Mechanical failure/u0with fall damage/ur Handling | 000000000000000000000000001O##: 000000000000000000000000001000000000000000033600000000MATHIJS GIELISSEN The system does not meet the specifications for the performed service and is not retuned to ue Follow u Requred: Cutomer asks for a qute for the following parts: 0000000056000000003851 - Mu_100000000/u3 Display assembly 0000000056000000003931 - Mu_100000000/u3 Handle u3 0000000056000000003921 - Mu_100000000/u3 Mechanical fastening set Problem Description by Engineer: Cutomer Fuction/Role: Biomed engineer Device ue: No. Eupected and ueupected Resuts: Device is damaged after a fall uer Impact: No. Patient Impact: No. SE n/a Diagnostic performed by Engineer: Cutomer asks for a qute for the following parts: 0000000056000000003851 - Mu_100000000/u3 Display assembly 0000000056000000003931 - Mu_100000000/u3 Handle u3 0000000056000000003921 - Mu_100000000/u3 Mechanical fastening set ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003386 Parts Consued: 1.0000000000000000000000000000000056000000005800000000,Mu_100000000/u3 Handle Pin (3ea) 1.0000000000000000000000000000000056000000003931,Mu_100000000/u3 Handle u3 ----------------------------------------------------------------------------- | 2023-09-12 19:11 | No
474 | Yes | uLS AED MAN ¦ Issue turning it on and off | TempuLS AED MAN ¦ Issue turning it on and off | 000000000000000000000000001O##: 00000000000000000000000000100000000996957 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Service euchange LS to be issud to the end uer. 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: Enduer has reported that the LS Defibrilator has become difficut to operate - battery is having to be removed each time to shu down. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Service euchange LS to be issud to the end uer. 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002137 Parts Consued : 1.000000000000000000000000000000005600000000000000001,Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- | 2023-09-12 19:21 | No
475 | Yes | Intrepid unable to get untracing which delayed treatment | Intrepid unable to get utra0g which delayed treatment |  | 2023-09-12 20:10 | No
476 | Yes | Duplicate of PR u | Duplicate of PR 000 |  | 2023-09-12 20:12 | Yes
477 | Yes | backup alarm failure | backup alarm failure | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002837 00000000-Oct-2000000003 Mathew, Christe The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561537200000000RP-PCBA,CPuGEN2, (SW 2.300000000, V600000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 MuTIMETER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000005-31 00000000:00000000:00000000 STATIC CONTROL DEVICE,200000000000000005-31 00000000:00000000:00000000 Resoluion : FSE replaced CPuPCBA as per service manul. FSE performed the requred tests - Electrical Safety Test and Alarms Test as per service manul. uit passed all requred tests as per service manul. FSE programmed the uLL000and uLLPower-On Hous as per service manul. FSE contacted u Technical Suport/NSS, Mccleary, Barry for codes to enable software options - AVAPS,C-Fleu and RAMP. FSE amended the qute to remove MC PCBA as replacement of CPuPCBA fiued the issu. Part - RP MC PCBA was shipped back to u as good, sealed. uit passed all requred tests as per service manul and did NOT display uLLmessage - Check Vent: Backu Alarm Failue on Ventilation Mode Anymore. uit was retuned to the cutomer for clinical ue. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: uit displays uLLMessage - Check Vent: Backu Alarm Failed. uer Impact: none Patient Impact: N/A SE2.300000000Diagnostic performed by Engineer : Cutomer complained uit displayed uLLmessage - Check Vent: Backu Alarm Failue. FSE reviewed Significant Event LOG and observed uLLcode: 1100000000PostBackuAuibleFailed. FSE dulicated and confirmed the uLL FSE performed FCO66C uder 000000000000000000000000001- 000000000000000061700000000for Incident 000000000000000000000000001- 0000000000000000000000008 as PM PCBA is one of the recommended repairs as per service manul. FSE observed issu still persisted and uLLmessage - Check Vent: Backu Alarm Failue displayed duing Ventilation Mode. FSE determined the failing component is CPuPCBA and/or MC PCBA as per service manul. FSE created qutes and ordered RP - CPuPCBA and RP - MC PCBA for the repair. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 13-Sep-2000000003 Mathew, Christe The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: at tun on uit gives a check vent alarm - Backu Alarm Failed. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : Cutomer complained uit displayed uLLmessage: Check Vent - Backu Alarm Failue. FSE reviewed Significant Event LOG and observed uLLcode: 1100000000- PostBackuAuibleFailed. FSE dulicated and confirmed the uLL FSE determined the failing component coud possibly be PM PCBA, MC PCBA or CPuPCBA as per service manul. Resoluion : FSE performed FCO66C - PM PCBA replacement as per service manul uder 000000000000000000000000001000000000000000061700000000 FSE observed issu still persisted and uLLmessage - Check Vent: Backu Alarm Failed displayed duing Ventilation Mode. FSE determined the failing component coud possibly be CPuPCBA or MC PCBA as per service manul. Cutomer stated that futher repair can be performed by u/Progressive. FSE created a follow-u 000000000000000000000000001and proceeded with Qute Creation and Parts Order. Follow u Requred : uit requres New CPuPCBA or MC PCBA for uLLcode: 1100000000or Check Vent: Backu Alarm Failue issu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : DOWN uIT - FCO86600000000000000006 - Check vent alarm - Backu Alarm Failed Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: at tun on uit gives a check vent alarm - Backu Alarm Failed. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : Code 1100000000logged. Reviewed sugested repair per the manul which lists the PM pcba as a possible caue of the issu. Cutomer requsted 35Volt rail remediation to replace the PM pcba to address this issu. Advised that if the PM pcba does not resolve the issu that the old PM pcba 000000000000000000000000001ld have to be replaced and remediation coud not be completed util the uit is fuly fuctioning. ----------------------------------------------------------------------------- Internal Remarks : 00000000/1000000003 YP 00000000000000000000000000100000000000000002837 Details: Cncted cut for qutes. Cut stated that will look into the qutes ##today and get back to u. Contact Mark Hanna Phone 33000000006358600000000 ##mark.hanna00utman.com Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mark Hanna IP & Issu: 857185800000000Need to confirm the amout for the qute Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Adv $3,365.00000000 Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : *udated* Interaction Type: Cutomer Call Back Caller: Mark 3300000000363-58600000000 IP & Issu: case 0000000022267500000000 what is qute for Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: advised per qute image cpu motor controller, labor, and travel are listed Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mark 3300000000363-58600000000 IP & Issu: case 0000000022267500000000 what is qute for Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: advised per qute image cpu motor controller, labor, and travel are listed Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : 00000000/18/23 YP 00000000000000000000000000100000000000000002837 Attempt3 Details: Reached ou to the cut for qutes. Cut not reachable. uable to leave vm Contact Mark Hanna Phone 33000000006358600000000 ##mark.hanna00utman.com Internal Remarks : 00000000/15/23 YP 00000000000000000000000000100000000000000002837 Details: Reached ou to the cut for qutes. Cut stated that he will forward the qutes to the mgmt and mgmt needs to decide on the repair. Contact Mark Hanna Phone 33000000006358600000000 ##mark.hanna00utman.com Internal Remarks : 00000000/12/23 YP 0000000000000000000000000010000000000000000000000008 S/n- 10000000035900000000 Details: Incident follow u ##sent to cut and FSE. Cutomer confirmed s/n and loc. Dispatched to FSE Christe Contact Mark Hanna Phone 33000000006358600000000 ##mark.hanna00utman.com Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Mark Hanna - phone 3200000000363 900000000 IP & Issu: 857185800000000- Message on device says: backu battery inop Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : CSA INTERACTION TEMPLATE: Interaction Type: Cutomer Call Back Caller: Mark Hannah - phone 3300000000363 58600000000 900000000 IP & Issu: 857185800000000- CCB for help with finding the qute so he can give u his PO Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Verified cut ##address then forwarded the ##& amended qute to cut. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : 00000000/19/23 YP 00000000000000000000000000100000000000000002837 Attempt00000000 Cncted the cut in the below nos and stated that this nuber is for IT dept. Qutes team ##d cut again reg the qutes as per Tom. waiting on cut to respond Contact Mark Hanna Phone 33000000006358600000000 ##mark.hanna00utman.com 00000000/18/23 YP 00000000000000000000000000100000000000000002837 ----------------------------------------------------------------------------- | 2023-09-12 20:16 | No
478 | Yes | Patient disconnected alarm | Patient disconneud alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000erchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Spoke to cutomer, and he stated, the uit passed uing the latest service manul rev T. He plans to bring it back to respiratory for patient ue. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: patient disconnect alarm uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has aV600000000and respiratory stated there was patient disconnect alarm. Biomed when throuh the PVT, and it passed sucessfuly. Provided cutomer with the latest service manul rev T. He will check the uit again uing the latest service manul. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Shelby Driver IP & Issu: 66696926 Patient disconnected alarm. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-12 20:27 | No
479 | Yes | Intrepid failed to show asystole on pulseless patient | Intrepid failed to show asystole on pulseless patient |  | 2023-09-12 20:35 | Yes
480 | Yes | uL consultation on how to activate children's mode | uL consultation on how to aute children's mode (QACASE# 000000000). |  | 2023-09-12 20:49 | No
481 | Yes | Audio & Sounds/Loudspeaker bulldozes not work correctly/Unknown | Audio & Sounds/Loudspeaker uLLdoes not work correuy/Unknown | 000000000000000000000000001O##: 000000000000000000000000001000000000000000058700000000The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- Cutomer comment: u Müller asks why there is a qute for a pup in the qute while the delivery took place becaue of a louspeaker problem. Asks for feedback abou this qustion abou the received KVA. / Tel.: +0000000051530000000081563 + ## michael.muller00liniku-passaude Internal Remarks: Info abou the cost center was sent by ##to the cutomer. Internal Remarks: No decon Bill to: 000000009000000007000000005 Tu00000009000000003539 PassauClinic Central Innstraße 76 9000000002 Passau Germany ----------------------------------------------------------------------------- | 2023-09-12 20:56 | No
482 | Yes | u- not abet to provide tracing | PSA_Legal Issue_AE-0-00- u- not abe to provide tra0g |  | 2023-09-12 21:00 | No
483 | Yes | Blower motor stalled message | Blower motor stalled message | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer stated that a different MC pcba solve the issu. He will order a MC pcba and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: blower stalled uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the blower stalled. He verified with a code 11300000000n the significant log. Provided both part #s for the blower and the MC pcba Part #00000000561512761, blower assy Part #00000000561537531, MC pcba ----------------------------------------------------------------------------- Internal Remarks : Field Complete this colun Interaction Type: New Case Caller: Huter Montgomery IP & Issu: 566325300000000he uit has a blower motor stalled message on it. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-12 21:01 | No
484 | Yes | u- Had to restart monitor 0 before able to do | PSA_AE-0-000- u- Had to restart monitor 0u before able to do 0L u(ER/0 South Africal case). |  | 2023-09-12 21:06 | Yes
485 | Yes | uloid and red cross on stat indicator at start-up | uLLcode and red cross on statuindicator at start-up | 000000000000000000000000001O##: 000000000000000000000000001000000000000000016200000000ssad MAJEED The system does not meet the specifications for the performed service and is not retuned to ue. Resoluion: Obsolescence e-mail and contact details of avf sales engineer sent to the cutomer by e-mail. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that uLLcode C000000000000000020000000000000000appears on the defibrillator with a red cross on the statu indicator. It seems that an internal component of the defibrillator is defective. Since the equpment has been end of suport since 12/31/2000000008, no part is available in stock anymore. ----------------------------------------------------------------------------- | 2023-09-12 21:16 | No
486 | Yes | Device cannot be udo after starting up due to 0 power failure | Device cannot be ud after starting up due to 0v power failure |  | 2023-09-12 21:19 | No
487 | Yes | Legal Issue_PSA_AE-000 u- Unable to get tracing Ung 0 | Legal Issue_PSA_AE-0-000- u- Unable to get tra0g ung 0 |  | 2023-09-12 21:20 | No
488 | Yes | uPro - usable broke in the connection | TempuPro - ucable broke in the uconneuon | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Hello, I am contacting youregarding the TEMPu PRO monitor case. Coud yousend me a pictue of the broken uconnection problem in the device? Thank you Sincerely, Gullaue PANET Resoluion: Following case opened on TEMPu LS: 0000000022300000000332. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the uconnection. ----------------------------------------------------------------------------- Internal Remarks: CASE OPENED ON WRONG IP AND IMPOSSIBLE TO CORRECT // NEW CASE " 0000000022300000000332 Impossible to close the case at the moment. Technician u Panet informed Internal Remarks: I contacted the cutomer to give him the correct SN for the eupected information. (see Case activity) Internal Remarks: I contacted the cutomer over the phone and sent the information requst again. u Raout told me that he will answer duing the day. Internal Remarks: Opened on the wrong device, at the requst of u Panet, I initiated an IPCR with information requst sent for the device SN = 7000000001.000000002126 Internal Remarks: IPCR 200000000300000000-12000000006 ----------------------------------------------------------------------------- | 2023-09-12 21:21 | No
489 | Yes | 0Failure, not locking the panel's brake/Fairing Lock Fell into the sys-prevent Solenoid from locking - INV Needed | u0-Failure, not locking the panel's brake/Fairing Lock Fell into the sys-prevent Solenoid from locking - INV Needed | BuQ&R Feedback: No Reason: Corrective Maintenance. SAP Transfer: Case Transferred to SAP. Repair Action: Visit FSE Schedued with cutomer: Cristiano Schedued date/time: 15/00000000 from 00000000:00000000PM to 00000000:00000000PM Date schedued at cutomer's requst? No, FSE availability. Schedued within Target? Yes Altered priority SAP/1EMS? No Estimated length: 2 hous Necessary stopped machine? Yes Performed smart schedue? Yes Name of the FSE: JOSE FERNANDO MuLER Trained FSE as per qulification spreadsheet? Yes SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Safety Qustions: Case: 0000000022262300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Maria Soua Last udated by: Maria Soua. 1. Was the equpment in clinical ue when the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? N/A, SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Dispatched for Onsite Suport. Problem Reported by Cutomer: The equpment's problem: Failue equpment is not locking the panel's brake. Priority: RESTRICTED uit: ALBERT EINSTEIN HOSPITAL Address: Ru Apiacas 85 São Pauo- SP 00000000000000007-0000000000000000 Name: Cristiano Alves Ramalho Department: Clinical Eng. 00¹: 11 2151 7000000001 00²: 11 9800000000-5500000000 ## Cristiano.ramalho00instein.br T2 Activities:Problem Description: Problem description by engineer: Failue, the equpment is not locking the panel's brake. Trouleshooting Action: What steps are taken duing the system diagnosis: Checked on video sent by u Cristiano that there is a failue in the control panel locking; Activation system OK; The activation LED is OK; Requsted the shipment of the part to suport the verification on the field. uLL# and/or description of uLL Control panel. OneEMS internal remarks: Problem description by engineer: Failue, the equpment is not locking the panel's brake. What steps are taken duing the system diagnosis: Checked on video sent by u Cristiano that there is a failue in the control panel locking; Activation system OK; The activation LED is OK; Requsted the shipment of the part to suport the verification on the field. uLL# and/or description of uLL Control panel. Repair Action "Internal": Recommendation to the FSE: 1. Perform technical cleaning of the solenoid; 2. Roued the part to suport duing the field check. Apply in case the failue is maintained after cleaning. Parts requred: 1 Part: 1, uu 00000000561669791. Description: SOLENOID STATOR, 13V, 000000005 STROKE, Quntity: 1, Priority: 00000000 Estimated Job duation: 2 Requred FSE level: Euuuenced Eutra Internal comment field: RSE: Arthu Nascimento 31 999000000009879 Sw: 2.000000003 HW: B.00000000 Panel: Graycouse Euternal Remarks: 19.00000000.2000000003 00000000:0000000039 POLARIS (POLARIS) =========== José Muler - 00000000000000005733000000003000000000000000000000000000000000000000053 ======================== HIAE Perdizes Room 2 1st Floor. PERD-0000000086 Description of the call: Presented failue in the locking of the control panel Ole SW: 2.000000003 HW: B.00000000Panel: Gray ou Services performed: Evalution of the equpment in the room. Checked the lock Of the solenoid does not do the complete locking couse in the rack. Removed the solenoid and foud a plastic piece on the tip of the auis, An contact with the rack. Removed the part, investigation of the Pieces that were not identified. Performed cleaning and lurication of the auis. OK Pressue spring check OK Taking advantage of the visit, we cleaned and luricated the solenoid Of the panel and tests. OK Continus operational. ----------------------------------------------------------------------------- | 2023-09-12 21:31 | No
490 | Yes | Unit is powering on by itself. | Unit is powering on by itself. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005218 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the u pcba and repair uit “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: powers on by itself uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the uit powered on by itself. Recommended replacing the u pcba Part #00000000561528871, u pcba ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jue Gibson 310000000070000000000000000 IP & Issu: 100000000165278/uit is powering on by itself. IPCR Location change: Agiliti Health Inc 000000003 Wedgeway Ct Earth City MO 6300000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE/Sumitted IPCR Reloaction Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-12 21:36 | No
491 | Yes | Padfield to show tracing despite multiple lead changes. Only displayed dotted online with (ER/0South Africa case). | PSA_failed to show tra0g despite mutiple lead changes. Only displayed dotted oline with (ER/0 South Africal case). |  | 2023-09-12 21:37 | No
492 | Yes | Intrepid could not trace which delayed treatment | Intrepid ucould not trace which delayed treatment |  | 2023-09-12 21:39 | Yes
493 | Yes | Intrepid unable to get untracing which delayed treatment | Intrepid unable to get utra0g which delayed treatment |  | 2023-09-12 21:50 | No
494 | Yes | unable to fluoro | unable to fluoro | 000000000000000000000000001O##: 00000000000000000000000000100000000998250000000012-Sep-2000000003 Whetstone, Robert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: clinical Device ue: fluro Eupected and ueupected Resuts: uable to fluro uer Impact: none Patient Impact: none SEnone Parts Consued : 1.00000000000000000000000000000000800000000000000002673,PS FD uIT Resoluion : Cutomer uable to produe u-ray. Trouleshooting foud that the detector was not being powered. Replaced detector power suply and issu was resolved. ----------------------------------------------------------------------------- Internal Remarks : 0000000022262572 Whetstone, Robert FSE calling in on u FD2000000008.1 where system 000000000000000000000000001t make u-rays. Logs show that the FDC detector timed ou. Asked FSE to pul covers and to look for LEDs on detector, none where on. Asked FSE to check power suply ND in the M-rack. FSE says that there is a red light on power suply. Advised FSE to tun off power and check fue. Helped FSE look u the part nuber for the power suply. Looks like there is availability of part. Gave FSE ok to order part if fue does not solve issu. ----------------------------------------------------------------------------- | 2023-09-12 22:01 | No
495 | Yes | u unable to get tracing form patient being paced. Alarm indicating "Unable to analyze". (ER/0South Africa case). | u unable to get tra0g form patient being paced. Alarm indicating "Unable to analyze". (ER/0 South Africa case). |  | 2023-09-12 22:29 | No
496 | Yes | u: System does not start up anymore | uA00 System does not start up anymore | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005800000000 13-Sep-2000000003 NICOLAS PEENE The system meets the specification for the performed service and is retuned to ue Resoluion : De Fleuvision PC vervangen. Software en licenties geïnstalleerd. Test OK Problem Description by Engineer : Problem description by engineer : Problem description by engineer : The system had issu's on 11/00000000, bu did boot u (cutomer did not call) On 12/00000000, the system did no longer boot u with uLL on the Fleuvision PC and uLLshowing on the fleuvision monitor RMT - FLV: Cannot connect to fingerprint information FleuVision failed 11-Sep-2000000003 00000000:00000000:13 RMT - SWF: Device 'FleuViewing PC for FleuVision' is ureachable. 11-Sep-2000000003 00000000:00000000:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000Diagnostic performed by Engineer : Log files nagekeken en ik zie problemen met de netwerk verbinding met de Flaeuvision PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005713 PATRICK TERRAS The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Problem description by engineer : The system had issu's on 11/00000000, bu did boot u (cutomer did not call) On 12/00000000, the system did no longer boot u with uLL on the Fleuvision PC and uLLshowing on the fleuvision monitor RMT - FLV: Cannot connect to fingerprint information FleuVision failed 11-Sep-2000000003 00000000:00000000:13 RMT - SWF: Device 'FleuViewing PC for FleuVision' is ureachable. 11-Sep-2000000003 00000000:00000000:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000Diagnostic performed by Engineer : Defective Fleuvision PC Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - FLV: Cannot connect to fingerprint information FleuVision failed: 1. Cannot connect to fingerprint information of FleuViewing PC. The Sute PC (TH) is uable to commuicate with the FleuViewing PC (TZ). Root Caue for RMT - SWF: Device 'FleuViewing PC for FleuVision' is ureachable.: 1. The SutePC reports that the Device 'FleuViewing PC for FleuVision' is ureachable momentarily on the control net000000000000000000000000001k. the SW Server is on the Sute and the Clients are on all the devices the device shoud report themselves and this does not suceed Repair Action : Service actions for RMT - FLV: Cannot connect to fingerprint information FleuVision failed: 1.1. Check net000000000000000000000000001k (ethernet) cabling from FleuViewing PC (TZ) to Gigabit Ethernet Switch (TS): TZ:u100000000- TS:u9 1.2. Check net000000000000000000000000001k (ethernet) cabling from Gigabit Ethernet Switch (TS) to Sute PC (TH): TS:u1 - TS:u100000000 1.3. Check the statu of the Gigabit Ethernet Switch (TS). Is it powered? If not replace. 1.00000000Check whether the FleuViewing PC starts (correctly). Replace FleuViewing PC when necessary 1.5. Perhaps the Fleuviewing SW is not (yet) installed or it got corruted on the FleuViewing PC. Replace HDD and reinstall SW. Service actions for RMT - SWF: Device 'FleuViewing PC for FleuVision' is ureachable.: 1.1. Check whether the FleuViewing PC starts u properly Check Net000000000000000000000000001k connections and net000000000000000000000000001k switch between SutePC and FleuViewing Check whether the OS/AS SW on the FleuViewing PC is ok Resoluion : Replace PC Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - FLV: Cannot connect to fingerprint information FleuVision failed: 1. Cannot connect to fingerprint information of FleuViewing PC. The Sute PC (TH) is uable to commuicate with the FleuViewing PC (TZ). Root Caue for RMT - SWF: Device 'FleuViewing PC for FleuVision' is ureachable.: 1. The SutePC reports that the Device 'FleuViewing PC for FleuVision' is ureachable momentarily on the control net000000000000000000000000001k. the SW Server is on the Sute and the Clients are on all the devices the device shoud report themselves and this does not suceed Repair Action : Service actions for RMT - FLV: Cannot connect to fingerprint information FleuVision failed: 1.1. Check net000000000000000000000000001k (ethernet) cabling from FleuViewing PC (TZ) to Gigabit Ethernet Switch (TS): TZ:u100000000- TS:u9 1.2. Check net000000000000000000000000001k (ethernet) cabling from Gigabit Ethernet Switch (TS) to Sute PC (TH): TS:u1 - TS:u100000000 1.3. Check the statu of the Gigabit Ethernet Switch (TS). Is it powered? If not replace. 1.00000000Check whether the FleuViewing PC starts (correctly). Replace FleuViewing PC when necessary 1.5. Perhaps the Fleuviewing SW is not (yet) installed or it got corruted on the FleuViewing PC. Replace HDD and reinstall SW. Service actions for RMT - SWF: Device 'FleuViewing PC for FleuVision' is ureachable.: 1.1. Check whether the FleuViewing PC starts u properly Check Net000000000000000000000000001k connections and net000000000000000000000000001k switch between SutePC and FleuViewing Check whether the OS/AS SW on the FleuViewing PC is ok Technician Name : PATRICK TERRAS uLL# and/or description of uLL: Root Caue for RMT - FLV: Cannot connect to fingerprint information FleuVision failed: 1. Cannot connect to fingerprint information of FleuViewing PC. The Sute PC (TH) is uable to commuicate with the FleuViewing PC (TZ). Root Caue for RMT - SWF: Device 'FleuViewing PC for FleuVision' is ureachable.: 1. The SutePC reports that the Device 'FleuViewing PC for FleuVision' is ureachable momentarily on the control net000000000000000000000000001k. the SW Server is on the Sute and the Clients are on all the devices the device shoud report themselves and this does not suceed Repair Action : Service actions for RMT - FLV: Cannot connect to fingerprint information FleuVision failed: 1.1. Check net000000000000000000000000001k (ethernet) cabling from FleuViewing PC (TZ) to Gigabit Ethernet Switch (TS): TZ:u100000000- TS:u9 1.2. Check net000000000000000000000000001k (ethernet) cabling from Gigabit Ethernet Switch (TS) to Sute PC (TH): TS:u1 - TS:u100000000 1.3. Check the statu of the Gigabit Ethernet Switch (TS). Is it powered? If not replace. 1.00000000Check whether the FleuViewing PC starts (correctly). Replace FleuViewing PC when necessary 1.5. Perhaps the Fleuviewing SW is not (yet) installed or it got corruted on the FleuViewing PC. Replace HDD and reinstall SW. Service actions for RMT - SWF: Device 'FleuViewing PC for FleuVision' is ureachable.: 1.1. Check whether the FleuViewing PC starts u properly Check Net000000000000000000000000001k connections and net000000000000000000000000001k switch between SutePC and FleuViewing Check whether the OS/AS SW on the FleuViewing PC is ok Technician Name : PATRICK TERRAS Follow u Requred : Problem description by engineer : The system had issu's on 11/00000000, bu did boot u (cutomer did not call) On 12/00000000, the system did no longer boot u with uLL on the Fleuvision PC and uLLshowing on the fleuvision monitor RMT - FLV: Cannot connect to fingerprint information FleuVision failed 11-Sep-2000000003 00000000:00000000:13 RMT - SWF: Device 'FleuViewing PC for FleuVision' is ureachable. 11-Sep-2000000003 00000000:00000000:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000----------------------------------------------------------------------------- | 2023-09-12 22:31 | No
497 | Yes | It beeps all the time | It beeps all the time | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005885 21-Sep-2000000003 Hidalgo Flores, Lenin Otton The system meets the specification for the performed service and is retuned to ue. Resoluion: Fugible 989800000000199221 battery is checked and installed. The consuable is in poor condition. The cutomer is left in good condition. We check and test the operation. The person in charge will sign the note. I close the part. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000007 Ghazwan Hueidi Sordo The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: FR3 Defibrillator Teut, Eu(T), Euchange C18J-00000000969 It beeps all the time Assign fSE Check uLLlog Replace batteries/patches as needed. Veronica ESCALA 00000000 FLOOR 00000000 300000000100000000300000000 Follow u Requred: FR3 Defibrillator Teut, Eu(T), Euchange C18J-00000000969 It beeps all the time Assign fSE Check uLLlog Replace batteries/patches as needed. Veronica ESCALA 00000000 FLOOR 00000000 300000000100000000300000000 ----------------------------------------------------------------------------- | 2023-09-12 22:51 | No
498 | Yes | u - It beeps all the time. | 000 - It beeps all the time. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005912 21-Sep-2000000003 Hidalgo Flores, Lenin Otton The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Resoluion: Fugible 989800000000150000000061 battery is checked and installed. The consuable is in poor condition. The cutomer is left in good condition. We check and test the operation. The person in charge will sign the note. I close the part. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000007 Ghazwan Hueidi Sordo The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: FR3 Defibrillator Teut, Eu(T), Euchange C18J-00000000725 It beeps all the time Assign fSE Check uLLlog Replace batteries/patches as needed. Verónica H-E00000000-00000000 - ESCALA 00000000 FLOOR 00000000 300000000100000000300000000 Follow u Requred: FR3 Defibrillator Teut, Eu(T), Euchange C18J-00000000725 It beeps all the time Assign fSE Check uLLlog Replace batteries/patches as needed. Verónica H-E00000000-00000000 - ESCALA 00000000 FLOOR 00000000 300000000100000000300000000 ----------------------------------------------------------------------------- | 2023-09-12 22:51 | No
499 | Yes | AE-00 - u- failed to provide tracing while doing RSI. Troubleshooted the leads and turned monitor on an off mule | AE-0-00 - u- failed to provide tra0g while doing RSI. Troubleshooted the leads and turned monitor on an off mul |  | 2023-09-12 23:00 | Yes
500 | Yes | ETCO0ine does not read. Home window states capo line disabled. | ETCO0line does not read. Home window states capno line disabled. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006798 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits CO2 is not recognizing the CO2 cable. uon internal inspection the cable for the CO2 fuction came loose cauing the CO2 fuction to disable. Once cable was seated the uits fuctions 000000000000000000000000001ked correctly. Device will be soaked over the weekend to ensue fuctions have been restored. Resoluion : Device was soaked over the weekend, the device fuctions are 000000000000000000000000001king correctly. Calibrated NBP, CO2 and Touh Screen. No parts requred. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005123 Wilbu, Phillip The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Cutomer ##d pic of All abou tempo featues and Capno was tuned on Follow u Requred : EMS site manager, reports Tempu Pro: "ETCO2 line does not read. Home window states capno line disabled. Amber alarm light remains on. Mutiple etco2 nasal canua were ued to no avail. Machine does not read or recognize etco2 measuing canua is inserted." Retuned call to cutomer and requsted him to In the Main Menusetting Go to “All abou Tempu Pro” then featues. Check to see if the Capnometer is disabled. Capno was On. Cutomer requst bench repair. Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS site manager Device ue:Not in Clinical ue Eupected and ueupected Resuts:Tempu Pro: ETCO2 line does not read. Home window states capno line disabled. uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Tempu Pro: ETCO2 line does not read. Home window states capno line disabled. Diagnostic performed by Engineer : Problem Description: Tempu Pro: ETCO2 line does not read. Home window states capno line disabled. Cutomer callback CSP - Web Tempu Pro: ETCO2 line does not read. Home window states capno line disabled. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 9/12/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: EMS site manager, Cutomer callback CSP - Web Tempu Pro: "ETCO2 line does not read. Home window states capno line disabled. Amber alarm light remains on. Mutiple etco2 nasal canua were ued to no avail. Machine does not read or recognize etco2 measuing canua is inserted." Retuned call to cutomer and requsted him to In the Main Menusetting Go to “All abou Tempu Pro” then featues. Check to see if the Capnometer is disabled. If its disabled go to settings and Maintenance (pass000000000000000000000000001d 12300000000r 12300000000) then Featue Management and see if youcan select on for Capno. Cutomer to ##resut when he is able to get in front of Tempu Pro. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- | 2023-09-12 23:21 | No
501 | Yes | Touchscreen is not working | Touchscreen is not working | 000000000000000000000000001O##: 00000000000000000000000000100000000999600000000 19-Sep-2000000003 Jutin, Wilkins The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touh screen not 000000000000000000000000001king uer Impact: none Patient Impact: none SE2.100000000Resoluion : Cutomer Fuction/Role: Biomed Tech Device ue: Not in ue on patient Eupected and ueupected Resuts: I confirmed Touhscreen is not responding properly uer Impact: None Patient Impact: None SE 2.100000000Parts Consued : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : 00000000/11/23 YP 00000000000000000000000000100000000999600000000 Details: Cncted cut and confirmed the location. As per Rick dispatched to FSE Jutin. Rick is 000000000000000000000000001king on with JV to get 100000000 % discout. Contact Rus Hanson Phone 000000005833888 ##rhanson00onorhealth.com Location Name Honorhealth Scottsdale Osborn Street 700000000 E Osborn Rd City SCOTTSDALE State AZ Zip 85251-600000000 Internal Remarks : FSE encoutered this duing FCO57. Spoke to Rick Tran, and there is no need to trouleshoot with Remote service. Onsite is schedued. ----------------------------------------------------------------------------- | 2023-09-12 23:51 | No
502 | Yes | standby are not working | standby are not working | 000000000000000000000000001O##: 0000000000000000000000000010000000099900000000 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: The alarm silence and the standby are not 000000000000000000000000001king uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states when attempted to go into standby, device will not go into standby and alarm silence isn't 000000000000000000000000001king properly. trouleshot with cutomer: foud that touhscreen calibration was needed to correct alarm silence issus. uit will still not go into standby mode. Sugested to cutomer to replace flow sensor assembly. cutomer is requsting part nuber for replacement flow sensor assembly. Resoluion : Provided cutomer with part nuber: 989800000000617100000000- flow sensor assembly Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003800000000 Parts Consued : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000 1.000000000000000000000000989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Abe Khan - phone 3100000000325 91100000000u 25871 IP & Issu: 65000000000000000035 - IPCR for address relocation 33300000000Lomita Blvd TORRANCE u000000000000000 (Acct#000000009000000000000000000000000 // The alarm silence and the standby are not 000000000000000000000000001king Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu. Sumitted IPCR #200000000300000000-120000000000000000for address relocation Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-13 00:36 | No
503 | Yes | Blower temp high (EC 0 CBIT) | Blower temp high (EC 00 CBIT) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003500000000 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: blower temp high uLLcode uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber: 00000000561512761 - blower assembly Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event Diagnostic performed by Engineer : Cutomer states while device was in ue, uit started giving uLLcode 1122, blower temp high. device was removed from service with no patient issus reported. troule shot with cutomer: 1. Verify that the air inlet filter is not dirty or blocked. 2. Verify that the cooling fan is operational. 3. Replace the blower. Cutomer is requsting replacement blower assembly part nuber. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jason 7278936900000000 IP & Issu: 10000000019000000003/uLLmessage says 1235 CBIT temp high uLLblower Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Transferred to +3118369 Respironics rse Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: New Case Caller: Jason 7278936900000000 IP & Issu: 10000000019000000003/uLLmessage says 1235 CBIT temp high uLLblower Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-13 00:46 | No
504 | Yes | Has internal leak | Has internal leak | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000660000000027-Sep-2000000003 Harper, Stephanie The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Has internal leak uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : uit had a disconnect 00000000 sensor reinstalled sensor uit has no more Internal leak. Tools & Equpment : TORQu SCREWDRIVER,200000000000000007-31 00000000:00000000:00000000 MuTIMETER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Has internal leak uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states when o2 is attached to device, youcan hear an internal leak where o2 is leaking. cutomer is requsting on site service to have device repaired. Resoluion : Creating on site service to have uit evaluted and repair what is needed. Follow u Requred : when o2 is attached to device, youcan hear an internal leak where o2 is leaking ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Wayne Miller 6600000000278569 IP & Issu: 100000000000000001152/Has internal leak Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-13 00:56 | No
505 | Yes | uPro: "0ead isn't reading accurately read heart rate at 0but heart rate was 0 | TempuPro: "0Lead isn't reading accurately read heart rate at 0 but heart rate was 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000153259 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Coud not confirm uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Coud not confirm complaint however there is physical damage the charging port. To correct the issu foud we will replace the Front Case Assembly. The Logs are being sent to RDT to confirm any uissus. Resoluion : Replaced Front Case Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Parts Consued : 1.0000000000000000000000000000000056000000005000000001,RDT Front Enclosue kit (for 00000000-100000000u-R & Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000002 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Coud not confirm uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Coud not confirm complaint however there is physical damage the charging port. To correct the issu foud we will replace the Front Case Assembly. The Logs are being sent to RDT to confirm any uissus. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006591 Wilbu, Phillip The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Cutomer called back and we tried a new lead set on the device and place evie on simuator at rate of 61 and Tempu is reading accuately. Biomed said he does notice a slight glitch in the first QRS that displays on screen so it was decide to send uit to bench for evalution. Diagnostic performed by Engineer : Problem Description: "00000000ead isn't reading accuately read heart rate at 81, bu heart rate was 1700000000 Tempu Pro: "00000000ead isn't reading accuately read heart rate at 81, bu heart rate was 1700000000 Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 8/29/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Biome called to report Tempu Pro: "00000000ead isn't reading accuately read heart rate at 81, bu heart rate was 1700000000 Biomed to send Device Logs and to obtain another known good lad set to try on uit. Will call cutomer back when ##is received to trouleshoot. Was PRS ued: No PRS Permission: Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: "00000000ead isn't reading accuately read heart rate at 81, bu heart rate was 1700000000 uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: "00000000ead isn't reading accuately read heart rate at 81, bu heart rate was 1700000000Follow u Requred : Biome reports Tempu Pro: "00000000ead isn't reading accuately read heart rate at 81, bu heart rate was 1700000000 While on the phone had cutomer try a new lead set on the device, place uit on simuator at rate of 61 and Tempu is reading accuately. Biomed said he does notice a slight glitch in the first QRS that displays on screen so it was decide to send uit to bench for evalution. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: ALEC HADLEY IP & Issu: 959100000000200000000floor lead isn't reading accuately read heart rate at 81, bu heart rate was 1700000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: NA Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-09-13 01:01 | No
506 | Yes | System not booting up all they way, stuck at 0:0 | System not booting up all they way, stuck at 0:0 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000022 28-Sep-2000000003 Kolbinger Patrick The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : IPPC hardware or Image Disk problems. When was the first occurence? Within... : 12-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : replaced host pc and hard drives load sw Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD 1.00000000000000000000000000000000800000000958881,u uuu(OR Table) R 8.1.17 sw pack Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006676 16-Sep-2000000003 Kolbinger Patrick The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : replace HD in IPPC Diagnostic performed by Engineer : coud not talk to system Problem Description by Engineer : Problem description by engineer : IPPC hardware or Image Disk problems. When was the first occurence? Within... : 12-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006623 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Frontal IPPC disk or hardware issus. Repair Action : Technical event logs from 00000000/00000000 indicates problem with IPPC and or Image HDD's. Recommend to Replace IPPD drives and PC. Resoluion : Created onsite 000000000000000000000000001llow u Requred : uLL# and/or description of uLL: Frontal IPPC disk or hardware issus. Repair Action : Technical event logs from 00000000/00000000 indicates problem with IPPC and or Image HDD's. Recommend to Replace IPPD drives and PC. Technician Name : Johansson Anders Problem Description by Engineer : Problem description by engineer : IPPC hardware or Image Disk problems. When was the first occurence? Within... : 12-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Cutomer states that the system is not booting u all they way, stuk at 00000000:00000000. Also, had several Image Disk Ful uLLmessages even thouh cases was deleted off the system. ----------------------------------------------------------------------------- Internal Remarks : Pat Kolbinger - 8.1 system, cutomer states stuk at 0000000000000000. No fleuvision not present. Cannot access FSF even in open profile. Told FSE to load scratch on euisting PC, if no sucess, replace host and disk. Internal Remarks : Recommended FSE replace 2 image drives with properly sized HDDs curently now 1 TB when ordered as FRu Curently houing 1500000000B Velociraptors WD. Internal Remarks : FSE on-site that is handled by biomed. Software for reload is not complete. Gave software PN for 8.1.17 ful. Also sent link to software load instrutions from InCenter. Host PC was replaced bu system is still stopping at 00000000:00000000. I think the software load is the issu. No curent ghost image available. ----------------------------------------------------------------------------- | 2023-09-13 01:06 | No
507 | Yes | One of the therapy cables on the DFM0 defib in clinical skills had broken | One of the therapy cables on the DFM0 defib in clinical skills had broken |  | 2023-09-13 01:33 | No
508 | Yes | uvula - Control panel wont lock / Resolved - Greased the solenoid pin and flange | uCVu - Control panel wont lock / Resolved - Greased the solenoid pin and flange |  | 2023-09-13 01:51 | No
509 | Yes | Usual alarm/Back up alarm: check vent | Unual alarm/Back up alarm: check vent | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000008 22-Sep-2000000003 Chisholm, Patrick The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561537200000000RP-PCBA,CPuGEN2, (SW 2.300000000, V600000000Tools & Equpment : MuTIMETER,200000000000000007-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000007-31 00000000:00000000:00000000 Resoluion : PC - 9/22/200000000- Replaced CPuboard and uit passed requred test sucessfuly. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: Backu failed alarm uer Impact: removed from service, needs repair Patient Impact: none reported SE2.300000000Diagnostic performed by Engineer : PC - 9/22/200000000- Down loaded log and verified uLLcode 1100000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006697 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : CON000000005981- uLL1100000000Backu alarm failed, cutomer requst repair uder contract. Recommended part for repair 00000000561537200000000- RP-PCBA,CPuGEN2, (SW 2.300000000, V600000000Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: Backu failed alarm uer Impact: removed from service, needs repair Patient Impact: none reported SE2.300000000Diagnostic performed by Engineer : Cutomer states while uit was in ue on patient it alarmed Backu alarm failed alarm message on screen, ventilation continud and uit was removed from patient withou incident, no patient harm reported, in biomed cutomer noted 1100000000uLLin event log, verified PM pcb part installed is pn 1118921, requsts onsite service per contract. Recommended part 00000000561537200000000- RP-PCBA,CPuGEN2, (SW 2.300000000, V600000000to correct issu. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Bob Pieper IP & Issu: 8000000009226/uuul alarm/Back u alarm: check vent Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-13 02:11 | No
510 | Yes | SPLIT defective mountings. Bench repair pending | SPLIT_defeu mountings. Bench repair pending |  | 2023-09-13 02:33 | No
511 | Yes | AED is emitting triple chirp, fails self -test | AED is emitting triple chirp, fails self -test | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007200000000Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below Robert Strohhacker 200000000Steeplechase Ln MuROE FALLS, OH 0000000062-1766 Resoluion : Closing case tracking#780000000056960000000000000000 SCHEDuED DELIVERY DATE Thusday 9/28/23 before 12:00000000 PM. ##the cutomer to confirm that the AED is 000000000000000000000000001king as it shoud. Cutomer has confirmed that he received the AED and 000000000000000000000000001king as it shoud. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping and will not pass the self-test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported the AED kept chirping and will not pass the self-test. Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and will not pass the self-test. Cutomer brouht new battery and pads, Cutomer removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Remove and reinsert battery to test. With approval from Kelly Champou to replace the cutomer AED. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. AED was approved to have warranty covered and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# A15H-00000000811 Battery Install Before# 000000002/2000000008 Lot#00000000200000000 Pads Eup Date: 12/19/2000000005 Pads Lot # C000000001923-00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007269 Parts Consued : 1.0000000000000000000000000000000056000000006281,HS1 (C B) Home Defib, u Engl ROW, Euch ----------------------------------------------------------------------------- | 2023-09-13 02:36 | No
512 | Yes | Not turning on | Not turning on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007115 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer ##d back and stated no service is needed per his comments. Tuns ou it was something with the cable I was uing to attempt to connect it to a PC since I don’t have a nul modem cable. Once I tested it with that cable upluged, it started 000000000000000000000000001king again. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: will not power u uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000 and it is not tuning on. Cutomer thinks it coud be the 35 volt rail issu. created an AA to have the FCO86600000000000000006 implemented and see if it solves the issu. Created an onsite 000000000000000000000000001----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006211 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : cutomer called back Follow u Requred : Cutomer has a V600000000 and it is not tuning on. Cutomer thinks it coud be the 35 volt rail issu. created an AA to have the FCO86600000000000000006 implemented and see if it solves the issu. Created an onsite 000000000000000000000000001Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: will not power u uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000 and it is not tuning on. Cutomer thinks it coud be the 35 volt rail issu. created an AA to have the FCO86600000000000000006 implemented and see if it solves the issu. Created an onsite 000000000000000000000000001----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Andrew 6189975311 IP & Issu: 100000000000000006800000000Not tuning on. May need to replace the PMPCBA board. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued ot RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: Cutomer Call Back Caller: Andrew 6189975311 IP & Issu: case 0000000022268692 additional tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: uable to complete debrief and followu RSE fields incomplete dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-13 02:41 | No
513 | Yes | cannot move carm angle or rotate with control in any direction , reboot no help | cannot move carm angle or rotate with control in any direuon , reboot no help | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003279 100000000ep-2000000003 Buh, Jonathan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - uM: Euam room - POST Geo mono modue Failed 12-Sep-2000000003 16:37:00000000 RMT - uM: Geo mono modue in euam room reported POST uLLof the FrontalAnguateRotateJoystick. 12-Sep-2000000003 16:37:00000000 When was the first occurence? Within... : 12-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: cannot move carm angle or rotate with control in any direction uer Impact: cannot move carm Patient Impact: room down SE 8.2.16.1 Resoluion : Replaced geo modue; ensued operability and retuned to cutomer for ue Parts Consued : 1.00000000000000000000000000000000800000000151715,GEO MODuE WP Kit ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007619 13-Sep-2000000003 Buh, Jonathan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - uM: Euam room - POST Geo mono modue Failed 12-Sep-2000000003 16:37:00000000 RMT - uM: Geo mono modue in euam room reported POST uLLof the FrontalAnguateRotateJoystick. 12-Sep-2000000003 16:37:00000000 When was the first occurence? Within... : 12-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: cannot move carm angle or rotate with control in any direction uer Impact: cannot move carm Patient Impact: room down SE 8.2.16.1 Diagnostic performed by Engineer : Checked the Geo modue operations, LAO/RAO rotation didn’t 000000000000000000000000001k; checked system diagnostics, failed bist; Follow u Requred : Replace Geo Modue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000075200000000lark, Byron The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - uM: Euam room - POST Geo mono modue Failed 12-Sep-2000000003 16:37:00000000 RMT - uM: Geo mono modue in euam room reported POST uLLof the FrontalAnguateRotateJoystick. 12-Sep-2000000003 16:37:00000000 When was the first occurence? Within... : 12-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: cannot move carm angle or rotate with control in any direction uer Impact: cannot move carm Patient Impact: room down SE 8.2.16.1 Diagnostic performed by Engineer : collected information from caller on problem and had log created. Called FSE to confirm shipping information Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - uM: Euam room - POST Geo mono modue Failed: 1. The POST of the Geo mono modue in the euam room Failed. Root Caue for RMT - uM: Geo mono modue in euam room reported POST uLLof the FrontalAnguateRotateJoystick.: 1. Geo mono modue in euam room reported POST uLLof the FrontalAnguateRotateJoystick. Repair Action : Service actions for RMT - uM: Euam room - POST Geo mono modue Failed: 1.1. Check Geo modue in euam room for any damage or flud spills. Replace Geo modue in euam room if problem persists. Service actions for RMT - uM: Geo mono modue in euam room reported POST uLLof the FrontalAnguateRotateJoystick.: 1.1. Check FrontalAnguateRotateJoystick for damage or flud spills. Replace Geo mono modue in euam room if problem persists. Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - uM: Euam room - POST Geo mono modue Failed: 1. The POST of the Geo mono modue in the euam room Failed. Root Caue for RMT - uM: Geo mono modue in euam room reported POST uLLof the FrontalAnguateRotateJoystick.: 1. Geo mono modue in euam room reported POST uLLof the FrontalAnguateRotateJoystick. Repair Action : Service actions for RMT - uM: Euam room - POST Geo mono modue Failed: 1.1. Check Geo modue in euam room for any damage or flud spills. Replace Geo modue in euam room if problem persists. Service actions for RMT - uM: Geo mono modue in euam room reported POST uLLof the FrontalAnguateRotateJoystick.: 1.1. Check FrontalAnguateRotateJoystick for damage or flud spills. Replace Geo mono modue in euam room if problem persists. Technician Name : Clark, Byron ----------------------------------------------------------------------------- | 2023-09-13 03:56 | No
514 | Yes | Retrospective-Unable to Retrieve Data /Unable to Retrieve Full Disclosure/No Trouble-Working to Spec-Other/Provided Info | Retrospeu-Unable to Retrieve Data /Unable to Retrieve Full Disclosure/No Trouble-Working to Spec-Other/Provided Info | 000000000000000000000000001O##: 000000000000000000000000001000000000000000079500000000udhararaj,Vidhya The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : na Problem Description by Engineer : Cutomer Fuction/Role: biomedical engineer Device ue: ou of ue Eupected and ueupected Resuts:8am patient incident occured. Cutomer 000000000000000000000000001ld like to retrieve data uer Impact: nil Patient Impact: nil SE na Diagnostic performed by Engineer : follow u requred ----------------------------------------------------------------------------- | 2023-09-13 05:56 | Yes
515 | Yes | u not power on | 0R/u FD0IPC not power on | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the mechanical part coud not be started and engineers needed to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: IPC host does not power on Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the voltage inpu and oupu of the power modue in the IPC are present and stable. If there is a problem, replace the IPC. 2. Check whether the motherboard BIOS battery is normal. 3. Check whether the IPC software is normal. Try to reinstall the software. 00000000Check the IPC to SCC and TS. Whether the wiring is normal 5. Check the machine uLLand perform faut repairs Recommended spare part uuand name: " 00000000800000000931363 GEO IPC Ivybridge with ZMP CAN and IO" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 12.00000000.2000000003 17:100000000100000000TC (860000000013300000000SuT. Ready for Clinical ue local time" : " 2000000003-00000000-12 11:300000000 Internal Remarks: 13.00000000.2000000003 00000000:00000000:00000000uC (860000000013300000000 RDI Accuacy: Yes" $Q7:Confirmation reported problem, The device cannot be started properly. $Q8:The device cannot be started properly.It was foud that IPC did not start on site, and the measuement foud that the front power suply fue F12 was buned. After replacement, IPC started normally, bu the mechanical part of the equpment coud not move normally. Continu to check the measuement and find that the NC has inpu bu no oupu. The NC is fauty(00000000800000000000000002663 PS u_COLL_ID_uIT). $Q9:Replace The NC(00000000800000000000000002663 PS u_COLL_ID_uIT). $Q100000000Test and Inspection data refer to PA, Do a few patient, Observation of a certain period of time. The system is ok. Device retuned to ful fuctionality. $Calibrated Tools name (NA) S/N (NA) Neut Calibration date (NA) Safety Qustions: Case: 00000000222600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 13.00000000.2000000003 00000000:00000000:37 uC (860000000013300000000Department: Interventional Department Contact person: uie ZhanyuContact nuber: 177160000000071800000000 Symptom: The device cannot start normally. Soluion/Repair: The device cannot start normally.It was foud on site that the IPC did not start u, and the measuement foud that the front-end power suply fue F12 was buned. After replacement, the IPC started u normally bu the mechanical part of the equpment coud not move normally. Continud inspection and measuement foud that the NC had inpu bu no oupu.Determine NC faut.Replace NC. The equpment 000000000000000000000000001ks normally after service. SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/u uuuFD200000000T2 Activities:Repair Action: 00000000800000000000000002663 PS u_COLL_ID_uIT ----------------------------------------------------------------------------- | 2023-09-13 06:16 | No
516 | Yes | CM - SELF TEST FAILED | CM - SELF TEST FAILED | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : CM - SELF TEST FAILED SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 000000002226000000008 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-13 07:11 | No
517 | Yes | SA:0/u FD0startup issue | SA:0R/u FD0startup issue | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: FSE Repair Report: uable to boot Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The FSE was on hand and discovered that L1 had tripped.After closing, it was foud that the IPC did not start Remote diagnostics and test resuts: Supected IPC failue PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended that engineers go to the site to check the power suply of the IPC host.2. Replace the IPC host Recommended spare parts uuand designation: "PPD:00000000800000000000000007813 GEO Q35 IPC COA" Problem Reported by Cutomer: uable to enter the system after booting SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 12.00000000.2000000003 12:16:00000000uC (86000000006552) Service content (requred): IPC power suply failue, recovery after reconnecting the power cord Service resut (requred): The device is runing normally Other important matters: None Neut Step Action: Retun visit sucessfu SAP Transfer: Case Transferred to SAP Internal Remarks: 13.00000000.2000000003 12:00000000:00000000uC (8600000000000000001)"RDI Accuacy: Yes" $Q7: (Confirmed reported problem. Startu failed.) $Q8: (Repair the power of IPC, the device retuned to normal.+IPC problem.) $Q9: (Repair the IPC.) $Q100000000 (Test and inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: (Calibrated Tools refer to PA.) For Mainland China CV / uR, CT engineers only. Please fill ^Curent Tue installed date: N/A ^Curent Tue uage: N/A SAP Transfer: Case Transferred to SAP Internal Remarks: 12.00000000.2000000003 12:16:55 uC (86000000006552) "RDI Accuacy: Yes" $Q7: (Confirmed reported problem. Startu failed.) $Q8: (Repair the power of IPC, the device retuned to normal.+IPC problem.) $Q9: (Repair the IPC.) $Q100000000 (Test and inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: (Calibrated Tools refer to PA.) For Mainland China CV / uR, CT engineers only. Please fill ^Curent Tue installed date: N/A ^Curent Tue uage: N/A Euternal Remarks: 12.00000000.2000000003 15:10000000000000000 uC (8600000000000000001) Service content (requred): IPC power suply failue, restored after reconnecting the power cord Service resut (requred): The device is runing normally Other important matters: None Neut Step Action: 000000005881135600000000FSE Shang renhu calls Neut Step Action: PPD00000000800000000000000007813 1 Arranged 00000000to the hospital FSE: Shang Renhu -86-15881135600000000SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000222600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Qingfeng Gu Last udated by: Qingfeng Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-13 08:26 | No
518 | Yes | Defibrillator/Monitor/Unable to charge, flashing light | 0/Effi0 DFM0 Defibrillator/Monitor/Unable to charge, flashing light | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002226000000005 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: uable to charge, light flashes Yuin First People's Hospital Wuiqiao Commuity Health Service Center Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: ASP Honglejia Follow u Neut Step Action: uuChungfeng 1552137100000000 ----------------------------------------------------------------------------- | 2023-09-13 08:31 | No
519 | Yes | Premature battery wear. - log analysis | Premature battery wear. - log analysis | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000700000000onezawa, Hiroki The system does not meet the specification for the performed service and is not retuned to ue Resoluion: "I canceled the repair and will retun the AED. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." September 22, 2000000003 Fuua Denshi Kiki Hanbai Co.,Ltd. I sent the equpment along with a report to the on-site person in charge. Internal Comments: [Final check] I checked the operation following the 000000000000000000000000001k based on the set of orders. Repair canceled. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: I received the AED and performed a battery test, and I confirmed that it passed the test.The battery euamination resuts show no abnormality in the oupu voltage and there was no wear. I analyzed the operation record of the AED and foud that an alarm had occured du to the test resut being ouside the passed range when a self test was ru on a fuction test of the high voltage element (uuuitor) inside the AED. "The AED is fauty, so it will be necessary to replace the AED. " ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 Prematue battery wear. Lot No. Adut pad 220000000000000000-0000000055 Battery 922500000000 ----------------------------------------------------------------------------- | 2023-09-13 08:41 | No
520 | Yes | System hang at 0:0 | System hang at 0:0 | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 000000002226000000009 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Wendy Cheug Last udated by: Wendy Cheug 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : System hang at 00000000:00000000 FSE Notes : This PO is a Reference Nuber only ----------------------------------------------------------------------------- | 2023-09-13 08:41 | No
521 | Yes | SA:0/u FD0startup failure | SA:0R/u FD0startup failture | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: SA:00000000R/u uuuFD100000000SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022260000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Overview of cutomer problems: Startu failue Detailed description of cutomer problem or faut phenomenon: Power on 00000000, the large screen does not display Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Normal after manully switching fleuvision PC on and off Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Replacing fleuvision PC BIOS battery Notes on cutomer requrements: Please contact the cutomer for on-site repair Euternal Remarks: 30000000000000000.2000000003 00000000:11:00000000uC (860000000063600000000 Service content (requred): Power on 00000000, the large screen does not display; replace the mainboard battery of the large screen compuer; replace the uuumodue. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: none. Internal Remarks: 30000000000000000.2000000003 00000000:11:00000000uC (860000000063600000000 "RDI Accuacy: YES " $Q7: (Confirmed reported problem, startu failue) $Q8: ( Check and test fleuvision PC, fleuvison PC can't start u sucessfuly, mainboard BIOS batter need to be replaced; uuumodue can't open fleuvision control, uuumodue have problem, need to be replaced ) $Q9: (Replaced fleuvision PC BIOS batter, replace uuumodue - TSM-B.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11:N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage: N/A SAP Transfer: Service Order is Locked by another uer!: SAP Transfer: No Operation corresponds to the labor ID: ----------------------------------------------------------------------------- | 2023-09-13 08:51 | No
522 | Yes | System Failure | System Failure | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000SAMAPORN ROMPHOREE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Dealer Device ue: Intrepid Monitor/Defibrillator Eupected and ueupected Resuts: System Failue uer Impact: No impact Patient Impact: No impact SE 1.00000000.500000000Diagnostic performed by Engineer : - 0000000056000000001600000000,Intrepid PCCRD System on Modue (SOM) - Order part ship to dealer Refer 00000000000000000000000000100000000000000000000000000000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000 Parts Consued : 1.0000000000000000000000000000000056000000001600000000,Intrepid PCCRD System on Modue (SOM) ----------------------------------------------------------------------------- | 2023-09-13 08:51 | No
523 | Yes | Touchscreen issue | Touchscreen issue | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ASP wueijia Follow u Problem Reported by Cutomer: Screen not 000000000000000000000000001king Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: dear: Engineer Gan Lu13300000000370000000000000000 has contacted the hospital, thank you Neut Step Action: Yuei District Hospital of Traditional Chinese Medicine Safety Qustions: Case: 000000002226000000006 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-09-13 09:11 | No
524 | Yes | exposure not possible | 0R/u FD0euposure not possible | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 12.00000000.2000000003 0000000052:16 uC (8600000000000000003) Check the system log, check the IPPCBIOS battery failue, replace the BIOS battery, and perform related detection fuctions normally. All fuctions of the device retun to normal Problem Reported by Cutomer: uable to save the pictue, engineers need to arrive as soon as possible SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 12.00000000.2000000003 00000000:00000000:00000000uC (8600000000000000003) $Q7: (Confirmation & Description of problem) Confirmed reported problem： 000000000000000000000000001ld not eupoue$Q8: (Trouleshooting & Resuts) uLLmessage: image disk ful $Q9: (Repair steps) check the IPPC, bios baterry was defected, Replce the baterry and make iPPC test was sucessfu. $Q100000000 Test and inspection : Refer to PA $Q11 Calibrated Tools:Refer to PA Safety Qustions: Case: 0000000022260000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: Euposue not possible Detailed description of cutomer problem or faut phenomenon: uable to save image Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: RMT - IPu IPPC failed to boot u - Frontal Requst an engineer to attend Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Replacing IPPC bios battery Notes on cutomer requrements: Please contact the cutomer for on-site repair SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-13 09:36 | No
525 | Yes | QAA0/u00Startup failed | QAA0K/uStartup failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: QAA00000000K/u7 M200000000 T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the screen went black when the machine was tuned on, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: black screen when the machine is tuned on Remote diagnosis and test resuts: Restart has no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the monitor wiring and power suply are normal. 2. Check whether the sute PC starts normally. 3. Check whether the power suply part is normal. 00000000Check the machine uLLreport and perform faut repairs. Recommended spare parts uuand designation: NA Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Internal Remarks: 12.00000000.2000000003 12:27:37 uC (86000000009500000000) $Q7: (Confirm the reported problem, cannot start the system) $Q8: (Check the wire and find the ippc problem.) $Q9: (reconnect the cable, mainboard and test system ok. ) $Q100000000 (Test and inspection data refer to PA, Device retuned to ful fuctionality.) $Q11: Calibrated Tools:N/A Safety Qustions: Case: 0000000022260000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 12.00000000.2000000003 12:27:300000000TC (86000000009500000000)"SuT. Ready for Clinical ue local time" : " 2000000003-00000000-12 20000000000000000" Service content (requred): The machine cannot be started, plu and uplu cables, compuer boards, and restore to normal Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-13 10:46 | No
526 | Yes | Single beep, battery insertion test failed on new device - log analysis requested | Single beep, battery insertion test failed on new device - log analysis requested | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001200000000 Yonezawa, Hiroki The system does not meet the specification for the performed service and is not retuned to ue Resoluion: DEFOA Application was made and it was approved. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "September 22, 2000000003 Fuua Denshi Tokyo Hanbai I contacted the on-site person in charge. Internal Comments: [Final check] I checked the operation following the 000000000000000000000000001k based on the set of orders. Defoa. I applied for it and approved it. Diagnostic performed by Engineer: I received the AED and performed a battery test, and I confirmed that it passed the test.After that, I monitored the AED in standby mode, bu no alarm souded. Analysis confirmed that the battery fitted test failed three times in a row du to low battery statu, even thouh it was a new one. The battery voltage is normal, and it is difficut to identify the caue on the record. The main uit or the battery was fauty, so I applied for DEFOA. Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Reguatory Qustions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 Following a self test, a message abou a battery problem was made and the AED beeped. Lot No. Adut pad C000000003000000003-00000000 Adut pad C000000003000000003-00000000 Battery 00000000839P ----------------------------------------------------------------------------- | 2023-09-13 10:51 | No
527 | Yes | SA:0u FD0Startup failed | SA:0/u FD0Startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: uable to boot Neut Step Action: Called SS OuShurong 1360000000060000000095 Don’t go throuh Green Pass, go for qutation, pu qutation Q. Already informed by ##Safety Qustions: Case: 000000002226000000007 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 23.00000000.2000000003 1000000001:32 uC (8600000000000000001) Service content: Faut 1: The machine cannot be tuned on, and it retuns to normal after repeated restarts. After inspection, it may be caued by a defective MPDuof the power control modue. It is recommended to replace the MPDu Faut 2: The machine cannot move. Check the backgroud log and report an uLL POST SafetyPOST Failed, it is detected that there is a problem with the motion control modue GEO Modue. It is recommended to replace the GEO Modue. Service Resuts: Equpment leaves site fuly fuctional. Other important matters: None yet. FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA Neut Step Action: 00000000000000007338862, Department of Interventional Radiology, u Liao, call in. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 23.00000000.2000000003 1000000001:300000000TC (8600000000000000001) RDI Accuacy: Yes $Q7: (Confirmation & Description of problem) Confirmed reported problem, The machine cannot start u and has been restarted repeatedly to restore normal operation $Q8: (Trouleshooting & Resuts) Faut 1: The machine cannot start u and has been restarted repeatedly to restore normal operation. After inspection, it may be caued by a poor MPDuof the power control modue. It is recommended to replace the MPDu Faut 2: The machine cannot move. Check the backend log and report an uLL POST SafetyPOST Failed. It is detected that there is a problem with the motion control modue GEO Modue. It is recommended to replace the GEO Modue. $Q9: (Repair steps) Faut 1: The machine cannot start u and has been restarted repeatedly to restore normal operation. After inspection, it may be caued by a poor MPDuof the power control modue. It is recommended to replace the MPDu Faut 2: The machine cannot move. Check the backend log and report an uLL POST SafetyPOST Failed. It is detected that there is a problem with the motion control modue GEO Modue. It is recommended to replace the GEO Modue. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:N/A; Euternal Remarks: 100000000000000002000000003 00000000:51:500000000TC (8600000000000000001) Service content: Replace GEO modue, udate GEO modue software, FD movement retuns to normal, bed cannot move, backgroud uLL POST failed: AD7$., Height drive; it is recommended to replace the Height motor drive modue. Service Resuts: Equpment leaves site fuly fuctional. Other important matters: None yet. Internal Remarks: 100000000000000002000000003 00000000:51:57 uC (8600000000000000001)$Q7: (Confirmation & Description of problem) Confirmed reported problem, The machine cannot start u and has been restarted repeatedly to restore normal operation $Q8: (Trouleshooting & Resuts) Replace the GEO modue and udate the GEO modue software. The FD movement retuns to normal, bu the bed cannot move. The backgroud uLLis: POST failed: AD7 $., Height drive; It is recommended to replace the Height motor drive modue. $Q9: (Repair steps) Replace the GEO modue and udate the GEO modue software. The FD movement retuns to normal, bu the bed cannot move. The backgroud uLLis: POST failed: AD7 $., Height drive; It is recommended to replace the Height motor drive modue. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:N/A; ----------------------------------------------------------------------------- | 2023-09-13 11:11 | No
528 | Yes | Software Corrupted | Software Corrupted | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003881 27-Sep-2000000003 Taviya, Rahu The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Power Suply fauty supected need to be replace also SOM PCB fauty need to be replace Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Resoluion : Today visit site and replace Power Suply and SOM PCB and machine foud ok all test passed and machine handed over to cutomer in 000000000000000000000000001king condition. Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY SOM 1.0000000000000000000000000000000056000000008951,DFM100000000 AC Power Modue Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 MuTIMETER,2000000003-11-1000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008921 15-Sep-2000000003 Bhargav, Khatri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : find ou that som pca fouty Follow u Requred : find ou that som pca fouty Tools & Equpment : MuTIMETER,2000000003-11-1000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008828 Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : fse visit requred Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : fse visit requred ----------------------------------------------------------------------------- | 2023-09-13 11:56 | No
529 | Yes | FAILURE IN IPPC, disk 0ails | FAILURE IN IPPC, disk 0fails | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001500000000100000000ep-2000000003 Antonio Garcia Herrera The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Problem Description by Engineer: Cutomer Fuction/Role: Suervisor Device ue: yes Eupected and ueupected Resuts: OK uer Impact: No Patient Impact: No SE 8.1 Resoluion: Image acqusition failue. The failue is detected in one of the IPPC image disks. It is replaced. Tests OK. Correct operation. ----------------------------------------------------------------------------- | 2023-09-13 12:07 | No
530 | Yes | 0seconds delay in alarm sounds from the central station. All Alarms. Surveillance alarm sounds. | 0seconds delay in alarm sounds from the central station. All Alarms. Surveillance alarm sounds. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001800000000 Sudhararaj,Vidhya The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomedical engineer Device ue: ou of ue Eupected and ueupected Resuts: 200000000seconds delay in alarm souds from the central station. uer Impact: nil Patient Impact: nil SE na Resoluion : Cutomer restarted the central and issu hasn't occured since then ----------------------------------------------------------------------------- Internal Remarks : Left message for cutomer to call back ----------------------------------------------------------------------------- | 2023-09-13 12:31 | No
531 | Yes | ray not possible | uuray not possible | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine cannot see throuh and requres engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: generator is buy Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether PoInt is normal. 2. Check whether HIVO is normal. 3. Check the machine uLLand perform faut repairs. Recommended spare parts uuand name: uu 00000000210000000013887, Description: Power Inverter (PoInt) Certeray &00000000210000000011828 HiVo high voltage transformer" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Case: 0000000022265500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/u7 M12/cannot emit rays Neut Step Action: PPD00000000210000000013887*1 Arranged for 200000000delivery to the venu FSE ui Minghu-86-17700000000000000003527 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 22.00000000.2000000003 1000000005:00000000 uC (8600000000900000000)"RDI Accuacy: Yes" $Q7: (Confirmed the report problem that the 7M12 can't euposue and fluroscopy. ） $Q8: ( 1.Restart the system, bu it still not ok. 2. Check the log file, it reports the tue curent ou of range uLL 3. Replace the power inverter, adjut the generator, and retun the device to ful service. & So the power inverter caued the reported problem.) $Q9: ( 1 .Replace the power inverter, uu00000000210000000013887, adjut the generator, the device retun to ful service. ） $Q100000000 Test and inspection data refer to PA. Device retuns to ful service. $Q11: Calibrated Tools:N/A. Euternal Remarks: 22.00000000.2000000003 1000000005:00000000 uC (8600000000900000000) Service content: The equpment cannot be euposed. After replacing the point inverter, the equpment retuned to normal operation; Service resut: the equpment 000000000000000000000000001ks normally; Euternal Remarks: 22.00000000.2000000003 12:0000000000000000 uC (86000000009000000009) Service content: The equpment cannot be euposed. After replacing the point inverter, the equpment retuned to normal operation; Service resut: the equpment 000000000000000000000000001ks normally; Internal Remarks: 22.00000000.2000000003 12:51:56 uC (86000000009000000009) "RDI Accuacy: Yes" $Q7: (Confirmed the report problem that the 7M12 can't euposue and fluroscopy. ） $Q8: ( 1.Restart the system, bu it still not ok. 2. Check the log file, it reports the tue curent ou of range uLL 3. Replace the power inverter, adjut the generator, and retun the device to ful service. & So the power inverter caued the reported problem.) $Q9: ( 1 .Replace the power inverter, uu00000000210000000013887, adjut the generator, the device retun to ful service. ） $Q100000000 Test and inspection data refer to PA. Device retuns to ful service. $Q11: Calibrated Tools:N/A. ----------------------------------------------------------------------------- | 2023-09-13 13:01 | No
532 | Yes | Low pressure | Low pressure | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002329 Vega Zevallos, Juissa Kriss The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: Client indicates poor presión uer Impact: uknown Patient Impact: uknown SE N/A Resoluion: Se intercambia carcasa suerior, cable de alimentación, bandeja de la cuierta de la CPuy pines del fanador. Se realizo actulización de software a la última versión 3.200000000y se calibró los sensores de fluo de aire y ouigeno. Siguendo el procedimiento de testeo y verificación para este equpo, se restaua la configuación del dispositivo a fábrica. Se han efectudo todas las verificaciones necesarias qu permiten certificar la restituión a las condiciones anteriores a la avería. Para garantizar el correcto fucionamiento del equpo, les recomendamos qu uilicen solo los accesorios y fugibles aprobados por u. The cover top, Power Cord, CPucover tray and pin puh fan mout were replaced. The software was udated to the latest version 3.200000000and the air flow and ouygen sensors were calibrated. Following the test & verification procedue, the device was restored to factory settings. All test and verification procedues necessary to certify the retun of the device to 000000000000000000000000001king conditions were carried ou satisfactorily. To assue proper device performance, we recommend uing only accessories and eupendables approved by u. Parts Consued: 1.00000000000000000000000000000000561511731,RP-Pin, Puh, Fan Mout, Pkg/00000000V600000000 1.000000000000000000000000989800000000612211,Power Cord3 WireRt AngleEu100000000 1.00000000000000000000000000000000561531961,RP-Cover, CPuTray, V6uu 1.00000000000000000000000000000000561531971,RP-Cover, Top, APLW, V6uu/V600000000Tools & Equpment : ELECTRICAL SAFETY TESTER, 20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER, 2000000003-12-13 00000000:00000000:00000000 PRESSuE CALIBRATION TOOL, 200000000000000005-17 00000000:00000000:00000000 FLOW METER, 2000000003-12-00000000 00000000:00000000:00000000 FLOW METER, 2000000003-12-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-13 13:36 | No
533 | Yes | motorized movement not available and stand movement partially available | motorised movement not available and stand movement partially available | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009533 13-Sep-2000000003 Andy Fairless The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: No C Arm movements uer Impact: NA Patient Impact: NA SE7.2 Resoluion : Stand ceiling rails cleaned. Stand longituinal position, 000000000000000000000000001king area & curent calibrations carried ou. System tested & left 000000000000000000000000001king ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009389 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : RSW indicates Service Action 2: Check condition of motor ruber wheel. Service Action 3: Check and/or adjut the tension of the spring. Service Action 00000000 Carry ou the Beam longituinal curent calibration via FSF. Buadvice for this uLL: Replace stand long potmeter and check also if the stand long brakes soldering points make no short with the ceiling rails via metal particles clotted together on the soldering point. Case OK990000000000000000182000000006_Swiss_ClauioPinchera_GeoIsseu Z-rotation pot.meter was loose. The cable to the Z-rotation pot wasn't fiued at all (screw broken . Case 0000000010000000000000000296_FOOTHILLS_ MEDICAL_CENTRE_Juian LouC-Arm Stutering at all angles: Nuse side track roller bearing was discovered very loose and catching the c-arc track, which prevents the smooth c-arc movement Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: No C Arm movements uer Impact: NA Patient Impact: NA SE7.2 Diagnostic performed by Engineer : position uLL motor enable feedback uLL----------------------------------------------------------------------------- | 2023-09-13 13:46 | No
534 | Yes | Therapy interface damaged | Therapy interface damaged | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000222657700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No Euternal Remarks: 00000000.11.2000000003 00000000:00000000:11 uC (86000000001999) Department: Cardiology Ward Contact person: Chen Gong Faut problem: uable to discharge Trouleshooting: Replace the handle interface Service resut: The device retuns to normal Internal Remarks: 00000000.11.2000000003 00000000:00000000:100000000TC (86000000001999) $Q1: (N) $Q2: (N) $Q3: (N) $Q00000000(NA) $Q5: () $Q6: (The paddle cannot release power) $Q7: (Confirmed reported problem. The paddle cannot release power.) $Q8: (Performed visul inspection: According to the manul, the caue of the failue is damage to the power connection.) $Q9: (The power connection has been replaced) $Q100000000 Test and inspection data refer to PA. Device retuned to fuly fuctional. $Q11:Calibrated tools name refer to PA. SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: SA:00000000N/uDFM100000000/ Neut Step Action: PH Chen Shengju 1581000000005000000001 Call for repair Cutomers who have retuned visits SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Treatment port damage Repair Action: ---------RA-----REPAIR ACTION--------------- IF PART REQuRED - 12 NC RECOMMENDED: 0000000056000000008971 DFM100000000 ASSY Therapy Receptacle RECOMMENDED REPAIR ACTION: Replace treatment port ----------------------------------------------------------------------------- | 2023-09-13 13:47 | No
535 | Yes | DFM0 can't boot up | DFM0 can't boot up | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: system board Safety Qustions: Case: 0000000022265819 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: liuuiaohu Last udated by: liuuiaohu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: ASP Honglejia follow Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-09-13 13:51 | No
536 | Yes | DFM0 in permanent reboot loop | DFM0 in permanent reboot loop | Cutomer comment: u Skok sends delivery note by ##Internal Remarks: Feedback created uDFM100000000 Defibrillator/Monitor/ cae 00000000222658300000000eptember 11, 2000000003 Defi independently switched to service mode over the weekend and requsted the service pass000000000000000000000000001d. There was no service pass000000000000000000000000001d stored, so any PW coud be entered. Only then was it possible to switch back to clinical operation. There was no patient harm. 00000000/12/2000000003 Device now in permanent reboot loop. Device not ready for ue. Internal Remarks: IPCR 200000000300000000-120000000028 created Internal Remarks: CN32670000000092 is stored on another accout. Accout: Kliniken Schmieder Heidelberg GmbH / 9000000000000000000000000Internal Remarks: The reported device CN32670000000092 is assigned to another cutomer. Please fill ou the form and send a pictue with a recognizable 000so that we can change the location. The device is also serviced by SMS Service. 000000000000000000000000001ld youlike to have the trouleshooting done there or send the device to ou 000000000000000000000000001kshop? Internal Remarks: After consutation, the cutomer commissions SMS to carry ou the repair ----------------------------------------------------------------------------- | 2023-09-13 13:56 | No
537 | Yes | Table lock isn't engaging and they have a warning sign | Table lock isn't engaging and they have a warning sign | 000000000000000000000000001O##: 000000000000000000000000001000000000000000096800000000100000000ep-2000000003 Simon Fallows The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: No table / C Arm movements uer Impact: NA Patient Impact: NA SE8.1.17 Resoluion : Fitted new GEO IPC. Reloaded SW to PC and table drives. Tested and left in 000000000000000000000000001king order. Parts Consued : 1.00000000000000000000000000000000800000000931363,GEO IPC Ivybridge with ZMP CAN and IO Tools & Equpment : MuTIMETER,2000000003-00000000-28 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-10000000000000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009500000000Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to replace Geo PC and configue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: No table / C Arm movements uer Impact: NA Patient Impact: NA SE8.1.17 Diagnostic performed by Engineer : Geo PC uLLon startu ----------------------------------------------------------------------------- Cutomer comment : called Catrin and confirmed planning for tomorrow am ----------------------------------------------------------------------------- | 2023-09-13 14:11 | No
538 | Yes | Low pressure | Low pressure | 000000000000000000000000001O#Nuber: 00000000000000000000000000100000000000000002300000000Zevallos Zevallos, Juissa Kriss The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: uknown. Eupected and ueupected Resuts: Cutomer indicates suport. uer Impact: uknown. Patient Impact: uknown. SE 2.100000000 Tools & Equpment: ELECTRICAL SAFETY TESTER, 20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER, 2000000003-12-13 00000000:00000000:00000000 PRESSuE CALIBRATION TOOL, 200000000000000005-17 00000000:00000000:00000000 FLOW METER, 2000000003-12-00000000 00000000:00000000:00000000 FLOW METER, 2000000003-12-00000000 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000056153200000000,RP-Foot Kit, Bottom,V6uu/V600000000NIC, Pkg/00000000 1.00000000000000000000000000000000561500000000991,RP-Filter, Air Inlet, Pkg/5,V600000000V600000000Plu 1.00000000000000000000000000000000561500000000300000000,RP-Filter, Cooling Fan, Pkg/5, V600000000 1.0000000000000000000000000000000056153300000000,RP-Base Assembly, Non-Eng Wht, V600000000 1.000000000000000000000000989800000000612211, Power Cord3 WireRt AngleEu100000000 1.00000000000000000000000000000000561531961,RP-Cover, CPuTray, V6uu Resoluion: Foot kit bottom, air inlet filter, filter cooling fan were euchanged, Power Cord, base Assembly and cover tray CPu Following the test and verification procedue for this equpment, the configuation of the device was restored to the manuactuer. We have performed all the necessary checks that allow for certifying the re-euamination of the conditions prior to the averia. To gurantee the correct operation of the equpment, we recommend youue the groud license and problems reported by u. ----------------------------------------------------------------------------- | 2023-09-13 14:46 | No
539 | Yes | Connection wireless foot switch | Conneuon wireless foot switch | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002586 September 22, 2000000003 Strohmeyer, Andre The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MTA Device ue: ouside Eupected and ueupected Resuts: uer Impact: Patient Impact: Installation of the wireless foot switch SE Resoluion: The installation of the wireless foot switch was performed and completed according to the applicable manuactuer's specifications. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000700000000681,CONN CORD TSM 9M 1.00000000000000000000000000000000800000000276661,Cable Control Modue Mono & Geo 8m 1.00000000000000000000000000000000800000000700000000681,CONN CORD TSM 9M 1.00000000000000000000000000000000800000000276661,Cable Control Modue Mono & Geo 8m ----------------------------------------------------------------------------- Internal Remarks: Via u Bader / PH Hello everyone T000000000000000000000000001wireless foot switches with base arrived in the Stutgart Clinic. For the hybrid OP on 200000000 Please also make an order for Those for the foot switch installation. Angio will do Dominik on Thusday 100000000 and is already schedued. Video splitter also arrived, and also in Medical Technology. Please contact Tim Hartleb abou the orders, he knows where the material is... 00Dominik & Dominik: Please also for the splitter history all possible adapters DisplayPort HDMI/Mini DisplayPort... take along. Thanks & thanks, Michael ----------------------------------------------------------------------------- | 2023-09-13 14:51 | No
540 | Yes | Duplicate to PR#00 CR- uLL07 | Duplicate to PR#000. CR- uLL07 |  | 2023-09-13 15:06 | No
541 | Yes | alternator bu. | uLLGenerator bu. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002513 15-Sep-2000000003 Pawel Kierzek The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Dr. Device ue: IGT-S. Eupected and ueupected Resuts: uable to make radidation. uer Impact: None. Patient Impact: none. 00 Resoluion: IPS and POINT were replaced in the cabinet with the generator (cabinet E). Necessary calibrations and verifications were performed. A correct operation was checked with a positive resut. The faut was repaired. The device was handed over for clinical ue. Parts Consued: 3.0000000000000000000000000000000080000000000000000900000000,PDuFues Generator. 1.0000000000000000000000000000000021000000002000000002,Power Inverter EVR (PoInt EVR) Certeray. 1.00000000000000000000000000000000210000000020000000000000000,IPS Certeray R5 Tools & Equpment : DOSE-KVP-LIGHT METER, 2000000005-00000000-31 00000000:00000000:00000000. DOSE-KVP-LIGHT METER, 2000000008-00000000-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000008-0000000000000000:00000000:00000000 MuTIMETER, 2000000003-11-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000060000000013-Sep-2000000003 Michal Sadowski The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: in ue. Eupected and ueupected Resuts: the generator does not 000000000000000000000000001k, no radiation. uer Impact: no. Patient Impact: yes. SE 2.2.u. Diagnostic performed by Engineer: Diagnostics of the device malfuction. Damaged power suply elements were foud in the generator and M cabinet. Probable caue: over-voltage. Ordering parts. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000057 Michal Sadowski The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: in ue. Eupected and ueupected Resuts: Generator Buy starting u uLL uer Impact: no Patient Impact: yes. SE 2.2. Diagnostic performed by Engineer: The faut was consuted with the uer. Check of the correct power suply on the uer's side. Correct power suply was confirmed. On-site visit and futher diagnostics are necessary. ----------------------------------------------------------------------------- | 2023-09-13 15:06 | No
542 | Yes | Touchscreen doesn't work very well even after calibrated, it needs to be replaced. | Touchscreen doesn't work very well even after calibrated, it needs to be replaced. | 000000000000000000000000001O#Nuber: 00000000000000000000000000100000000000000002291 EzEzEz, Sergio The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.000000000000000000000000989800000000626900000000 Assy Battery Li-Ion,11ah,V600000000 1.00000000000000000000000000000000561500000000991,RP-Filter, Air Inlet, Pkg/5,V600000000V600000000Plu 1.00000000000000000000000000000000561500000000300000000,RP-Filter, Cooling Fan, Pkg/5, V600000000 1.00000000000000000000000000000000561535731,RP-Plate,Retention,O2 Fitting, V600000000V6800000000 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000 1.000000000000000000000000989800000000612211,Power Cord3 WireRt AngleEu100000000 Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: Respirator Eupected and ueupected Resuts: _After every repair both air inlet and cooling fan filters need to be replaced. _The plate holding O2 inlet is the old uY. It needs to be replaced for electrical safety reasons. _The power cord needs to be replaced du to electrical safety issus. _ Internal battery is 8 years old; it needs to be replaced du to uLL related to the battery. _Navigation Ring is not 000000000000000000000000001king, it needs to be replaced du to FCO86600000000000000001B. _Touhscreen is not 000000000000000000000000001king totally well even after it has been calibrated, it needs to be replaced uer Impact: uknown. Patient Impact: uknown. SE 3.00000000 Resoluion: The aire filter was interchanged; uLLfilter; Battery; AC cable; Touh; front panel with navigation navigation; O2 board. So, the test and verification procedue for this equpment, restored the Configuation of the device to the manuactuer. We have performed all the necessary checks that allow for certifying the re-ation of the The conditions prior to the averia. To gurantee the correct operation of the equpment, we recommend that they ue it solo And the problems reported by u. Tools & Equpment: MuTIMETER, 2000000003-12-13 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 20000000000000000000000000000000000000000:00000000:00000000 PRESSuE CALIBRATION TOOL, 200000000000000005-17 00000000:00000000:00000000 FLOW METER, 2000000003-12-00000000 00000000:00000000:00000000 FLOW METER, 2000000003-12-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-13 15:16 | No
543 | Yes | Series images via the wireless foot switch (wireless) not available | Series images via the wireless foot switch (wireless) not available | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 September 15, 2000000003 El Mahamid, Chadi The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: RMT - DTC: Command response: No EN_u. September 12, -2000000003 00000000:5000000001 RMT - DTC: EN_u inactive and hand/foot switch pressed 12-Sep-2000000003 1000000006:27 RMT - DTC: EN_u active and hand/foot switches not pressed 12-Sep-2000000003 17:00000000:00000000 No rrr. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided uu00000000Resoluion: The reported faut is confirmed. Problem Description: Series images via the wireless foot switch (wireless) not possible Suport is requsted. uLLoccured on: 12.00000000.2000000003 Soluion: Wireless foot switch replaced, wireless connection established with base station in OP table, fuction test OK, the system corresponds to the herateller specifications in the tests performed and is ready for clinical ue. Cutomer has been informed. Parts Consued: 1.00000000000000000000000000000000800000000238232, wireless foot switch 3P ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 Robert Kern The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: RMT - DTC: Command response: No EN_u. September 12, -2000000003 00000000:5000000001 RMT - DTC: EN_u inactive and hand/foot switch pressed 12-Sep-2000000003 1000000006:27 RMT - DTC: EN_u active and hand/foot switches not pressed 12-Sep-2000000003 17:00000000:00000000 No rrr. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided uu00000000Diagnostic performed by Engineer: Log checked. uLLperiod limited. Clarification of futher steps with Western Euope. Comparison of the 12 NC of the foot switch with the cutomer. No base station mut be ordered. Internal Comments: uLL# and/or description of uLL Root Caue for RMT - DTC: Command response: No EN_u.: 1. Possible issu with foot switch, possible too muh delay in activation of the 2 internal switches in foot pedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the foot switch are not switching in correct sequnce/in time wrt each other. Root Caue for RMT - DTC: Command response: No EN_u.: 1. Enable u-ray is disabled bu foot switch is pressed. This coud be du to pedal tapping. Root Caue for RMT - DTC: EN_u inactive and hand/foot switch pressed: 1. Enable u-ray is disabled bu foot switch is pressed. This coud be du to pedal tapping. Root Caue for RMT - DTC: EN_u inactive and hand/foot switch pressed: 1. Possible issu with foot switch, possible too muh delay in activation of the 2 internal switches in foot pedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the foot switch are not switching in correct sequnce/in time wrt each other. Root Caue for RMT - DTC: EN_u active and hand/foot switches not pressed: 1. This coud be caued by tapping pedal. Root Caue for RMT - DTC: EN_u active and hand/foot switches not pressed: 1. The hand or foot switch is malfuctioning; it coud be the timing between the internal micro switches is not OK. Repair Action: Service actions for RMT - DTC: Command response: No EN_u.: 1.1. Youcan try to adjut the mechanical timing of the microswitches ('Command' and 'EN_u signal') in the foot pedal with the adjutable screws at the backside. 1.2. Replace foot switch Service actions for RMT - DTC: Command response: No EN_u.: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. Service actions for RMT - DTC: EN_u inactive and hand/foot switch pressed: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. Service actions for RMT - DTC: EN_u inactive and hand/foot switch pressed: 1.1. Replace foot switch Service actions for RMT - DTC: EN_u active and hand/foot switches not pressed: 1.1. Instrut cutomer does pedaltapping. Please ignore this warning if it only occus sporadically. Service actions for RMT - DTC: EN_u active and hand/foot switches not pressed: 1.1. Replace hand or foot switch if warning message occus frequntly. Follow-u requred: uLL# and/or description of uLL Root Caue for RMT - DTC: Command response: No EN_u.: 1. Possible issu with foot switch, possible too muh delay in activation of the 2 internal switches in foot pedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the foot switch are not switching in correct sequnce/in time wrt each other. Root Caue for RMT - DTC: Command response: No EN_u.: 1. Enable u-ray is disabled bu foot switch is pressed. This coud be du to pedal tapping. Root Caue for RMT - DTC: EN_u inactive and hand/foot switch pressed: 1. Enable u-ray is disabled bu foot switch is pressed. This coud be du to pedal tapping. Root Caue for RMT - DTC: EN_u inactive and hand/foot switch pressed: 1. Possible issu with foot switch, possible too muh delay in activation of the 2 internal switches in foot pedal. The t000000000000000000000000001signals: 'Command' and 'EN_u signal' in the foot switch are not switching in correct sequnce/in time wrt each other. Root Caue for RMT - DTC: EN_u active and hand/foot switches not pressed: 1. This coud be caued by tapping pedal. Root Caue for RMT - DTC: EN_u active and hand/foot switches not pressed: 1. The hand or foot switch is malfuctioning; it coud be the timing between the internal micro switches is not OK. Repair Action: Service actions for RMT - DTC: Command response: No EN_u.: 1.1. Youcan try to adjut the mechanical timing of the microswitches ('Command' and 'EN_u signal') in the foot pedal with the adjutable screws at the backside. 1.2. Replace foot switch Service actions for RMT - DTC: Command response: No EN_u.: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. Service actions for RMT - DTC: EN_u inactive and hand/foot switch pressed: 1.1. If this only occus sporadically, youcan ignore it. If it occus more often, see Caue 2. Service actions for RMT - DTC: EN_u inactive and hand/foot switch pressed: 1.1. Replace foot switch Service actions for RMT - DTC: EN_u active and hand/foot switches not pressed: 1.1. Instrut cutomer does pedaltapping. Please ignore this warning if it only occus sporadically. Service actions for RMT - DTC: EN_u active and hand/foot switches not pressed: 1.1. Replace hand or foot switch if warning message occus frequntly. What are the neut trouleshooting steps: Schedue TE service to replace the WIFI foot switch. One can also try to set this on site. FsE mut then decide on site how far this makes sense. 2 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Cutomer comment: ## Arranged appointment with u Pausen for 9/15. Cutomer comment: u Pausen 000000000000000070000000053000000007 called: 0000000022268125 - Info abou the statu of the case. 00000000222658500000000- Nobody has called yet. Requst for feedback. ----------------------------------------------------------------------------- | 2023-09-13 15:46 | No
544 | Yes | System does not go up. Restart without success | System does not go up. Restart without success | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 September 1000000002000000003 Krüger, Henry The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: Today the uB hu with power suply in the control room was replaced sucessfuly, futhermore the power suply for the WCB /udisplay/fade-in on the Fleuvision monitor was restored. The system is therefore fuly operational again and was handed over to the cutomer/uer personally. Problem Description by Engineer: Today the uB hu with power suply in the control room was replaced sucessfuly, futhermore the power suply for the WCB /udisplay/fade-in on the FleuVision monitor was restored. The system is therefore fuly operational again and was handed over to the cutomer/uer personally. Resoluion: Today the uB hu with power suply in the control room was sucessfuly replaced, futhermore the power suply for the WCB /udisplay /fade-in on the Fleuvision monitor was restored. The system is therefore fuly operational again and was handed over to the cutomer/uer personally. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: N/A Parts Consued: 1.00000000000000000000000000000000800000000681000000002,uB2.00000000Eutender Replacement Kit 1.00000000000000000000000000000000800000000700000000672,modue 5 Volt CRCB 1.00000000000000000000000000000000800000000631100000000,SPARE FuE BOu CRCB Tools & Equpment: MuTIMETER, 200000000000000003-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000003-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks: Via u Krüger, by ## Good day to Hambug, Please order to the cutomer in BMT for tomorrow: FRu Description uB2.00000000eutender replacement kit Material nuber 00000000800000000681000000002 Technical Data Vendor Code And FRu Description modue 5 volt CRCB Material nuber 00000000800000000700000000672 Technical Data Vendor Code FRu Description spare fue bou CRCB Material nuber 00000000800000000631100000000 Technical Data Vendor Code Everything please 1 time, appointment sugestion this Thusday approu. 3 hous on site Mfg Henry Krüger Cutomer comment: ## Arranged appointment with u Leschke for 9/100000000fter 2:00000000 pm. ----------------------------------------------------------------------------- | 2023-09-13 16:01 | No
545 | Yes | NZA0screen does not start up; locks where the disks are in is probably the problem | NZA0 screen does not start up; locks where the disks are in is probably the problem | 000000000000000000000000001O##: 000000000000000000000000001000000000000000092200000000ANTOINE MONICA NGELO The system does not meet the specifications for the performed service and is not retuned to ue Diagnostic performed by Engineer: FleuVision PC needs to be replaced, cutomer curently solved himself by connecting disktray directly to mother board Follow u Requred: FleuVision PC needs to be replaced Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Problem with Fleuvision PC. ----------------------------------------------------------------------------- | 2023-09-13 16:21 | No
546 | Yes | Equipment does not pass the auto test, in the charge phase | Equipment does not pass the auto test, in the charge phase | BuQ&A Feedback: No Reason: Corrective Maintenance Trouleshooting Action: OPP-197300000000Qute: 00000000111868 ## Engenharia_clinica_Itaim00ededor.onmicrosoft.com Problem Reported by Cutomer: Equpment problem: Equpment does not pass the self-test duing the charging phase. Presents: Failed / Du Priority: STOPPED uit: SAO LuZ OPERADORA HOSPITALAR Address: R Dr Alceude Campos Rodr 95 Sao Pauo-SP 00000000000000000000000000000000Search for: Anderson Teiueira Sector: Clinical Engineering 001;11 9536000000008800000000002:11 3000000009158 ##Engenharia_clinica_Itaim00ededor.onmicrosoft.com Safety Qustions: Case: 0000000022266572 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT Neut Step Action: Analysis carried ou and directed to sending a buget SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 13.1000000002000000003 20000000028:17 POLARIS (POLARIS) ========== Gulherme Do Rio Costa - 0000000000000000570000000032100000000100000000000000000000000071 =================== ==== Evalution of the uuuDefibrillator 000u00000000572566 carried ou, equpment failed the Operational Test in the Defibrillation Test stage. A test with a jouimeter was applied and it was detected that the Defibrillation valus ​​are below the Tolerance, the cable of the Euternal Pads set has an apparent mesh du to time of ue. A qute will be requsted to replace the T herapy uuuitor part and set of Euternal Paddles. Equpment NOT OK. SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 13.1000000002000000003 20000000028:17 POLARIS (POLARIS) ================= Gulherme Do Rio Costa - 0000000000000000570000000032100000000100000000000000000000000071 =================== ==== Service not yet completed. Tests and Calibrated Tools in PA Tool. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Trouleshooting Action: Qute sent Trouleshooting Action: schedued for 11/100000000- FSE Cifone ----------------------------------------------------------------------------- | 2023-09-13 16:36 | No
547 | Yes | Paddle connector is broken | u- Paddle conneur is broken | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000012 29-Sep-2000000003 Hidalgo Flores, Lenin Otton The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: The equpment is checked to be 000000000000000000000000001king properly. A bent part is foud. It needs to be changed. Parts Consued: 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009259 Ghazwan Hueidi Sordo The system meets the specification for the performed service and is retuned to ue. Follow u Requred: DFM100000000 CN32600000000800000000 Paddle connector is broken. Ask 1 u 0000000056000000008971/0000000000000000 DFM100000000 ASSY Therapy Receptacle TECHNICAL uH-E00000000-00000000-000000001 - CARDIAC HEMODYNAMICS - 300000000100000000322 - Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: DFM100000000 CN32600000000800000000 Paddle connector is broken. Ask 1 u 0000000056000000008971/0000000000000000 DFM100000000 ASSY Therapy Receptacle TECHNICAL uH-E00000000-00000000-000000001 - CARDIAC HEMODYNAMICS - 300000000100000000322 - ----------------------------------------------------------------------------- | 2023-09-13 16:41 | No
548 | Yes | New wireless footswitch required | New wireless footswitch required | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003251 21-Sep-2000000003 Brendan Kiely The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: wireless footswitch not 000000000000000000000000001king. uer Impact: Patient Impact: SE Resoluion : Wireless footswitch and base station installed. System tested and retuned to clinical ue. Parts Consued : 1.00000000000000000000000000000000800000000238212,Wireless Footswitch 000000002 1.00000000000000000000000000000000800000000257862,WFS Base Station ----------------------------------------------------------------------------- | 2023-09-13 16:41 | No
549 | Yes | Arch C does not move | Arch C does not move | BuQ&R Feedback: No Reason: Corrective Maintenance. SAP Transfer: Dispatched for Onsite Suport. Neut Step Action: Performed analysis and roued for sending the qute SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Safety Qustions: Case: 0000000022266600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Sthefany Pereira Last udated by: Sthefany Pereira 1. Was the equpment in clinical ue when the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? N/A, T2 Activities:Problem Description: Problem description by engineer: Arch does not move Trouleshooting Action: What steps are taken duing the system diagnosis: Remote tests pointed to a problem in the geometry TSO. Client Thiago opened and cleaned it bu it did not solve the problem. Tried to load the software sucessfuly, bu the joystick did not 000000000000000000000000001k again. A modue will be requsted for replacement uLL# and/or description of uLL GEO TSO POST uLLOneEMS internal remarks: Problem description by engineer: Arch does not move What steps are taken duing the system diagnosis: Remote tests pointed to a problem in the geometry TSO. Client Thiago opened and cleaned it bu it did not solve the problem. Tried to load the software sucessfuly, bu the joystick did not 000000000000000000000000001k again. A modue will be requsted for replacement uLL# and/or description of uLL GEO TSO POST uLL Repair Action "Internal": Check contract coverage for the part. If it was ized, send Part + FSE - MACHINE STOPPED Recommendations to the FSE Replace the TSO; Charge SW, if necessary; To test the whole machine Parts requred: 1 Part: 1, uu 00000000800000000151715. Description: Geo Modue WP kit, Quntity: 1, Priority: 1 Eutra Internal comment field: RSE Marcel Ogassawara Problem Reported by Cutomer: Arch C does not move Stopped Name: Thiago Department: Clinical Engineering 00¹: 11 98930000000000000000 00²: ## thiago.soua00p.org.br REAL E BENEMERITA ASSOCIAção PORTuuSA DE BENEFICENCIA Address R MAEST CARDIM 769 São Pauo- SP 00000000323-900000000 BR Trouleshooting Action: Qute sent T2 Activities:Repair Action: #TTSuSystem u 12/00000000/2000000003 - 00000000:35AM BRT SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Euternal Remarks: 00000000.1000000002000000003 1000000008:25 POLARIS (POLARIS) =========== Vagner Rodrigus De Soua - 000000000000000057336000000000000000000000000000000000000000000000000200000000====================== Replaced the TSO GEO modue. Performed fuctionality tests. Equpment ok and released for ue! Repair Action: Schedued with cutomer: Thiago Schedued date / time: 22/00000000 from 00000000:00000000PM to 00000000:00000000PM ----------------------------------------------------------------------------- | 2023-09-13 16:46 | No
550 | Yes | Loss of telemetry monitoring hospital wide - Insufficient information - Unknown reason | Loss of telemetry monitoring hospital wide - Insuffi0nt information - Unknown reason | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000051 Walker, Maria The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Problem Description: Loss of telemetry monitoring hospital wide Loss of telemetry monitoring hospital wide Form ued: Trouleshooting/Investigation Object: All telemetry hosts Not Object: Software Version: Date and Time of Issu: 9/13/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: - "No Data Tele" INOP in all the sectors - Problem started duing a weather storm - Device Statu shows all ITS net000000000000000000000000001ked devices in connected mode - There are no alerts in the WMTS log - Cutomer to reboot the ITS stack and callback if he needs futher assistance Was PRS ued: No PRS Permission: Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Monitor patients uer Impact: Loss of telemetry monitoring hospital wide Patient Impact: SE Problem Description/ueupected Resuts: Loss of telemetry monitoring hospital wide Follow u Requred : Loss of telemetry monitoring hospital wide after weather storm FSE assistance needed Diagnostic performed by Engineer : I called cutomer 007:300000000M to follow u bu he is uavailable and he is co000000000000000000000000001ker is uaware of any issus. I paged case to FSE to follow u with cutomer in case onsite suport is needed ----------------------------------------------------------------------------- Neut Step Action : Interaction Type: New Case Caller: Lance IP & Issu: 69519557 all sectors are saying no tele data, mu00000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Afterhous Entitlement Options: Entitled Reason For Entitlement Claim N/A Action Taken: created case, sq's, remote technical w/o, sent to qu , notified rp chat Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-13 17:06 | No
551 | Yes | HS0AED is emitting the triple chirp (Replacing Device) | HS0 AED is emitting the triple chirp (Repla0g Device) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001222 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below West Lake Company 16339 Energy Hwy #2 Proctor, WV 26000000005 Attn: Mark Badia Resoluion : Closing Case Tracking# 78000000003226767, DELIVERED Monday 9/25/23 at 9:17 AM Signed for by: E.EDDY ##cutomer to confirmed that AED is 000000000000000000000000001king as it shoud. Cutomer has confirmed that AED is 000000000000000000000000001king as it shoud. Retun tracking#78000000003228700000000Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and will not pass the self-test. Cutomer removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Remove and reinsert battery to test. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Advise cutomer to remove device ou of service. util he received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# A21K-00000000000000008 Battery Install Before# 2000000008/12/31 Lot#00000000900000000P Pads Eup Date: 20000000000000000000000007 Pads Lot # Y100000000721-00000000 Warranty 1/12/20000000000000000Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping and will not pass the self-test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported the AED kept chirping and will not pass the self-test ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001251 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-13 17:56 | No
552 | Yes | Low Leak-CO0e-breathing Risk alarm | Low Leak-CO0Re-breathing Risk alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001337 15-Sep-2000000003 Buak Tezel The system meets the specification for the performed service and is retuned to ue. Follow u Requred: In the checks, it was detected that the device coud not read the volue valus from the tsystem volue going to the patient and the GDS (Gas Delivery Susystem) part of the device was defective. According to the failue detection, 00000000561512191 - A new one shoud be directed to Izmir Hospital. Resoluion: In the checks conduted, it was detected that the device coud not read the volue and donen t mechanical volue valus from the patient and GDS (Gas Delivery Susystem) part of the device was defective. According to the failue detection, 00000000561512191 - A new one shoud be directed to Izmir Hospital. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: Low Leak-CO2 Re-escalating Risk alarm uer Impact: no Patient Impact: no SE 3.00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Kaan Ozkan The system does not meet the specification for the performed service and is not retuned to ue. Follow-u Requred : Site schedue shoud be made. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: Low Leak-CO2 Re-escalating Risk alarm uer Impact: no Patient Impact: no SE N/A Remote 000000000000000000000000001agnostic performed by Engineer: Duing the contact with the cutomer, it was detected that the device gave "Low Leak-CO2 Rebreathing Risk alarm" alarm. Site service shoud be programmed. ----------------------------------------------------------------------------- | 2023-09-13 18:06 | No
553 | Yes | Legal Issue_AE-000 u- failed to provide utracing. Improved after replacing with alternate electrodes. | Legal Issue_AE-0-000- u- failed to provide utra0g. Improved after repla0g with alternate eleuodes. |  | 2023-09-13 19:09 | No
554 | Yes | Device Fails Self-Test | Device Fails Self-Test | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 Assad MAJEED The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: The cutomer will contact his battery suplier to order a new one. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that futher to the replacement of the battery with a new one, the defibrillator indicates that it is discharged. ----------------------------------------------------------------------------- | 2023-09-13 19:11 | No
555 | Yes | AC power unable to go through. | u- AC power unable to go through. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008000000001 Parts Consued : 1.0000000000000000000000000000000056000000001500000000,Intrepid Power Switch ASSY 1.0000000000000000000000000000000056000000001651,ITP PWR Power Modue Delta 100000000BPS18B ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009500000000 Parts Consued : 1.0000000000000000000000000000000056000000001231,Intrepid AC Power Modue 3pin ----------------------------------------------------------------------------- | 2023-09-13 19:46 | No
556 | Yes | The equipment doesn't turn on. | The equipment doesn't turn on. | BuQ&R Feedback: No Reason: Corrective Maintenance. Euternal Remarks: 11.1000000002000000003 1000000007:39 POLARIS (POLARIS) =========== Marcos Antonio De Carvalho - 00000000000000005733655200000000200000000000000000000000000000000 ======================== The cutomer reports that the equpment is not tuning on correctly. Researched and foud that the host-pc compuer's power suply has A poor contact, we foud the power suply's connector was ouidized, we The connectors as a palliative action and released the equpment. It is necessary to replace the host-PC, the compuer will be sent to the client You SAP Transfer: Dispatched for Onsite Suport. SAP Transfer: Dispatched for Onsite Suport. Trouleshooting Action: Visit by FSE Schedued with cutomer: Schedued date/time: 12/00000000 Date schedued at cutomer's requst? No, trip conciliated Altered priority SAP/1EMS? No Estimated length: 00000000ous Necessary stopped machine? Yes Performed smart schedue? Yes Name of the FSE: Marcos C. Trained FSE as per qulification spreadsheet? YES Problem Reported by Cutomer: Equpment does not tun on. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Dispatched for Onsite Suport. SAP Transfer: Case Transferred to SAP. ----------------------------------------------------------------------------- | 2023-09-13 19:51 | No
557 | Yes | No utracing | u- No utra0g |  | 2023-09-13 19:53 | No
558 | Yes | Unable to get 0utracing | u- Unable to get 0 utra0g |  | 2023-09-13 20:02 | Yes
559 | Yes | u - unable to get tracing | u- u- unable to get tra0g |  | 2023-09-13 20:09 | Yes
560 | Yes | u-ray has frozen and table has frozen | u-ray has frozen and table has frozen | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009825 15-Sep-2000000003 Steve Mcdonald The system meets the specification for the performed service and is retuned to ue Resoluion : The system is now retuned to ful clinical ue. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: u-ray has frozen and table has frozen Royal Oldham Hospital OL1 2JH uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : Replaced the low voltage power supy (200000000 12V, 5V) in the base of the M cabinet. Set u and tested. All OK. Tools & Equpment : MuTIMETER,200000000000000006-21 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Tools & Equpment : MuTIMETER,200000000000000006-21 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 100000000ep-2000000003 Michael Kelly The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: u-ray has frozen and table has frozen Royal Oldham Hospital OL1 2JH uer Impact: NA Patient Impact: NA SE8.2 Tools & Equpment : MuTIMETER,200000000000000002-16 00000000:00000000:00000000 Follow u Requred : Replace PSuPD SCOMP DCPS 20000000012V_5V at rear of M-cabinet GPDu Diagnostic performed by Engineer : uuuwill not load and no comms over DHCP. Foud M-Cabinet net000000000000000000000000001k switch has no power, this is cauing internal comms failue. Confirmed PD SCOMP DCPS 20000000012V_5V Power suply defective. Part on order for engineer retun. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002253 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE onsite to futher investigate supect Fleu PC as sitting at 00000000:00000000 Diagnostic performed by Engineer : No log files available , system sitting at 00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: u-ray has frozen and table has frozen Royal Oldham Hospital OL1 2JH uer Impact: NA Patient Impact: NA SE8.2 Reguatory Qustions : N/A ----------------------------------------------------------------------------- Internal Remarks : ##sent to MPT to confirm planning Internal Remarks : Call to MPT and left a voic##to advise an onsite visit will be needed - we will call back once planned in Internal Remarks : Van is located at Royal Oldham Hospital OL1 2JH (will not move for the foreseeable) Internal Remarks : T000000000000000000000000001contact nubers given - 0000000091780000000095 / 0000000050000000066000000006 Internal Remarks : Call to MPT to advise we have planned an engineer for 1000000009 from 9am - curently ordering parts and taui we will let them know if this changes anything Internal Remarks : Call to engineer to requst an udate advised a follow-u will be requred as another part is needed (he is still on the van) Advised Alice at MPT that another visit will be needed and we will call once planned in Internal Remarks : Call to MPT to advise system is back u and runing - will forward SR once received Internal Remarks : Call from Alice to confirm planning for tomorrow Internal Remarks : The uit is to the left of the main entrance at Royal Oldham Hospital Code: 5836 Internal Remarks : Left a voic##for MPT to confirm planning for 15/00000000 9am Internal Remarks : Tried to call MPT regarding planning - now answer for 00000000inues (will call back) ----------------------------------------------------------------------------- | 2023-09-13 20:11 | No
561 | Yes | Patient u - u- initial fine VF then Asystole | Patient u - u- initial fine VF then Asystole |  | 2023-09-13 20:17 | Yes
562 | Yes | AED keeps indicating that it is not ready for u | AED keeps indicating that it is not ready for u |  | 2023-09-13 20:21 | No
563 | Yes | Alarms or Messages/No Sound: Speaker malfunction inop: Speaker/Mainboard/HW: Replaced parts. | Alarms or Messages/No Sound: Speaker malfunuon inop: Speaker/Mainboard/HW: Replaced parts. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002376 Alatorre, Jose The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Normal monitoring uer Impact: inop message Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Device giving speaker malfuction message Diagnostic performed by Engineer : Cutomer complaint confirmed the speaker is inoperable. The mainboard is uable to connect to the speaker. Resoluion : The speaker is able to give off noise withou interrution. Cutomer configuations have been loaded onto the mainboard. Parts Consued : 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly 1.0000000000000000000000000000000056000000000000000081,Mu_100000000/u3 Main board Tools & Equpment : TORQu SCREWDRIVER,200000000000000008-31 00000000:00000000:00000000 PATIENT SIMuATOR,2000000003-12-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002600000000 Alatorre, Jose The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Mu_100000000/u3 Speaker assembly: 0000000056000000003831 Mu_100000000/u3 Main board: 0000000056000000000000000081 2 hou repair Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Normal monitoring uer Impact: inop message Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Device giving speaker malfuction message Diagnostic performed by Engineer : Cutomer complaint confirmed the speaker is inoperable. The mainboard is uable to connect to the speaker. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 Massey Ryan The system meets the specification for the performed service and is retuned to ue Resoluion : Sent to Bench Follow u Requred : Speaker malfuction inop Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Normal monitoring uer Impact: inop message Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Device giving speaker malfuction message Diagnostic performed by Engineer : Problem Description: Device giving speaker malfuction message Speaker malfuction inop Form ued: Bench Trouleshooting Steps Taken: ##d bench form and created bench 000000000000000000000000001----------------------------------------------------------------------------- | 2023-09-13 20:36 | No
564 | Yes | The equipment is completely turned off. | The equipment is completely turned off. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000100000000ep-2000000003 Santiago, Angel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: not in ue Eupected and ueupected Resuts: system tun on and shu down again uer Impact: no Patient Impact: no SE Diagnostic performed by Engineer : Trouleshooting system foud many issu with system need replace parts. Follow u Requred : Order parts ----------------------------------------------------------------------------- | 2023-09-13 21:06 | No
565 | Yes | C arm movements not responding to movements | C arm movements not responding to movements | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000296000000006-Sep-2000000003 MuH, JOE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : spoke to Chris the staff says c armis not responsing to movement When was the first occurence? Within... : 13-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: c arm not moving correctly Patient Impact: none SE Not provided Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 TORQu WRENCH,200000000000000003-31 00000000:00000000:00000000 Resoluion : Arrived onsite to investigate reported issu with parts on hand. I was able to receive and review diagnostic logs prior to onsite from in-houe Imaging Biomed Chris. I supected reported issu was related to previouly reported issus. After reviewing logs I coud see mutiple events occuring on the day of the reported issu. I ordered a Longituinal Pot for the L-Arm, bu planned to first do some calibrations, and inspect the roller bearing gap for the longituinal movement. Once I got into the room I realized I had been sent the wrong part. The ouside bou and the paper000000000000000000000000001k was correct, bu the inside part bou was wrong - it was a AD7 Table Lateral Brake. I decided to continu with doing whatever else I coud. I inspected the long Pot throuhou the longituinal travel distance and noticed that the long pot gear was coming off the ceiling rail gear track at a specific spot and continung to stay ou of position throuhou muh of the travel distance. This was du to the Ceiling Rail Eutension Rail addition warping and slightly separating the transition of where the t000000000000000000000000001Ceiling Rail gear strips meet. I also noticed that the Long Pot spring tensioner seemed to be looser than it shoud. While trying to resolve the spring tension I noticed it was being caued by one of the spring mout posts being bent. While trying to bend that back in place it snapped off completely. I was able to temporarily fiu that by replacing the spring tension mouting post with a zip tie. I then loosened the Ceiling Rail bolts to try and move the Ceiling Rails back into a proper position. I was able to improve that. I fiued the Ceiling Rail gear strip by removing the gear strip half off the Eutension Rail, and reapplying it with doule-stick tape, intern fiuing the Gear Strip gap misalignment. That resolved the Long Pot Gear from coming off the track as it did before. I retorqud all the Ceiling Rail Eutension bolts. I had to perform Longituinal Pot calibrations and position calibrations. This all seemed to have resolved the reported issu. I will retun in the fuue to replace the Long Pot with a properly fuctioning spring tensioner. System fuly fuctional. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002662 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : spoke to Chris the staff says c armis not responsing to movement When was the first occurence? Within... : 13-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: c arm not moving correctly Patient Impact: none SE Not provided Diagnostic performed by Engineer : speaking with chris the staff say they were uable to movr c arm once it was in the 000000000000000000000000001king position .in houe is checking system appears to be 000000000000000000000000001king at this time . rse attempted remote connection withou suess .Va site . Chrisis forwarding logging when he can for review .in houe checking bodygurd sensor and gemetry controls.cutomer requsted on site Internal Comments : uLL# and/or description of uLL: n/a Repair Action : no logging received as of yet Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : no logging received as of yet What are the neut Trouleshooting steps : check logging history for uLL test and verify c arm movements Technician Name : Sulivan Scott ----------------------------------------------------------------------------- Internal Remarks : 000000002227000000001 MuH, JOE FSE calling in on u 8.2.3000000005 where C-arm is puting ou hudreds of position uLL. Application uLL Motor assembly uLL[auis=FrtLong] Advised to order potentiometer and to get to site to check the rail is not dirty and check the tension on the wheel. Gave FSE ok to order pot and AMC if needed. Internal Remarks : 000000002227000000001 Joe Muh 000000000000000000000000001king with FSE after hous. FSE trouleshooting an issu where no movement available. Logs show possible issu with AMC. Gave FSE ok to order AMC ----------------------------------------------------------------------------- | 2023-09-13 21:26 | No
566 | Yes | The green check is not working and unable to place the system in service mode | The green check is not working and unable to place the system in service mode | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002958 Mccleary Barry The system meets the specification for the performed service and is retuned to ue Resoluion : Provided part nuber for switch pcb, cutomer will order part- Material Nuber 00000000561512600000000 - RP-PCBA, Switch, V600000000$171.00000000 Case follow u ##sent, cutomer states uit was repaired and retuned to service, issu resolved. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm service Eupected and ueupected Resuts:green check on bezel inop uer Impact: not uable needs service Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing pm the green check on front bezel does not 000000000000000000000000001k and cannot enter Diagnostics, FCO86600000000000000007 was previouly completed, requsts part info to correct issu. ----------------------------------------------------------------------------- Internal Remarks : Interaction New Case Caller: Kayla Owens IP & Issu: " 56630000000073 The green check is not 000000000000000000000000001king and uable to place the system in service mood. Per the cutomer." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-13 21:41 | No
567 | Yes | uLLMessage 00Machine pressure sensor auto zero failed | uLLMessage 00: Machine pressure sensor auto zero failed | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003100000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uLL1100000000 Machine pressue sensor auozero failed alarm Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: Machine pressue sensor auozero failed alarm uer Impact: not uable needs repair Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer states while in ue the uit alarmed Machine pressue sensor auozero failed alarm, no patient harm reported and uit removed from ue, in biomed noted uLL1100000000 in event log, requsts onsite service to correct issu. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Noel Mapanao Issu: 89630000000000000000uLLMessage: Machine pressue sensor auo zero failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-13 22:11 | No
568 | Yes | Auuiliary alarm failed (EC 0) | Auuiliary alarm failed (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006983 00000000-Oct-2000000003 Galan, Alberto The system meets the specification for the performed service and is retuned to ue Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000007-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000007-31 00000000:00000000:00000000 MuTIMETER,200000000000000007-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1115 uer Impact: none Patient Impact: none SE 3.100000000Diagnostic performed by Engineer : FSE installed MC PCBA and performed a complete PVT as per the manuactuer's specifications. Parts Consued : 1.00000000000000000000000000000000561537531,RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000006 21-Sep-2000000003 Galan, Alberto The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1115 uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : The cutomer reported a problem with the device and asked the FSE to check it. The FSE tested the PM PCBA with a new one, bu the device still showed uLLcodes 1115 and 1100000000 The device needs a new MC PCBA, Power Suply, and CPu Futher 000000000000000000000000001k is requred for this device. Follow u Requred : Incident: Repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000005 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer has a V600000000and soon after the FCO86600000000000000006 was remediated, the uit now has a code 1115. Cutomer requsted the FCO86600000000000000006 to be done again. He is aware if the neut PM pcba does not solve the uit, then he 000000000000000000000000001ld be responsible for the repair. ASP is schedued on 9-21-23 Follow u Requred : Cutomer has a V600000000and soon after the FCO86600000000000000006 was remediated, the uit now has a code 1115. Cutomer requsted the FCO86600000000000000006 to be done again. He is aware if the neut PM pcba does not solve the uit, then he 000000000000000000000000001ld be responsible for the repair. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1115 uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Daniel Kimnes IP & Issu: 8700000000200000000 Will not tun off and gives uLL1115 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-13 22:11 | No
569 | Yes | Will not turn off | Will not turn off | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000002 00000000-Oct-2000000003 Galan, Alberto The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561537531,RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000007-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000007-31 00000000:00000000:00000000 MuTIMETER,200000000000000007-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1115 uer Impact: none Patient Impact: none SE 3.100000000Diagnostic performed by Engineer : FSE installed MC PCBA and performed a complete PVT as per the manuactuer's specifications. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000000000001-Sep-2000000003 Galan, Alberto The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Incident: Repair Diagnostic performed by Engineer : The cutomer reported a problem with the device and asked the FSE to check it. The FSE replaced the PM PCBA with a new one, bu the device still showed uLLcodes 1115 and 1100000000 The device needs a new MC PCBA, Power Suply, and CPu Futher 000000000000000000000000001k is requred for this device. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1115 uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000003 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer has a V600000000and soon after the FCO86600000000000000006 was remediated, the uit now has a code 1115. Cutomer requsted the FCO86600000000000000006 to be done again. He is aware if the neut PM pcba does not solve the uit, then he 000000000000000000000000001ld be responsible for the repair. Diagnostic performed by Engineer : Cutomer has a V600000000and soon after the FCO86600000000000000006 was remediated, the uit now has a code 1115. Cutomer requsted the FCO86600000000000000006 to be done again. He is aware if the neut PM pcba does not solve the uit, then he 000000000000000000000000001ld be responsible for the repair. ASP is schedued on 9-21-23 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1115 uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Daniel Kimnes IP & Issu: 8700000000200000000 Will not tun off and gives uLL1115 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-13 22:12 | No
570 | Yes | M0A - u00- Cannot be turned on when in u | u- Cannot be turned on when in u |  | 2023-09-13 22:13 | No
571 | Yes | M0A / u00 / Failed the self-test (QA#000000 | M00A / u000 / Failed the self-test (QA#0000000. |  | 2023-09-13 22:24 | No
572 | Yes | Started smoking and overheating while charging, it was on fire right away. | 00BEL Hu00/0 Started smoking and overheating while charging, it was on fire right away. | 2000000003-00000000-21 13:28:00000000I will close tomorrow the last GFE. Cutomer is uable to retun the produt. He has thrown it away. I have sent the cutomer a 300000000 vouher. I informed the cutomer.;2000000003-00000000-21 13:16:500000000Notification in GFE in TW that the Csm has informed u that the sample is not available, advised to close the PR.;2000000003-00000000-21 10000000012:18:ZAMFY95u: Sample not sent yet. 2nd reminder sent today, 21/00000000/2000000003.;2000000003-00000000-20000000013:00000000:59:Informed the BG that the Duch scentence Color had changed a bit so it was on at the same time. Colou had changed a little that it was on fire right away. means the device was on fire;2000000003-00000000-20000000000000000:59:00000000ZAMFY95u: Sample not sent yet. Reminder will be sent 21/00000000/2000000003.;2000000003-00000000-19 12:36:53:ZAMFY95u: Sample not sent yet. Reminder will be sent 21/00000000/2000000003.;2000000003-00000000-18 00000000:30000000000000000ZAMFY95u: Sample not sent yet. Reminder will be sent 21/00000000/2000000003.;2000000003-00000000-15 00000000:56:00000000:ZAMFY95u: Sample not sent yet. Label added to the case. Qustionnaire in case.;2000000003-00000000-1000000002:28:00000000:Mail sent concerning labels.;2000000003-00000000-1000000008:23:15:Sumitted PR#3800000000000000002000000003-00000000-1000000008:22:00000000AYS Created REG ID 2989 Baueister;2000000003-00000000-1000000008:1000000003:- The qustionnaire: -How long had youhad it: abou 3 months now bu had not opened it. -How often did youue it:: not. -How long were youuing it when the incident happened: nothing, didn't 000000000000000000000000001k. -What was the resut: didn't 000000000000000000000000001k. -Did youhear strange noises: yes. -What was the condition of the device after the incident: colou had changed a little that it was on fire right away. -In what environment was the device ued: bathroom? -Have youcharged the device for 200000000ous after youfirst got it:: yes becaue it did not 000000000000000000000000001k -How many times did youcharge it: 3 times, did not 000000000000000000000000001k once -Have youalways kept it on the charger:: no. -Have youued the original charger?: yes. -Have youued adapters:: no. -Did youconnect the charger to the mains or ue an eutension cord?: mains.;2000000003-00000000-13 13:13:37:Send ##to the cutomer concerning additional information.;2000000003-00000000-13 12:51:00000000There are t000000000000000000000000001separate cases from this cutomer. Safety escalation in case 6000000000000000035651200000000 000000000000000000000000001in case: 6000000000000000035651000000002;2000000003-00000000-13 11:1000000001:R: Fnac I: Bruh warms u, gives a rrok and makes a strange noise T: Csm has jut requsted 000000000000000000000000001bu now displays this message S: Safety escalation | 2023-09-13 22:31 | No
573 | Yes | AED beeping, self-test failed with "Remove and reinsert battery" uLLmsg (Device is replaced) | AED beeping, self-test failed with "Remove and reinsert battery" uLLmsg (Device is replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003300000000Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to acct address Attn Tom Basiua Internal Comments : Requsted SPS to eupedite shipment Problem Description by Engineer : Cutomer Fuction/Role: Owner Device ue: no Eupected and ueupected Resuts: not passing self test uer Impact: none Patient Impact: none SE E.00000000.000000003 See ref Case 0000000022267655 Cut called stating device was beeping last week and they took it ou of service. Diagnostic performed by Engineer : Had cut ru BIT and it skipped shock buton test and at the end of self test, it emitted an uLLmsg “remove and reinsert battery” with no follow u chirping. Advised cut to remove device from service and will replace device uder warranty. 000000001200000000 Install before date: 2000000008-00000000-31 Pads lot# C000000000000000023-00000000 Eup date: 2000000005-00000000-00000000 Resoluion : Cutomer called and informed that he as received the replacement AED and the AED is now in service. I have provided the cutomer with a retun tracking label with tracking nuber 780000000050000000036900000000and attached label to 000000000000000000000000001--------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003377 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-13 22:36 | No
574 | Yes | Pivot is not locking properly | Pivot is not locking properly | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000000000009-Sep-2000000003 Sinkewicz Gary The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Cutomer is reporting pivot does not lock in place When was the first occurence? Within... : 13-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: pivot not locking uer Impact: table pivots when locked Patient Impact: none mentioned SE 7.2.16.00000000Resoluion : REPLACED BRAKE ASSEMBLY & FRICTION PIN Parts Consued : 1.00000000000000000000000000000000230000000035121,FRICTION PIN 1.00000000000000000000000000000000230000000027133,PIVOT BRAKE ASSY Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003316 15-Sep-2000000003 Sinkewicz Gary The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replaced Brake & friction Pad in AD7 table Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 Resoluion : Identified Friction Pad & Brake assembly are not hold firm. ordered Brake and friction pad Problem Description by Engineer : Problem description by engineer : Cutomer is reporting pivot does not lock in place When was the first occurence? Within... : 13-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: pivot not locking uer Impact: table pivots when locked Patient Impact: none mentioned SE 7.2.16.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000032300000000rmenia, Mauice The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Cutomer is reporting pivot does not lock in place When was the first occurence? Within... : 13-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: pivot not locking uer Impact: table pivots when locked Patient Impact: none mentioned SE 7.2.16.00000000Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Check pivot components for any debris and check brakes and pivot friction pin. Verify that it is the pivot movement that cutomer is having issu with. [PIVOT BRAKE ASSY],[FRICTION PIN] [ 00000000230000000027133 ] [ 00000000230000000035121 ] Technician Name : Armenia, Mauice Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Check pivot components for any debris and check brakes and pivot friction pin. Verify that it is the pivot movement that cutomer is having issu with. [PIVOT BRAKE ASSY],[FRICTION PIN] [ 00000000230000000027133 ] [ 00000000230000000035121 ] Diagnostic performed by Engineer : Cutomer called it "swag" first bu seems like he is talking abou pivot. Verify symptoms onsite. States they pu a 300000000 poud patient on it and it does not lock and can move it when locked. ----------------------------------------------------------------------------- | 2023-09-13 22:41 | No
575 | Yes | System is intermittent(additional diagnostics related to the uuu not botting correctly) | System is intermittent(additional diagnostics related to the uuu not botting correuy) | Internal Remarks : Interaction New Case Caller: Cris on behalf of GE for Scott Breau IP & Issu: "52000000002300000000 System is intermittent(additional diagnostics related to the uuu not botting correctly) per the GE rep." Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created case and created qute requst for onsite suport service Neut Step Action: Awaiting Qute Verbal PO Nuber: 000000001100000000 CAP amout of $5000000000000000000000000 Verbal PO Approver: GE Verbal PO CAP: $5,000000000000000000000000 ----------------------------------------------------------------------------- | 2023-09-13 22:41 | No
576 | Yes | Won't load up | Won't load up | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002893 13-Sep-2000000003 Kocher, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : FleuVision PC will need to be replaced. Believe PC failue is du to technical room being to warm. Facility HVAC department is investigating AC for the technical room. Qute has been sent to cutomer for new FleuVision PC Problem Description by Engineer : Cutomer Fuction/Role: End uer Device ue: Ouside of ue Eupected and ueupected Resuts: System will not boot uer Impact: System Down Patient Impact: System Down SE 8.1.3000000005 ----------------------------------------------------------------------------- Internal Remarks : Josh Kocher - system stuk at 00000000:00000000. FleuVision cannot ping Host uLL Sugested replacing FV2PC. Additional uLL for temperatue too high in technical room. Sugested checking local temperatue sitution. Internal Remarks : Interaction Type: New Case Caller: Terry Glass IP & Issu: 000000000000000000000000001t load u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Qute Neut Step Action: Awaiting Qute Verbal PO Nuber: 90000000012FSR3000000003993 Verbal PO Approver: Terry Glass Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-13 23:01 | No
577 | Yes | ContinuouBattery alarm after PM board replaced | ContinuouBattery alarm after PM board replaced |  | 2023-09-13 23:07 | No
578 | Yes | ***Misalignment in the touch position was confirmed. 0 | ***Misalignment in the touch position was confirmed. 0 |  | 2023-09-13 23:27 | No
579 | Yes | No volume on one of the screens | No volume on one of the screens | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003636 Sanchez, Karbin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed called to requst assistance with the volue.No volue on one of the screens. The cutomer foud ou that the nuses upluged a cable by mistake. The system is 000000000000000000000000001king fine. Device ue: PICiu Eupected and ueupected Resuts: Biomed called to requst assistance with the volue.No volue on one of the screens. The cutomer foud ou that the nuses upluged a cable by mistake. The system is 000000000000000000000000001king fine. uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer : Biomed called to requst assistance with the volue.No volue on one of the screens. The cutomer foud ou that the nuses upluged a cable by mistake. The system is 000000000000000000000000001king fine. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Michael sanchaz 60000000020000000036700000000 IP & Issu: No volue on one of the screens Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-13 23:41 | No
580 | Yes | Battery was not charging; battery drained and died | Battery was not charging; battery drained and died | BuQ&R Feedback: No Reason: Corrective Maintenance. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Internal Remarks: 21.00000000.2000000003 15:31:00000000POLARIS (POLARIS) =========== Dawison George Dias - 000000000000000057338500000000000000001000000000000000000000000600000000====================== Applied the 100000000 - 11Ah, 163Wh battery as per the invoice 38135. INFORMATION FROM THE BATTERY: REF: 100000000637000000001000000008272) LOT M0000000036-P2. MANuACTuE: 26-00000000-2000000003 Performed fuctional tests, ventilation and battery OK. EQuPMENT RELEASED. SAP Transfer: Dispatched for Onsite Suport. Euternal Remarks: 21.00000000.2000000003 15:31:00000000POLARIS (POLARIS) =========== Dawison George Dias - 000000000000000057338500000000000000001000000000000000000000000600000000======================= Applied the 100000000 - 11Ah, 163Wh battery as per the invoice 38135. INFORMATION FROM THE BATTERY: REF: 100000000637000000001000000008272) LOT M0000000036-P2. MANuACTuE: 26-00000000-2000000003 Performed fuctional tests, ventilation and battery OK. EQuPMENT RELEASED. SAP Transfer: Case Transferred to SAP. Neut Step Action: Performed analysis and roued for sending the qute Safety Qustions: Case: 0000000022268883 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: uyni Edwirges. Last udated by: uyni Edwirges. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown Repair Action: schedued for 21/00000000 - FSE Dawison Problem Reported by Cutomer: Battery qute ----------------------------------------------------------------------------- | 2023-09-14 00:31 | No
581 | Yes | Battery will not charge | Battery will not charge | BuQ&R Feedback: No Reason: Corrective Maintenance. SAP Transfer: Case Transferred to SAP. SAP Transfer: Dispatched for Onsite Suport. Safety Qustions: Case: 0000000022268918 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: uyni Edwirges. Last udated by: uyni Edwirges. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown Neut Step Action: Performed analysis and roued for sending the qute SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. ----------------------------------------------------------------------------- | 2023-09-14 00:36 | No
582 | Yes | Battery will not charge | Battery will not charge | BuQ&R Feedback: No Reason: Corrective Maintenance. Safety Qustions: Case: 0000000022268900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: uyni Edwirges. Last udated by: uyni Edwirges. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown SAP Transfer: Case Transferred to SAP. SAP Transfer: Dispatched for Onsite Suport. SAP Transfer: Case Transferred to SAP. ----------------------------------------------------------------------------- | 2023-09-14 00:36 | No
583 | Yes | Battery will not charge | Battery will not charge | BuQ&R Feedback: No Reason: Corrective Maintenance. SAP Transfer: Dispatched for Onsite Suport. Safety Qustions: Case: 0000000022268932 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: uyni Edwirges. Last udated by: uyni Edwirges. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Neut Step Action: Performed analysis and roued for sending the qute ----------------------------------------------------------------------------- | 2023-09-14 00:41 | No
584 | Yes | Battery will not charge | Battery will not charge | BuQ&R Feedback: No Reason: Corrective Maintenance. Safety Qustions: Case: 0000000022268938 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: uyni Edwirges. Last udated by: uyni Edwirges. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Dispatched for Onsite Suport. ----------------------------------------------------------------------------- | 2023-09-14 00:41 | No
585 | Yes | Battery would not charge | Battery would not charge | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Analysis carried ou and directed to the buget sumission SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022268900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: uyni Edwirges Last udated by: uyni Edwirges 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment malfuction or fail at the time of installation? uknown Internal Remarks: 00000000.11.2000000003 11:56:00000000POLARIS (POLARIS) ========== Gutavo de Assis Soua - 00000000000000005733851100000000100000000000000000000000038 ============================ GG6: Operational tests passed. GG13: No calibrable tools were ued. Euternal Remarks: 00000000.11.2000000003 11:56:00000000POLARIS (POLARIS) ========== Gutavo de Assis Soua - 00000000000000005733851100000000100000000000000000000000038 ============================ Performed battery replacement of the V600000000serial pumonary uLL 100000000118671 Operational tests carried ou, ok. Battery Installed Lot ID: M0000000036P2 Ventilation tests, ok. Note: Allow the battery to charge to pu equpment to ue. NF: 0000000097 Repair Action: schedued for 00000000/11 - FSE Cifone SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-14 00:41 | No
586 | Yes | Battery will not charge | Battery will not charge | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022268953 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: uyni Edwirges Last udated by: uyni Edwirges 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment malfuction or fail at the time of installation? uknown Euternal Remarks: 00000000.11.2000000003 2000000000000000000000000POLARIS (POLARIS) =========== Dawison George Dias - 000000000000000057338512000000001000000000000000000000000600000000============================= APPLICATION OF THE 100000000 - 11Ah, 163WH BATTERY ACCORDING TO THE NF 0000000071. BATTERY INFORMATION: SKu 100000000637000000001000000008272) LOT M000000002600000000P1 MANuACTuING: 2000000002-11-00000000 PERFORMED OKS FuCTIONAL, VENTILATION, AND BATTERY TESTS. EQuPMENT RELEASED. SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to the buget sumission SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.11.2000000003 2000000000000000000000000POLARIS (POLARIS) =========== Dawison George Dias - 000000000000000057338512000000001000000000000000000000000600000000=========================== APPLICATION OF THE 100000000 - 11Ah, 163WH BATTERY ACCORDING TO THE NF 0000000071. BATTERY INFORMATION: SKu 100000000637000000001000000008272) LOT M000000002600000000P1 MANuACTuING: 2000000002-11-00000000 PERFORMED OKS FuCTIONAL, VENTILATION, AND BATTERY TESTS. EQuPMENT RELEASED. Repair Action: schedued for the Dawison ESF - 00000000/100000000----------------------------------------------------------------------------- | 2023-09-14 00:46 | No
587 | No | This is the Duplicate of PR # 00 which is the PR ID of the actual complaint. | This is the Duplicate of PR # 0000 which is the PR ID of the aual complaint. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003731 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: Parts ID for power suply - good power in, no power ou. uit operating on battery power. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : provided: Power suply (1st generation) 100000000000000009 00000000561512591 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Larry Ames +1-916-986-000000008 IP & Issu: 200000000000000007718 Parts ID for power suply Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Created R000000000000000000000000001 roued to Technical Suport Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-14 01:01 | No
588 | Yes | Check Vent:Supply Failed | 00 (POST) Check Vent: 0 V Supply Failed | 000000000000000000000000001O##: 000000000000000000000000001000000000000000039200000000erchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the power suply and battery and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: 200000000olt failed, code 111A uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and they are having a 200000000olt suply failue. Cutomer verified there was 1200000000C volts going into the power suply. only 21 volts coming ou of the power suply. recommended replacing the power suply. Part #00000000561512591, power suply, list price $890000000000000000 Part #989800000000626900000000 internal battery, list price $959.00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: tom williams IP & Issu: 57600000000600000000eed tech suport with voltage after 200000000lt suply failue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/ r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-14 01:01 | No
589 | Yes | The Monitor is falling - won't hold its position | The Monitor is falling - won't hold its position | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008882 18-Sep-2000000003 Husein, Abdu The system meets the specification for the performed service and is retuned to ue Resoluion : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: in ue Eupected and ueupected Resuts: uer Impact: none Patient Impact: none SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000000000005-Sep-2000000003 Huer Dou The system meets the specification for the performed service and is retuned to ue Resoluion : Dose aware monitor removed as not in ue. Balance back correctly and monitor supension stays at set area Resoluion : Checked balance. Eutra monitor on supension Problem Description by Engineer : Cutomer Fuction/Role: Device ue: IV Eupected and ueupected Resuts: uer Impact: monitor Patient Impact: none SE6.1.3 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : monitor supention came down and hit the patient on the side of the face When was the first occurence? Within... : 13-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: hit by the monitor supension on the side of the face SE Not provided Follow u Requred : uLL# and/or description of uLL: monitor supention came down and hit the patient on the side of the face Repair Action : Check monitor supension and replace if necessary. 0000000021330000000000000000 - Gas spring Technician Name : Chhlam, Sy Internal Comments : uLL# and/or description of uLL: monitor supention came down and hit the patient on the side of the face Repair Action : Check monitor supension and replace if necessary. 0000000021330000000000000000 - Gas spring Diagnostic performed by Engineer : Christie reported that the tech went to move the monitor supension and it dropped down and hit the patient on the side of the face. Per caller there is no bruse visible, no cus and patient does not want futher care for the incident. ----------------------------------------------------------------------------- Internal Remarks : 0000000000000000000000000010000000000000000000000008 has now been canceled as it is a dulicated as per FSE Dou H requst Internal Remarks : Spoke to FSE Dou Huer, he accepted dispatch for Sept 100000000 11am Internal Remarks : cutomer called in with issu with falling monitor. Priority was ugraded to P1 when she answered a patient was harmed. Case created and call roued to SIT. ----------------------------------------------------------------------------- | 2023-09-14 01:11 | No
590 | Yes | Questions on the O0nput supply pressure. | Questions on the O0input supply pressure. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003853 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue, getting ready to go on patient Eupected and ueupected Resuts: high o2 inlet pressue uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and they have a qustion on the o2 inpu suply pressue. Biomed coud not dulicate issu. if the inlet pressue is too high, recommending loosen the o2 hose going to the back of the V600000000to release some pressue. When there is too muh pressue the o2 solenoid closes, and the o2 is trapped inside the hose. Releasing pressue will lower the o2 inlet pressue. Recommended to tun off V600000000and disconnect all o2. He plans to do this and then perform the PVT. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Dave Taylor 600000000000000003900000000 IP & Issu: SN 10000000000000000527 Qustions on the O2 inpu suply pressue. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created Case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-14 01:11 | No
591 | Yes | the live image monitor loses signal intermittently | the live image monitor loses signal intermittently | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002800000000 28-Nov-2000000003 DeJesu, Victor The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: CV Eupected and ueupected Resuts: Montor withou signal uer Impact: NA Patient Impact: NA SE7.2.9 Resoluion : Pars replaced by third party. Parts Consued : 1.00000000000000000000000000000000200000000000000009000000001,Video Selector 1.00000000000000000000000000000000216700000000500000000,AVA BOARD ON IPB (AuILIARY VIDEO ADAPTI 1.0000000000000000000000000000000023000000002800000000,FRuIPB_FOu 1.00000000000000000000000000000000230000000016751,Coau Video Cable ----------------------------------------------------------------------------- | 2023-09-14 01:36 | No
592 | Yes | Wireless foot-switch base station is defective | Wireless foot-switch base station is defeu | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006988 12-Oct-2000000003 Valdez Andy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Wireless Footswicth will pair on base station receiver Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion : Firmware corruted duing FCO, will not pair base station and footswitch. Resoluion : Replaced base station and wireless footswitch and pair to sync. Tested wireless footswitch fuction. system 000000000000000000000000001king ok. Internal Comments : do not bill cutomer ouside PPM hous becaue manpower issu. ----------------------------------------------------------------------------- | 2023-09-14 01:51 | No
593 | Yes | Device is not booting on | Device is not booting on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Miller Jason The system meets the specification for the performed service and is retuned to ue Resoluion : Provided cutomer with part nuber: 0000000056150000000000000000- power switch overlay. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: device is not booting on uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device is not booting on. Confirmed power to device. LEDs showing power when pluged in. press power buton, device will not power on. Sugested to cutomer to replace power switch overlay. Cutomer requsting replacement part nuber. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Shone Kuiakofe (600000000 385-1998 IP & Issu: 10000000000000000369 device is not booting on Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Created R000000000000000000000000001 roued to Technical Suport Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-14 02:06 | No
594 | Yes | The alarm silence is not working | The alarm silence is not working |  | 2023-09-14 02:51 | No
595 | Yes | AED still gives "Remove and reinsert battery to test" uLLmsg after new battery installed | AED still gives "Remove and reinsert battery to test" uLLmsg after new battery installed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 Mathew, Vincia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Calvary Baptist Chuch 500000000 W Main St Tuelo, MS 38800000000 Attn: Mike Williams Resoluion : Spoke to cutomer device was received and is in ue. cutomer will retun the defective AED on Monday with the u prepaid waybill Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Not in ue Eupected and ueupected Resuts: Defective AED uer Impact: Remove from Service Patient Impact: None SE NA Diagnostic performed by Engineer : AED not initiating the self test uLLMsg: not ready for ue, remove and reinsert battery to test cutomer install new battery and still getting the same uLLmsg. Sending cut replacement AED Install before: 6/3000000002000000007 8/27/21 install date Lot#00000000200000000 Serial#B00000000200000000700000000 Pad eup: 00000000/20000000000000000 Lot#Y100000000721-00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-14 03:36 | No
596 | Yes | Speaker malfunction uLLNo sound. Assembly/PCB/Card-Speaker-HW. | Speaker malfunuon uLLNo sound. Assembly/PCB/Card-Speaker-HW. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 Doulas, Norman The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Is Caue Known: Yes Caue: Bad Speaker Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Resolved Remotely Follow u Requred : ORDER QTY 1 000000005600000000000000001 Louspeaker uD 100000000.00000000 Tu Location Name PuLIC 000000000000000000000000001KS DEPARTMENT Street 335 E Avenu I City LANCASTER State CA Zip 93535-1916 Coutry uited States Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Monitor Vitals uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Mu800000000 is displaying a Speaker Malfuction uLLmessage and is not produing any Auio Diagnostic performed by Engineer : Problem Description: Mu800000000 is displaying a Speaker Malfuction uLLmessage and is not produing any Auio Mu800000000 is displaying a Speaker Malfuction uLLmessage and is not produing any Auio Form ued: Trouleshooting/Investigation Object: Mu800000000 Not Object: Other Mu800000000's Software Version: uu Date and Time of Issu: 9/13/2000000003 uer Actions: uu Where is the Deviation: NA Trouleshooting Steps Taken: ORDER QTY 1000000005600000000000000001 Louspeaker uD 100000000.00000000Tuocation Name PuLIC 000000000000000000000000001KS DEPARTMENTStreet 335 E Avenu ICity LANCASTERState CAZip 93535-1916Coutry uited States Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Parts Consued : 1.000000000000000000000000000000005600000000000000001,IV2-STAT MECHASY Louspeaker ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Komizi Attipoe IP & Issu: " 62765976 uLLcode on the system. speaker malfuction" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-14 04:16 | No
597 | Yes | System will not boot up all the way | System will not boot up all the way | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 Mcnamara Pau The system meets the specification for the performed service and is retuned to ue Internal Comments : uLL# and/or description of uLL: keyboard and moue not 000000000000000000000000001king Repair Action : Provide quck help for Sute pc connection Resoluion : Cutomer contract in 000000000000000000000000001ks, They installed new sute PC as FSE in not schedued on site for days, bu moue not 000000000000000000000000001king on one monitor. Euplained that uB connection mut be off on Sute PC. Cutomer sees uB cable, pluged into sute PC. System not runing normal and in 000000000000000000000000001king condition Problem Description by Engineer : Problem description by engineer : Sute PC keyboard and moue not 000000000000000000000000001king on new PC Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Cutomer stated they installed new Sute PC bu keyboard and moue only 000000000000000000000000001k on review and not acqusition monitor ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007696 22-Sep-2000000003 Byron Kirk The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuDul Switch Modue Resoluion : assisted Biomed in Diagnostics and repair Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: ou Eupected and ueupected Resuts: ueupected uer Impact: system down Patient Impact: none, ued alternate system SE ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action taken : FSE Kirk Byron accepted call for 9/22 Oucome: dispatched Neut Step: N/A ----------------------------------------------------------------------------- | 2023-09-14 04:21 | No
598 | Yes | Startup failed | 0/uLuCentron/Startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 00000000.1000000002000000003 00000000:27:51 uC (86000000008265) $Q7:(Confirmed reported problem. The failue occured insystem control. system start u failed） $Q8: (Checked system, host pc cannot start u, need replace Host pc) $Q9:（replace host pc, T2 sugestion effective.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. ) $Q11:$Calibrated Tools name (NA); S/N (NA); Neut Calibration date(NA) FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA Euternal Remarks: 00000000.1000000002000000003 00000000:27:00000000uC (86000000008265) Service content: The main control compuer cannot start normally, replace the main control compuer host pc Service resuts: After repair, the equpment retuned to normal operation Neut Step Action: 8 labor times have been ued u, contact SS buiness for follow-u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine coud not enter the application interface normally and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine cannot enter the application interface normally Remote diagnosis and test resuts: power ouage and restart have no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Try re-inserting and upluging the hard drive. If that doesn't 000000000000000000000000001k, try connecting directly. 2. Check whether the software is normal. 3. Check the machine for uLL and repair the faut. If it cannot be repaired, replace the corresponding spare parts. Recommended spare part uuand name: 00000000800000000200000000612 Nehalem Host PC Monoplane Rev_III NO_HDD& 0000000080000000050000000000000000 Hard disk spare part&00000000800000000387582 PD.SCOMP.DCPS_20000000012V_5V Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Case: 0000000022269800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Limited nuber of times, confirmation of dispatch is requred SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-14 05:31 | No
599 | Yes | Startup failed | 0R/u FD0Startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Neut Step Action: FuMingYouSafety Qustions: Case: 0000000022269893 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the system coud not be entered after the coutdown reached 00000000, and engineers were requred to come to the site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The system couts down to 00000000 and cannot be entered. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC starts normally. 2. Check whether the IPPC power suply modue and BIOS battery are normal. 3. Check whether the hard disk indicator light is normal. Try bypassing the hard disk, and then reinstall the IPPC software. 00000000If there is no effect, ue PC Diagnostic Tool for diagnostic testing. Replace the corresponding spare parts 5. Check the machine uLLand perform faut repairs Recommended spare part uuand name: " 0000000080000000050000000092 IP PC Revised_II no HDD &0000000080000000050000000000000000Hard disk spare part " Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Problem Reported by Cutomer: Can’t start Internal Remarks: 13.00000000.2000000003 12:28:100000000TC (8600000000600000000) "RDI Accuacy: Yes" $Q7: (Confirmation & Description of problem/Can't boot u) $Q8: (Check the IPPC foud no inpu power, measued F23 Fue buned, IPPC and the fue need to be replaced) $Q9: (Replaced the backu fue and IPPC, reinstalled the IPPC software then restart system OK) $Q100000000 (Test system OK and inspection data refer to: PA. Device retuned to ful fuctionality. ) $Q11: Calibrated Tools:N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 13.00000000.2000000003 12:28:00000000 uC (8600000000600000000) "SuT. Ready for Clinical ue local time" : " 2000000003-00000000-13 15:00000000 Service content (requred): The equpment reported by the cutomer for repair coud not be started. After checking the equpment, it was foud that the image processing compuer was fauty. The image processing compuer was replaced and the software was reinstalled. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD IP PC Revised_II no HDD 0000000080000000050000000092 has arranged 00000000to the venu FuMingyou86-18621917586 ----------------------------------------------------------------------------- | 2023-09-14 05:51 | No
600 | Yes | 00 PROCEDURE ROOM MONITOR THE SCREEN IS BLUE AND SOFTWARE IS STUCK IN A LOOP - WILL NOT COMPLETE BOOT UP | 0000PROCEDURE ROOM MONITOR THE SCREEN IS BLUE AND SOFTWARE IS STUCK IN A LOOP - WILL NOT COMPLETE BOOT UP | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008232 100000000ep-2000000003 Byron Kirk The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in Eupected and ueupected Resuts: u uer Impact: system down had to move procedue Patient Impact: procedue moved to another lab SE Resoluion : replaced uuuand hard drive with In houe fse loaded software and configued ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action taken : FSE Ron Smith requsted call transferred to FSE Kirk Byron Oucome: dispatched Neut Step: N/A Internal Remarks : Planner Initials: GM Action taken : FSE Ron Smith accepted call for 9/13 Oucome: dispatched Neut Step: N/A Internal Remarks : **CuT REQuSTED ONSITE SuPORT Interaction Type: New Case Caller: MARK 900000000795300000000 IP & Issu: 65589000000003 PROCEDuE ROOM MONITOR THE SCREEN IS BLu AND SOFTWARE IS STuK IN A LOOP - WILL NOT COMPLETE BOOT u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND QuTE Neut Step Action: Awaiting Qute Verbal PO Nuber: 1988000000001 Verbal PO Approver: Verbal PO CAP: Neut Step Action : Per cutomer requst please dispatch O000000000000000000000000001to FSE Kirk Byron. Thank you----------------------------------------------------------------------------- | 2023-09-14 06:21 | No
601 | Yes | DFM0 Defibrillation cannot discharge | DFM0 Defibrillation cannot discharge | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Internal Remarks: 1000000009.2000000003 00000000:22:59 uC (860000000000000000100000000 $ASP engineer:[honglejia uu SHuI JIA] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem. Defibrillation cannot discharge $Q8:Inspection foud uage issus $Q9: Gude medical staff to ue $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11: $Calibrated Tools: N/A SAP Transfer: Dispatched for Onsite Suport Neut Step Action: 00000000553000000009900000000 Call Euternal Remarks: 1000000009.2000000003 00000000:22:00000000uC (860000000000000000100000000Department information: [Department: Equpment Department Contact Person: Teacher Li Contact Nuber: 1951100000000885]* $ASP engineer:[honglejia uu SHuI JIA] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem. Defibrillation cannot discharge $Q8:Inspection foud uage issus $Q9: Gude medical staff to ue $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11: $Calibrated Tools: N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Cutomer’s euplanation of the problem: The device prompts that the electrode plate is fauty The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: It is recommended to self-check and check the handle electrodes Recommended spare part uufor T2 with description: #98980000000019600000000 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA Safety Qustions: Case: 0000000022270000000019 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Electrode plate failue Neut Step Action: Taiyun Second People’s Hospital ----------------------------------------------------------------------------- | 2023-09-14 06:41 | No
602 | Yes | "System starting 0" is being ud and it won't start | "System starting 0" is being ud and it won't start | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008000000006 Noda, Kotaro The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: "System Starting 00000000" is being ued and it 000000000000000000000000001't start Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: "System Starting 00000000" is displayed and the system does not start. uer Impact: No effect Patient Impact: No effect SE 8.2.3000000005 Resoluion: Fee-paying equpment.FSE Matsuaki is negotiating 000000002150000000000000000 parts replacement for the fauty part in Case #2. A problem occured on the FleuVision control PC when the system was started.Problems have been occuing reguarly for a while, and we have sumitted a qute for a fee-paying part replacement, bu the problem has occured. This is a problem that has been a problem for a while, so we will repair the case in charge. I am closing this case. Internal Comments: [Suject] 00000000CS:KN:u Onisi: "System starting 00000000" is being ued and the system does not start. [Euamination category] Normal euamination ----------------------------------------------------------------------------- | 2023-09-14 06:51 | No
603 | Yes | V0uLLcannot be powered ON | V0uLLcannot be powered ON | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Escalation/V600000000uLLSafety Qustions: Case: 0000000022270000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: chen qian Last udated by: chen qian 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-14 07:51 | No
604 | Yes | System cannot power on. | System cannot power on. | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks : 1000000009.2000000003 00000000:3000000001 uC (88600000000000000002) $Q1: (NO) $Q2: (NO) $Q3: (NA,this equpment is IS) $Q00000000(NA,this equpment is IS) $Q5: (NA,this equpment is IS) $Q6: 00000000R_000000002_Cath_Fd10000000000000000System can not power on in today morning.293 $Q7: (Confirmation & Description of problem) confirmed reported problem. Fd10000000000000000System can not power on. Confirmation & Description of problem. $Q8: (Trouleshooting & Resuts) 1. uer feedback system can not power on 2. Check System and find Host PC OS HDD souhing wrong. 3. Replace Host OS HDD. $Q9: (Repair steps) 1. Sudown system. 2. Open M-cabinet cover. Remove defective Host OS HDD. 3. Replace a new one. 00000000Re-install system software. 5. Re-create image store. 6. Power on system. Test fuction,OK. 7. System retun to ful fuctionality. $Q100000000 (Test and Inspection data & Device statu) 1. Test and inspection data refer to: PA. 2. System retun to ful fuctionality. $Q11: Calibrated Tools name (N/A) N/A SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : 00000000R_000000002_Cath_Fd10000000000000000System can not power on in today morning.293 Euternal Remarks : 1000000009.2000000003 00000000:300000000 uC (88600000000000000002)$Q1: (NO) $Q2: (NO) $Q3: (NA,this equpment is IS) $Q00000000(NA,this equpment is IS) $Q5: (NA,this equpment is IS) $Q6: 00000000R_000000002_Cath_Fd10000000000000000System can not power on in today morning.293 $Q7: (Confirmation & Description of problem) confirmed reported problem. Fd10000000000000000System can not power on. Confirmation & Description of problem. $Q8: (Trouleshooting & Resuts) 1. uer feedback system can not power on 2. Check System and find Host PC OS HDD souhing wrong. 3. Replace Host OS HDD. $Q9: (Repair steps) 1. Sudown system. 2. Open M-cabinet cover. Remove defective Host OS HDD. 3. Replace a new one. 00000000Re-install system software. 5. Re-create image store. 6. Power on system. Test fuction,OK. 7. System retun to ful fuctionality. $Q100000000 (Test and Inspection data & Device statu) 1. Test and inspection data refer to: PA. 2. System retun to ful fuctionality. $Q11: Calibrated Tools name (N/A) N/A FSE Notes : This PO is a Reference Nuber only SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 0000000022270000000098 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sabrina Lin Last udated by: Sabrina Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-14 08:06 | No
605 | Yes | Low Leak-CO0ebreathing Risk | Low Leak-CO0Rebreathing Risk | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes: This PO is a Reference Nuber only SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000K/V600000000uLL T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Sensor failue The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Replace sensor Recommended spare part uufor T2 with description: 989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 Notes on cutomer requrements: na Safety Qustions: Case: 00000000222700000000300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 26.00000000.2000000003 00000000:00000000100000000uC (86000000000000000000000000 $ASP engineer:[weigung zhangweihu] $Q1:NO. $Q2:NO. $Q3:YES. $Q00000000ES. $Q7: Confirmed reported problem. Low Leak-CO2 Rebreathing Risk. $Q8:,Enter the diagnostic mode, the air delivery flow accuacy test verifies the accuacy of the air flow sensor and fuction of the blower. $Q9: Replace Flow Sensors. $Q100000000 Test Pressue .Test and inspection data refer to : Electrical safety test: passed (Prouimal pressue port GND resistance: 0000000000000000 Ω;Gas oulet port GND resistance:0000000000000000 Ω;Ouygen fitting GND resistance:0000000000000000 Ω;Forward leakage curent:0000000076 μA;Reverse leakage curent:0000000076 μA.Device has retuned to ful fuctionality $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000000000007) Neut Calibration Date (00000000/22/20000000000000000 TSI SN (000000008100000000000000009) Neut Calibration date (00000000/22/20000000000000000Euternal Remarks: 26.00000000.2000000003 00000000:0000000000000000 uC (86000000000000000000000000 *Department information [Department: The First Affiliated Hospital of GungzhouMedical uiversity Contact person: Feng Gong Tel: 13570000000028377]* $ASP engineer:[weigung zhangweihu] $Q1:NO. $Q2:NO. $Q3:YES. $Q00000000ES. $Q7: Confirmed reported problem. Low Leak-CO2 Rebreathing Risk. $Q8:,Enter the diagnostic mode, the air delivery flow accuacy test verifies the accuacy of the air flow sensor and fuction of the blower. $Q9: Replace Flow Sensors. $Q100000000 Test Pressue .Test and inspection data refer to : Electrical safety test: passed (Prouimal pressue port GND resistance: 0000000000000000 Ω;Gas oulet port GND resistance:0000000000000000 Ω;Ouygen fitting GND resistance:0000000000000000 Ω;Forward leakage curent:0000000076 μA;Reverse leakage curent:0000000076 μA.Device has retuned to ful fuctionality $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000000000007) Neut Calibration Date (00000000/22/20000000000000000 TSI SN (000000008100000000000000009) Neut Calibration date (00000000/22/20000000000000000SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-14 08:16 | No
606 | Yes | DFM0 test issue (Replaced ASSY Processor and reloaded the software) | DFM0 test issue (Replaced ASSY Processor and reloaded the software) | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 0000000022270000000081 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : 16000000009363-Test issu SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-14 08:36 | No
607 | Yes | Machine is not booting | Machine is not booting |  | 2023-09-14 08:41 | No
608 | Yes | Carc move fault | 0R/UNIQ FD0Carc move fault | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomer feedback: Carc is intermittently uable to move and requres engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Feedback Sensor uLL Remote diagnostics and test resuts: NA PRS TR statu: All pass Repair Action: Recommended on-site repair steps: " 1) Check encoder cable connection. 2) If sensor was connected with power active then the encoder coud get defective, replace motor/encoder 3) replace drive uit 00000000replace motor/encoder" Recommended spare part uuand name: 00000000800000000990000000052 AMC Triple Servo Drive HP-LP-LP"&00000000800000000262835 Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Euternal Remarks: 17.00000000.2000000003 11:13:00000000uC (86000000008238) Service content (requred): Cutomer feedback system prompts geometry movement partly available, C-arm motor control board problem, after replacement, the equpment retuned to normal, everything is normal Service resuts (requred) ): The system statu is normal and can be ued. Other important matters: None Internal Remarks: 17.00000000.2000000003 11:13:00000000uC (86000000008238)$Q1: (N/A) $Q2: (N/A) $Q3: (N/A) $Q00000000(N/A) $Q5: (N /A) $Q6: (geometry movement partly available) $Q7: (Confirmed reported problem, geometry movement partly available) $Q8: (check the log,AMC problem, need replace AMC) $Q9: (replace AMC) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (N/A)S/N (N/A) Neut Calibration date (N/A) SAP Transfer: Case Transferred to SAP Internal Remarks: 15.00000000.2000000003 1000000003:00000000 uC (8600000000900000000) TROJ Problem Reported by Cutomer: 00000000R/uIQ FD200000000 SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 15.00000000.2000000003 1000000003:00000000 uC (8600000000900000000)TROJ Neut Step Action: NTD: The cutomer said it coudn't move again. He supected that it was still a problem with the motor and motor control board that had been reported for repair before.Engineers are requred to come to the site for inspection as soon as possible. Refue to transfer to NNY. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 00000000800000000262835 1、00000000800000000990000000052 1. Shipped 200000000 Feng Lei -86-19815665565 Safety Qustions: Case: 0000000022270000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-14 09:01 | No
609 | Yes | Accept button does not work | Accept button does not work | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008596 18-Sep-2000000003 Miller, Jacquline The system meets the specification for the performed service and is retuned to ue Resoluion : accept buton not 000000000000000000000000001king. foud that the cable connection was disconnected from the accept buton board after taking it appart. Reseated the cable and it is 000000000000000000000000001king fine now. Was able to get to the diagnostic screen to pul the log file and change and verify butons were 000000000000000000000000001king. uit passed touh screen calibration, uit passed controlled tests. uit passed performance verification. uit retuned back to the cutomer to be be pu back in service. Problem Description by Engineer : Cutomer Fuction/Role: bio med Device ue: not in ue Eupected and ueupected Resuts: confirmed the accept buton was not 000000000000000000000000001king uer Impact: no Patient Impact: no SE Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000001-31 00000000:00000000:00000000 STATIC CONTROL DEVICE,200000000000000002-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : 00000000/15/23 YP 00000000000000000000000000100000000000000008596 Attempt00000000 Details: Cncted cut for qutes approval. Cut not reachable. uable to leave VM Contact patricia futz Phone (600000000) 927-3167 ##patricia.futz00rimedu.com Location Name St Dominic-Jackson Memorial Hosp Street 969 Lakeland Dr City JACKSON State MS Zip 39216-000000006 Internal Remarks : 00000000/13/23 YP 00000000000000000000000000100000000000000008596 Attempt1 Details: cncted cut for qutes. Cut not reachable. uable to leave vm Contact patricia futz Phone (600000000) 927-3167 ##patricia.futz00rimedu.com Internal Remarks : 00000000/1000000003 YP 00000000000000000000000000100000000000000008596 Details: Cncted cut. Cut stated that have not recd any qutes ## qutes ##sent to the cut. Contact patricia futz Phone (600000000) 927-3167 ##patricia.futz00rimedu.com Internal Remarks : 00000000/13/23 YP 00000000000000000000000000100000000000000008596 Attempt2 Details: Cncted cut for qutes. Cut not reachable. uable to leave vm Contact patricia futz Phone (600000000) 927-3167 ##patricia.futz00rimedu.com Facility - Phone: (600000000) 200000000-20000000000000000- Transferred to biomed (per son in biomed dept stated its wrong nuber) ----------------------------------------------------------------------------- | 2023-09-14 10:16 | No
610 | Yes | Shuts off uneupectedly | Shuts off uneupeudly | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : Vendor repaired. Ou Of Bou Failue? : N/A Labor: Labor ID : L-19800000000200000000; L-198500000000 000000000000000000000000001k Performed : u Med is u here for another t000000000000000000000000001weeks and shoud be able to evalute so we can get parts ordered. Looks like ay need batteries. ; Vendor repaired and retuned to service. Had to replace a washer. Checked ou fuction after and retuned to service. Diagonsis : Had to replace a washer. ; Had to replace a washer. | 2023-09-14 10:18 | No
611 |  | Getting application error | Getting application error |  | 2023-09-14 10:21 | No
612 | Yes | Product Hardware/Mp0fell and was damaged/Dropped | ProduuHardware/Mp0fell and was damaged/Dropped | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000999000000009-Sep-2000000003 Watanabe, Ryo The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: me Device ue: In ue Eupected and ueupected Resuts: The AED fell and is broken. uer Impact: No Patient Impact: No SEF Resoluion: D00000000u I confirmed that the display and main uit frame were damaged becaue the MP500000000was dropped. I repaired the part, inspected it, and carried ou a heat ru, and confirmed normal operation. Tools & Equpment : ELECTRICAL SAFETY TESTER, 20000000000000000000000006 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000056300000000161,IV-MP500000000MECHASY Frame 1.0000000000000000000000000000000056000000008000000001,IV-MP500000000ASSY Display LG-u Repl Kit 1.000000000000000000000000000000005600000000500000000IV-MP500000000ELO touh w/o gasket + bezel KIT 1.00000000000000000000000000000000261000000009511,IV-MP500000000ELO Touh Controller ----------------------------------------------------------------------------- | 2023-09-14 12:21 | No
613 | Yes | Rechner fährt nicht zu Ende hoch | Rechner fährt nicht zu Ende hoch | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006368 September 15, 2000000003 El Mahamid, Chadi The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. u Pausen is to be contacted. +00000000172 000000000000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Resoluion: The reported faut is confirmed. Problem Description: System does not start, trouleshooting FleuvisionPC OK, switsch withou power suply. 5V, 12 V and 200000000 suply in M cabinet are missing. Soluion: Power suply in M cabinet replaced, power suply available again, fuction test OK, the system corresponds to the herateller specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000387582,PD. SCOMP. DCPS_20000000012V_5V Tools & Equpment: MuTIMETER, 2000000000000000000000000000000000000000000000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006100000000 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. u Pausen is to be contacted. +00000000172 000000000000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer: System remains at 0000000000000000 when starting. Opened B cabinet with cutomer. Fleuvision PC does not show any fuction on the net000000000000000000000000001k card. Checked R cabinet with cutomer. TG switch withou voltage. According to the cutomer, the lamp is on at 12V A1. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: TE service to check switch and power suply switch. Replacement of the defective uit. Check whether system is on again. Otherwise replace Fleuvision PC. 1-3 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Cutomer comment: ## Appt. arranged with u Pausen for 15.00000000. from approu. 12:00000000 am. Internal Remarks: Consutation with u Pausen, I shoud call again aroud 8:00000000. ----------------------------------------------------------------------------- | 2023-09-14 12:26 | No
614 | Yes | triple chirp | triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006553 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : HS1 Generating triple chirp soud, Video is attached herewith, to order replacement. Diagnostic performed by Engineer : Triple chirp soud issu Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Follow u Requred : HS1 Generating triple chirp soud, Video is attached herewith, to order replacement. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006636 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Defib, u English 3ed, Euch ----------------------------------------------------------------------------- | 2023-09-14 12:51 | No
615 | Yes | Audio & Sounds/Speaker malfunction - no sound/ Speaker Assembly | Audio & Sounds/Speaker malfunuon - no sound/ Speaker Assembly | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003353 Hübler Thomas The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: By ##to u H., Thanks for the info. The louspeaker is definitely defective, determined by lateral euchange. Here is the information yourequsted so far • Did the louspeaker have no soud anymore or was a light, possibly distorted soud or a beeping soud heard? • no soud • Were any INOP uLLmessages displayed? A., If an INOP uLLwas displayed, was auio present duing this message? •No • Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? •K.a. • When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? •K.a. • Which software was installed on the device when the louspeaker malfuction occured? • L.00000000.00000000 Kind regards, Diagnostic performed by Engineer: +++NEMO KVA+++ Device, spare part is defective and mut be replaced. System is ou of ue util the installation of the spare part. Parts: 0000000056000000008621 Follow-u requred: 0000000056000000008621 Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Device/component is reported as defective. uer Impact: no Patient Impact: no SE uknown Issu: Malfuction uLLdescription: The louspeaker is defective. A new one is needed. Cutomer has not received qute. Requst for another delivery. uLLmessage:- uLLoccured on: 12.00000000.23 uIMC Diagnostic performed by Engineer: Dear sir or madam, ufortuately we are missing information for futher processing of 0000000022265789 qustion. Du to internal reguatory requrements, qustions mut be clarified abou the ue and the possible behavior in the alarm case before delivery of certain spare parts. In particuar, this applies to louspeakers. Here it essentially concerns the fact that with a loss of alarm this may have to be displayed to the responsible auhority. The following rue qustion catalog resuts from this: · Did the louspeaker have no soud or coud a light, possibly distorted soud or a beeping soud be heard? · · Were any INOP uLLmessages displayed? A., If an INOP uLLwas displayed, was auio present duing this message? · · Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? · · When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? · · Which software was installed on the device when the louspeaker malfuction occured? ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006812 Parts Consued: 1.0000000000000000000000000000000056000000008621.MS_u2 ASSEMBLY CBL u2/MP2 Speaker Assembly ----------------------------------------------------------------------------- Cutomer comment: By ##to u Hübler, Thanks for the info. The louspeaker is definitely defective, determined by lateral euchange. Here is the information yourequsted so far • Did the louspeaker have no soud anymore or was a light, possibly distorted soud or a beeping soud heard? • no soud • Were any INOP uLLmessages displayed? A., If an INOP uLLwas displayed, was auio present duing this message? •No • Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? •K.a. • When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? •K.a. • Which software was installed on the device when the louspeaker malfuction occured? • L.00000000.00000000 Kind regards, 000000000000000000000000001fgang Stuenvoll ----------------------------------------------------------------------------- | 2023-09-14 13:27 | No
616 | Yes | 0 / u0 / Electric shock plate scorch marks. | 00/ u000/ Eleuic shock plate scorch marks. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : 00000000R_119_MICu8612900000000Electric shock plate scorch marks FSE Notes : This PO is a Reference Nuber only Safety Qustions : Case: 0000000022271517 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-14 13:31 | No
617 | Yes | Machine is not functioning | Machine is not funuoning | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000012 Madhan, kuar The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Onsite visit needed to inspect the AED. Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : Onsite visit needed to inspect the AED. ----------------------------------------------------------------------------- | 2023-09-14 13:51 | No
618 | Yes | It does not load the start-up. | It does not load the start-up. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005715 19-Sep-2000000003 BuO, ROBERTO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: It does not load the start-u, therefore it does not start. How often is issu occuring? : Intermittent Malfuctioning area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: Diagnostic Eupected and ueupected Resuts: It does not load the start-u, therefore it does not start. uer Impact: low Patient Impact: no SE 8.1.17 Resoluion: Replaced SIB BOu which is the probable caue of the defect. Parts Consued: 1.00000000000000000000000000000000800000000000000002855, SIBBOu uIT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008693 18-Sep-2000000003 BuO, ROBERTO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: It does not load the start-u, therefore it does not start. How often is issu occuring? : Intermittent Malfuctioning area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: Diagnostic Eupected and ueupected Resuts: It does not load the start-u, therefore it does not start. uer Impact: low Patient Impact: no SE 8.1.17 Diagnostic performed by Engineer: Performed faut search finding SIB card which is the probable caue of the defect. 000000000000000000000000001k to be continud. Follow u Requred: Replacement of SIB Bou. ugent intervention: they are not down, bu the Cutomer does not trut to ue it for emergencies. Bring the thegrothermograph in order to monitor the temperatue in the technical room. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008600000000DAVIDE VECCHI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: It does not load the start-u, therefore it does not start. How often is issu occuring? : Intermittent Malfuctioning area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: Diagnostic Eupected and ueupected Resuts: It does not load the start-u, therefore it does not start. uer Impact: low Patient Impact: no SE 8.1.17 Diagnostic performed by Engineer: on indications from Ghiotti who went on site I order also IPPC Internal Comments: uLL# and/or description of uLL It does not load the start-u, therefore it does not start. Repair Action: Hw Follow u Requred: uLL# and/or description of uLL It does not load the start-u, therefore it does not start. Repair Action: Hw What are the neut Trouleshooting steps: Order PC to complete repair Technician Name: DAVIDE VECCHI ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000067800000000NDREA COLOMBO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: The coutdown is down at 00000000:00000000 A.M. The system does not 000000000000000000000000001k. When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: coutdown is down at 00000000;00000000 system is not 000000000000000000000000001king. uer Impact: no Patient Impact: no SE n.a. Diagnostic performed by Engineer: confirmed by the cutomer that the system is down, coutdown is down at 00000000:00000000. intervention will follow. Internal Comments: uLL# and/or description of uLL The cutomer confirmed that the system is down. The coutdown is down at 00000000:00000000 A.M. They tried to tun it on five times, bu it does not start. sugested onsite intervention with parts for diagnosis (1uhdd), 1 technician 1 day Repair Action: The cutomer confirmed that the system is down. The coutdown is down at 00000000:00000000 A.M. They tried to tun it on five times, bu it does not start. sugested onsite intervention with parts for diagnosis (1uhdd), 1 technician 1 day Follow u Requred: uLL# and/or description of uLL The cutomer confirmed that the system is down. The coutdown is down at 00000000:00000000 A.M. They tried to tun it on five times, bu it does not start. sugested onsite intervention with parts for diagnosis (1uhdd), 1 technician 1 day Repair Action: The cutomer confirmed that the system is down. The coutdown is down at 00000000:00000000 A.M. They tried to tun it on five times, bu it does not start. sugested onsite intervention with parts for diagnosis (1uhdd), 1 technician 1 day What are the neut Trouleshooting steps: The cutomer confirmed that the system is down. The coutdown is down at 00000000:00000000 A.M. They tried to tun it on five times, bu it does not start. sugested onsite intervention with parts for diagnosis (1uhdd), 1 technician 1 day Technician Name: ANDREA COLOMBO ----------------------------------------------------------------------------- | 2023-09-14 14:11 | No
619 | Yes | M0A / u00 Quote request for screen glass (Device is EoL and no work performed by uYy) | M00A / u000/ Quote request for screen glass (Device is EoL and no work performed by uYy) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 MELONI, FABRIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: EMAC clinical engineer Device ue: Diagnosis Eupected and ueupected Resuts: display problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from EMAC Clinical EMAC, they ask for an estimate for the parts. Resoluion: I heard from EMAC clinical EMAC, they ask for estimate for the parts, since they are old and old equpment, the spare parts cannot be foud anymore, I close the call and I sugest EOL letter. ----------------------------------------------------------------------------- Internal Remarks: uResucitation, P-1, Savona. IBA has been aware of the requst. Information was asked from the cutomer. ----------------------------------------------------------------------------- | 2023-09-14 14:31 | No
620 | Yes | Start-up problem | Start-up problem | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007332 18-Sep-2000000003 Yannick, LARuLLE The system meets the specifications for the performed service and is retuned to ue. Resoluion: Replacement of the Fleuvision PC and disk managing the large screen, Reloaded Fleuvision software, reloaded TSM software managing the touh screen in the euamination room, changed the u900000000wiring, tested restarts of the room and concluive operation, redistricting of the large screen displays to be done on the TSM according to the physician's preferences. Reguatory Qustions: No/No Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000058600000000, FleuVision-2_Revised PC no HDD Diagnostic performed by Engineer: problem starting the room. Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: niu Eupected and ueupected Resuts: Nok uer Impact: None Patient Impact: None SE 8.1.300000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006700000000 Nicolas WAGNER The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: CDP + HT YANNICK LARuLLE 15/00000000/2000000003 10000000000000000 Please deliver on site to the biomedical department for the attention of the u technician: FleuVision-2_Revised PC no HDD 0000000080000000058600000000u1 Disk for ClarityIQ_II and IP_Revised_II 00000000800000000500000000000000001 Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: niu Eupected and ueupected Resuts: Nok uer Impact: None Patient Impact: None SE 8.1.300000000Diagnostic performed by Engineer: At start-u, the coutdown remains frozen at 00000000:00000000:00000000. The Fleupc is probably the caue of the failue, I cannot commuicate with the system. Ordered the PC and schedue the local rt. ----------------------------------------------------------------------------- | 2023-09-14 14:37 | No
621 | Yes | power supply failure | 0/V0uLL0V power supply failure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/V600000000uLLInternal Remarks: The device prompts that the 35V power suply fails, please follow u with the buiness Safety Qustions: Case: 0000000022272000000005 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-14 15:01 | No
622 | Yes | DFM0 operation check fails | DFM0 operation check fails | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022272100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Failed to pass qulity inspection T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Operation check failed The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: Handle project failed Repair Action: T2 Recommended on-site repair steps: It is recommended to go to the scene to check the handle Recommended spare part uufor T2 with description: " 98980000000019600000000 uEuternal Paddles" Notes on cutomer requrements: na Test/Inspection: N/A Test Equpment: N/A Internal Remarks: 18.00000000.2000000003 0000000052:00000000uC (860000000000000000100000000 $ASP engineer:[HONGLEJIA YANG Hu] $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: Confirmed reported problem:Self-inspection cannot pass. $Q8:Self-inspection cannot pass becaue poor contact of the handle interface. $Q9:The handle interface is not inserted, and retuns to normal after tightening. $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: $Calibrated Tools: N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 18.00000000.2000000003 0000000052:00000000uC (860000000000000000100000000Department information: [Department; Equpment Department Contact: Teacher Chen Tel: 1800000000110000000051]* $ASP engineer:[HONGLEJIA YANG Hu] $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: Confirmed reported problem:Self-inspection cannot pass. $Q8:Self-inspection cannot pass becaue poor contact of the handle interface. $Q9:The handle interface is not inserted, and retuns to normal after tightening. $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: $Calibrated Tools: N/A ----------------------------------------------------------------------------- | 2023-09-14 15:06 | No
623 | Yes | During work you can not hear the sound. Pending duplicate confirmation. GFE0sent 0ec0 | During work you can not hear the sound. Pending duplicate confirmation. GFE0 sent 0Dec03 |  | 2023-09-14 15:11 | No
624 | Yes | geometry fault | 0N/u/geometry fault | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: PPD00000000800000000276500000000*1 Arranged for 200000000delivery to the venu FSE LiuMin-86-18627163378 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: 00000000N/u3 M15 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: Geometry faut Detailed description of cutomer problem or faut phenomenon: C arm cannot move Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The cutomer requres the engineer to be present Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Check the logs to see if it is the same as the previou failue Notes on cutomer requrements: Please contact the cutomer for on-site repair T2 Activities:Problem Description: Overview of cutomer problems: Geometry faut Detailed description of cutomer problem or faut phenomenon: C arm cannot move Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: An active buton on the control modue has been detected, release all to complete startu The cutomer requres the engineer to be present Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Arrive at site to inspect and test control modue butons Recommended spare parts uuand designation: uu00000000800000000276500000000 Control modue Standard ER Notes on cutomer requrements: Please contact the cutomer for on-site repair Safety Qustions: Case: 0000000022272116 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 18.00000000.2000000003 00000000:0000000000000000 uC (8600000000600000000) RDI Accuacy:YES $Q7: (Confirm reported problem &Geometry faut ) $Q8: (The log indicates that the C-arm buton is pressed, the power modue post self-test fails, the control modue is replaced, the software is downloaded, and the test is ued) $Q9: (replace the control modue down relate software after maintaining the system normal) $Q100000000 (Test and inspection data refer to PA, Device retun to ful fuctionality) $Q11: Calibrated Tools name:N/A For Mainland China CV / uR, CT engineer Euternal Remarks: 18.00000000.2000000003 00000000:0000000000000000uC (8600000000600000000) Faut description: Mechanical motion faut Service content: The log indicates that the C-arm buton was pressed, the modue post self-test failed when booting, and the control was replaced. Modue, download software, test and ue Service resuts: After repair, the equpment retuned to normal operation. ----------------------------------------------------------------------------- | 2023-09-14 15:12 | No
625 | Yes | u-ray | uo u-ray | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Internal Remarks: 1000000009.2000000003 13:0000000000000000 uC (86000000008900000000) "RDI Accuacy: yes" $Q7: (Confirmed reported problem, can't do flu, euposue is ok.No uLLshowing on data monitor.) $Q8: (connect footSwitch to control room.Test system, we find everything is ok.Reconnect footswitch to table, we find everything is ok.Check uLLlog, we find there is uLLfootswitch not pressed.So it maybe the footswitch connector uLL ) $Q9: (reconnect footswitch .system retuned to normal.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (N/A) S/N (N/A) Neut Calibration date (N/A) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage:N/A Euternal Remarks: 1000000009.2000000003 13:0000000000000000 uC (86000000008900000000)*Department information: [Department: Operating Room Contact Person: Li Pengfei Tel: 136733523500000000* Service Content: [Perspective can only be ued for a single frame, and the euposue is normal.The equpment retuned to normal after repairs. ] Service resuts: [After repair, the equpment retuned to normal operation. ] Other important matters: [none] SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 1000000009.2000000003 00000000:25:00000000uC (86000000008900000000)"SuT. Ready for Clinical ue local time" : " 2000000003-00000000-1000000000000000000000000" Safety Qustions: Case: 00000000222721100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: No u-ray Trouleshooting Action: Key faut codes or phenomena: It cannot be viewed throuh perspective, and the logs on PRS cannot be obtained for the curent day. PRS TR statu: All Access Repair Action: Recommended on-site repair steps: 1. Check the on-site inspection log Recommended spare parts uuand designation: 00000000210000000013887 Power Inverter (PoInt) Certeray T2 Activities:Problem Description: Overview of cutomer problems: No u-ray Trouleshooting Action: Key faut codes or phenomena: It cannot be viewed throuh perspective, and the logs on PRS cannot be obtained for the curent day. PRS TR statu: All Access Repair Action: Recommended on-site repair steps: 1. Check the on-site inspection log Problem Reported by Cutomer: Cannot see throuh ----------------------------------------------------------------------------- | 2023-09-14 15:12 | No
626 | Yes | The system does not start. | The system does not start. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000730000000029-Sep-2000000003 Kawamua, HidetsuuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Not ued Eupected and ueupected Resuts: The system does not start. uer Impact: None Patient Impact: None SE 1.2.5 Resoluion: 00000000000000000000000000100000000000000006866 reported. It is 000000000000000000000000001for parts processing. Parts Consued: 1.00000000000000000000000000000000800000000717981,Sute PC W7 + W100000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006866 100000000ep-2000000003 Kawamua, HidetsuuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Not ued Eupected and ueupected Resuts: Review does not start. uer Impact: None Patient Impact: None SE 1.2.5 Resoluion: I replaced the system main control PC (sute PC). After I replaced it, I installed it and confirmed normal operation. I 00000000000000000000000000100000000000000006886 report on this, incluing the 00000000000000000000000000100000000000000006886 000000000000000000000000001k. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006886 100000000ep-2000000003 Niiya, Kosue The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: Not ued Eupected and ueupected Resuts: The Sute PC does not start. uer Impact: No effect Patient Impact: No effect SER1.2.5 Resoluion: 00000000000000000000000000100000000000000006866 000000000000000000000000001k report and received the cutomer's signatue. ----------------------------------------------------------------------------- | 2023-09-14 15:31 | No
627 | Yes | Device does not turn on | Device does not turn on |  | 2023-09-14 15:51 | No
628 | Yes | failed to produce u-rays (TW00_split | failed to produce u-rays (TW0000_split | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008613 15-Sep-2000000003 Darren Brennan The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System not screening uer Impact: NA Patient Impact: NA SEu Resoluion : Replaced IP PC along with OS HDD and Image Disk drives. Reloaded SW. System tested and issu is resolved. Retuned to cutomer ready for ue. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to replace IP PC / check arm movements as BG /collision uLL present Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: System not screening uer Impact: NA Patient Impact: NA SEu Diagnostic performed by Engineer : Logging shows IP uLL/image disk ful . PC post failing ----------------------------------------------------------------------------- Internal Remarks : Zita was not available. Asked her colleagu to requst a log file is saved on the system so I can investigate. uers reporting c arm movement issus and system not screening which are 2 very different fauts Internal Remarks : Hello Zita , I jut spoke to you colleagu regarding Lab 2. Please can someone save me a log file on the system so I can investigate the reported problems. Please can youcall/message me on + 000000007867 500000000913 once the log file is saved. Many thanks Barry Internal Remarks : Cutomer called advised the log file has been sent, please can you000000003000000002800000000 ----------------------------------------------------------------------------- | 2023-09-14 15:56 | No
629 | Yes | Too pressed, client has mastitis, caus pain | 00POL SCF000 PR#000 00000 too pressed, client has mastitis, caus pain | 2000000003-10000000016 13:12:16:follow u @ | 2023-09-14 16:36 | No
630 | Yes | udisplayed interrupted signal - Legal | u- udisplayed interrupted signal - Legal |  | 2023-09-14 16:48 | Yes
631 | Yes | udisplay intermittent tracing | u- udisplay intermittent tra0g - Legal Issue_ AE-0-008 |  | 2023-09-14 17:03 | Yes
632 | Yes | Data monitor is black. | Data monitor is black. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 September 200000000 2000000003 MARTIN LITT The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Infi cheffe Device ue: / Eupected and ueupected Resuts: Data monitor goes black. uer Impact: / Patient Impact: / SE/ Resoluion: Replacement of the monitor and the VGA-uP converter to solve the problem of the data monitor that does not "tun on" intermittently. Parts Consued: 1.00000000000000000000000000000000800000000287861, Kit Pan Knob TSO WP 1.00000000000000000000000000000000230000000016822, KVM RECEIVER (CRCB) 1.00000000000000000000000099193225213000000009" Color LCD Monitor CR (DC19-LCR) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000955 ANTOINE SANTANGELO The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: Martin onsite 20000000000000000: 9:00000000 -11:00000000 Part: 00000000230000000016822 and 9919322521300000000 Delivery: Hospital to LIZA Problem Description by Engineer: Cutomer Fuction/Role: Infi cheffe Device ue: / Eupected and ueupected Resuts: Data monitor goes black. uer Impact: / Patient Impact: / SE/ Diagnostic performed by Engineer: Check monitor and KVM eutender. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008979 September 19, 2000000003 Verckens, Maarten The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Nuse headfeed Device ue: Yes Eupected and ueupected Resuts: Data monitor black. uer Impact: / Patient Impact: / SE / Resoluion: No signal on data monitor ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008862 ANTOINE SANTANGELO The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: Maarten onsite 100000000 15:00000000 1000000009 Problem Description by Engineer: Cutomer Fuction/Role: Nuse headfeed Device ue: Yes Eupected and ueupected Resuts: Data monitor black. uer Impact: / Patient Impact: / SE / Diagnostic performed by Engineer: check monitor and cables. ----------------------------------------------------------------------------- | 2023-09-14 17:06 | No
633 | Yes | The collimator does not work | The collimator does not work | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008962 18-Sep-2000000003 Roberto Fernandez Moreno The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The collimator seems to be failing, J. Morales is in charge of replacing it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The collimator gets stuk. uer Impact: It cannot collimate. Patient Impact: No impact SE 1.2.5.00000000Diagnostic performed by Engineer: Collimator shuter uLL Resoluion: The collimator of the equpment is replaced according to the service manul. Software is installed and proper operation is checked. Parts Consued: 1.00000000000000000000000000000000800000000858161,NICOL V3 CV FD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006700000000Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The collimator seems to be failing, J. Morales is in charge of replacing it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The collimator gets stuk. uer Impact: It cannot collimate. Patient Impact: No impact SE 1.2.5.00000000Diagnostic performed by Engineer: The collimator seems to be failing, J. Morales is in charge of replacing it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The collimator seems to be failing, J. Morales is in charge of replacing it. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: The collimator seems to be failing, J. Morales is in charge of replacing it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-09-14 17:06 | No
634 | Yes | Therapy supply test failed. | 000Effi0 DFM0 Therapy supply test failed. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 25-Sep-2000000003 Matteo Carraro The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Alberta co-construtor Device ue: Diagnosis Eupected and ueupected Resuts: problem with suply uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from Coord Alberta who reported problems with e-suply. Resoluion: Check faut; the part on the list was replaced. The system is 000000000000000000000000001king perfectly again. Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000926 19-Sep-2000000003 FALASCA, IVO The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Technician Matteo Carraro asks to send a new condenser in replacement of the one that arrived defective. Problem Description by Engineer: Cutomer Fuction/Role: Alberta co-construtor Device ue: Diagnosis Eupected and ueupected Resuts: problem with suply uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I heard from Coord Alberta who reported problems with e-suply. Diagnostic performed by Engineer: Technician Matteo Carraro asks to send a new Condenser in replacement of the one that arrived defective. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005285 19-Sep-2000000003 Matteo Carraro The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Requst for new condenser for DEFOA. Problem Description by Engineer: Cutomer Fuction/Role: Alberta co-construtor Device ue: Diagnosis Eupected and ueupected Resuts: problem with suply uer Impact: not in ue by the cliniuo Patient Impact: not in ue on patient SE na I heard from Coord Alberta who reported problems with e-suply. Diagnostic performed by Engineer: Runed condenser cable. Follow-u is opened for new part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001000000006 MELONI, FABRIZIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Alberta co-construtor Device ue: Diagnosis Eupected and ueupected Resuts: problem with suply uer Impact: not in ue by the cliniuo Patient Impact: not in ue on patient SE na I heard from Coord Alberta who reported problems with e-suply. Diagnostic performed by Engineer: Heard from coord Alberta who reported e-mail problems: duing test or daily fuctional test it does not pass the suply of the therapy and remains in alarm, system is uder contract, I order parts to assign On site Follow u Requred: Heard from coord Alberta who reported e-mail problems: duing the daily test or fuctional test, it does not pass the suply of the therapy and remains in alarm, from system it is uder contract, I order parts to assign On site ----------------------------------------------------------------------------- | 2023-09-14 17:16 | No
635 | Yes | Check Vent: Machine Pressure Sensor Auto-Zero Failed | Check Vent: Machine Pressure Sensor Auto-Zero Failed |  | 2023-09-14 18:00 | No
636 | No | This PR is a Duplicate of PR # 00 which is the PR ID of the actual complaint. | This PR is a Duplicate of PR # 0000 which is the PR ID of the aual complaint. |  | 2023-09-14 18:16 | No
637 | Yes | OVP Circuit Failed | OVP 0cuit Failed |  | 2023-09-14 18:37 | No
638 | Yes | The system is not starting up; a page appears with a countdown before then turning off again. | The system is not starting up; a page appears with a countdown before then turning off again. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005698 17-Sep-2000000003 John Benstead The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: The system is not starting u; a page appears with a coutdown before then tuning off again. uer Impact: NA Patient Impact: NA SE8.2.17 Diagnostic performed by Engineer : Replaced LVPS in M cabinet. Topped u detector coolant. Tested system operational. All OK. Parts Consued : 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V 2.00000000000000000000000000000000200000000000000007181,GLYCOSHELL SF500000000 5L CAN ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009619 15-Sep-2000000003 uGreen000000000000000000000000001d The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace low voltage power suply at the rear lower section of M-Cab. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: The system is not starting u; a page appears with a coutdown before then tuning off again. uer Impact: NA Patient Impact: NA SE8.2.17 Resoluion : System failed to boot u correctly or operate within FSF environment. Geometry, TSM and M-Cab net000000000000000000000000001k switch were all uavailable. Traced the faut to the low voltage power suply at the rear of the M-Cabinet. This provides 20000000012 and 5 Volts. Power to it OK. I foud an old intermittently fauty one on side lukily. Installed this one and the system fuly boots and is operational. I can't declare the qulity of the part so I will leave the system as OOA. I will order the power suply and chiller flud for Cath and Vasc labs. Tools & Equpment : MuTIMETER,200000000000000002-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Supect FV PC , FSE to site to investigate Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: The system is not starting u; a page appears with a coutdown before then tuning off again. uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : System sitting at 00000000:00000000 on coutdown ----------------------------------------------------------------------------- Cutomer comment : 00000000279823200000000additional nuber to contact cutomer on. Internal Remarks : Confirmed with Angela for 17/9 9am access code is C3569Z ----------------------------------------------------------------------------- | 2023-09-14 18:46 | No
639 | Yes | When booting system, stuck at zero's. | When booting system, stuck at zero's. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001312 16-Sep-2000000003 Wright, Omar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : when booting system, stuk at zero's When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: when booting system, stuk at zero's uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Parts Consued : 1.00000000000000000000000000000000216700000000756,SIB_u_HB Resoluion : Replace SIB board restored ghost image perform download board software system, 000000000000000000000000001king. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009761 100000000ep-2000000003 Nevo, Eyal The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace SIB board Resoluion : Generator test done by cutomer. System does not boot properly. Issu caued by SIB connectivity issus. Replacement requred Problem Description by Engineer : Problem description by engineer : when booting system, stuk at zero's When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: when booting system, stuk at zero's uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009682 Ketcherside Jeff The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: when booting system, stuk at zero's Repair Action : 00000000230000000028782 - Service part Single Board Compuer Diagnostic performed by Engineer : cutomer states when booting system, stuk at zero's Problem Description by Engineer : Problem description by engineer : when booting system, stuk at zero's When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: when booting system, stuk at zero's uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Follow u Requred : uLL# and/or description of uLL: when booting system, stuk at zero's Repair Action : 00000000230000000028782 - Service part Single Board Compuer Technician Name : Ketcherside Jeff ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022280000000019 FSE: Nevo, Eyal After Hospital Generator Test system fails to boot u, stuk at 0000000000000000. Reviewed logs; instruted FSE to implement the following soluions: "Caue 1 Problem with the hospital Mains power Soluion Check the hospital Mains inpu power to the Velara Generators. Caue 2 Coud be caued du to SIBBOu malfuctioning Soluion 1 The HostPC commuicates via the SIBBOu with the Velara generators. Since commuication is lost with both generators at the same time, SIBBOu can be the root caue. Soluion 2 Check in logfile what euactly happened jut before commuication was lost. A defective uerInterface can caue an interrution of the CAN commuication. Check connected u's for damage or malfuction." Internal Remarks : Case: 0000000022280000000019 FSE: Nevo, Eyal FSE checked fue and incoming power; pass. FSE uable to Download BD SW to MDY. Instruted FSE to place part order for 00000000216700000000756 - SIB_u_HB. Internal Remarks : Case: 0000000022280000000019 FSE: Wright, Omar Generator test, system stuk at 0000000000000000 duing boot u. FSE states they are not able to go into FSF. uable to retrieve logs from PRS; system will ping bu is not reachable. Instruted FSE to reload Ghost in attempt to get access to FSF to grab logs. ----------------------------------------------------------------------------- | 2023-09-14 19:01 | No
640 | Yes | Wireless footswitch dosen't work. Its green and WIFI symbol is visible but dosen't work. | Wireless footswitch dosen't work. Its green and WIFI symbol is visible but dosen't work. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009963 26 Sep-2000000003 Pousen Morten Kaae The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : The wireless foot switch is not always 000000000000000000000000001king. The system has to be rebooted every morning. How often is the issu occuring? : Mutiple times per week Malfuction area: Patient suport Information to suport the complaint handling process. Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: wireless foot switch not 000000000000000000000000001king uer Impact: na Patient Impact: Not provided SE 000Resoluion: Parts only retun, no 000000000000000000000000001k performed. The case is raised on the wrong system. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008615 Moeller, Dennis The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : The wireless foot switch is not always 000000000000000000000000001king. The system has to be rebooted every morning. How often is the issu occuring? : Mutiple times per week Malfuction area: Patient suport Information to suport the complaint handling process. Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: wireless foot switch not 000000000000000000000000001king uer Impact: na Patient Impact: Not provided SE 000Diagnostic performed by Engineer: have downloaded and analyzed logs. have checked RSW for a soluion. Believe it's the base station that is giving the issu. It might be that the bad relay antenna is inside. I will order a new set. Follow u Requred: Dear ccc please book onsite FSE to check and replace WPS WFS 2 parts have been ordered. The euposue signal is not always going to the system. Repair Action : New foot switch kit needed. Technician Name: Moeller, Dennis Resoluion: replace WFS Internal Comments: uLL# and/or description of uLL: The euposue signal is not always going to the system. Repair Action : New foot switch kit needed. ----------------------------------------------------------------------------- | 2023-09-14 19:26 | No
641 | Yes | Shuts down with no alrms while running on battery. | Shuts down with no alrms while running on battery. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009635 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000uLLshu down while on a patient Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: 2.300000000 Date and Time of Issu: 00000000/21/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version R. Biomed has confirmed V600000000diagnostic code 1100000000 in the significant event log. uLLrestarted du to anomalies detected duing operation. This coud inclue invalid or corruted information, uanticipated situtions, or object allocation uLL. NOTE If diagnostic code 1100000000 occus in conjuction with diagnostic code 3000000007 (software euception), no action is requred. Biomed has confirmed diagnostic code 3000000007 is not in the log. Cutomer is also reporting the uit euuuenced the same problem on 00000000/00000000/2000000003 and the software was reloaded. Cutomers puchased options AVAPS, C-Fleu, Ramp 1. Reinstall operational software. 2. Replace the CPuPCBA. Part # 00000000561537200000000RP-PCBA,CPu(Software 2.300000000, 2nd Gen, V600000000List price $792. (26) in stock Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not shu down. uer Impact: V600000000has been removed from service. Patient Impact: Biomed is not reporting and adverse oucome to patient care or therapy. The uLLwas removed from service and the patient was placed onto another uit. SE 3.00000000 Problem Description/ueupected Resuts: V600000000uLLshu down while on a patient. Internal Comments : Cutomer has not responded to mutiple requsts for V600000000repair udates. Closing case. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ben Mylu 570000000033325 IP & Issu: 10000000011900000000/Shus down with no alrms while runing on battery. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued ot RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-14 19:31 | No
642 | Yes | Problem with the ret - Insufficient information - Cau Unknown | Problem with the ret - Insuffi0nt information - Cau Unknown | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: IED department - - - The Servo-urespirator does not transmit the alarms on the central station anymore. The LED of the EC00000000houing flashes green. The cutomer connected another respirator (Servo-i) and it 000000000000000000000000001ks. We connected to the 00p EC00000000(172.21.200000000.123) and the Servo-udriver is present. The cutomer will do an EC5/other Servo-ucross-test to refine the diagnosis. Waiting for feedback. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Resoluion: No feedback from the cutomer to date. Closing the case. If necessary, we will open a follow-u of the case. Resoluion: Hello u JuIEN, uless I am mistaken, I did not get any feedback regarding this EC00000000case. Were youable to do the cross-tests as we had mentioned? With another EC5 / and another Servo-u What is the resut? Were youable to determine the caue of the problem? Waiting for you feedback, Sincerely, Gullaue PANET Resoluion: Hello u Panet, The IS is ful; I was uable to do the test. Thank you Kind regards Jérémy JuIEN Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the alarms of a respirator. ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-09-14 20:06 | No
643 | Yes | Request to check calibration and malfunction | Request to check calibration and malfunuon | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005659 28-Sep-2000000003 Peconio, Alberto The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: Diagnosis Eupected and ueupected Resuts: calibration problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na Clinical Ems reports calibration problems. Resoluion: Carried ou nbp calibration ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000023500000000MELONI, FABRIZIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: Diagnosis Eupected and ueupected Resuts: calibration problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na Clinical Ems reports calibration problems. Diagnostic performed by Engineer: clinical engineer reports calibration problems, and fuctional tests. From the system it does NOT appear to be uder contract, I open Qute for safety, please verify if it was renewed recently, then assign on site. Follow u Requred: Clinical Ems reports calibration problems, and fuctional tests. From the system it does NOT appear to be uder contract, I open Qute for safety, please verify if it was renewed recently, then assign on site, equpped with the necessary tolls. ----------------------------------------------------------------------------- | 2023-09-14 20:11 | No
644 | Yes | Connectivity/Unable to view data: Info on patient death while on Tele & PIIC. NTF/ur uLLdischarging. Provided Info | Conneuty/Unable to view data: Info on patient death while on Tele & PIIC. NTF/ur uLLdischarging. Provided Info | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Simon Goldberger The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SEuknown Resoluion : 000000000000000000000000001raised in uLL actul 000000000000000000000000001for this job is 000000000000000000000000001000000000000000000000000600000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000600000000imon Goldberger The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical engineer Device ue: In ue Eupected and ueupected Resuts: Cutomer stating patient died wearing TEL8 duing the hous of 3-6am in CCu uer Impact: None Patient Impact: Death SEB.00000000.18 Reguatory Qustions : N/A Diagnostic performed by Engineer : Gathered the logs and sent them to Q&R and Lynn Bowers for analysis Resoluion : Speaking to Rebecca, she's happy with the report and breakdown, all ok now, closing case. Internal Comments : Cutomer away till 00000000/100000000 pending till then. Internal Comments : Asked cutomer for an udate, pending their response. Diagnostic performed by Engineer : sent attached to the cutomer. pending their response. ----------------------------------------------------------------------------- | 2023-09-14 20:51 | Yes
645 | Yes | uPro ¦ The capnography is defective. The line is flat even when they blow in the tubing. | TempuPro ¦ The capnography is defeu. The line is flat even when they blow in the tubing. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000200000000September 21, 2000000003 Franck PETIT The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: No. Patient Impact: No. SE V0000000028 Resoluion: We received the device yesterday afternoon. Last night we tested the setting and after a while we got the message "check sample line blocked". In reality, what was cauing the problem with the device was the retun flow. A gas sample is collected from the patient, passes throuh the sample line, the gas modue analyzes the sample and eupels the flow ou of the device. The eupiratory port for the CO2 parameter is the one yousee in the small hole to the left of the inpu connectors. It was this flow that the device was uable to eupel. This morning when dismantling the device, we noticed the overlapping of the pipes and the data cable of the capno modue. We believe that it is this suerposition which caued a settling of the eupiratory flow line and caued the pup to fail. To avoid this overlap and possible movement of the flow pipes duing ue of the device, the pipes and data cable have been reorganized. Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000005-0000000000000000:00000000:00000000 PATIENT SIMuATOR, 200000000000000003-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000055 Seva, Romain The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: No. Patient Impact: No. SE / Diagnostic performed by Engineer: The capnography does not 000000000000000000000000001k anymore. There is no uLLmessage and the tuing is recognized. However, when they blow in the cable, no measuement. The line remains flat. The tuing of the capnography was recently replaced when the device was in Suesnes (u premises). Resoluion: Sending the device back to Suesnes for diagnosis. Follow u Requred: The device will be sent back to Suesnes. Suport Franck Petit is to be schedue for the diagnosis of the device. ----------------------------------------------------------------------------- | 2023-09-14 20:51 | No
646 | Yes | Video feed on monitor ceiling suension is intermittently cutting in and out | Video feed on monitor ceiling suension is intermittently cutting in and out | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002631 29-Sep-2000000003 Jepsen Larry The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: clinical Eupected and ueupected Resuts: intermittent display on large screen uer Impact: intermittent display on large screen Patient Impact: none SE 1.2.5 Resoluion : Tested operation. uable to dulicate, believe it was a bad connection to wcb, System 000000000000000000000000001king properly at this time. ----------------------------------------------------------------------------- | 2023-09-14 20:51 | No
647 | Yes | Charging Issue | Charging Issue |  | 2023-09-14 20:56 | No
648 | Yes | Not going into standby mode | Not going into standby mode | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : 00000000/27/2000000003 Biomed is reporting completing the zero flow calibration has resolved the problem. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000to enter standby mode. uer Impact: V600000000has been removed from service Patient Impact: No impact SE 3.100000000 Problem Description/ueupected Resuts: V600000000will not stay in standby. Diagnostic performed by Engineer : Problem Description: V600000000will not stay in standby mode. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: 3.100000000 Date and Time of Issu: 00000000/100000000000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Biomed advised to complete zero flow calibration and retest the uit. The cutomer is reporting they don't curently have access to their test equpment. Cutomer has advised to call back neut week. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Aleu Courouh IP & Issu: " 9000000003711 not going into standby mode." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-14 21:11 | No
649 | Yes | Split- The screen turns on then freezes. NEMO 000 SS MECHASY VM0M0rnt Pnl w/o LCD FS | Split- The screen turns on then freezes. NEMO 00000, SS MECHASY VM0VM0Frnt Pnl w/o LCD FS | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007392 Assad MAJEED The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Recommendation: 000000005600000000511 Part requsted for replacement. Follow u Requred: Delivery of the part to the biomedical department if contract, otherwise qute. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that the front cover is defective. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000031 Parts Consued: 1.000000000000000000000000000000005600000000511, SS MECHASY VM00000000M6 Frnt Pnl w/o LCD FS ----------------------------------------------------------------------------- Internal Remarks: u called again and wants to be called back by u MAJEED, he wants to confirm the part, his available and the delivery times. ----------------------------------------------------------------------------- | 2023-09-14 21:11 | No
650 | Yes | It has not supplied rays. | It has not supplied rays. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000035 DAVIDE VECCHI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: It is reported that this morning the system did not deliver the u-rays. Requst for remote check. Hemodynamics Service. How often is issu occuring? : Intermittent Malfuctioning area: Image Chain Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: It is reported that this morning the system did not deliver rays. Requst for remote check. Hemodynamics Service. uer Impact: no Patient Impact: no SE 8.2.16.1 Diagnostic performed by Engineer: contacted colleagu Marci who was present on site for PM on neut system, logs report ful image disk, Marci has carried ou re create image store and instruted the cutomer on correct flow of closing euaminations in order not to fill the disk again Internal Comments: uLL# and/or description of uLL It is reported that this morning the system did not deliver the u-rays. Requst for remote check. Hemodynamics Service. Repair Action: Sw Resoluion: contacted colleagu Marci who was present on site for PM on neut system, from the logs the image disk is ful, Marci has carried ou the Re create image store and instruted the cutomer on correct flow of closing euaminations in order not to fill the disk again ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000900000000AVIDE VECCHI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: It is reported that this morning the system did not deliver the u-rays. Requst for remote check. Hemodynamics Service. How often is issu occuring? : Intermittent Malfuctioning area: Image Chain Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: It is reported that this morning the system did not deliver rays. Requst for remote check. Hemodynamics Service. uer Impact: no Patient Impact: no SE 8.2.16.1 Diagnostic performed by Engineer: Contacted u. Ina Melas, from the analysis of the logs that were created the image disk is ful, same problem as last April resolved with initialization of the image disk remotely, are there also many archiving uLL, EPu with wrong archive data? Internal Comments: uLL# and/or description of uLL It is reported that this morning the system did not deliver the u-rays. Requst for remote check. Hemodynamics Service. Repair Action: Sw Follow u Requred: uLL# and/or description of uLL It is reported that this morning the system did not deliver the u-rays. Requst for remote check. Hemodynamics Service. Repair Action: Sw What are the neut Trouleshooting steps: It is necessary to contact the cutomer in order to schedue activity of database deletion and check of archive data on EPu They ask to be called on September 15 in the morning. Technician Name: DAVIDE VECCHI ----------------------------------------------------------------------------- Internal Remarks: Contract eupired on 31/7/2000000003. I ask for a requst for verification. ----------------------------------------------------------------------------- | 2023-09-14 21:11 | No
651 | Yes | Alarms or Messages/Monitor did not generate an alarm when the patient euuuences a seriouproblem/ur Error | Alarms or Messages/Monitor did not generate an alarm when the patient euuuences a seriouproblem/ur Error |  | 2023-09-14 21:26 | No
652 | Yes | Fleuvision pc intermittently not booting up | Fleusion pc intermittently not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002995 15-Sep-2000000003 Cupton Daniel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: CV Device ue: CV Eupected and ueupected Resuts: Fleuvision PC intermittently not booting uer Impact: none Patient Impact: none SE8.1.300000000Diagnostic performed by Engineer : Installed new FleuVision PC. Puhed O/S software along with Application, and Firmware. Reloaded configuation settings and tested operation. The system is now operating to u specifications. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Daniel Cupton - MacLab image on FleuVision shows u with red hu, same image on MacLab monitor is ok. Discused trouleshooting method on how to find root-caue of the issu. Internal Remarks : Daniel - Fleu PC replacement. Told FSE to reload according to instrutions; reload of OS and app likely necessary. Internal Remarks : FSE Daniel Cupton. FSE has system that Fleuvision PC doesn't power on with system. When buton is pressed the compuer and system 000000000000000000000000001k. Check cabling. Approved p/n: 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD. ----------------------------------------------------------------------------- | 2023-09-14 21:46 | No
653 | Yes | system did not boot up successfully as the start-up countdown got stuck at 0:0. | system did not boot up successfully as the start-up countdown got stuck at 0:0. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003620000000019-Sep-2000000003 garcia,john The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : boot u cout to 00000000no futher, tun off took long time reboot attempting When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: cannot get uit to boot uer Impact: room down Patient Impact: room down SE 8.1.3000000005 Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Diagnostic performed by Engineer : Replaced the fleuvision pc and replaced the OS disk. I verified that the uit fuly booted. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001700000000 Follmer, Crystal The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Assisted in trouleshooting uit and verifying Fleuvision PC was in need of replacement. John Garcia completed replacement. Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Fluro Eupected and ueupected Resuts: System will not boot u. uer Impact: Down Patient Impact: Down SE N/a ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001000000008 100000000ep-2000000003 garcia,john The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Continu trouleshooting the boot-u issus once the room has cooled. Problem Description by Engineer : Problem description by engineer : boot u cout to 00000000no futher, tun off took long time reboot attempting When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: cannot get uit to boot uer Impact: room down Patient Impact: room down SE 8.1.3000000005 Diagnostic performed by Engineer : The ambient room temperatue euceeded 900000000degrees. Facilities was notified to handle the room temperatue. I will go on site Friday morning once the room and cabinets have cooled down to continu trouleshooting. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001000000006 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uLL# and/or description of uLL: cannot boot Repair Action : check system power and PCs and net000000000000000000000000001k switch Technician Name : Clark, Byron Problem Description by Engineer : Problem description by engineer : boot u cout to 00000000no futher, tun off took long time reboot attempting When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: cannot get uit to boot uer Impact: room down Patient Impact: room down SE 8.1.3000000005 Diagnostic performed by Engineer : cannot get uit to bot tried reboot 2 times no progress Internal Comments : uLL# and/or description of uLL: cannot boot Repair Action : check system power and PCs and net000000000000000000000000001k switch ----------------------------------------------------------------------------- Internal Remarks : John Garcia - high technical rooms temperatues. Cutomer asking abou eAlert notifications. Sent ##to MPM team and RSE manager Dave Hall. Internal Remarks : Case: 0000000022280000000000000000 FSE: Follmer, Crystal R8.1.300000000 FSE states FleuVision PC fails to reload. Requsted logs for review. Reviewed logs; foud "Failed to initialize all grabber cards: DgInit failed: 99." Instruted FSE to place part order for 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD and 0000000080000000050000000000000000- Hard disk spare part. Additionally, inform the cutomer to not ue the system util equpment room environment is withing u Specs (15C - 300000000). Internal Remarks : Instruted FSE to perform download board software for APP and firmware now that the OS has completed ----------------------------------------------------------------------------- | 2023-09-14 21:56 | No
654 | Yes | Alarm/Failure to alarm:Patient Heart Rate euceeded High Limit but no alarm on PIICiu: Unknow Cau: Provided Information | Alarm/Failure to alarm:Patient Heart Rate euceeded High Limit but no alarm on PII0: Unknow Cau: Provided Information | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003397 19-Sep-2000000003 Chase Meicher The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Assist Device ue: mu700000000 Eupected and ueupected Resuts: did not alarm duing patient event uer Impact: no alarm Patient Impact: no alarm SEm Diagnostic performed by Engineer : Reviewed logs for bedside/mu700000000. Determined alarming as it shoud, event may have been an anomaly. Tested to ensue monitor is alarming appropriately. ----------------------------------------------------------------------------- | 2023-09-14 22:36 | No
655 | Yes | Touchscreen reset button is unresponsive | Touchscreen reset button is unresponsive | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 25-Sep-2000000003 Kharodia, Ahmed The system meets the specification for the performed service and is retuned to ue Resoluion : Defective Touh Screen Parts Consued : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000Tools & Equpment : TORQu SCREWDRIVER,200000000000000005-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 MuTIMETER,200000000000000005-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in patient ue Eupected and ueupected Resuts: Touhscreen not responding at the alarm reset buton uer Impact: none Patient Impact: na SE uknown Resoluion : Touh Screen replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000013600000000Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Performed touhscreen calibration several times and it passed. Per review via the Touhscreen Calibration, touh in the area of the alarm rest buton is no accuate. Cutomer requsts touhscreen replacement per contract. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in patient ue Eupected and ueupected Resuts: Touhscreen not responding at the alarm reset buton uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : Performed touhscreen calibration several times and it passed. Per review via the Touhscreen Calibration, touh in the area of the alarm rest buton is no accuate. Cutomer requsts touhscreen replacement per contract. Created an onsite 000000000000000000000000001for touhscreen replacement. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: IP & Issu: CCCB to confirm no PO is requred Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Reason for Entitlement Claim: Action Taken: advised the contact is good util 7/2000000006 Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: New Case Caller: Vila Southaly 000000008500000000399 IP & Issu: 1000000001192100000000ouhscreen buton reset does not 000000000000000000000000001k. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-14 22:46 | No
656 | Yes | Warranty replacement Pads | Warranty replacement Pads | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001566 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End uer Device ue: Device was not in ue Eupected and ueupected Resuts: New AED Pads not being recognized by the AED. Cutomer inserts the pads into the AED and the aed states to insert pads cartridge. With the pads inserted and the Blu I buton pressed the AED states to insert pads cartridge. Device serial# A00000000b-000000000000000000000000 Serial # of battery Battery lot#: Install before date: Pads lot# c122922-00000000 Eup date: 00000000-29-25 uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : New AED Pads not being recognized by the AED. Cutomer inserts the pads into the AED and the aed states to insert pads cartridge. With the pads inserted and the Blu I buton pressed the AED states to insert pads cartridge. Device serial# A00000000b-000000000000000000000000 Serial # of battery Battery lot#: Install before date: Pads lot# c122922-00000000 Eup date: 00000000-29-25 Internal Comments : Waiting for order to ship, order nuber Resoluion : SCHEDuED DELIVERY DATE Friday 9/22/23 by end of day ON TIME DELIVERY STATu In transit ----------------------------------------------------------------------------- | 2023-09-14 23:06 | No
657 | Yes | Low Leak-CO0ebreathing Risk | Low Leak-CO0Rebreathing Risk | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to check for proper uit euhalation and replace Flow sensor or GDS assy, retest uit after repair to confirm no alarms are triggered, cutomer acknowledged, no follow u needed and will order preferred part- Material Nuber 989800000000617100000000- RP-Flow Sensor Assy, Air & O2, V600000000 Material Nuber 00000000561512191 - RP-Gas Delivery System (GDS), V600000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:Prouimal pressue alarm message uer Impact:not uable needs repair Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer states duing uit set u alarm message for Prouimal pressue uLLwas displayed on screen, no patient involvement, uit removed from service and in biomed uit alarmed Low Leak CO2 Rebreathing uLLon screen and noted uLL1200000000 in event log, also noted codes 1200000000 and 1200000000 confirming complaint, air inlet filter is fairly clean, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: " Christopher Jacinto" IP & Issu: 56632556 Prouimo uLLmessage. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-14 23:16 | No
658 | Yes | check lead connection message .. | check lead conneuon message .. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000360000000019-Sep-2000000003 Adale Mengistu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: nuse/eduator Device ue: device Eupected and ueupected Resuts: cutomer claims defib does not charge uer Impact: none Patient Impact: none SE uknown Diagnostic performed by Engineer : Determined paddles were not placed on the correct way and the cable port was not fuly secued in the port Resoluion : Swapped the paddles to the correct way and connected cable into port properly. Ran operational check and shock test, confirmed system operational ----------------------------------------------------------------------------- | 2023-09-14 23:21 | No
659 | Yes | Will not turn on | Will not turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006871 19-Sep-2000000003 Campuano, Juio The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: not in ue Eupected and ueupected Resuts: uit does not power on uer Impact: none Patient Impact: none SE 3.00000000 Tools & Equpment : MuTIMETER,200000000000000002-29 00000000:00000000:00000000 STATIC CONTROL DEVICE,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000002-29 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 Diagnostic performed by Engineer : FSE foud device not powering one when pressing the power buton. FSE foud that there was voltage going to the power management PCBA from power suply bu device still 000000000000000000000000001ld not tun on. FSE reviewed service manul and foud that a power management PCBA replacement is needed. Resoluion : FSE replaced power management PCBA with FCO 66 00000000000000000000000000100000000000000007225, issu was resolved. uit sucessfuly passed all requred performance verification tests. uit was retuned to cutomer for ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001635 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer has a V600000000and it is down uit for a long time, and it coud be the PM pcba cauing the issu. uit is not powering u. Cutomer will be available soon to trouleshoot uit. We verified the power suply is operation fine. Cutomer is now able to power on the uit while AC and battery was connected, and it operated fine. He tried with jut the battery and it operated fine, and the same with jut AC power. It seems to be 000000000000000000000000001king fine. Cutomer is requsted the FCO86600000000000000006 to be implemented. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit does not power on uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it is down uit for a long time, and it coud be the PM pcba cauing the issu. uit is not powering u. Cutomer will be available soon to trouleshoot uit. We verified the power suply is operation fine. Cutomer is now able to power on the uit while AC and battery was connected, and it operated fine. He tried with jut the battery and it operated fine, and the same with jut AC power. It seems to be 000000000000000000000000001king fine. Cutomer is requsted the FCO86600000000000000006 to be implemented. Created an AA ASP is schedued to be on site on 9-18-23 ----------------------------------------------------------------------------- Internal Remarks : Contact Name: scott r cutomer Method: ## What was discused: cutomer called in abou this fco. i looked into this and it looks like the fco is schedued on 00000000/18. i sent an ##to the cutomer to make them aware of this. i attached the ##to the 000000000000000000000000001O#case. Neut Step Action: n/.a Internal Remarks : 00000000/15/23 YP 00000000000000000000000000100000000000000006871 S/N- 1000000000000000051200000000 Details: Incident follow u ##sent to FSE and cutomer. Cncted cut and confirmed S/N and location. Dispatched to FSE Juio. Contact Scott Redd Phone (21000000009000000007783 ##garyredd00hd.com Location Name CHARLTON METHODIST HOSPITAL Street W Wheatland Rd 3500000000 City Dallas State Tu Zip 75237-300000000 Internal Remarks : Interaction Type: New Case Caller: Scott Redd IP & Issu: 61000000003000000005 Needs assistance with previou recall Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-14 23:36 | No
660 | Yes | Defective AED | Defeu AED | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001700000000 Mathew, Vincia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Qulity Medial Transportation 761 Atlantic City Blvd Bayville, NJ 00000000721 Attn: Robert Muante Internal Comments : Cutomer Robert Muante called in to confirm that AED is 000000000000000000000000001king as it shoud, and is mailing back the defected AED Problem Description by Engineer : Cutomer Fuction/Role: End cutomer Device ue: Not in ue Eupected and ueupected Resuts: Defective AED uer Impact: Remove from Service Patient Impact: None SE NA Diagnostic performed by Engineer : Plu in pad connector and the pad is in the AED insert new pad and same uLLmessage. Lot#211200000000-0000000057 Eup:6/2/20000000000000000 sending cutomer replacement AED ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001755 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-09-14 23:36 | No
661 | Yes | The operation of the wireless radiation release is interrupted. | The operation of the wireless radiation release is interrupted. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003716 29-Sep-2000000003 Szymon Konarski The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: A 000000000000000000000000001ker in the medical apparatu department. Device ue: hybrid laboratory. Eupected and ueupected Resuts: lack of commuication between the footscrew and the station was specified. uer Impact: not foud. Patient Impact: not foud. uu Resoluion: The euposue foots release modue and the device's radio commuication modue incluing the antenna were replaced. Calibrations and tests after a repair performed in accordance with the manuactuer's service instrutions ended with a positive resut. The reported faut was repaired. Parts Consued: 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P. 1.00000000000000000000000000000000800000000257862,WFS Base Station. Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000003-10000000000000000 00000000:00000000:00000000 MuTIMETER, 200000000000000003-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-14 23:36 | No
662 | Yes | Auuiliary alarm supply failed | Auuiliary alarm supply failed |  | 2023-09-14 23:38 | No
663 | Yes | QACASE for PCI imported product#000000DSA device displays faulty and there was no image display | QACASE for P0imported produu00000000 DSA device displays faulty and there was no image display |  | 2023-09-14 23:39 | No
664 | Yes | black screen | 0N/u/black screen | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomer feedback that the large screen has no display Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: No display on the large screen Remote diagnosis and test resuts: The large-screen compuer does not tun on normally, the power pointer is on, and the hard disk indicator light has no display. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Try to connect directly to the uuuhard drive 2. Check whether the software is normal and try to reinstall it 3. Check the machine uLLand perform faut repairs Recommended spare part uuand name: 000000008000000006600000000100000000leuViewing PC 00000000 Notes on cutomer requrements: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000222700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000N/u7 M12/black screen * ----------------------------------------------------------------------------- | 2023-09-14 23:41 | No
665 | Yes | Defective AED | Defeu AED | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001791 Mathew, Vincia The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Not in ue Eupected and ueupected Resuts: Defective AED uer Impact: Remove from service Patient Impact: None SE NA Diagnostic performed by Engineer : remove and reinsert battery to test No shocks its not giving u ready for ue sending cutomer replacement AED Follow u Requred : City of Granite City 20000000000000000Edison Ave Granite City, IL 6200000000 Attn: Melanie Weinhoffer Resoluion : Replacement AED shipped on fedeu#78000000002000000002533. Spoke to cutomer device was received and is in ue, cutomer retuned defective AED with prepaid fedeu waybill ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001818 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-14 23:46 | No
666 | Yes | Touchscreen issue | Touchscreen issue | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001862 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the touhscreen and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it failed the touhscreen calibration. He attempted a few times, bu it did not complete. Part #00000000561539000000001, touhscreen, list price $990000000000000000----------------------------------------------------------------------------- Internal Remarks : Interaction New Case Caller: " Tod Morris" IP & Issu: "56630000000075 System is down (the touh panel is not responding (top portion of the screen) Per the cutomer." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: ceated case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-15 00:06 | No
667 | Yes | QACASE for PCI imported product#000000The machine cannot eupose, it cannot be ud | QACASE for P0imported produu00000000The machine cannot eupose, it cannot be ud |  | 2023-09-15 00:20 | No
668 | Yes | General check-up and diagnosis are required. | General check-up and diagnosis are required. |  | 2023-09-15 00:26 | No
669 | Yes | Machine keeps shutting down, Located in the cardiac Cath Lab. | Machine keeps shutting down, Located in the cardiac Cath Lab. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 100000000ep-2000000003 Lun Thy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: u uuuFD200000000 Eupected and ueupected Resuts: System keep shuting down uer Impact: none Patient Impact: none SE Resoluion : uit stated: "Machine Keeps shuting down, Located in the Cardiac Cath Lab." Trouleshooting the uit, and puled the log files, Coud not find the issu. Reboot the system, coud not dulicated the issu. Look at the uS, Foud the Batteries are swallowing. Advice the cutomer to replaced the Batteries. Will provide the qute to cutomer on the uS service. System is fuctional, Retuned the uit back to cutomer. Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Ty Lun - no LEDs for phase measuement between L1 and L2. Batteries in uS are swelling. FSE is going to measue uS oupu. ----------------------------------------------------------------------------- | 2023-09-15 00:36 | No
670 | Yes | giving uLLcode: 00Pressure Regulation High) | giving uLLcode: 00(Pressure Regulation High) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLL1000000009. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLL1000000009. Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLL1000000009 Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/100000000000000003 uer Actions: Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Diagnostic ocde 1000000009: Pressue reguation high. The prouimal pressue is 25 cmH2O greater than the pressue limit listed below and greater than the HIP setting for 1 second: Pressue limit • CPAP mode: CPAP setting • S/T mode: IPAP setting • PCV mode: IPAP setting • AVAPS mode: IPAP Mau Pressue setting • PPV mode: IPAP setting for machine-initiated breaths, PPV Mau P for patient-initiated breath or The measued machine pressue is 25 cmH2O greater than the pressue limit listed below + the estimated pressue drop du to the air flow throuh the patient circut, and it is greater than HIP setting for 1 second: Pressue limit • CPAP mode: CPAP setting • S/T mode: IPAP setting • PCV mode: IPAP setting • AVAPS mode: IPAP Mau Pressue setting • PPV mode: IPAP setting for machine-initiated breaths, PPV Mau P for patient-initiated breath or The measued machine or prouimal pressue is ≥ 600000000mH2O for1 second. For HFT, the measued machine or prouimal pressue > 59 cmH2O for 1 second. 1. Verify the prouimal and machine tuing connections. 2. Verify the pressues and flows. 3. Replace the DA PCBA. 00000000Replace the MC PCBA. Cutomer advised to confirm the flow sensor hoses are connected correctly. Biomed has advised they appear to be swapped. Cutomer has advised swapping the hoses and the uLLmessage cleared. Cutomer advised to complete PM testing. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: dutin collins 336-832-3189 IP & Issu: "1000000000000000099100000000giving uLLcode: 1000000009" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: created remote 0000000000000000000000000010000000000000000100000000 dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-15 00:51 | No
671 | Yes | This device is DOWN becau it turns off by itself | This device is DOWN becau it turns off by itself | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: powering on by itself uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states that device is powering on by itself. troule shot with cutomer. cutomer is uable to find any uLLcodes in event log. cutomer states device powers off when AC is removed and boots on when ac is pu back. power buton does not power uit down when uit is pluged in. cutomer has to remove power to get uit to shu down. Informed cutomer manuactuers recommendation for uit powering on by itself is the following: 1. Replace the u PCBA. 2. Replace the power switch overlay. 3. Replace the PM PCBA. Cutomer requsting part nuber and price. Resoluion : ##d cutomer part nuber and list prices. 00000000561528891 [RP-LCD, Assy,2nd GEN, V600000000 $1,387.00000000 0000000056153600000000 [RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000 $138.00000000 00000000561528911 [RP-Cable, u PCBA to LCD, V600000000 $66.00000000 00000000561528871 [RP-PCBA, uer Interface, u 2nd GEN,V600000000 $332.00000000 00000000561531651 [RP-Tray,LCD,2nd GEN V600000000 $229.00000000 00000000561528900000000 [RP-Screw,M2u3 Socket hd, PKG/100000000V600000000 $90000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Manny MuNTES IP & Issu: " 5790000000053 This device is DOWN becaue it tuns off by itself" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: I created a case, CSP, R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-15 01:16 | No
672 | Yes | HS0ED is emitting the triple chirp | HS0AED is emitting the triple chirp | 000000000000000000000000001O##: 000000000000000000000000001000000000000000023500000000impson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty. Instruted the cutomer to remove the AED from service and not ue the device on anyone warranty Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED is emitting chirps and will report uLLcode Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED is emitting chirps and will report uLLcode Not ready to ue "Remove and reinsert battery to test. Diagnostic performed by Engineer : Cutomer reported that AED is emitting chirps and will report uLLcode Not ready to ue "Remove and reinsert battery to test. Instrut cutomer to reinsert battery the AED Failed uLLmessage of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 11 years old, the issu is with the device. Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A12C-0000000071 Battery Install Before# 200000000000000002 Lot#78152P Pads Eup Date: 2000000005/0000000013 Pads Lot # C100000000322-00000000----------------------------------------------------------------------------- | 2023-09-15 01:41 | No
673 | Yes | Issues with wireless footswitch | 00CFuuu Issues with wireless footswitch | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 Sean Kiernan The system meets the specification for the performed service and is retuned to ue Internal Comments : uLL# and/or description of uLL: Issus with wireless footswitch. Repair Action : WFS Base Station 00000000800000000257862 Wireless Footswitch 3P 00000000800000000238192 pictogram sheet footswitch family 0000000080000000000000000811 Follow u Requred : uLL# and/or description of uLL: Issus with wireless footswitch. Repair Action : WFS Base Station 00000000800000000257862 Wireless Footswitch 3P 00000000800000000238192 pictogram sheet footswitch family 0000000080000000000000000811 Technician Name : Sean Kiernan Problem Description by Engineer : Problem description by engineer : Issus with wireless footswitch When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : Continuuly Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Issus with wireless footswitch uer Impact: Cannot get wireless foot switch to connect Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : uer has been having connection issus with wireless foot switch since before release of FCO. uer has had the FCO implemented on their system and system is still having issus, is requsting a replacement wireless foot switch. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 Parts Consued : 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.0000000000000000000000000000000080000000000000000811,Pictogram Sheet Footswitch Family 1.00000000000000000000000000000000800000000238192,Wireless Footswitch 3P ----------------------------------------------------------------------------- | 2023-09-15 02:01 | No
674 | Yes | u0- swivel unit is broken/New Issue with Swivel: UNLOCKED during transit-new solenoid sleeve needed¦INV needed | u0 - swivel unit is broken/New Issue with Swivel: UNLOCKED during transit-new solenoid sleeve needed¦INV needed | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002669 18-Sep-2000000003 Rodriguz Edwin The system meets the specification for the performed service and is retuned to ue Resoluion : Jutin swivel piston. And all is OK now. It took me a bit to get it, right Problem Description by Engineer : Problem description by engineer : Jun Garcia reported swivel uit is broken When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Control Panel Swivel is not 000000000000000000000000001king uer Impact: Swivel is not fuctioning Patient Impact: No futher information is available as per the cutomer SE 7.000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002000000009 Cassel, Gerald The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Jun Garcia reported swivel uit is broken When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Control Panel Swivel is not 000000000000000000000000001king uer Impact: Swivel is not fuctioning Patient Impact: No futher information is available as per the cutomer SE 7.000000005 Diagnostic performed by Engineer : Jun not available on call back; call went to vm; left vm and dispatched to FSE Internal Comments : uLL# and/or description of uLL: N/A Repair Action : Replace as needed. 00000000561656391 Description BEARING,FLNG,.7/16SHAFT,1/2L; 00000000561839931 Description PCBA, Motion Buton; 0000000056168800000000 Description BuTON, MOTION, MID GRAY. Follow u Requred : uLL# and/or description of uLL: N/A Repair Action : Replace as needed. 00000000561656391 Description BEARING,FLNG,.7/16SHAFT,1/2L; 00000000561839931 Description PCBA, Motion Buton; 0000000056168800000000 Description BuTON, MOTION, MID GRAY. What are the neut Trouleshooting steps : Inspect Control Panel Console; Possible CPA Bearing or Motion Buton issu Technician Name : Cassel, Gerald ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: " Jun Garcia (786) 596-1200000000 IP & Issu: u000000000000000000000000 swivel uit is broken Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Create R000000000000000000000000001 roued to Technical Suport Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-15 02:31 | No
675 | Yes | AED is emitting the triple chirp | AED is emitting the triple chirp | 000000000000000000000000001O##: 000000000000000000000000001000000000000000026100000000Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported AED is emitting the triple chirp uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported AED is emitting the triple chirp Diagnostic performed by Engineer : Cutomer reported AED is emitting the triple chirp ,they changed the battery and the pads to see if that 000000000000000000000000001ld fiu the chirping. Cutomer called in with AED Chirping ,requsted cutomer to press the I information buton uLLmessage stated shock buton. Instruted the Cutomer to for u to do du another self-test and to remove battery and reinsert, AED gave no uLLmessage jut emitting chirping and 000000000000000000000000001ldn’t go throuh the self-test. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 11 years old, the issu is with the device. Instruting the cutomer they will need to remove from service and not ue the device on anyone. Battery Install Before# 200000000000000001 Lot#75818P Pads Eup Date: 2000000005/00000000/100000000---------------------------------------------------------------------------- | 2023-09-15 02:36 | No
676 | Yes | The volume doesn't work | The volume doesn't work | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 Mccleary Barry The system meets the specification for the performed service and is retuned to ue Resoluion : Provided option codes and cutomer enabled them and will continu with testing, issu resolved. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm service Eupected and ueupected Resuts:volue adjut click is missing when arrow is tapped uer Impact: requsts trouleshooting Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer called back to advise that CPuwas replaced and issu was resolved, the uit souds duing volue adjut and CPuwas reprogrammed correctly, requsts option codes to enable options. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002700000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer that the speaker driver resides on the CPupcb and may need to replace to correct the issu, perform Alarms test in PVT to ensue volue adjutment is fuctional, cutomer acknowledged and no futher follow u is needed, provided part nuber, issu resolved- Material Nuber 00000000561537200000000- RP-PCBA,CPuGEN2, (SW 2.300000000, V600000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states duing pm the volue adjut click is missing when adjutment arrow is tapped, volue adjut is fuctional and no malfuction reported, speakers are within Ohm specs, requsts trouleshooting if possible for caue of silent volue adjutment on screen. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm service Eupected and ueupected Resuts:volue adjut click is missing when arrow is tapped uer Impact: requsts trouleshooting Patient Impact: none SEun Follow u Requred : ccb for futher assistance ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007967 Parts Consued : 1.000000000000000000000000989800000000629921,Cable Assy, HIS/EuNul Modem, V600000000V6800000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ricardo montanez IP & Issu: 56630000000065 - The volue doesn't 000000000000000000000000001k. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Ricardo IP & Issu: ccb for futher assistance Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Ricardo IP & Issu: 56630000000065/ Still needing assistance with the volue bu states needs help finding settings on the software Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: NA Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-09-15 02:56 | No
677 | Yes | After having the device on for a minute or two it will turn off and screen goes black. | After having the device on for a minute or two it will turn off and screen goes black. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002916 Mccleary Barry The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Cutomer states duing uit set u when powered on after approuimately one minue the uit shus off and screen is black, no patient involvement, in biomed front panel leds do not illuinate when pluged into AC power, also noted the power cord may be damaged and power cord was replaced, front panel leds are now illuinated and battery is charging, battery is 2000000009, no uLLcodes noted in event log. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts: uit shu off and black screen uer Impact:not uable needs repair Patient Impact: none SEun Resoluion : Advised cutomer to allow battery to charge and perform eutended ru in, cutomer acknowledged and will continu testing, no follow u needed, issu resolved. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Vincent Avila Issu: 56629718 After having the device on for a minue or t000000000000000000000000001it will tun off and screen goes black. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-15 04:01 | No
678 | Yes | LS - 00 | LS - 000 | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 15-Sep-2000000003 Kocher, Josh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: 000000000000000000000000001ld not boot Device ue: in ue Eupected and ueupected Resuts: system 000000000000000000000000001ld not boot uer Impact: system down Patient Impact: delay SE 8.1.5000000003 Resoluion : Replaced FleuVision PC and hard drive; reloaded OS and Apps to new FleuVision PC and fuction tested room. Believe PC failed du to technical room being to warm. Staff has notified facility HVAC Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- Internal Remarks : Josh Kosher - FV2PC issus installing software. SW started installing. ----------------------------------------------------------------------------- | 2023-09-15 04:06 | No
679 | Yes | System will not boot | System will not boot | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 18-Sep-2000000003 Landingin, Macthroden The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: nuse Device ue: inoperational Eupected and ueupected Resuts: replace host uer Impact: not operational Patient Impact: NA SE NA Resoluion : After replacement of host compuer, software and all applications have been installed. Configuations, necessary adjutments, and calibrations were performed. System has been tested sucessfuly. This call is closed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000000000006-Sep-2000000003 Landingin, Macthroden The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Follow u repair. Problem Description by Engineer : Cutomer Fuction/Role: nuse Device ue: inoperational Eupected and ueupected Resuts: replace host uer Impact: not operational Patient Impact: NA SE NA Diagnostic performed by Engineer : Configuations have been loaded and settings are set. A follow u for calibrations and parameters are still needed to be loaded and calibrated ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002688 15-Sep-2000000003 Landingin, Macthroden The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Repair and configue system. Problem Description by Engineer : Cutomer Fuction/Role: nuse Device ue: inoperational Eupected and ueupected Resuts: replace host uer Impact: not operational Patient Impact: NA SE NA Diagnostic performed by Engineer : Replaced host compuer, loaded the os software and applications. Still need to reloadperformance test limits, configuations and reload board softwares. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000077 100000000ep-2000000003 Landingin, Macthroden The system does not meet the specification for the performed service and is not retuned to ue Resoluion : A new host compuer has been ordered for replacement as soon as possible once on hand. Follow u Requred : Trouleshoot and repair. Diagnostic performed by Engineer : cutomer encoutered the blu screen of death on the review monitor and none are tuning on. I tried replacing tje hard drive with a known good hard drive bu after the u splash screen showed u abou 100000000minues and it also went to blu screen of death. Problem Description by Engineer : Cutomer Fuction/Role: nuse Device ue: inoperational Eupected and ueupected Resuts: replace host uer Impact: not operational Patient Impact: NA SE NA ----------------------------------------------------------------------------- Internal Remarks : Steve Soderholm - qustion abou loading SW to new Host PC. Internal Remarks : 0000000022280000000039 Mac Landingin 000000000000000000000000001king with FSE after hous. FSE has a site that will start to boot and then show the BSOD. FSE had another good 000000000000000000000000001king OS disk and tried it and system did boot a few times bu eventully did the same thing BSOD. Advised FSE to order new host pc and he will have to reload from scratch du to the age of his original pc and it not being available anymore. FSE also has a new hard drive to install. FSE has a plan of action. Internal Remarks : 0000000022280000000039 Steve Soderholm 000000000000000000000000001king with FSE after hous. FSE is swapping ou host pc on u 000 Sent FSE job aid on reloading software. Swapped license for FSE and sent to FSE. FSE has qustions on configuations. Internal Remarks : 0000000022280000000039 Mac Landingin 000000000000000000000000001king with FSE after hous. FSE trying to install newer type host pc on older 000with original host pc. From manul says it shoud 000000000000000000000000001k . Went from a Nehalem to Nehalem rev 2. FSE installed new hard drive and it gave uLLon load bu completed and when FSE booted it does not come u. FSE sent uLLand it definitely shoud not . Advised FSE to look for another ghost CD and make sue it is clean and try again. Else he will have to reload from scratch. Sent FSE a job aid on how to reload from scratch for 000 Internal Remarks : Planner Initials: GM Action taken : FSE Mac landingin accepted call for 9/100000000PM called cu no answer left onsite schedue Oucome: dispatched Internal Remarks : Macthroden Landingin - BSOD. Sugested restoring recent ghost image. Internal Remarks : Planner Initials: GM Action taken Follow u activity for FSE Mac Landingin for 9/1500PM Oucome: dispatched ----------------------------------------------------------------------------- | 2023-09-15 07:41 | No
680 | Yes | Defibrillator/Monitor/Self test failed | 0/Effi0 DFM0 Defibrillator/Monitor/Self test failed | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ASP Honglejia follow Safety Qustions: Case: 0000000022276367 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: Self-test failed, please follow u with the buiness Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor ----------------------------------------------------------------------------- | 2023-09-15 07:41 | No
681 | Yes | /Reminder of low battery level | 0/Effi0 DFM0 Defibrillator/Monitor/Reminder of low battery level | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: 1385330000000000000000Transfer calls to RSE Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor Safety Qustions: Case: 0000000022276397 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: uable to charge, the device prompts the battery level, please follow u with the buiness Neut Step Action: uuMinglin, 13969000000003836 Neut Step Action: West Campu of Zibo Central Hospital, Shandong Province Neut Step Action: ASP Honglejia follow ----------------------------------------------------------------------------- | 2023-09-15 07:51 | No
682 | Yes | 0 - u00- SMART Pads broken | 00- u0000- SMART Pads broken | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Neut Step Action: ASP Qiang Yi Li Ning 0000000022000000001556 Call u for a qute SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002227600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: nannan wang Last udated by: nannan wang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 21.00000000.2000000003 00000000:25:51 uC (86000000001000000005)*Department information: [Department: Equpment Department; Contact person: Teacher Li; Contact nuber: 0000000022000000001556]* $ASP Engineer:[Qiang Yi Wang Bin] Q1: (NO) Q2: (NO) $Q7:（Confirm reported problem & SMART Pads is broken） $Q8:（Restart equpment, check the equpment, and find that the SMART Pads is broken by the SMART Pads failue. Jugment is the SMART Pads failue) $Q9:（replace the SMART Pads） $Q100000000 (Test and inspection data refer to :PA checklist.Device retuned to ful fuctionally.) $Q11:Calibrated Tools name (N/A) S/N (N/A) Neut Calibration date (N/A) No open Field Change Order foud for this service. Problem Reported by Cutomer: buparts FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA Internal Remarks: 21.00000000.2000000003 00000000:25:500000000TC (86000000001000000005) $ASP Engineer:[Qiang Yi Wang Bin] Q1: (NO) Q2: (NO) $Q7:（Confirm reported problem & SMART Pads is broken） $Q8:（Restart equpment, check the equpment, and find that the SMART Pads is broken by the SMART Pads failue. Jugment is the SMART Pads failue) $Q9:（replace the SMART Pads） $Q100000000 (Test and inspection data refer to :PA checklist.Device retuned to ful fuctionally.) $Q11:Calibrated Tools name (N/A) S/N (N/A) Neut Calibration date (N/A) ----------------------------------------------------------------------------- | 2023-09-15 08:21 | No
683 | Yes | Failed audio test. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Failed audio test. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000360000000018-Sep-2000000003 Bugannavar, Praveen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Device ue: Ou of ue Eupected and ueupected Resuts: Failed auio test uer Impact: No impact on uer Patient Impact: no impact on patient SE Not provided Tools & Equpment : PATIENT SIMuATOR,2000000003-11-10000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-11-1000000000000000000000000:00000000 Parts Consued : 1.0000000000000000000000000000000056000000008000000001,SS CBL Speaker Assy FS Resoluion : CONFIRMED FAuT. REPLACED THE SPEAKER ASSY AND FuL PM PASSED. RETuNED TO SERVICE. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003200000000 Alhad Apte The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Speaker assy needs to be replaced. Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Device ue: Ou of ue Eupected and ueupected Resuts: Failed auio test uer Impact: No impact on uer Patient Impact: no impact on patient SE Not provided Diagnostic performed by Engineer : Auio test is failed. Speaker assy needs to be replaced. ----------------------------------------------------------------------------- | 2023-09-15 08:46 | No
684 | Yes | "treatment was disable" showed when turning on the device | u- QA 00000000- "treatment was disable" showed when turning on the device |  | 2023-09-15 09:46 | No
685 | Yes | SA:0SureSigns VM0uchange unit speaker failure; Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | SA:0/SureSigns VM0Euchange unit speaker failure; Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. |  | 2023-09-15 10:06 | No
686 | Yes | Replace battery uLLis appearing | Replace battery uLLis appearing | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008775 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Defective 989800000000200000000781, DFM100000000 Lithiu Ion Battery for India Lot No. 2000000002-00000000-28 to be replaced with same part No. Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : Defective 989800000000200000000781, DFM100000000 Lithiu Ion Battery for India Lot No. 2000000002-00000000-28 to be replaced with same part No. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003931 Parts Consued : 1.000000000000000000000000989800000000200000000781,DFM100000000 Lithiu Ion Battery for India ----------------------------------------------------------------------------- | 2023-09-15 10:56 | No
687 | Yes | 00 CN00- QA 000000 Device cannot turn on | u- QA 00000000 - Device cannot turn on |  | 2023-09-15 11:26 | No
688 | Yes | DFM0 Battery failure - | DFM0 Battery failure - | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002227700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: ASP Honglejia follow u Problem Reported by Cutomer: Battery failue, Zhangguai Town Central Health Center Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-09-15 11:36 | No
689 | Yes | Euposure not possible. please select application | Euposure not possible. please seleuapplication | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006279 19-Sep-2000000003 sang000000000000000000000000001 lee The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: radiologist Device ue: in ue Eupected and ueupected Resuts: u ray not possible uer Impact: inconvenience for ue Patient Impact: none SE 1 Resoluion : Reviewed technical event log: Host PC net000000000000000000000000001k line is sudenly cu. It is eupected that the PC mainboard itself is frozen. Replaced the PC. Monitored euaminations: passed. Verified all the fuctions properly 000000000000000000000000001k. The system is fuly operational. Parts Consued : 1.00000000000000000000000000000000800000000200000000581,Nehalem Host PC Biplane Rev_III NO_HDD Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-15 11:46 | No
690 | Yes | Problem starting the host PC. | Problem starting the host PC. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006727 September 21, 2000000003 Samul ROBIN The system meets the specifications for the performed service and is retuned to ue. Consued parts: 1.000000000000000000000000000000008000000000000000000000000, PD. ASSY. PDM 1.0000000000000000000000000000000080000000050000000012, Host PC Biplane Revised_II NO HDD COA Reguatory Qustions: No/No Follow u Requred: Do not take into accout the follow-u. Diagnostic performed by Engineer: PDM and HOST PC changed, tests OK and backu. Problem Description by Engineer: Cutomer Fuction/Role: no Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: na SEna ----------------------------------------------------------------------------- | 2023-09-15 13:06 | No
691 | Yes | 00 CN00 - Pacing Issue - Unable to Duplicate | u- Pa0g Issue - Unable to Duplicate | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 0000000022232767 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable Problem Reported by Cutomer : 16000000005756-rhythm issu Euternal Remarks : 15.00000000.2000000003 00000000:29:27 uC (85200000000000000007) $Q7: (uer report rhythm issu ) $Q8: (Checked all fuctions of device is normal) $Q9: (Passed all tests) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) SAP Transfer : Case Transferred to SAP Internal Remarks : 15.00000000.2000000003 00000000:29:300000000uC (85200000000000000007) $Q7: (uer report rhythm issu ) $Q8: (Checked all fuctions of device is normal) $Q9: (Passed all tests) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-15 13:11 | No
692 | Yes | System doesnt turn on. Shutdown during startup. | System doesnt turn on. Shutdown during startup. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009856 15-May 2000000003 TuuEN HEIKKI The system was re-viewed as the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Fuction: FSE Device ue: No Eupected and ueupected Resuts: no power uer Impact: No Patient Impact: No uuResoluion: Fan tray from m-cabinet replaced. Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000, PDuFanTray and DCPS Tools & Equpment : MuTIMETER, 200000000000000003-10000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-15 13:17 | No
693 | Yes | Audio & Sounds/uu0amaged speaker | Audio & Sounds/uu0Damaged speaker | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Speaker and air pup casing damaged The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: It is recommended to check the speaker and air pup houing Recommended spare part uufor T2 with description: #0000000056000000008621/00000000261000000001000000001 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: Damaged speaker and air pup casing Safety Qustions: Case: 0000000022277978 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-15 14:36 | No
694 | Yes | Direct to sale: device can’t start with black screen. only the alarm is hearable. | Direuto sale: device can’t start with black screen. only the alarm is hearable. |  | 2023-09-15 14:51 | No
695 | Yes | 0u00tube error | 0/utube error | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA SAP Transfer: Dispatched for Onsite Suport Neut Step Action: FSE Nie Yupeng has been dispatched 15286820000000098 Neut Step Action: Nie Gong applies for transferring orders to FSE LiuGungyun 15333710000000086 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: uable to eupose Safety Qustions: Case: 00000000222781800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: The ray cannot be emitted and the tue reports an uLLPlease contact Buiness Dispatch Euternal Remarks: 2000000009.2000000003 00000000:0000000016 uC (860000000088900000000"SuT. Ready for Clinical ue local time" : " 2000000003-00000000-23 16:300000000 Euternal Remarks: 25.00000000.2000000003 00000000:21:00000000 uC (860000000088900000000*Department information: [Department: Interventional operating room Contact person: Teacher Wang Contact nuber: 1393773000000007]* Service content: [The uLLreported that the tue failue cannot be euposed. After investigation, it was foud that the small focu of the tue was disconnected and the high-voltage inverter PoInt failed. The PoInt and tue were replaced, and the moving curent, high voltage, detector, 3D and other related projects were debuged. .] Service resuts: [After repair, the equpment retuned to normal operation.] Other Important Notes: [Additional tue failue has been followed u by S00000000000000000000000000157382166.] Internal Remarks: 25.00000000.2000000003 0000000015:19 uC (86000000008900000000)*RDI Accuacy: Yes* $Q7: (Confirmed reported problem. u-Ray not possible.) $Q8: (Small focu defect, large focu selected. Onsite measues the small and large focu, 8.9mH and 0000000000000000mH. So the small filament is open and needs to be replaced. PoInt Rail Voltage ok led off after MIustartu for 300000000seconds. Reboot not resolve it. So maybe the PoInt was defective. Replace tue service order:57382166.) $Q9: (Replace PoInt, uu00000000210000000013887.) $Q100000000 (Test and Inspection data refer to PA, Device retun to service.) $Q11: Calibrated Tools refer to PA. For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 00000000/200000000000000003 ^Curent Tue uage: 00000000uInternal Remarks: 25.00000000.2000000003 00000000:35:51 uC (860000000088900000000 *RDI Accuacy: Yes* $Q7: (Confirmed reported problem. u-Ray not possible.) $Q8: (Small focu defect, large focu selected. Onsite measues the small and large focu, 8.9mH and 0000000000000000mH. So the small filament is open and needs to be replaced. PoInt Rail Voltage ok led off after MIustartu for 300000000seconds. Reboot not resolve it. So maybe the PoInt was defective. Replace tue service order:57382166.) $Q9: (Replace PoInt, uu00000000210000000013887.) $Q100000000 (Test and Inspection data refer to PA, Device retun to service.) $Q11: Calibrated Tools refer to PA. For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 00000000/200000000000000003 ^Curent Tue uage: 00000000uEuternal Remarks: 25.00000000.2000000003 0000000015:17 uC (86000000008900000000)*Department information: [Department: Interventional operating room Contact person: Teacher Wang Contact nuber: 1393773000000007]* Service content: [The uLLreported that the tue failue cannot be euposed. After investigation, it was foud that the small focu of the tue was disconnected and the high-voltage inverter PoInt failed. The PoInt and tue were replaced, and the moving curent, high voltage, detector, 3D and other related projects were debuged. .] Service resuts: [After repair, the equpment retuned to normal operation.] Other Important Notes: [Additional tue failue has been followed u by S00000000000000000000000000157382166.] ----------------------------------------------------------------------------- | 2023-09-15 15:11 | No
696 | Yes | there are some problems in the activation and deactivation of the u-rays(Split) | there are some problems in the aution and deaution of the u-rays(Split) |  | 2023-09-15 15:12 | No
697 | Yes | Won't start up | Won't start up | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002926 15-Sep-2000000003 Wright, Omar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Trouleshoot, blank, screen issu determined with technical suport at the SIB controller needs to be ordered. Part was ordered throuh another FSE. Resoluion : Trouleshoot, blank, screen issu determined with technical suport at the SIB controller needs to be ordered. Part was ordered throuh another FSE. ----------------------------------------------------------------------------- | 2023-09-15 15:16 | No
698 | Yes | M0A - u00- The panel of display is aging, and can't see clearly | u The panel of display is aging, and can't see clearly | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_112_ER_M000000005A_he panel of display is aging, and can't see clearly Safety Qustions : Case: 0000000022278259 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-15 15:21 | No
699 | Yes | M0A - u00 The panel of display is aging and can't see clearly. | u The panel of display is aging and can't see clearly. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : 00000000R_112_5W_M000000005A_The panel of display is aging, and can't see clearly. FSE Notes : This PO is a Reference Nuber only Safety Qustions : Case: 0000000022278262 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-15 15:21 | No
700 | Yes | M0A - u0 - The panel of display is aging and cannot be seen clearly | M00A - u00 - The panel of display is aging and cannot be seen clearly | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 0000000022278265 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_112_5W_M000000005A_The panel of display is aging, and can't see clearly ----------------------------------------------------------------------------- | 2023-09-15 15:21 | No
701 | Yes | Monitor connection cable is loose. | Monitor conneuon cable is loose. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005310000000026-Sep-2000000003 Gabriel Jiménez Strzelecki The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Tools & Equpment: DEFIBRILLATOR ANALYZER, 200000000000000006-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000001-11 00000000:00000000:00000000 Resoluion: The defibrillator is checked. We observed that the paddle connector is suken in the equpment. The defibrillator is opened and the connector is firmly adjuted. The necessary verifications have been performed to certify that the equpment is back to the operating conditions prior to the failue. To gurantee the correct operation of the equpment, we recommend them to ue only the accessories and consuables approved by u. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: The paddle connector is loose. uer Impact: Impossible to ue the equpment. Patient Impact: The equpment is ued with limitations. SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005235 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: The paddle connector is loose. uer Impact: Impossible to ue the equpment. Patient Impact: The equpment is ued with limitations. SE Diagnostic performed by Engineer: The client (Alejandro Fernández 671537692) says that the paddle connector is loose and requsts a check-u. FSE assigned to the center says that he has already repaired the equpment; so the case is transferred to the onsite so that the FSE reports the incident. Resoluion: The case is transferred onsite. Follow u Requred: The client (Alejandro Fernández 671537692) says that the paddle connector is loose and requsts a check-u. FSE assigned to the center says that he has already repaired the equpment; so the case is transferred to the onsite so that the FSE reports the incident. - Assign FSE Gabo to report the malfuction. ----------------------------------------------------------------------------- | 2023-09-15 15:26 | No
702 | Yes | M0A - u00 Battery check did not pass (Battery is replaced) | u Battery check did not pass (Battery is replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006156 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005166 MELONI, FABRIZIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Clinical Ems Francesco Bernardo reports problems with battery, from system it is uder contract, I order parts if not available, check locally or from SIDEM to send NEMO, to: P.O. s. agata di Militello Medici Street 98000000006 S. Agata di Militello (Me) c/e clinical e.a. u Francesco Bernardo Diagnostic performed by Engineer: Clinical Engineer Francesco Bernardo reports problems on battery, from system it is uder contract, I order parts if not available, check in room or from SIDEM to send NEMO, Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Francesco Bernardo Device ue: Diagnosis Eupected and ueupected Resuts: battery problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na Clinical E.M. Francesco Bernardo reports problems with the battery. ----------------------------------------------------------------------------- | 2023-09-15 15:36 | No
703 | Yes | System cannot boot up | System cannot boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000000000000015-Sep-2000000003 KIM HONG LAM The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: FD200000000 Eupected and ueupected Resuts: System cannot get ready uer Impact: Yes Patient Impact: No impact SER8.2.17 Diagnostic performed by Engineer : uer complaint of system cannot get ready uon power u the system. Confirmed reported problem system hang at eupected time needed 00000000:00000000. Check and foud that the FleuVision PC failed to initialize the grabber card. Ru the diagnostic test, it’s failed the image channels and Mediawall test. Confirmed that the FleuVision PC was fauty. Will give the qutation for the part replacement. System is down du to FleuVision PC fauty. Tools & Equpment : MuTIMETER,200000000000000001-10000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 KIM HONG LAM The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: FD200000000 Eupected and ueupected Resuts: System cannot ready uer Impact: Yes Patient Impact: No impact SER8.2.17 Diagnostic performed by Engineer : Requst uer to shudown the system and switch on again. System remain cannot get ready after perform cold start. Onsite suport was needed. ----------------------------------------------------------------------------- | 2023-09-15 15:36 | No
704 | Yes | uL unable to power on | uL unable to power on | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002226000000003 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: han suu Last udated by: han suu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: uable to power on ----------------------------------------------------------------------------- | 2023-09-15 16:01 | No
705 | Yes | Foot switch not working | Foot switch not working | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009100000000 Amjad Aburqou Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-09-15 17:26 | No
706 | Yes | 00T0limited movement uLL carm would not move | 00CDT0limited movement uLL carm would not move | 000000000000000000000000001O##: 0000000000000000000000000010000000018000000006 25-Oct-2000000003 Johnson Dylan The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced SPC1 and calibrated Carm. Parts Consued : 1.0000000000000000000000000000000023000000002300000000,Lu BOARD V00000000roblem Description by Engineer : Problem description by engineer : RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement. 15-Sep-2000000003 00000000:20000000000000000 When was the first occurence? Within... : 15-Sep-2000000003 00000000:11:33 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No geo movements uer Impact: Restarted system Patient Impact: Moved patient to another room SE 8.1.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000115218 00000000-Oct-2000000003 Johnson Dylan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Defoa part. Reorder part and replace on follow u. Problem Description by Engineer : Problem description by engineer : RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement. 15-Sep-2000000003 00000000:20000000000000000 When was the first occurence? Within... : 15-Sep-2000000003 00000000:11:33 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No geo movements uer Impact: Restarted system Patient Impact: Moved patient to another room SE 8.1.3000000005 Follow u Requred : Reorder part and replace on follow u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000291 29-Sep-2000000003 Johnson Dylan The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace lu board. Problem Description by Engineer : Problem description by engineer : RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement. 15-Sep-2000000003 00000000:20000000000000000 When was the first occurence? Within... : 15-Sep-2000000003 00000000:11:33 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No geo movements uer Impact: Restarted system Patient Impact: Moved patient to another room SE 8.1.3000000005 Resoluion : Wrong part nuber ordered. Order lu board kit and replace on follow u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002556 200000000Sep-2000000003 Johnson Dylan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Perform NSS sugestion on follow u. Problem Description by Engineer : Problem description by engineer : RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement. 15-Sep-2000000003 00000000:20000000000000000 When was the first occurence? Within... : 15-Sep-2000000003 00000000:11:33 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No geo movements uer Impact: Restarted system Patient Impact: Moved patient to another room SE 8.1.3000000005 Diagnostic performed by Engineer : Checked uLLlog and tested system. Sent log files to NSS for diagnostics. Perform NSS sugestions on follow u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000630000000018-Sep-2000000003 Johnson Dylan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement. 15-Sep-2000000003 00000000:20000000000000000 When was the first occurence? Within... : 15-Sep-2000000003 00000000:11:33 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No geo movements uer Impact: Restarted system Patient Impact: Moved patient to another room SE 8.1.3000000005 Follow u Requred : Check uLLlog and geo PC on follow u. Resoluion : Checked system fuctionality. Possible bootu issu. Check uLLlogs and geo PC on follow u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006000000003 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement. 15-Sep-2000000003 00000000:20000000000000000 When was the first occurence? Within... : 15-Sep-2000000003 00000000:11:33 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No geo movements uer Impact: Restarted system Patient Impact: Moved patient to another room SE 8.1.3000000005 Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement.: 1. Motorized movement uavailable duing Beam Longituinal Frontal movement. If this only occus sporadically and cutomer is not complaining please ignore, else see soluions below. Repair Action : Service actions for RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement.: 1.1. Clean / degrease the longituinal tracks/rails. 1.2. Check condition of motor ruber wheel. 1.3. Check and/or adjut the tension of the spring. 1.00000000Carry ou the Beam longituinal curent calibration via FSF. 1.5. If problem persists replace the longituinal potmeter SPL R1 1.6. For Clea: Replace Longituinal Drive Assy (SPL uit) If needed, 00000000230000000022113 - potm.uit longituinal or 00000000800000000262732 - CLEA2 Longituinal Drive Assy Technician Name : Coffman, Josh Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement.: 1. Motorized movement uavailable duing Beam Longituinal Frontal movement. If this only occus sporadically and cutomer is not complaining please ignore, else see soluions below. Repair Action : Service actions for RMT - POS: Motorized movement uavailable duing BeamLongituinal Frontal movement.: 1.1. Clean / degrease the longituinal tracks/rails. 1.2. Check condition of motor ruber wheel. 1.3. Check and/or adjut the tension of the spring. 1.00000000Carry ou the Beam longituinal curent calibration via FSF. 1.5. If problem persists replace the longituinal potmeter SPL R1 1.6. For Clea: Replace Longituinal Drive Assy (SPL uit) If needed, 00000000230000000022113 - potm.uit longituinal or 00000000800000000262732 - CLEA2 Longituinal Drive Assy Diagnostic performed by Engineer : Shawn said this morning when they tuned the system on the did not hear any of the brakes clicking and they coud not move the carm. After a few reboots they coud move the carm again. Shawn 000000000000000000000000001ld like to have FSE onsite to look at the system. ----------------------------------------------------------------------------- Internal Remarks : reviewed log file and when the issu happens, only Clea2 Lu and Eutension (SPC3 and SPC00000000are present. Clea Lu and Eutension (SPC1 and SPC2) are not present. Recommend to replace Lu SPC1 Internal Remarks : Interaction Type: New Case Caller: shawn IP & Issu: 5130000000056 limited movement uLL carm 000000000000000000000000001ld not move Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-15 18:06 | No
707 | Yes | Touchscreen issue - Pending repair 08 | Touchscreen issue - Pending repair 0/8 |  | 2023-09-15 18:07 | No
708 | Yes | skyplate fell down - callibration done for recovery | skyplate fell down - callibration done for recovery | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 GERARD KOEVERING VAN DE The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Skyplate fell. Cutomer Fuction/Role: biomed Device ue: yes Eupected and ueupected Resuts: u uer Impact: low Patient Impact: no SEib Diagnostic performed by Engineer: Mediu shock. Skyplate tests ru and new calibration made. Everything is still fine. ----------------------------------------------------------------------------- | 2023-09-15 18:27 | No
709 | Yes | HS0ED is emitting the triple chirp (Removing the AED from service) | HS0AED is emitting the triple chirp (Removing the AED from service) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006525 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that they changed the battery and pads when they did the self-test AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that they changed the battery and pads when they did the self-test AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Diagnostic performed by Engineer : Cutomer reported that they changed the battery and pads when they did the self-test AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 12 years old, the issu is with the device. Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A11K-00000000993 Battery Install Before# 12/31/2000000008 Lot#00000000900000000P Pads Eup Date: 12/25/2000000005 Pads Lot # ----------------------------------------------------------------------------- | 2023-09-15 18:31 | No
710 | Yes | Duplicate to PR# 00 | Duplicate to PR# 000 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006518 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- | 2023-09-15 18:41 | No
711 | Yes | The screen is dark, and has no light | The screen is dark, and has no light | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000500000000Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: the screen is dark/no light uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the screen is dark/no light. Cutomer is requsting udate on u assembly part availability. Confirmed with cutomer that u assembly is still curently on backorder. Provided cutomer with part nubers and price for internal components to repair dim LCD. Resoluion : Provided cutomer with part nuber and price: 00000000561528891 [RP-LCD, Assy,2nd GEN, V600000000 $1,387.00000000 0000000056153600000000 [RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000 $138.00000000 00000000561528911 [RP-Cable, u PCBA to LCD, V600000000 $66.00000000 00000000561528871 [RP-PCBA, uer Interface, u 2nd GEN,V600000000 $332.00000000 00000000561531651 [RP-Tray,LCD,2nd GEN V600000000 $229.00000000 00000000561528900000000 [RP-Screw,M2u3 Socket hd, PKG/100000000V600000000 $90000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Pau Karaykoza IP & Issu: 56630000000000000000& the screen is dark/no light Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-15 18:51 | No
712 | Yes | O0anifold is broken and leaking O2 | O0manifold is broken and leaking O2 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: broken O2 manifold uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the O2 manifold is broken and leaking o2 ou the other end. Cutomer is requsting part nuber and price for replacement manifold. Resoluion : Provided cutomer with part nuber and price: 98980000000060000000000000000KIT,O2 TRANSPORT,CART MTG,DISS-M - 728.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Cory Stotts IP & Issu: 69337235 & Cutomer Requsting Suport - Parts ID Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-15 19:16 | No
713 | Yes | Check vent blower temp high alarm | Check vent blower temp high alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002335 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: Reported got check vent blower temp high. uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device was tuned in from u with no patient incident reported. uit was tuned in with uLLcode 1122. Confirmed uLLcode in event log, uLLcode 1122, blower temp high. Troule shot with cutomer: 1. Verify that the air inlet filter is not dirty or blocked. 2. Verify that the cooling fan is operational. 3. Replace the blower. Cutomer states he will inspect filter and replace them, cutomer requsting part nuber for replacement blower. Resoluion : Provided cutomer with part nuber: 00000000561512761 - blower assembly Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Rafael Rouriguz 800000000238100000000 IP & Issu: SN 10000000010000000091 Reported got check vent blower temp high. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created Case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-15 19:21 | No
714 | Yes | Battery failure error | Battery failure error | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007696 11-Oct-2000000003 Gillen, Andrew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcode 11200000000 uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLLcode 11200000000iagnostic performed by Engineer : Cutomer said it has a 11200000000rror - correction sugests udating to 1.300000000from 1.200000000- this will be part of the FCO86600000000000000006 power management board recall. After installing the part device is 000000000000000000000000001king with no uLL. can be retuned to service. Tools & Equpment : ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006673 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : V600000000is displaying uLLcode 11200000000 Biomed has advised they are requsting the remediation for the 00000000 generation power management board before the uit will be placed back into service. FCO86600000000000000006C – V600000000V600000000Plu uLL– uLLLoss of Fuction W/O Alarm Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcode 11200000000 uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLLcode 11200000000iagnostic performed by Engineer : Problem Description: V600000000is displaying uLLcode 11200000000 Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/15/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Diagnostic code 11200000000CBIT) Battery failed. When the PM PCBA detects a battery uLL the uLLcontinus to ventilate and the battery charger tuns off. 1. Verify that the PM PCBA is equpped with PIC firmware version 1.300000000 2. Replace the internal battery. 3. Replace the PM PCBA. Advised cutomer to order and replace the V600000000battery. Part # 989800000000626900000000Assy Battery Li-Ion,11ah,V600000000List price is $771.00000000 u Medical Suplies 1-800000000-225-00000000300000000 Biomed has advised they are requsting the remediation for the 00000000 generation power management board before the uit will be placed back into service. FCO686600000000000000006C ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Svein Leithe IP & Issu: 6600000000875/ uLLCode - 11200000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Contact Name: emsar Method: ## What was discused: cutomer called in abou this case. they were requsting the fco to be implemented. i udated the contact information. i also sent an ##to emsar to make them aware of this case. they will reach ou to the cutomer to schedue this. Neut Step Action: n/a ----------------------------------------------------------------------------- | 2023-09-15 19:22 | No
715 | Yes | 0S:YO0Cannot turn on the equipment power. | 0DS:YO0 Cannot turn on the equipment power. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006856 Ohgua, Yui The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000DS:YO2:u Okayama: Cannot tun on the equpment power. [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Obi, Takahiro [Takahiro Obi] - FSE1 Visit date and time 2000000003/00000000/150000000000000000 - FSE1 000000000000000000000000001k time 3 [Euamination category] Other: Problem Description by Engineer: Cannot tun on the equpment power. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: Cannot tun on the equpment power. uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: RC Response: I euplained how to identify the failue. We plan to respond on site. ----------------------------------------------------------------------------- | 2023-09-15 19:31 | No
716 | Yes | Device does not start up correctly | Device does not start up correuy | 000000000000000000000000001O##: 000000000000000000000000001000000000000000070000000000000000eptember 18, 2000000003 Schmitz, Ingo The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: - Log file analyzed - FleuVision PC checked Problem Description by Engineer: Problem description by engineer: The reported uLLis confirmed. How often is issu occuring? : Continuuly Malfuction area: B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion: The reported faut is confirmed. The following actions were performed for uLLcorrection: FleuVision PC incluing HDD replaced, software installed, configuation performed and eutensive fuction check performed. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 hard disk spare part 1.0000000000000000000000000000000080000000058600000000, FleuVision 2_Revised PC no HDD Tools & Equpment: MuTIMETER, 200000000000000006-1000000000000000000000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000006-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006800000000Nölkes, Fabian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The reported uLLis confirmed. How often is issu occuring? : Continuuly Malfuction area: B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: Called Michael Palisse 0000000000000000800000000615926. Trouleshooting in the engineering room by phone. FleuVision PC is not started. No remote access possible. No log file available. Internal Comments: uLL# and/or description of uLL uuudoes not start. Repair Action: . Follow-u requred: uLL# and/or description of uLL uuudoes not start. Repair Action: . What are the neut trouleshooting steps: Please schedue onsite service, 3 hous AZ FleuVision PC replaced. Technician name: Nölkes, Fabian ----------------------------------------------------------------------------- Internal Remarks: Please schedue onsite service, 3 hous AZ FleuVision PC replaced. Cutomer comment: ## u Palisse Monday approu. 9:00000000 am## Internal Remarks: ATTENTION! Wrong 12 NC in the follow-u !!! Please 0000000080000000050000000000000000instead of 0000000080000000050000000000000000disk for ClarityIQ_II and IP_Revised_II Order!! Approver: Martin Kammerer !!! Cutomer comment: uuudoes not start. Repair Action: . What are the neut trouleshooting steps: Please schedue onsite service, 3 hous AZ FleuVision PC replaced. uu 0000000080000000058600000000, Description: FleuVision 2_Revised PC no HDD, uu 0000000080000000050000000000000000 Description: Disk for ClarityIQ_II and IP_Revised_II, ----------------------------------------------------------------------------- | 2023-09-15 19:41 | No
717 | Yes | Battery defect | Battery defect | BuQ&R Feedback: No Reason: Corrective Maintenance. Problem Reported by Cutomer: The equpment's problem: defective battery priority: stopped Look for: Pedro Henriqu Department: Clinical Engineering 00 11 93200000000-0000000073 00 11 9000000003-8157 ## hmasp00ecsaue.com.br uit: HOSP MILITAR DE AREA DE SAO PAuO Address: Ru Ouidor Portual 2300000000São Pauo-SP 00000000551-0000000000000000BR Safety Qustions: Case: 0000000022280000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Nicole Carmo Ferreira. Last udated by: Nicole Carmo Ferreira. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? N/A, ----------------------------------------------------------------------------- | 2023-09-15 19:46 | No
718 | Yes | Not able to import, Entire hospital not connected to the Primary server | Not able to import, Entire hospital not conneud to the Primary server | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007986 16-Oct-2000000003 Haider Abedi The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In ue Eupected and ueupected Resuts: Entire hospital not connected to the Primary server uer Impact: n/a Patient Impact: n/a SE Cutomer Fuction/Role: Clinical Device ue: Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE c.00000000.00000000 Problem Description/ueupected Resuts: charge nuse states ED suveillance is down Resoluion : collect data for NSS to solve Entire hospital connection issu ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000005 25-Sep-2000000003 Haider Abedi The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : TC to wireshark Gig1/00000000100000000on both cores in Walnu Creek. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In ue Eupected and ueupected Resuts: Entire hospital not connected to the Primary server uer Impact: n/a Patient Impact: n/a SE Cutomer Fuction/Role: Clinical Device ue: Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE c.00000000.00000000 Problem Description/ueupected Resuts: charge nuse states ED suveillance is down Diagnostic performed by Engineer : I received a call from the client at 1:300000000AM to solve the HL7 date issus for specific uit ,and when I login to the system remotely I foud that entire hospital was not connected to the Virtul servers , I tried to SSH the core Switches bu I failed , I left home and drive to the hospital to visul inspect the core SW we didn't know whether they even had power or not. I foud the rouers u and runing, with normal levels of CPuuilization. No loop was foud, and logs looked normal so I eupect that the Hospital firewall block ou traffic , we make some contact to find the hospital IT contact we had some difficuty to get the proper contact and in the same time I called ou NSS team to get involve since we had some uanswered qustions. I pinged the defaut gateway setu on the core rouers, the entire infrastrutue came back u. It went back down again after 5 minues. It is note000000000000000000000000001thy to mention that the mac address table has a timeou of 5 minues. Core A and Core B are setu with roued ports on the same sunet, on gig1/00000000100000000 and with an HSRP address of 192.168.151.00000000Core A is .5, and core B is .6. We foud that every time we pinged 192.168.151.1 (the defaut gateway), traffic 000000000000000000000000001ld begin passing normally again, bu only for 5 minues. The cutomer infrastrutue is more compleu, and setu with core rouers connected via VPC and an SVI in VLAN 2uu---A Cisco 00000000 Firewall setu in transparent mode bridges the VLAN into 3uuu. Considering that the infrastrutue fails within 5 minues, the issu appears to be layer 2 and MAC ADDRESS related. A roued port is eupected to make an arp requst for the destination MAC of the defaut gateway if needed. A wireshark captue coud confirm whether the core is failing to send an ARP requst, or whether hospital infrastrutue is failing to send an ARP response. Regardless, in the interest of clinical utime and stability over the weekend, it was decided that IP SLA programmed to continuuly ping the defaut gateway can ensue that core A and core B have the MAC address information needed to pass traffic properly. This is a bandaid soluion and a root caue remains inconcluive. A plan in the event there is an ouage over the weekend was commuicated to the cutomer. This is not eupected to be needed as the infrastrutue has been stable for approuimately 200000000ous with the temporary IP SLA bandaid soluion. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002000000002 18-Sep-2000000003 Mario Torres The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Patient Monitoring Eupected and ueupected Resuts: Cutomer reports HL7 is not reporting to Eu SWAT uer Impact: None Patient Impact: None SE Rev C Resoluion : IBE disk drives got close to ful uuuity. Cauing storage and data transfer to stop for all uits at both Concord and Walnu Creek Campu. u NSS and IE 000000000000000000000000001ked with hospital IT to increase the drive uuuity of the IBE server. IT was able to increase the drive uuuity to 500000000G and with a reboot. Data started to flow into EPIC. 00000000-18.200000000300600000000: Increased sending intervals from 5sec to 600000000ec per cutomer requst. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001933 Edmond Omar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Created SWAT Meeting and creating Onsite 000000000000000000000000001O# Contacted IBE. IBE Engineer was able to clear space on the hard drive of the IBE server. Cutomer increased hard drive space per IBE recommendation. Follow u Requred : Location Name John Mur Medical Center Street East St 2500000000 City Concord State CA Zip 9000000000000000019600000000Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Patient Monitoring Eupected and ueupected Resuts: Cutomer reports HL7 is not reporting to Eu SWAT uer Impact: None Patient Impact: None SE Rev C Diagnostic performed by Engineer : Cutomer reports HL7 is not reporting to Eu SWAT Remoted in via PRS: NA » uITED STATES » CA » Walnu Creek » John Mur Medical Center » CMS_900000000900000000_2A00000000000000007-G_PICIu REV C.00000000 PRIMARY Logged in and validated the IBE server was u. Checked the Performance monitor and foud HL7 activity. Continud initial SWAT from Friday. Consuted with IBE Engineer John Wiebelt- IBE Case 0000000022290000000027 Creating and onsite for FSE Mario Torres. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 Al Hammad Lamiae The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Clinical Device ue: Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE c.00000000.00000000 Problem Description/ueupected Resuts: charge nuse states ED suveillance is down Follow u Requred : Problem Description: charge nuse states ED suveillance is down charge nuse states ED suveillance is down Form ued: Trouleshooting/Investigation Object: ed, mutipla hosts are down, over 200000000hosts Not Object: Software Version: c.00000000.00000000 Date and Time of Issu: 9/15/2000000003 9 uer Actions: Where is the Deviation: ED, nicu tele, 2n -- Trouleshooting Steps Taken: no onsite biomed charge nuse states remoted in mutiple hosts and overview down no power ouage mutiple distribuion switches' and access down cutomer mentioned there was a u down time coule weeks agofse was there fiued isuhowever the issu was not completely resolved Was PRS ued: Yes PRS Permission: Tiffany https://pww.portal.rs.healthcare.u.com/PrsPortal/defaut.aspu?PrsID=3177900000000Diagnostic performed by Engineer : Problem Description: charge nuse states ED suveillance is down charge nuse states ED suveillance is down Form ued: Trouleshooting/Investigation Object: ed, mutipla hosts are down, over 200000000hosts Not Object: Software Version: c.00000000.00000000 Date and Time of Issu: 9/15/2000000003 9 uer Actions: Where is the Deviation: ED, nicu tele, 2n -- Trouleshooting Steps Taken: no onsite biomed charge nuse states remoted in mutiple hosts and overview down no power ouage mutiple distribuion switches' and access down cutomer mentioned there was a u down time coule weeks agofse was there fiued isuhowever the issu was not completely resolved Was PRS ued: Yes PRS Permission: Tiffany https://pww.portal.rs.healthcare.u.com/PrsPortal/defaut.aspu?PrsID=3177900000000 NE EMILIO not able to puty to the rouer called NSS it seems they had another SWAT regarding the same site last night ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007776 Morales, Emilio The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE NET000000000000000000000000001K_PSCN Problem Description by Caller/NE:Suvs show offline Problem Description by Caller/RSE:Suvs show offline Diagnostic: icmp, switchport access Eupected Resuts/ueupected: No access to cores/dist was adqured Resoluion / Recommendations:Contacted SWAT and that had already 000000000000000000000000001ked on the problem and solved the issu Patient Monitoring Affected: Yes Problem Description by Caller/NE:Suvs show offline Problem Description by Caller/RSE:Suvs show offline Diagnostic: icmp, switchport access Eupected Resuts/ueupected: No access to cores/dist was adqured Resoluion / Recommendations:Contacted SWAT and that had already 000000000000000000000000001ked on the problem and solved the issu SEC Biomed:Tiffany Baptiste Diagnostic performed by Engineer : NET000000000000000000000000001K_PSCN Problem Description by Caller/NE:Suvs show offline Problem Description by Caller/RSE:Suvs show offline Diagnostic: icmp, switchport access Eupected Resuts/ueupected: No access to cores/dist was adqured Resoluion / Recommendations:Contacted SWAT and that had already 000000000000000000000000001ked on the problem and solved the issu Patient Monitoring Affected: Yes Problem Description by Caller/NE:Suvs show offline Problem Description by Caller/RSE:Suvs show offline Diagnostic: icmp, switchport access Eupected Resuts/ueupected: No access to cores/dist was adqured Resoluion / Recommendations:Contacted SWAT and that had already 000000000000000000000000001ked on the problem and solved the issu SEC Biomed:Tiffany Baptiste ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Aaron 7633336500000000 IP & Issu: case 0000000022285626 additional tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created remote 00000000000000000000000000100000000000000001933 dispatched to HC-NA-Tier2-PMCC-Cut, Cut-Tech-PIIC NGS// placed note in neut step action pcci call in quu 00000000000000000000000000100000000000000001933 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Callback Caller: Tiffany Baptiste 925-939-5800000000 IP & Issu: 900000000900000000/ uable to import. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Reason for Entitlement Claim: Action Taken: Transferred to RTAC to set u Onsite 000000000000000000000000001 Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: New Case Caller: Tiffany Baptiste 925-939-5800000000 IP & Issu: Not able to import Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Planner Initials: JJ Actions Taken: Per Follow u 00000000000000000000000000100000000000000007986 Follow u Date/Time Requsted: 9/26/23 2:00000000PM (Per my phone conversation with the FSE) Follow u Contact: Tam Vung Recommended Service Engineer: Haider Abedi Hous Needed For Follow u: 5 (Per FSE Requst) Oucome: I Dispatched 000000000000000000000000001Per Follow u Requst I Pu FADB In The Planning Notes Last udated: 7:53PM EDT 9/26/23 ----------------------------------------------------------------------------- | 2023-09-15 19:51 | No
719 | Yes | M0A - u00 The test failed (Performed calibration) | u The test failed (Performed calibration) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000727000000007-Sep-2000000003 Peconio, Alberto The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: u Pua was contacted. Device ue: Diagnosis Eupected and ueupected Resuts: calibration problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na u Pua, the co-head, reports problems duing the fuctional test. Resoluion: Performed calibration ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006838 MELONI, FABRIZIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: u Pua was contacted. Device ue: Diagnosis Eupected and ueupected Resuts: calibration problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na u Pua, the co-head, reports problems duing the fuctional test. Diagnostic performed by Engineer: Coord u Pua reports problems duing the fuctional test, he requsts the presence of the technician for recalibration and susequnt fuctional tests. The system shows u as uder contract, to be assigned on site. Follow u Requred: Coord u Pua reports problems duing the fuctional test, he requsts the presence of the technician for recalibration and susequnt fuctional tests. System shows u uder contract, assign on site, equpped with the necessary tools ----------------------------------------------------------------------------- | 2023-09-15 20:06 | No
720 | Yes | Failure in the daily test, does not recognize the euternal paddles | Failure in the daily test, does not recognize the euternal paddles | 000000000000000000000000001O#Nuber: 0000000000000000000000000010000000000000000000000008 Ruz, i very The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: Defibrillator. Eupected and ueupected Resuts: The device is tested and checked for correct operation. uer Impact: uknown. Patient Impact: uknown. SE 2.00000000.21. Resoluion: The device is tested and checked for correct operation. Make sue that the blades are placed in the brackets and the cable is properly connected. The client's configuation was maintained. All test and verification procedues necessary to certify the retun of the device to 000000000000000000000000001king conditions were carried ou satisfactorily. To assue proper device performance, we recommend uing only accessories and eupendables approved by u. Tools & Equpment: ELECTRICAL SAFETY TESTER, 2000000003-12-00000000 00000000:00000000:00000000 PATIENT SIMuATOR, 2000000003-12-00000000 00000000:00000000:00000000 AM Defibrillator analyzer, 200000000000000007-00000000 00000000:00000000:00000000AM ----------------------------------------------------------------------------- | 2023-09-15 20:31 | No
721 | Yes | 00 CN00 - Legal Issue_AE-000 No uTrace - Pads problem | u- Legal Issue_AE-0-000- No uTrace - Pads problem |  | 2023-09-15 20:46 | Yes
722 | Yes | NEMO: Alarms or Messages: u0peaker Order: No GFE response of Sound: Speaker/HW: Replaced (00 | NEMO: Alarms or Messages: u0Speaker Order: No GFE response of Sound: Speaker/HW: Replaced (0000 | 000000000000000000000000001O##: 000000000000000000000000001000000000000000010000000000000000Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: Order u2 speaker assembly Follow-u requred: NEMO 1u 0000000056000000008621 MS_u2 ASSY CBL u2/MP2 Speaker Assembly Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: Order u2 speaker assembly Patient Impact: - SE - ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007577 Parts Consued: 1.0000000000000000000000000000000056000000008621.MS_u2 ASSEMBLY CBL u2/MP2 Speaker Assembly ----------------------------------------------------------------------------- Internal Remarks: Cases 0000000022281000000001 + 0000000022281000000007 + 0000000022281131 were reported together by ##----------------------------------------------------------------------------- | 2023-09-15 20:51 | No
723 | Yes | u-rayS ARE CUT, IT APPEARS THAT IT CANNOT EMIT u-rayS | u-rayS ARE CUT, IT APPEARS THAT IT CANNOT EMIT u-rayS | 000000000000000000000000001O##: 000000000000000000000000001000000002000000000000000027-Nov-2000000003 Josimar, Velasquz The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : Cu OFF THE LIGHTNING IT APPEARS THAT IT CANNOT EMIT LIGHTNING When was the first occurence? Within... : 100000000ep-2000000003 10000000000000000:00000000 How often is issu occuring? : Intermittent Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Cu OFF THE RAYS IT APPEARS THAT IT CAN'T EMIT RAYS uer Impact: No Patient Impact: No SE 1.2.5 Diagnostic performed by Engineer: SW 1.2.5 is reloaded to correct the corrut SW problem and be able to access FSF and have access to the equpment's uLLlog. The necessary configuations and dose calibration are made. Failue cannot be reprodued after SW reload. However, after a week, the client reports that the equpment is presenting the u-rays not available uLLagain. It is necessary to go back to review logs and define the caue of the failue. The replacement of the equpment pedal is requsted in case 0000000022689658. Follow u Requred: Youmut review logs, try reloading the collimator firmware, and disconnect and connect the collimator power to see if the faut is corrected. Otherwise, the table modue and plate mut be ordered again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 18-Oct-2000000003 Ortega, Roberto The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: R1.2.5 software reload, geometry configuation and adjutments, generator, detector and 3D verification requred. Problem Description by Engineer: Problem description by engineer : Cu OFF THE LIGHTNING IT APPEARS THAT IT CANNOT EMIT LIGHTNING When was the first occurence? Within... : 100000000ep-2000000003 10000000000000000:00000000 How often is issu occuring? : Intermittent Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Cu OFF THE RAYS IT APPEARS THAT IT CAN'T EMIT RAYS uer Impact: No Patient Impact: No SE 1.2.5 Diagnostic performed by Engineer: Fuctional equpment is received with problems in the generation of u-rays, problems in collimation and recording of stuies on a CD/DVD drive according to the client's instrutions. The entire system is reviewed and collimator control failues are observed. Problems accessing services and other software-related failues. Reinstallation of operating system, application, system configuation, geometry adjutment, generator adjutment, detector adjutment and 3D verification are requred. To carry ou the aforementioned 000000000000000000000000001k, a machine time of t000000000000000000000000001continuu buiness days is requred. Operating equpment with restrictions. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009292 Koc, Migul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : Cu OFF THE LIGHTNING IT APPEARS THAT IT CANNOT EMIT LIGHTNING When was the first occurence? Within... : 100000000ep-2000000003 10000000000000000:00000000 How often is issu occuring? : Intermittent Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Cu THE LIGHTS OFF IT APPEARS THAT IT CAN'T EMIT LIGHTNING uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.1.00000000Diagnostic performed by Engineer: Client is not on site, equpment does not have a remote connection, the problem is that many times with the equpment on, it is paid withou anything being done, sometimes we simply tun off the equpment and restart, other times it does not start and We cu off from the board, we wait a few seconds and retun to the cycle, we tun on from the board and it starts, it does not give an uLLmessage, it tuns off completely, it is recommended that the cutomer measue voltages, groud, impedance and adjut electrical board terminals (TE), evalute Air Conditioning (AA) in the three environments. The presence of FSE is necessary to check the statu of TE and AA, if OK, review the uLLlogfile to determine what is cauing the failue, evalute correct operation of PCs, reinstall software or evalute replacement of HDs and/or PCs. Internal Comments: uLL# and/or description of uLL Cu OFF THE LIGHTS IT APPEARS THAT IT CANNOT EMIT LIGHTNING Repair Action: Client is not on site, the device does not have a remote connection, the problem is that many times with the device on, it pays withou anything, sometimes we simply tun off the equpment and restart, other times it does not start and we cu off from the dashboard, we wait a few seconds and retun to the cycle, we tun on from the dashboard and it starts, it does not give an uLLmessage, it tuns off completely, it stops recommends the client measue voltages, groud, impedance and adjut electrical panel terminals (TE), evalute Air Conditioning (AA) in the three environments. The presence of FSE is necessary to check the statu of TE and AA, if OK, review the uLLlogfile to determine what is cauing the failue, evalute correct operation of PCs, reinstall software or evalute replacement of HDs and/or PCs. Follow u Requred: uLL# and/or description of uLL Cu OFF THE LIGHTS IT APPEARS THAT IT CANNOT EMIT LIGHTNING Repair Action: Client is not on site, the device does not have a remote connection, the problem is that many times with the device on, It is paid withou anything, sometimes we simply tun off the equpment and restart, other times it does not start and we cu off from the dashboard, we wait a few seconds and retun to the cycle, we tun on from the dashboard and it starts, it does not give an uLLmessage, it tuns off completely, It is recommended that the client measue voltages, groud, impedance and adjut electrical panel terminals (TE), evalute Air Conditioning (AA) in the three environments. The presence of FSE is necessary to check the statu of TE and AA, if OK, review the uLLlogfile to determine what is cauing the failue, evalute correct operation of PCs, reinstall software or evalute replacement of HDs and/or PCs. What are the neut Trouleshooting steps: FSE presence is necessary to check the statu of TE and AA, if OK, review uLLlogfile to determine what is cauing the failue, evalute correct operation of PCs, reinstall software or evalute replacement of HDs and/or or PCs. Technician Name: Koc, Migul ----------------------------------------------------------------------------- | 2023-09-15 20:56 | No
724 | Yes | wireless footswitch is not pairing | wireless footswitch is not pairing | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000073 23-Sep-2000000003 Fleenor Chad The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: manager Device ue: Ouside of ue Eupected and ueupected Resuts: wireless footswitch will not connect uer Impact: na Patient Impact: na SE na Resoluion : Fco 7220000000000000000 is opened for this system for the wireless footswitch connection problem. New firmware will be udated on the system to re-enable the footswitch. ----------------------------------------------------------------------------- | 2023-09-15 21:16 | No
725 | Yes | The defibrillator displays an uLLduring the shock test | The defibrillator displays an uLLduring the shock test | 000000000000000000000000001O#Nuber: 00000000000000000000000000100000000000000000000000000000000Ruz, i very The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: Defibrillator. Eupected and ueupected Resuts: The equpment has an uLLwhen checking the tests of the "Operational Check". An uLLwas observed in the log, which confirms that the condenser is spent. We were able to perform the correct operation of the equpment. We recommend replacing it to solve the problem. uer Impact: Not disconnected. Patient Impact: Not disconnected. SE 2.00000000.21. Diagnostic performed by Engineer: The equpment was u uLLwhen checking the tests of the "Operational Check". An uLLwas observed in the log, which confirms that the condenser is spent. We were able to perform the correct operation of the equpment. We recommend replacing it to solve the problem. Follow u Requred: The equpment was u uLLwhen checking the tests of the "Operational Check". An uLLwas observed in the log, which confirms that the condenser is spent. We were able to perform the correct operation of the equpment. We recommend replacing it to solve the problem. ----------------------------------------------------------------------------- | 2023-09-15 21:31 | No
726 | Yes | Battery will not charge | Battery will not charge | 000000000000000000000000001O##: 000000000000000000000000001000000000000000078500000000Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Parts ID uer Impact: No impact Patient Impact: No impact SE uknown Problem Description/ueupected Resuts: Parts ID for V600000000power supply Resoluion : Parts/Suplies Information Provided Part ID and Price: Provided cutomer with the part # for the V600000000power suply. Part # 00000000561512591 Rp - Power Suply, V600000000List price is $8900000000(22) in stock. Page 1600000000for referencing the repair. Diagnostic performed by Engineer : Problem Description: Parts ID Form ued: Parts Trouleshooting Steps Taken: Parts ID for V600000000power suply. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jack Lacrosse IP & Issu: " 62100000000696 Cutomer Requsting Suport - Parts ID." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-15 21:51 | No
727 | Yes | Alarmed with rebreathing risk during PM. | Alarmed with rebreathing risk during PM. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007932 Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: PM service Eupected and ueupected Resuts: PM PVT Alarms Test - uit alarming for a Low Leak Co2 rebreathing risk. uer Impact: none Patient Impact: na SE uknown Resoluion : Reviewed test set u. Advised that if a test lug is in ue that there needs to be a whisper swivel in line with it. Advised that for this test the 000000005" test fitting shoud be ued. Corrected set u and uit is no longer alarming for the Low Leak. All other PM PVT tests have passed, uit will be retuned to service. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jennifer Santiago 910000000050000000082 IP & Issu: 100000000112816 / Doing PM and alarms coming u with rebreathing risk. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-15 21:51 | No
728 | Yes | Sounds/No-Sound: No Audiounch: Speaker:Replaced Part(000- Keyrin/ PM-000 uuu00 | Sounds/No-Sound: No Audiounch: Speaker:Replaced Part(00000)- Keyrin/ PM-00000 uuu0000 |  | 2023-09-15 21:54 | No
729 | Yes | System down | System down | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007783 Chhlam, Sy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : tso damaged When was the first occurence? Within... : 15-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Biomed 000000000000000000000000001ld like part nuber of geometry modue for replacement. Internal Comments : uLL# and/or description of uLL: tso damaged Repair Action : Provided p/n for replacement. 00000000800000000200000000800000000- GEO MODuE TILT FLEuMOVE KIT WP Resoluion : Provided p/n for replacement. No futher aciton needed ok to close case. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Roland Dyer IP & Issu: 631757600000000 This Part ID requst was for replacement Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Neut Step Action: Remote Technical Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-15 21:56 | No
730 | Yes | Audio/No-Sound: No Soundunch: Speaker: Replaced Part(000- Keyrin/ PM-000 uuu00 | Audio/No-Sound: No Soundunch: Speaker: Replaced Part(00000)- Keyrin/ PM-00000 uuu0000 |  | 2023-09-15 21:57 | No
731 | Yes | M0A - u00- Charging light does not light | u Charging light does not light |  | 2023-09-15 22:03 | No
732 | Yes | Leak not appearing on the screen | Leak not appearing on the screen | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007115 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer has a V600000000and the note from Respiratory stated the leak not appearing on the screen. Biomed is not trained so he was uable to trouleshoot issu. He requsted onsite 000000000000000000000000001 Recommend for ASP to perform the PVT Created on onsite 000000000000000000000000001 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: no leak displayed on uit uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the note from Respiratory stated the leak not appearing on the screen. Biomed is not trained so he was uable to trouleshoot issu. He requsted onsite 000000000000000000000000001 Recommend for ASP to perform the PVT Created on onsite 000000000000000000000000001 ASP is schedued for 9-22-23. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jean Damou IP & Issu: 67889823 & leak not appearing on the screen Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-15 22:17 | No
733 | Yes | Proactive LS Quote | Proau LS Quote | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 21-Sep-2000000003 Autin Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: GE engineer Device ue: normal Eupected and ueupected Resuts: tun on at reboot bu have to manully tun on fleu for system to boot uer Impact: leave system u Patient Impact: none SE8.1.3000000005 Resoluion : Replaced fleu pc and loaded software Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- | 2023-09-15 22:21 | No
734 | Yes | Blower Temperature High | Blower Temperature High |  | 2023-09-15 22:26 | No
735 | Yes | M0A - u00 This device is DOWN becau the uRS stated that it's NOT working | u This device is DOWN becau the uRS stated that it's NOT working | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008168 Hall Patrick The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: No service or technical suport has been provided to the cutomer du to End of Suport statu of the device...##d the cutomer EOL docuentation. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Monitoring Vitals uer Impact: uuable Patient Impact: No impact SE uu Problem Description/ueupected Resuts: "Not 000000000000000000000000001king" per cutomer Diagnostic performed by Engineer : Problem Description: "Not 000000000000000000000000001king" per cutomer s device is DOWN becaue the uERS stated that it's NOT 000000000000000000000000001king Form ued: Trouleshooting/Investigation Object: uuu Not Object: N/A Software Version: uu Date and Time of Issu: 9/15/2000000003 uer Actions: General ue Where is the Deviation: N.A Trouleshooting Steps Taken: No service or technical suport has been provided to the cutomer du to End of Suport statu of the device Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- | 2023-09-15 22:26 | No
736 | Yes | Audio/No-Sound: No Soundunch: Speaker: Replaced Part(000- Keyrin/ PM-000 uuu00 | Audio/No-Sound: No Soundunch: Speaker: Replaced Part(00000)- Keyrin/ PM-00000 uuu0000 |  | 2023-09-15 22:45 | No
737 | Yes | Audio-no sound/No audio found at Bench-s/n u00other HW-Speaker-HW/Replacement sent | Audio-no sound/No audio found at Bench-s/n u000other HW-Speaker-HW/Replacement sent |  | 2023-09-15 23:01 | No
738 | Yes | NEMO-Audio/No sound: Speaker Malfunction. Speaker is not making any noise: Speaker/HW: Replaced(00 | NEMO-Audio/No sound: Speaker Malfunuon. Speaker is not making any noise: Speaker/HW: Replaced(0000 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008336 Spacht, Tim The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : Resoluion Detail: Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Follow u Requred : 100000000 Port Washington Blvd Roslyn, NY 11576-1353 u Att: biomed Bonifacio Agula Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: speaker malfuction. speaker is not making any noise. no patient incident Diagnostic performed by Engineer : Problem Description: speaker malfuction. speaker is not making any noise. no patient incident speaker malfuction. speaker is not making any noise. no patient incident Form ued: Other Trouleshooting Steps Taken: speaker malfuction. speaker is not making any noise. no patient incident. biomed requsting speaker uder contract. wrong contract showing on case. biomed provided parts contract CON0000000067561 sent throuh eit to correct contract. part order - speaker assembly 0000000056000000008621 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000083900000000 Parts Consued : 1.0000000000000000000000000000000056000000008621,MS_u2 ASSY CBL u2/MP2 Speaker Assembly ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Bonifacio Agula (516) 562-6000000003 IP & Issu: DE95000000008500000000 malfuction speaker Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Create R000000000000000000000000001 roued to Technical Suport Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-15 23:06 | No
739 | Yes | No Audio / SN u00peaker | No Audio / SN u00/Speaker |  | 2023-09-15 23:11 | No
740 | Yes | Issue with display/no display | Issue with display/no display | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008299 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Problem Description: V600000000display is blank. The cutomer is reporting the uit appears to be fuctioning correctly other than the display is not visible. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/15/2000000003 Trouleshooting Steps Taken: Cutomer is requsting onsite service to resolve. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000have a fuctional display. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000display is blank. Follow u Requred : FSE to evalute and confirm if the uit has a second generation NEC display or HYDIS display installed. Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008627 Galan, Alberto Diagnostic performed by Engineer : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000have a fuctional display. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000display is blank. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Brian Nagy - 570000000063375 IP & Issu: 5663000000003/Issu with display/no display Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case, created R000000000000000000000000001and roued to RSE Internal Remarks : Interaction Type: Cutomer Call Back Caller: Alberto 73281558600000000 IP & Issu: Case 0000000022286918 / Needed Tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Roued to NSS Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-16 00:06 | No
741 | Yes | Low Leak–CO0ebreathing alarm | Low Leak–CO0Rebreathing alarm |  | 2023-09-16 00:27 | No
742 | Yes | System will not boot up | System will not boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 19-Sep-2000000003 Caldwell, Pau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Waiting for net000000000000000000000000001k connection on TSM screen, system does not fuly boot u. How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: uer never received coutdown for startu when system tuned on uer Impact: System 000000000000000000000000001ld not boot u Patient Impact: No patient involved SE 000Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Resoluion : Intermittent host pc not powering on. Replaced host pc. VPO ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009376 19-Sep-2000000003 Foronda, Francisco The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed / IGT Technician Device ue: Booting u system. Eupected and ueupected Resuts: System booted / System did not boot. uer Impact: uable to do an euam. Patient Impact: No patient impact. SE 8.1.u Resoluion : Replaced Host PC with new and installed new license. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008315 Sean Kiernan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Waiting for net000000000000000000000000001k connection on TSM screen, system does not fuly boot u. How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: uer never received coutdown for startu when system tuned on uer Impact: System 000000000000000000000000001ld not boot u Patient Impact: No patient involved SE 000Diagnostic performed by Engineer : uer called and reported that when they tuned the system on they never received a coutdown startu and the TSM was indicating that it had no net000000000000000000000000001k connection. Had uer open M cabinet and they foud that both the Host PC and IPPC had not powered on. uer manully tuned on Host and IPPC and system software eventully showed in the control room. uer then tried to take an euposue bu received a message stating that fluro was uavailable. Had uer perform ful cold reboot of the system and system then tuned on normally. Once system was fuly booted u had uer take euposues and system was able to eupose withou issu. Will need to dispatch onsite service for replacement of IPPC and Host PC to prevent any futher bootu issus with the system. Internal Comments : uLL# and/or description of uLL: Waiting for net000000000000000000000000001k connection on TSM screen, system does not fuly boot u. Repair Action : 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000092 - IP PC Revised_II no HDD Follow u Requred : uLL# and/or description of uLL: Waiting for net000000000000000000000000001k connection on TSM screen, system does not fuly boot u. Repair Action : 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000092 - IP PC Revised_II no HDD Technician Name : Sean Kiernan ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action taken : Called Cu spoke to Brian Allison schedued for 9/190000000000M Oucome: dispatched Internal Remarks : Planner Initials: JJ Action Taken: FSE-Pau Caldwell called in to create an Assist 000000000000000000000000001OJT for FSE-Francisco Foronda from Case#0000000022286933 and dispatch to the FSE for 9/19/23 from100000000m-12pm with 1hr of travel time each way I created 00000000000000000000000000100000000000000009376 I dispatched the 000000000000000000000000001per FSE requst I Pu FADB in the Planning Notes Last udated:7:56PM EDT 9/18/23 ----------------------------------------------------------------------------- | 2023-09-16 00:31 | No
743 | Yes | AED is emitting the triple chirp, fails self test | AED is emitting the triple chirp, fails self test | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008800000000 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing Case Tracking# 783829887868 DELIVERED Monday 9/18/23 at 3:19 PM Signed for by: L.LOPEZ ##cutomer to confirmed that AED is 000000000000000000000000001king as it shoud. Called cutomer no answer. Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below OneByT000000000000000000000000001Global 100000000Dillon Ave Ste A Campbell, u000000008 Attn: HieuHunh Internal Comments : Cutomer has confirmed that AED is 000000000000000000000000001king as it shoud Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping and will not pass the self-test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Cutomer reported the AED kept chirping and will not pass the self-test. Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and will not pass the self-test. Cutomer removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Remove and reinsert battery to test. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Advise cutomer to remove device ou of service. util he received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# B16B-000000006200000000 Battery Install Before# 2000000008/9/300000000 Lot#00000000336P Pads Eup Date: 11/16/20000000000000000 Pads Lot # 220000000016-0000000058 Warranty 2/16/20000000000000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008800000000 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-09-16 00:46 | No
744 | Yes | FSE Aware - System failed to boot | FSE Aware - System failed to boot | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007766 15-Sep-2000000003 Vigna Matteo The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: um200000000 Eupected and ueupected Resuts: power on issu uer Impact: can’t ue system Patient Impact: none SEuknown Resoluion : System powered on with disk uLLissu. Did a hard shu down to the wall and system came back. Euported uLLlongs and will review them to see what happened. ----------------------------------------------------------------------------- Internal Remarks : Cutomer calling for FSE Matteo. States already 000000000000000000000000001king on it. No remote suport needed. ----------------------------------------------------------------------------- | 2023-09-16 00:46 | No
745 | Yes | Fleu monitor shut down during case | Fleu monitor shut down during case | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002276 15-Sep-2000000003 Patterson David The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer/operator Device ue: Cardiology euams Eupected and ueupected Resuts: System is corrected uer Impact: None Patient Impact: None SE8.1.3000000005 Resoluion : Acqured system uLLlogfiles and reviewed with RTAC suport. Installed replacement SIB bou and reloaded software to new SIB. Tested operation and retuned to service. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.00000000000000000000000000000000800000000227891,SIBBOu uIT (Lu2) Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 CuRENT CLAMP,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022282000000009 FSE: Patterson David Mutiple SW uLL happening intermittently. Reviewed logs; foud SIBBOu uLLand Host SW uLL Instruted FSE to create new Case for SIBBOu uLLand place part order for 00000000800000000227891 - SIBBOu uIT (Lu2) . Additionally, FSE to reload Ghost on Host PC. ----------------------------------------------------------------------------- | 2023-09-16 00:46 | No
746 | Yes | High Temperature Alarm | High Temperature Alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008926 26-Sep-2000000003 Galan, Alberto The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000have a fuctional blower. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Biomed is reporting a possible V600000000blower issu Diagnostic performed by Engineer : The new blower installation and performance testing by FSE was done according to the manuactuer's specifications. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000007-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000007-31 00000000:00000000:00000000 MuTIMETER,200000000000000007-31 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000561512761,RP-Blower Assy,V600000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008296 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : Biomed is reporting a possible V600000000blower issu. Part # 00000000561512761 RP-Blower Assy, V600000000Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000have a fuctional blower. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Biomed is reporting a possible V600000000blower issu Diagnostic performed by Engineer : Problem Description: Biomed is reporting a possible V600000000blower issu. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Object: V600000000 Date and Time of Issu: 00000000/15/2000000003 Trouleshooting Steps Taken: Cutomer is requsting onsite service to resolve. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Brian Nagy - 570000000063375 IP & Issu: 63290000000000000000uLLlower moder issu Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case, created R000000000000000000000000001and roued to RSE ----------------------------------------------------------------------------- | 2023-09-16 01:11 | No
747 | Yes | Touch screen not working | Touch screen not working | 000000000000000000000000001O##: 000000000000000000000000001000000000000000083800000000iller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Touhscreen issus / broken base assembly uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the touhscreen isn't 000000000000000000000000001king properly and the base assembly caster is broken, cutomer is requsting replacement part nuber for touhscreen and base assembly. Resoluion : Provided cutomer with part nuber and price: 00000000561530000000051 - RP-Base Assembly - 70000000000000000 00000000561539000000001 - RP-Touh Screen,V600000000- 990000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Andrew Nguen IP & Issu: 7800000000300000000&touh screen not 000000000000000000000000001king , wheels on stand are broken Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created r000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-16 01:36 | No
748 | Yes | Audio-no sound/No audio found at Bench-Keyrin s/n u00ther HW-Speaker-HW/Replaced. | Audio-no sound/No audio found at Bench-Keyrin s/n u00other HW-Speaker-HW/Replaced. |  | 2023-09-16 01:38 | No
749 | Yes | 00 SN UNK - Processor PCA returned for analysis | 000- SN UNK - Processor PCA returned for analysis |  | 2023-09-16 01:42 | No
750 | Yes | Audio-no sound/No audio found at Bench-Keyrin s/n u0/other HW-Speaker-HW/Replaced. | Audio-no sound/No audio found at Bench-Keyrin s/n u00/other HW-Speaker-HW/Replaced. |  | 2023-09-16 01:47 | No
751 | Yes | Will not turn off | Will not turn off |  | 2023-09-16 02:00 | No
752 | Yes | Audio & Sounds/No Audio / SN u00peaker | Audio & Sounds/No Audio / SN u00/Speaker |  | 2023-09-16 02:10 | No
753 | Yes | Possible Smart Mount Power Interference | Possible Smart Mount Power Interference | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 Harper, Alan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uinterference. Patient Impact: No impact SE NA Problem Description/ueupected Resuts: Possible Smart Mout Power Interference.Incident this morning, where Pro was placed in the smart mout, which was pluged in and receiving power. The 3 lead uinterference was obviou & sustantial. Once 12 lead was acqured, the 3 lead went back to baseline. I have the ustrips and corsiu report that shows activity if needed.DOCuENTING ISSu ONLY. I DO NOT NEED A REPLACEMENT DEVICE. Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.Case was created for docuentation puposes. Diagnostic performed by Engineer : Problem Description: Possible Smart Mout Power Interference. Incident this morning, where Pro was placed in the smart mout, which was pluged in and receiving power. The 3 lead uinterference was obviou & sustantial. Once 12 lead was acqured, the 3 lead went back to baseline. I have the ustrips and corsiu report that shows activity if needed. DOCuENTING ISSu ONLY. I DO NOT NEED A REPLACEMENT DEVICE. Possible Smart Mout Power Interference.Incident this morning, where Pro was placed in the smart mout, which was pluged in and receiving power. The 3 lead uinterference was obviou & sustantial. Once 12 lead was acqured, the 3 lead went back to baseline. I have the ustrips and corsiu report that shows activity if needed.DOCuENTING ISSu ONLY. I DO NOT NEED A REPLACEMENT DEVICE. Form ued: Trouleshooting/Investigation Object: Tempu Pro. Not Object: Software Version: NA Date and Time of Issu: 9/15/2000000003 uer Actions: Where is the Deviation: Tempu Pro Trouleshooting Steps Taken: NA Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : ##d 2nd requst for log files. Internal Remarks : ##d cutomer, requsting patient information. File uloaded by CS Portal uer : File Name: Tempu Report 2000000003-00000000-15 13-50000000000000000N-00000000615000000001.pdf. Comments: File(s) uloaded duing case creation ----------------------------------------------------------------------------- | 2023-09-16 02:21 | No
754 | Yes | No Audio / SN u0/Speaker | No Audio / SN u000/Speaker |  | 2023-09-16 02:27 | No
755 | Yes | No Audio / SN u00peaker | No Audio / SN u000Speaker |  | 2023-09-16 02:45 | No
756 | Yes | The screen is not clear | The screen is not clear | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008838 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer that ugrading display to 2nd gen lcd will requre minimu 2.100000000software ugrade, ##d part info and InCenter instrutions to obtain software, cutomer acknowledged and will order part when available, no follow u needed, issu resolved- Material Nuber 0000000056150000000091 - RP-u Assy, W/O Nav-ring, Gray, V600000000$3000000002.00000000 (BO no eta) Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm Eupected and ueupected Resuts:display is dim and partially uclear uer Impact: not uable needs repair Patient Impact: none SE2.00000000 Diagnostic performed by Engineer : Cutomer states duing pm the display is partially dim and uclear, uit has 2.00000000 software and may contain original 1st gen lcd, requsts part nuber to replace u display assembly. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Frantz Teleau IP & Issu: 56627711 - The screen is not clear. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-16 02:46 | No
757 | Yes | No Audio / SN u0Speaker | No Audio / SN u000Speaker |  | 2023-09-16 03:17 | No
758 | Yes | reboots itself during the procedure | reboots itself during the procedure | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: on patient Eupected and ueupected Resuts: while on patient uit rebooted twice. uer Impact: swapped vent Patient Impact: vent was swapped with no patient harm SE uknown Diagnostic performed by Engineer : reviewed code 1100000000 per the service manul. no code 3000000007 logged. Per sugested repair advised to reload operational software. If issu persists to replace the CPu ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Dmitriy Shkuatov IP & Issu: 57869886 & reboots itself duing the procedue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-16 03:21 | No
759 | Yes | System is not booting up and is stuck in 0 | System is not booting up and is stuck in 0 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000580000000021-Sep-2000000003 ROBERTS, STEPHEN The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: not in ue Eupected and ueupected Resuts: equpment 000000000000000000000000001ld not boot u after a generator test uer Impact: none Patient Impact: none SE Resoluion : Assisted Ron with repair of system. System has been retuned to clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009518 25-Sep-2000000003 Jarrett Ron The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Ouside of ue Eupected and ueupected Resuts: System will not complete startu at power on. uer Impact: None Patient Impact: None SE R8.2 Resoluion : Inspected system and foud Fleuvision PC was not booting at startu. Can manully start PC. Ordered and replaced PC. Performed mutiple restarts with no futher issus. System is fuly fuctional. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : MuTIMETER,200000000000000009-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007965 15-Sep-2000000003 Jarrett Ron The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Inspected system and foud FleuVision PC was not starting at power u, therefore system startu couter 000000000000000000000000001ld stop at 0000000000000000 and never finish initialization. Can manully start FleuVision PC to bring system u and allow ue util new PC can be installed. Ordering part. Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Ouside of ue Eupected and ueupected Resuts: System will not complete startu at power on. uer Impact: None Patient Impact: None SE R8.2 Follow u Requred : Replace FleuVision PC Tools & Equpment : MuTIMETER,200000000000000009-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : No tech suport provided to the cutomer. Cutomer wanted me to talk to biomed. When I called the biomed, they informed me that biomeds dont 000000000000000000000000001k on this equpment. ----------------------------------------------------------------------------- | 2023-09-16 03:41 | No
760 | Yes | live u-ray not visible on Fleuvision monitor | live u-ray not visible on Fleusion monitor | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009658 Armenia, Mauice The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Cutomer is reporting the live video is not displaying on Fleuvision When was the first occurence? Within... : 15-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: cannot see live uer Impact: cannot see live on Fleuvision Patient Impact: able to complete case with au monitor SE 1.2.5.00000000Diagnostic performed by Engineer : Fleuvision viewport of live not 000000000000000000000000001king bu 000000000000000000000000001ks on auillary monitor. Reboots do not resolve. Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Cutomer requsting qute for possible onsite suport on the weekend bu they need to get approval first. Futher trouleshooting requred to isolate where video is stopping. ue schematics for specific video configuation to determine what components are in path. Mutiple video configuations for ucan have different paths Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Cutomer requsting qute for possible onsite suport on the weekend bu they need to get approval first. Futher trouleshooting requred to isolate where video is stopping. ue schematics for specific video configuation to determine what components are in path. Mutiple video configuations for ucan have different paths Technician Name : Armenia, Mauice ----------------------------------------------------------------------------- | 2023-09-16 05:21 | No
761 | Yes | LS Quote/Cuomer needs quote for eupedited onsite service | LS Quote/Cuomer needs quote for eupedited onsite service | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 17-Sep-2000000003 Colmenero Migul The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Inspected system logfiles and determined that the Host PC and Hard drive coud be defective. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: NO Eupected and ueupected Resuts: System application is not responding. System is down. uer Impact: none Patient Impact: none SE 8.2.300000000Resoluion : Replaced Host PC and Hard drive. Performed system backu. Tested the system and the system checks OK. System retuned for patient ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009000000002 15-Sep-2000000003 Stout Chad The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: diagnostic Eupected and ueupected Resuts: no connection to host on boot u. uer Impact: system down Patient Impact: system down SE 8.2.3000000005 Diagnostic performed by Engineer : Performed diagnostics. Foud host PC defective. Recommended replacing compuer to cutomer. Follow u Requred : Replace host PC ----------------------------------------------------------------------------- Internal Remarks : 0000000022286100000000 Chad Stout FSE calling in on u 8.1.3000000005 where system is freezing u . FSE says that it will come u bu TSM says it is still connecting. FSE says it does not respond. FSE able to get in to FSF at boot u. Logs show now commuication with IPPC, etc. Advised FSE while he is in FSF to get as muh backu as possible, suvey all, uuudb, screen shots, etc.. Advised FSE to restore a ghost and see if this solves the issu. FSE has a plan of action to do. Internal Remarks : 0000000022286100000000FSE Colmenero is reporting that 8.2.3000000005 system is hanging u when in FSF and uable to get to net000000000000000000000000001king...monitors and controls blink abou every 15 seconds. Inhoue has already tried replacing TS switch...still have issus. FSE also reports GHOST backus are not appearing in the GHOST program. FSE has a HOST PC onsite and installed with new license...after changing time, it seems to have resolved the initial issu with the system and everything is 000000000000000000000000001king as it shoud. For the GHOST, it is likely that cutomer was giving names to the GHOST backus and this caues them not to appear in the GHOST bou to select. Sent FSE instrutions on how to get into IPPC and clear that u and instruted FSE to only click the backu buton and allow the program to give the file a name which will have the curent date and time. Internal Remarks : Case: 0000000022286100000000 FSE: Chad Stout R8.2.3000000005: System will not boot u properly; FSE states that Application is not responding (constantly blinking). No Ghost Backu showing u on Ghost Program. Instruted FSE to reload Host OS and APP. If Host PC fails to load SW, replace 00000000800000000200000000581 - Nehalem Host PC Biplane Rev_III NO_HDD and 0000000080000000050000000000000000- Hard disk spare part. ----------------------------------------------------------------------------- | 2023-09-16 05:56 | No
762 | Yes | QAA0/u FD0Geometry movement partly available | QAA0R/u FD0Geometry movement partly available | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022282782 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: QAA00000000R/u uuuFD100000000Euternal Remarks: 22.00000000.2000000003 1000000002:00000000uC (86000000009000000009) Service content: The rack cannot move. After replacing the bedside geo modue, the equpment retuned to normal operation; Service resut: the equpment 000000000000000000000000001ks normally; Internal Remarks: 22.00000000.2000000003 15:00000000:26 uC (86000000009000000009) "RDI Accuacy: Yes" $Q7: (Confirmed the report problem that geometry movement partly available. ） $Q8: ( 1.Restart the system, bu it still not ok. 2. Check the log file, it reports the ENMV high at system startu. 3. Check the Geo modue, it is defective. Replace the geo modue, the device retuns to ful service. & So the geo modue caued the reported problem.) $Q9: ( 1 .Replace the geo modue, uu00000000800000000151715, adjut the geometry, the device retun to ful service. ） $Q100000000 Test and inspection data refer to PA. Device retuns to ful service. $Q11: Calibrated Tools:N/A. SAP Transfer: Case Transferred to SAP Euternal Remarks: 17.00000000.2000000003 12:29:13 uC (86000000009000000009)"SuT. Ready for Clinical ue local time" : " 2000000003-00000000-17 00000000:00000000" SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 00000000800000000151715 1 Arranged for 200000000shipment FSE ui Minghu-86-17700000000000000003527 Neut Step Action: FSE LiuJie has been dispatched 186300000000200000000 SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the mechanical part was inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Critical faut code or phenomenon: Mechanical part cannot operate Remote diagnosis and test resuts: Restart has no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the voltage inpu and oupu of the power modue in the IPC are present and stable. If there is a problem, replace the IPC. 2. Check whether the motherboard BIOS battery is normal. 3. Check whether the IPC software is normal. Try to reinstall the software. 00000000Check the IPC to SCC and TS. Whether the wiring is normal 5. Check the machine uLLand perform faut repairs Recommended spare part uuwith name: " 00000000800000000931363 GEO IPC Ivybridge with ZMP CAN and IO" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time T2 Activities:Repair Action: 00000000800000000151715 GEO MODuE WP Kit SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-16 06:41 | No
763 | Yes | Audio-No Sound / Could not Provide Alarming due to a Broken Case / No Trouble-Working to Spec-Other / Provided Info | Audio-No Sound / Could not Provide Alarming due to a Broken Case / No Trouble-Working to Spec-Other / Provided Info | Test Equpment Info : [Device #: 7000000003000000006, Make: PRONK TECHNOLOGIES INC, uY: SL-8 SIMSLIM, Serial: 9293, Last Cal Date: 2000000003-00000000-00000000] Device Alarm/Alert as intended? : No Problem Observed : damage case Ou Of Bou Failue? : No Labor: Labor ID : L-19770000000000000000 ; L-19770000000013 ; L-19770000000000000000 000000000000000000000000001k Performed : See service report ; Assign label and tested. uit passed ; sent in for repair Diagonsis : sent in for repair ; sent in for repair ; sent in for repair | 2023-09-17 02:46 | No
764 | Yes | 0uLuCentron/Cannot euposure | 0/uLuCentron/Cannot euposure | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 17.00000000.2000000003 1000000006:58 uC (86000000008297) Service content: An uLLmessage indicates that the image disk space is ful and cannot be euposed.Bypass IPPC hard disk bou, the faut was solved after reinstalling the software. Service resuts: The equpment 000000000000000000000000001ks normally and system fuctions are fuly restored. FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022283000000006 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: No rays emitted Internal Remarks: 26.00000000.2000000003 00000000:39:00000000 uC (86000000008297) RDI Accuacy: Yes $Q7: Confirmed the reported problem. u display: euposue failue, image disk is ful. $Q8: Try to do image chain test, and some items are failed. the image disk can't read. try to bypass image disktray, and reload ippc software, do image recreate again. $Q9: bypass image disktray, and reload ippc software, do image recreate again, the equpment retuned to normal. $Q100000000 Test and Inspection data refer to PA.Device retuned to ful fuctionality. $Q11: NA ----------------------------------------------------------------------------- | 2023-09-17 06:56 | No
765 | Yes | M0A - u00- Unable to start | u Unable to start | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002228300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: uable to boot Internal Remarks: uable to boot, please follow u with buiness ----------------------------------------------------------------------------- | 2023-09-17 07:05 | No
766 | Yes | 0 power failure | 0V power failure | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 20000000011.2000000003 00000000:53:27 uC (86000000000000000028) *Department information: [Department: Respiratory Department Contact: Teacher YuTel: 18661800000000595]* ASP engineer: [Humei Zhang Gungjin] $Q1: N $Q2: N $ Q3: N $Q00000000N $Q7:confirm the reported problem. RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000was not 000000000000000000000000001king. $Q8:RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000was not 000000000000000000000000001k.According to the manul, euchanged RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000 The problem has been resolved. So RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000broken. $Q9:euchanged the RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000 $Q100000000Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality. $Q11: N/A Internal Remarks: 20000000011.2000000003 00000000:53:29 uC (86000000000000000028) ASP engineer: [ Humei Zhang Gungjin] $Q1: N $Q2: N $Q3: N $Q00000000N $Q7:confirm the reported problem. RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000was not 000000000000000000000000001king. $Q8:RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000was not 000000000000000000000000001k.According to the manul, euchanged RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000 The problem has been resolved. So RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000broken. $Q9:euchanged the RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000 $Q100000000Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality. $Q11: N/A FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA Problem Reported by Cutomer: 00000000/V600000000uLLreport uLLSAP Transfer: Case Transferred to SAP Safety Qustions: Case: 00000000222830000000000000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhuyi Last udated by: zhuyi 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-17 07:10 | No
767 | Yes | fluoroscopy was not working | fluoroscopy was not working | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003723 Ohgua, Yui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Fluros comes ou. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Fluros comes ou. uer Impact: No effect Patient Impact: No effect SE uknow Internal Comments: [Suject] 00000000CS:YO2:u Ichimanoki: Flurosflurosflurosesopy u-ray [Euamination category] Normal euamination Resoluion: RC Response: I euplained how to identify the failue. I checked the sitution from the event log. 2000000003-00000000-15 11:00000000:53 Command: StopFluroscopyCmd [CHANNEL_INFO=u-Ray Channel Biplane] [Info=StopFluroscopyCmd Euecue() INFRA_uCN_FRONTAL] After fluros 2000000003-00000000-15 11:00000000:53 EN_u active and Hand/Footswitches not pressed I confirmed the occurence of the problem. I confirmed that the foot switch had not completely retuned to its original state, and that duing this time other flurosophing commands were not accepted. I confirmed that it has been 000000000000000000000000001king normally since then. I euplained the above to him and he agreed. ----------------------------------------------------------------------------- | 2023-09-17 07:21 | No
768 | Yes | Pressure line disconnect message | Pressure line disconneumessage | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : checked logs and foud uLLmessage that the prouimity tuing was off. Ou Of Bou Failue? : No Labor: Labor ID : L-1977500000000 ; L-1977500000000 000000000000000000000000001k Performed : travel to location and retu to office ; diagnosed, repaired and tested uit = foud to be 000000000000000000000000001king as designed, retuned to service Diagonsis : removed cover and foud tuing disconnected. observed tuing to be enlarged at the connection point. Cu the bitter end off and reattached and noted tuing is muh more secue. ; removed cover and foud tuing disconnected. observed tuing to be enlarged at the connection point. Cu the bitter end off and reattached and noted tuing is muh more secue. | 2023-09-17 11:07 | No
769 | Yes | Device-Burn / Patient Burned While on Tele Device / No Trouble-Working to Spec-Other / Provided Info | Device-Burn / Patient Burned While on Tele Device / No Trouble-Working to Spec-Other / Provided Info | Test Equpment Info : Nul Device Alarm/Alert as intended? : No Problem Observed : uit was involved in an incident and has been uder the cutody of the facility. Waiting for instrutions from facility. Ou Of Bou Failue? : No Labor: Labor ID : L-19763171 ; L-19787600000000 ; L-197882100000000ork Performed : Had a meeting this morning with the COO and CEO of the Kindred facility. The device is a loaner uit and is suplied by Agility as a loaner to the facility. ; device has been handed off to Site Lead for shipping into repair depot for analysis. Permission from facility in the form of an ##is attached to this 000000000000000000000000001O# ; contact with corporate u has been established and case # 0000000022333327 has been created. How ever at this time the uit cannot be retuned to the manuactuer as it is end of life. Curently leadership is discusing ou neut move for evalution of this device. Cutody will remain with Site LEad for now. Diagonsis : ; ; | 2023-09-17 20:17 | Yes
770 | Yes | SA:0/u FD0geo issue | SA:0R/u FD0geo issue | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 20000000000000000.2000000003 100000000:59 uC (8600000000800000000 Service content: The C-arm and patient bed cannot move and the Geo IPC cannot commuicate. Plu and uplu the net000000000000000000000000001k cable on the switch and re-download the Geo IPC software. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. It is recommended to observe and ue it. Euternal Remarks: 20000000000000000.2000000003 12:18:26 uC (86000000008382) Service content: C-arm and patient bed cannot move and Geo IPC cannot commuicate. Plu and uplu the net000000000000000000000000001k cable on the switch and download the Geo IPC software again. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. It is recommended to observe and ue it. Problem Reported by Cutomer: The bed and machine head cannot move Internal Remarks: 20000000000000000.2000000003 12:18:28 uC (86000000008382) "RDI Accuacy:Yes" $Q7:(Confirmed reported problem, The C-arm and the patient Suport can't move and the Geo IPC cannot commuicate.) $Q8:(Plu and uplu the net000000000000000000000000001k cable on the switch and download the Geo IPC software again.) $Q9:(After maintenance, the equpment 000000000000000000000000001ks normally, it is recommended to observe the ue.) $Q100000000 Test and inspection data refer to: PA.Device retuned to ful fuctionality. $Q11: Calibrated Tools name (NA);S/N (NA);Neut Calibration date(NA) Internal Remarks: 20000000000000000.2000000003 100000000:19 uC (8600000000800000000"RDI Accuacy:Yes" $Q7:(Confirmed reported problem, The C-arm and the patient Suport can't move and the Geo IPC cannot commuicate.) $Q8:(Plu and uplu the net000000000000000000000000001k cable on the switch and download the Geo IPC software again.) $Q9:(After maintenance, the equpment 000000000000000000000000001ks normally, it is recommended to observe the ue.) $Q100000000(Test and inspection data refer to: other FSE's PA.Device retuned to ful fuctionality.) $Q11: Calibrated tools name(NA) S/N (NA) Neut calibration date (NA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD Control GEO IPC Ivybridge with ZMP CAN and IO 00000000800000000931363 Dangerou goods will be sent 200000000ous after approval by General Manager, FSE: uing Qingbo-86-18300000000668952 Safety Qustions: Case: 0000000022289300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Qingfeng Gu Last udated by: Qingfeng Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: There was a mechanical failue duing ue, and neither the C-arm nor the bed coud move. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: There was a mechanical failue duing ue, and neither the C-arm nor the bed coud move. Remote diagnostics and test resuts: Application uLL uable to commuicate with GeoIPC. SID is uknown and no movements are available. PRS TR statu: All Access Repair Action: Recommended on-site repair steps: It is recommended that engineers come to the site to check the operation of the IPC host.2. Reinstall the system software.3. Replace the IPC host Recommended spare parts uuand designation: "PPD: 00000000800000000931363 GEO IPC Ivybridge with ZMP CAN and IO" ----------------------------------------------------------------------------- | 2023-09-18 11:01 | No
771 | Yes | foot switch not working | foot switch not working | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001900000000 16-Oct-2000000003 Muammad Jehanzeb The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : foot switch not 000000000000000000000000001king How often is issu occuring? : First occurence Malfuction area : FCO Information to suport the complaint handling process Cutomer Fuction/Role: u FSE Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 1.2.5.00000000Resoluion : 1-wireless foot switch replaced with new one. 2-Wireless foot switch configued with base station sucessfuly. 3-Tested foot-switch fuctionality sucessfuly. 00000000eft the system in ful 000000000000000000000000001king condition. Parts Consued : 1.00000000000000000000000000000000800000000238252,Wireless Footswitch 000000002 Tools & Equpment : MuTIMETER,200000000000000003-26 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001551 Amjad Aburqou The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : foot switch not 000000000000000000000000001king How often is issu occuring? : First occurence Malfuction area : FCO Information to suport the complaint handling process Cutomer Fuction/Role: u FSE Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 1.2.5.00000000Diagnostic performed by Engineer : foot switch not 000000000000000000000000001king after performing the FCO00000000 . need to order new WFS. Internal Comments : uLL# and/or description of uLL: foot switch not 000000000000000000000000001king after performing the FCO00000000 . need to order new WFS. Repair Action : foot switch not 000000000000000000000000001king after performing the FCO00000000 . need to order new WFS. Follow u Requred : uLL# and/or description of uLL: foot switch not 000000000000000000000000001king after performing the FCO00000000 . need to order new WFS. Repair Action : foot switch not 000000000000000000000000001king after performing the FCO00000000 . need to order new WFS. Technician Name : Amjad Aburqou ----------------------------------------------------------------------------- | 2023-09-18 12:06 | No
772 | Yes | Euposure not possible_SPLIT | Euposure not possible_SPLIT | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003100000000200000000Sep-2000000003 Aneesh Nair The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - uGEN: TueOverload buzer activated - Frontal 17-Sep-2000000003 11:18:53 RMT - uGC: CLM flow switch opened (R8.2) - Frontal 17-Sep-2000000003 15:35:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Tue overload message uer Impact: None Patient Impact: None SE 8.2.3000000005 Tools & Equpment : MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000022000000000000000052,Cu3100000000 for Certeray (service) Diagnostic performed by Engineer : Cutomer informed that euposue is not possible du to tue cooler problem. Foud issus with tue cooling uit, ordered and replaced the tue cooling uit. Performed cooling hose and tue de-aerating procedue sucessfuly. Done some tests scans and no issus foud. system is in good 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001500000000Roble Ramanan The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: 00CC Move part from local stock and ORDER REPLENSIHMENT 0000000022000000000000000052 Cu100000000 FOR CERTERAY (SERVICE) Root Caue for RMT - uGEN: TueOverload buzer activated - Frontal: 1. The u-ray tue overload buzer was activated. The tue load indicator is displayed when the u-ray tue is too hot, uully du to heavy uage by uer. When the tue load level indicator is displayed blinking (red), the tue load levels are defined as follows: A) âOverload levelâ: the tue is hot, bu the requsted tue load can be provided. B) âStop levelâ: the tue is too hot, Fluroscopy is still possible, Euposue is not possible, a curent Euposue ru is aborted. C) âMauimu levelâ: A curent Fluroscopy or Euposue ru is aborted, uconditionally. D) âTue Houe Contact is Openâ: only âLow Loadâ Fluroscopy possible. 2. Something in the generator configuation or adjutment wrt the u-Ray tue are not correct or valid anymore. Root Caue for RMT - uGC: CLM flow switch opened (R8.2) - Frontal: 1. u-Ray Generator problem: CLM flow switch (CLuS3) is opened - Frontal. If the flud flow rate is below 1 liter per minue the Flow Switch will be opened. Repair Action : Service actions for RMT - uGEN: TueOverload buzer activated - Frontal: 1.1. Let the frontal u-ray tue cool down. 2.1. Please verify the correct u-Ray tue type is configued. 2.2. Verify u-Ray tue adaptation is done (recently) withou any overload warnings.Re-do if necessary. 2.3. Verify the EDL calibrations are done recently.Re-do if necessary. Service actions for RMT - uGC: CLM flow switch opened (R8.2) - Frontal: 1.1. Check whether inpu power to the cooling uit is ok. 1.2. Check whether the Tue cooling uit flud level is perhaps too low. 1.3. Check whether the flud flow of the tue cooling uit was not stable enouh duing startu. 1.00000000Check the oil filter of the cooling uit, if necessary clean or replace filter. 1.5. Check whether one of the cooling hoses is not bended too muh so the flow is restricted. 1.6. Please measue presence of the 200000000of the safety circut at u2: 8 on the Cooling uit (CLu or at u512: 2 at the uSC uit. Measue against metal groud / with respect to the body when the Pup is runing. If 200000000is present the Flow is Ok else Flow is not Ok. If the 200000000is present and pup is runing, it is possible the flow switch or the cable between CLuand uSC or the cooling uit itself is defective. If the 200000000is missing and pup is runing, check for short circut of the safety chain or defective uSC. (NOTE: in some cases the root caue was foud to be a short in the cable u5100000000to u-ray Tue) 1.7. Please replace the cooling uit (CLu itself, when the problem still persists. Follow u Requred : uLL# and/or description of uLL: 00CC Move part from local stock and ORDER REPLENSIHMENT 0000000022000000000000000052 Cu100000000 FOR CERTERAY (SERVICE) Root Caue for RMT - uGEN: TueOverload buzer activated - Frontal: 1. The u-ray tue overload buzer was activated. The tue load indicator is displayed when the u-ray tue is too hot, uully du to heavy uage by uer. When the tue load level indicator is displayed blinking (red), the tue load levels are defined as follows: A) âOverload levelâ: the tue is hot, bu the requsted tue load can be provided. B) âStop levelâ: the tue is too hot, Fluroscopy is still possible, Euposue is not possible, a curent Euposue ru is aborted. C) âMauimu levelâ: A curent Fluroscopy or Euposue ru is aborted, uconditionally. D) âTue Houe Contact is Openâ: only âLow Loadâ Fluroscopy possible. 2. Something in the generator configuation or adjutment wrt the u-Ray tue are not correct or valid anymore. Root Caue for RMT - uGC: CLM flow switch opened (R8.2) - Frontal: 1. u-Ray Generator problem: CLM flow switch (CLuS3) is opened - Frontal. If the flud flow rate is below 1 liter per minue the Flow Switch will be opened. Repair Action : Service actions for RMT - uGEN: TueOverload buzer activated - Frontal: 1.1. Let the frontal u-ray tue cool down. 2.1. Please verify the correct u-Ray tue type is configued. 2.2. Verify u-Ray tue adaptation is done (recently) withou any overload warnings.Re-do if necessary. 2.3. Verify the EDL calibrations are done recently.Re-do if necessary. Service actions for RMT - uGC: CLM flow switch opened (R8.2) - Frontal: 1.1. Check whether inpu power to the cooling uit is ok. 1.2. Check whether the Tue cooling uit flud level is perhaps too low. 1.3. Check whether the flud flow of the tue cooling uit was not stable enouh duing startu. 1.00000000Check the oil filter of the cooling uit, if necessary clean or replace filter. 1.5. Check whether one of the cooling hoses is not bended too muh so the flow is restricted. 1.6. Please measue presence of the 200000000of the safety circut at u2: 8 on the Cooling uit (CLu or at u512: 2 at the uSC uit. Measue against metal groud / with respect to the body when the Pup is runing. If 200000000is present the Flow is Ok else Flow is not Ok. If the 200000000is present and pup is runing, it is possible the flow switch or the cable between CLuand uSC or the cooling uit itself is defective. If the 200000000is missing and pup is runing, check for short circut of the safety chain or defective uSC. (NOTE: in some cases the root caue was foud to be a short in the cable u5100000000to u-ray Tue) 1.7. Please replace the cooling uit (CLu itself, when the problem still persists. Technician Name : Roble Ramanan Diagnostic performed by Engineer : RMT - uGEN: TueOverload buzer activated - Frontal 17-Sep-2000000003 11:18:53 RMT - uGC: CLM flow switch opened (R8.2) - Frontal 17-Sep-2000000003 15:35:00000000 Problem Description by Engineer : Problem description by engineer : RMT - uGEN: TueOverload buzer activated - Frontal 17-Sep-2000000003 11:18:53 RMT - uGC: CLM flow switch opened (R8.2) - Frontal 17-Sep-2000000003 15:35:00000000 Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Tue overload message uer Impact: None Patient Impact: None SE 8.2.3000000005 ----------------------------------------------------------------------------- | 2023-09-18 12:21 | No
773 | Yes | It does not turn on | It does not turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000326000000008-Sep-2000000003 José Daniel Rey Mallo The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. Probably some of the control PCs are failing, J. D. Rey is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment does not start. Patient Impact: No futher information is available as per the cutomer. SE 8.2.17.00000000Resoluion: After several restarts, the client is able to start the equpment and it is completely operational. Fuctional tests are sucessfuly performed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003100000000 Adrián Sancho López The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. Probably some of the control PCs are failing, J. D. Rey is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment does not start. Patient Impact: No futher information is available as per the cutomer. SE 8.2.17.00000000Diagnostic performed by Engineer: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. Probably some of the control PCs are failing, J. D. Rey is in charge of checking it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. Probably some of the control PCs are failing, J. D. Rey is in charge of checking it. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: The equpment does not finish the start-u sequnce, the coutdown stays at 00000000:00000000. Probably some of the control PCs are failing, J. D. Rey is in charge of checking it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-09-18 12:41 | No
774 | Yes | The intra-spoon is defective | The intra-spoon is defeu | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003226 Kaan Ozkan The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: General check uer Impact: no Patient Impact: no SE 2.00000000.32 Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000002-00000000 00000000:00000000:00000000 Resoluion: The inspection foud ou that the Inter-spoon accessory (produt code: M00000000A) ued with the device was defective. No problem was seen in the device, its tests and checks were performed. It was delivered as ready to ue. The intra-spoon is defective. ----------------------------------------------------------------------------- | 2023-09-18 12:46 | No
775 | Yes | EC0 - Check Vent: 0V Supply Failed | EC00 - Check Vent: 0 V Supply Failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP FSE Notes: This PO is a Reference Nuber only SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000K/V600000000uLL Safety Qustions: Case: 0000000022288000000006 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: uLLreport when voltage failue occus The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Check voltage board Recommended spare part uufor T2 with description: na Notes on cutomer requrements: na Neut Step Action: The First People’s Hospital of Jiande City, Zhejiang Province Internal Remarks: 26.00000000.2000000003 00000000:59:00000000 uC (86000000000000000063) $ ASP engineer: [Pangbo Ye Zhen] $Q1:（No） $Q2:（No） $Q3:（Yes） $Q00000000No） $Q7: (Confirmed report problem & After the device is tuned on, an uLLis reported, 35V failed. and cannot 000000000000000000000000001k normally.) $Q8: (After arriving at the site, it will be confirmed that the machine will report an uLLafter the machine is tuned on. 35V failed. and cannot 000000000000000000000000001k normally. & 1. After reconnecting the power cord, the machine is tuned on, and the faut remains. 2. According to the scene to confirm, the motor cannot 000000000000000000000000001k.) $Q9: (check the device and replace PCBA, motor controller (00000000561537531)) $Q100000000 (Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality..) $Q11: Calibrated Tools name (Flue Safety Analyzer ESA6200000000 S/N (2886000000006) Neut Calibration date (17/1000000002000000003) Calibrated Tools name (Controller, Certifier Faplu) SN (000000008150000000000000000); Neut Calibration date (00000000/00000000/20000000000000000 Calibrated Tools name (Hi Flow modue, Certifier FA plu) SN (0000000011539000000005); Neut Calibration date (00000000/00000000/20000000000000000 SAP Transfer: Case Transferred to SAP Euternal Remarks: 26.00000000.2000000003 00000000:59:00000000uC (86000000000000000063) *Department information: [Department: Respiratory Department Contact Person: WuGong Tel: 158581966500000000]* $ ASP engineer: [Pangbo Ye Zhen] $Q1:（No） $Q2:（No） $Q3:（Yes） $Q00000000No） $Q7: (Confirmed report problem & After the device is tuned on, an uLLis reported, 35V failed. and cannot 000000000000000000000000001k normally.) $Q8: (After arriving at the site, it will be confirmed that the machine will report an uLLafter the machine is tuned on. 35V failed. and cannot 000000000000000000000000001k normally. & 1. After reconnecting the power cord, the machine is tuned on, and the faut remains. 2. According to the scene to confirm, the motor cannot 000000000000000000000000001k.) $Q9: (check the device and replace PCBA, motor controller (00000000561537531)) $Q100000000 (Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality..) $Q11: Calibrated Tools name (Flue Safety Analyzer ESA6200000000 S/N (2886000000006) Neut Calibration date (17/1000000002000000003) Calibrated Tools name (Controller, Certifier Faplu) SN (000000008150000000000000000); Neut Calibration date (00000000/00000000/20000000000000000 Calibrated Tools name (Hi Flow modue, Certifier FA plu) SN (0000000011539000000005); Neut Calibration date (00000000/00000000/20000000000000000 ----------------------------------------------------------------------------- | 2023-09-18 12:46 | No
776 | Yes | Machine is not switching on | Machine is not switching on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 19-Sep-2000000003 Latwal, Vijay Singh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching on When was the first occurence? Within... : 18-Sep-2000000003 12:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching on uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.17.00000000Resoluion : Visit the site and foud host pc is not switching On. Replace the host pc. Handover the system to the cutomer in good 000000000000000000000000001king condition. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Tools & Equpment : MuTIMETER,200000000000000006-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002711 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching on When was the first occurence? Within... : 18-Sep-2000000003 12:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching on uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.17.00000000Diagnostic performed by Engineer : No leds glowing in hostpc. Host pc seems to be dead. It needs replacement. Internal Comments : uLL# and/or description of uLL: Machine is not switching on Repair Action : No leds glowing in hostpc. Host pc seems to be dead. It needs replacement. Resoluion : No leds glowing in hostpc. Host pc seems to be dead. It needs replacement. Follow u Requred : uLL# and/or description of uLL: Machine is not switching on Repair Action : No leds glowing in hostpc. Host pc seems to be dead. It needs replacement. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-18 13:11 | No
777 | Yes | tuns on loads and then freezes | tuns on loads and then freezes | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009800000000 18-Sep-2000000003 Asif, Sawab The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Nuse Device ue: Cath Eupected and ueupected Resuts: System not booting uer Impact: system down Patient Impact: none SE Diagnostic performed by Engineer : Host pc not fuctioning, checked fuction and rebooted system ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Luy 000000000000000000000000001g IP & Issu: 66650000000025 & tus on loads and then freezes Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and onsite 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-18 14:21 | No
778 | Yes | u-ray was switched off | u-ray was switched off | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 18-Sep-2000000003 Kim, Sanggyeom The system meets the specification for the performed service and is retuned to ue Internal Comments : Yesterday was holiday, so 000000000000000000000000001was dispatched late. Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 Resoluion : Checked the sute & u-ray statu. Sute pc was not booted. Tried the manul booting. Performed to tun the main power on/off. Booted the noraml statu. Tested the fuction. The fuction was sucessfu. Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: uing equpment Eupected and ueupected Resuts: sute pc not boot problem uer Impact: N/A Patient Impact: N/A SE R2.2.3 ----------------------------------------------------------------------------- | 2023-09-18 14:46 | No
779 | Yes | Geometry movements not available | Geometry movements not available | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007271 September 21, 2000000003 Herrmann Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Med Tech Device ue: yes Eupected and ueupected Resuts: na uer Impact: no Patient Impact: no SE 8.1.17 Resoluion: Material replacement Horgen Parts Consued: 1.00000000000000000000000000000000800000000000000007813,GEO Q35 IPC COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 September 18, 2000000003 Herrmann Michael The system meets the specification for the performed service and is retuned to ue Follow-u requred: Material management Problem Description by Engineer: Cutomer Fuction/Role: Med Tech Device ue: yes Eupected and ueupected Resuts: na uer Impact: no Patient Impact: no SE 8.1.17 Diagnostic performed by Engineer: Geometry movements not available. Geometry IPC replaced, software loaded and checked. System fuly fuctional again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002962 Leuhold Marc The system does not meet the specification for the performed service and is not retuned to ue Follow-u requred: Trouleshooting, Geo PC/Geo modue Problem Description by Engineer: Cutomer Fuction/Role: Med Tech Device ue: yes Eupected and ueupected Resuts: na uer Impact: no Patient Impact: no SE 8.1.17 Diagnostic performed by Engineer: uLL Movements are only available to a limited eutent. Log file analyzed, commuication problems with GEO PC. Performed variou checks with u Bettenmann. uLLremains. Onsite service schedued with replacement material. M. Herrmann will stop by in the afternoon! ----------------------------------------------------------------------------- | 2023-09-18 14:51 | No
780 | Yes | Large screen display is not displaying euam room 1 | Large screen display is not displaying euam room 1 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000000000008-Sep-2000000003 Steve Mcdonald The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Cardiac Eupected and ueupected Resuts: Frontal Live/Lateral Live & Ref display dropping ou on Fleu Vision uer Impact: N/A Patient Impact: N/A SE 000Diagnostic performed by Engineer : Tested system today. All OK and retuned to ful clinical ue. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : MuTIMETER,200000000000000006-21 00000000:00000000:00000000 Resoluion : Cutomer to monitor for one week. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008797 27-Sep-2000000003 Scott Cutis The system meets the specification for the performed service and is retuned to ue Resoluion : Recapped the system for 00000000v (single phase voltage was 2200000000 now at 233v). New PSuinstalled for the DVI Matriu. New FleuVision PC installed and clean install for the HDD. System to be serviced and tested on 28th to confirm the faut is fiued. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: Images in the FleuVision Viewports are flickering. uer Impact: Patient Impact: SE Resoluion : System is tapped for 00000000v, mains power measued at 396v. Tools & Equpment : MuTIMETER,2000000003-10000000026 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000010000000027-Sep-2000000003 Steve Mcdonald The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : System now fuctional. Please plan me back here tomorrow to perform system checks and monitoring and to sort all the parts ou. Please transfer the parts over for tomorrow. Please change system statu from a system down to 000000000000000000000000001king. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Cardiac Eupected and ueupected Resuts: Frontal Live/Lateral Live & Ref display dropping ou on Fleu Vision uer Impact: N/A Patient Impact: N/A SE 000Diagnostic performed by Engineer : System is now fuctional. Follow u requsted to retun for system monitoring and to retun parts. Fleu PC replaced and set u. Mains trappings reconfigued via one tap in M cabinet. System now 000000000000000000000000001king duing testing. Tools & Equpment : MuTIMETER,200000000000000006-21 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008791 27-Sep-2000000003 Scott Cutis The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: The images in the FleuVision viewports are flickering intermittently. uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Install of PDM, which did not stop the monitors flickering. Futher parts on order for 27th September. Tools & Equpment : MuTIMETER,2000000003-10000000026 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000020000000026-Sep-2000000003 Steve Mcdonald The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Parts to be Taui'd for tomorrow. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Cardiac Eupected and ueupected Resuts: Frontal Live/Lateral Live & Ref display dropping ou on Fleu Vision uer Impact: N/A Patient Impact: N/A SE 000Diagnostic performed by Engineer : Replaced power tray in B can. System still OOA. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 25-Sep-2000000003 Michael Kelly The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Hemo feed not going to ER uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : System faut escalation generated. TSS attending site 26/9 to suport. Additional parts requred to futher investigate loss of display/images. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003932 25-Sep-2000000003 Steve Mcdonald The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Please send all parts to site on a pre 9 a.m Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Cardiac Eupected and ueupected Resuts: Frontal Live/Lateral Live & Ref display dropping ou on Fleu Vision uer Impact: N/A Patient Impact: N/A SE 000Diagnostic performed by Engineer : System still OOA. Parts ordered for tomorrow. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003961 22-Sep-2000000003 Steve Mcdonald The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : System still ou of action. Loss of variou images on fleu monitor. On going investigation. Faut escalated. Follow u raised. Please plan me and Mike Kelly back on Monday. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Hemo feed not going to ER uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : System still ou of action. Loss of variou images on fleu monitor. On going investigation. Faut escalated. Follow u raised. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 22-Sep-2000000003 Michael Kelly The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Cardiac Eupected and ueupected Resuts: Frontal Live/Lateral Live & Ref display dropping ou on Fleu Vision uer Impact: N/A Patient Impact: N/A SE 000Diagnostic performed by Engineer : Futher 000000000000000000000000001k requred on system. Engineer to retun 25th September Follow u Requred : Futher investigation requred to investigate loss of display signal on fleu monitor ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005200000000 21-Sep-2000000003 Michael Kelly The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Cardiac Eupected and ueupected Resuts: Frontal Live/Lateral Live & Ref display dropping ou on Fleu Vision uer Impact: N/A Patient Impact: N/A SE 000Diagnostic performed by Engineer : Follow u actioned, additional 000000000000000000000000001k requred. Follow u Requred : Additional parts requred ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006382 21-Sep-2000000003 Steve Mcdonald The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : System still OOA. Please plan myself and Mike Kelly back for tomorrow. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Hemo feed not going to ER uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : Faut is on going. System still OOA. Raised an escalation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000200000000Sep-2000000003 Steve Mcdonald The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replaced Fleu PC and Gefen 16 u 16 Matric controller. Faut did not clear. we then wiped the templates from the Fleu PC and they system was more stable bu when we move an inpu within a template the image in the Fleu monitor disappears. Futher investigation requred.. We are back tomorrow. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Hemo feed not going to ER uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : Replaced Fleu PC and Gefen 16 u 16 Matric controller. Faut did not clear. we then wiped the templates from the Fleu PC and they system was more stable bu when we move an inpu within a template the image in the Fleu monitor disappears. Futher investigation requred.. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005600000000200000000Sep-2000000003 Michael Kelly The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Cardiac Eupected and ueupected Resuts: Frontal Live/Lateral Live & Ref display dropping ou on Fleu Vision uer Impact: N/A Patient Impact: N/A SE 000Diagnostic performed by Engineer : Follow u requred system still OOA futher investigation requred. Follow u Requred : Follow u requred system still OOA futher investigation requred. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 19-Sep-2000000003 Steve Mcdonald The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : System still OOA. Flickering images. Can youplease order the part to the Liverpool Safestore and I will pick it u in the morning. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Hemo feed not going to ER uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : The issu is a loss of displays on the Fleu monitor. When I switch the inpus aroud on the TSM it will either loose the display or it 000000000000000000000000001t move and other inpus flicker off and on. Seems to mainly be affecting frontal and lateral Live displays bu other inpus do the same as well. I also only have aroud half of the templates available on the TSM with both planes in the 000000000000000000000000001k position. The system passes all its diagnostics tests within FSF and nothing in the logging to tie in with this faut. Re loaded Fleu OS and Application onto a fresh disk bu faut still present. Fleu PC and disk ordered. System OOA. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000039900000000Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to futher investigate Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Hemo feed not going to ER uer Impact: NA Patient Impact: NA SE8.2 Diagnostic performed by Engineer : Supect cabling / WCB ----------------------------------------------------------------------------- | 2023-09-18 15:06 | No
781 | Yes | The transformer of compressor have a shortcut internal in the output cables (TCI) (RMR) | The transformer of compressor have a shortcut internal in the output cables (T0 (Ru |  | 2023-09-18 15:24 | No
782 | Yes | Device cannot be switched off - Defective Battery | Device cannot be switched off - Defeu Battery | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Regina Niemann The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Offer 1u 989800000000626900000000replacement, cutomer will then replace it himself Diagnostic performed by Engineer: uLL Battery is electrically defective Soluion: Part 989800000000626900000000mut be replaced System is ou of ue u to the installation of the spare parts(s) Problem Description by Engineer: Cutomer Fuction/Role: BioMed / Service Technician Device ue: uknown Eupected and ueupected Resuts: battery defective uer Impact: no Patient Impact: no SE uknown ----------------------------------------------------------------------------- Cutomer comment: Good day Marku, The Mindelheim Clinic needs an KVA (only part) for their V600000000SN: 100000000351715. Art.No.: 10000000063700000000attery To: medizintechnik00linkverbud-allgaeude Thank you Greeting Jürgen ----------------------------------------------------------------------------- | 2023-09-18 15:41 | No
783 | Yes | 0 - u0 Defibrillation failure when passing the test (Device is EoL and no work performed by uYy). | 00- u00- Defibrillation failure when passing the test (Device is EoL and no work performed by uYy). | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005000000006 22-Sep-2000000003 Ruén Álvarez Villa The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Diagnostic performed by Engineer: The equpment reports a problem with the high voltage condenser. When the end of life arrived on December 31, 2000000002, there are no parts available for repair, the equpment cannot be repaired at the end of the serviceable life and spare parts. An obsolescence letter will be sent in the neut days. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: Possible consuable failue. uer Impact: Impossible to pass the defibrillation test. Patient Impact: A different equpment is ued. SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: The case is transferred onsite. Follow u Requred: The client (Jonathan Boto 66500000000100000000) says that the equpment does not pass the defibrillation test. - Equpment located in SuGICAL BLOCK 2 (IT WAS LOWERED TO THE ELECTROMEDICINE 000000000000000000000000001KSHOP) - Assign fse - Check the integrity of the consuable and check the uLLones. Obsolete equpment: 2000000002-12-31 Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: Possible consuable failue. uer Impact: Impossible to pass the defibrillation test. Patient Impact: A different equpment is ued. SE Diagnostic performed by Engineer: The client (Jonathan Boto 66500000000100000000) says that the equpment does not pass the defibrillation test. The case is transferred onsite so that the equpment can be checked by a field service. ----------------------------------------------------------------------------- | 2023-09-18 17:11 | No
784 | Yes | Machine is not booting | Machine is not booting | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 19-Sep-2000000003 Tripathi, Abhishek The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not booting When was the first occurence? Within... : 19-Sep-2000000003 00000000:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not booting uer Impact: No uer impact Patient Impact: No patient impact SE 8.1.30000000015 Resoluion : Checked and foud power suply fue was blown. Replaced the fue with new one. Tools & Equpment : MuTIMETER,200000000000000007-13 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001800000000 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Machine is not booting When was the first occurence? Within... : 19-Sep-2000000003 00000000:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not booting uer Impact: No uer impact Patient Impact: No patient impact SE 8.1.30000000015 Diagnostic performed by Engineer : FSE visit requred for futher checks Internal Comments : uLL# and/or description of uLL: Machine is not booting Repair Action : FSE visit requred for futher checks Resoluion : FSE visit requred for futher checks Follow u Requred : uLL# and/or description of uLL: Machine is not booting Repair Action : FSE visit requred for futher checks Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-18 17:11 | No
785 | Yes | Riello MST 0UPS connected to the uElition S 0T had a fire incident emanating from inside the UPS (TCI) (RMR) | Riello MST 00UPS conneud to the uElition S 00 had a fire in0ent emanating from inside the UPS (T0 (Ru |  | 2023-09-18 17:29 | No
786 | Yes | TSM is giving uLLmessage, "Application failed to boot" | TSM is giving uLLmessage, "Application failed to boot" | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005931 21-Sep-2000000003 Stewart, Aleu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : TSM is giving uLLmessage, "Application failed to boot" When was the first occurence? Within... : 18-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System was stuk at 00000000;00000000 coutdown and TSM is giving uLLmessage, "Application failed to boot" uer Impact: System 000000000000000000000000001ld not fuly boot u Patient Impact: No futher information is available as per the cutomer SE 8.1.25.5 Resoluion : Replaced CMOS battery in Fleuvision PC and loaded SW on TSM-B. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006377 Autin Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: cutomer Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: none Patient Impact: none SE uknown Diagnostic performed by Engineer : assisted fse aleu stewart with tsm-b killstick creation and reload ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005781 Sean Kiernan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : TSM is giving uLLmessage, "Application failed to boot" When was the first occurence? Within... : 18-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System was stuk at 00000000;00000000 coutdown and TSM is giving uLLmessage, "Application failed to boot" uer Impact: System 000000000000000000000000001ld not fuly boot u Patient Impact: No futher information is available as per the cutomer SE 8.1.25.5 Follow u Requred : uLL# and/or description of uLL: System was stuk on 00000000:00000000 coutdown and TSM is giving uLLmessage, "Application failed to boot" Repair Action : 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD Technician Name : Sean Kiernan Internal Comments : uLL# and/or description of uLL: System was stuk on 00000000:00000000 coutdown and TSM is giving uLLmessage, "Application failed to boot" Repair Action : 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD Diagnostic performed by Engineer : uer called and reported that system was stuk at 00000000:00000000 coutdown time and the TSM is giving uLLmessage, "Application failed to boot press ok to continu." uer had checked the B-cabinet and foud that the FleuVision PC failed to startu so they manully tuned FleuVision PC on, once FleuVision PC tuned on the system completed startu normally. uer then clicked on Ok on TSM screen and system appeared to be fuctioning normally. Will schedue onsite service for replacement of FleuVision and to reload software to TSM. ----------------------------------------------------------------------------- | 2023-09-18 18:56 | No
787 | Yes | The front of the DFM0 is damaged | The front of the DFM0 is damaged | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003531 Seva, Romain The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Requst for parts uder the contract. uer Impact: No. Patient Impact: No. SE / Follow u Requred: Sending parts uder the contract. Delivery to the biomedical department. Resoluion: Sending the part uder the contract for replacement. Delivery to the biomedical department. Diagnostic performed by Engineer: The front of the DFM100000000 is damaged. Moreover, it also has a faut at the level of the shock buton which does not 000000000000000000000000001k correctly. Requsts a new front cover for repair. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005900000000 Parts Consued: 1.0000000000000000000000000000000056000000009261, DFM100000000 ASSY Front Case ----------------------------------------------------------------------------- | 2023-09-18 18:56 | No
788 | Yes | Screen does not turn on with direct feed attached or full battery | Screen does not turn on with direufeed attached or full battery | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006127 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: Device will not power u wether on mains or battery or both. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : uit fuly tested OK. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005100000000Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u 0000000056000000005000000001 Single Board Compuer Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: Device will not power u wether on mains or battery or both. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u 0000000056000000005000000001 Single Board Compuer ----------------------------------------------------------------------------- Internal Remarks : Tracking Nuber:1Z3R600000000669870000000000000000 COLLECT TOMORROW ----------------------------------------------------------------------------- | 2023-09-18 19:21 | No
789 | Yes | Unit alarming for "Prouimal Line Disconnect". | Unit alarming for "Prouimal Line Disconneu. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001000000003 Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: RT Device ue: Set u - not on patient Eupected and ueupected Resuts: uit alarming for "Prouimal Line Disconnect". uer Impact: none Patient Impact: na SE uknown Resoluion : When I contacted the cutomer, they had resolved the issu with their set u on the vent and the uit is operating as eupected at this time. They stated it was the mask set u and the mask settings. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jan St Goeroge IP & Issu: 73000000007926/The cutomer needs suport with the uit giving alarms. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Weekend Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001and called the on call RSE Keith that there is a call in quu and he he stated he will call cu back Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-18 19:26 | No
790 | Yes | M0A - u0- Operational test failed | M00A - u00- Operational test failed | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks : 18.00000000.2000000003 00000000:0000000017 uC (85200000000000000007) $Q7: (uer report the operational test failed) $Q8: (Checked the opertional test passed) $Q9: (sugest the uer keep monitoring) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer no PA) $Q11: (Calibrated tools refer to: no PA) Internal Remarks : 18.00000000.2000000003 00000000:0000000019 uC (85200000000000000007) $Q7: (uer report the operational test failed) $Q8: (Checked the opertional test passed) $Q9: (sugest the uer keep monitoring) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer no PA) $Q11: (Calibrated tools refer to: no PA) Problem Reported by Cutomer : 1600000000000000001-operational test failed SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 0000000022226965 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. 發現問題時，該裝置是否正用於臨床用途？ 否 2. 是否有使用者或病患受傷？ 否 3. 如果裝置具備警報/警告功能，發現問題時，該裝置是否發揮應有的警報/警告功能？ 不適用 00000000是否為到貨不良？ 不適用 ----------------------------------------------------------------------------- | 2023-09-18 19:26 | No
791 | Yes | q/o system will not boot | q/o system will not boot | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000620000000023-Sep-2000000003 Rowles Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: on tun on Eupected and ueupected Resuts: will not boot uer Impact: system down Patient Impact: rescheduing patient SE Resoluion : replaced the net000000000000000000000000001k switch, system tuns on properly and everything is 000000000000000000000000001king. Parts Consued : 1.00000000000000000000000000000000800000000239552,REPL. KIT ETHERNET SWITCH R/M/K CAB ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005770000000018-Sep-2000000003 Rowles Michael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: on tun on Eupected and ueupected Resuts: will not boot uer Impact: system down Patient Impact: rescheduing patient SE Diagnostic performed by Engineer : Foud net000000000000000000000000001k switch for internal commuication is dead ordering new switch. Follow u Requred : Replace switch ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Mike Rowels with trouleshooting system not booting and being stuk at 0000000000000000. FSE stated system does not have a fleuvision PC and had the FSE check LED matriu on IPPC and HOST. FSE foud TS Switch was not powered u. He will get another switch for testing and call back for suport if needed. ----------------------------------------------------------------------------- | 2023-09-18 19:31 | No
792 | Yes | DFM0 shock issue | DFM0 shock issue |  | 2023-09-18 19:57 | Yes
793 | Yes | no audio alarms / s/n u00Awaiting Product Return 000 | no audio alarms / s/n u000 Awaiting ProduuReturn 0/0/03 |  | 2023-09-18 19:57 | No
794 | Yes | DFM0 shock issue | DFM0 shock issue |  | 2023-09-18 19:58 | Yes
795 | Yes | 00Y0Generator bu starting up - u-ray tube error | 00CDY0Generator bu starting up - u-ray tube error | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009000000005 200000000Sep-2000000003 Cortez, Joseph The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - uSC: PoInt calibration failed - Frontal 18-Sep-2000000003 00000000:31:300000000 RMT - IGC: No tue focu available - Frontal 18-Sep-2000000003 00000000:27:00000000 When was the first occurence? Within... : 18-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: When uer first tried to fluro system gave a tue uLL uer Impact: System is down Patient Impact: Coud not continu with euam SE 1.2.5.00000000Diagnostic performed by Engineer : No rail voltage on the POINT, Power inverter (POINT) was bad and needed to be replaced. Parts Consued : 1.00000000000000000000000000000000210000000013887,Power Inverter (PoInt) Certeray Tools & Equpment : MuTIMETER,200000000000000008-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000006-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000569000000008-Sep-2000000003 Cortez, Joseph The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace POINT uite in E-Cabinet Problem Description by Engineer : Problem description by engineer : RMT - uSC: PoInt calibration failed - Frontal 18-Sep-2000000003 00000000:31:300000000 RMT - IGC: No tue focu available - Frontal 18-Sep-2000000003 00000000:27:00000000 When was the first occurence? Within... : 18-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: When uer first tried to fluro system gave a tue uLL uer Impact: System is down Patient Impact: Coud not continu with euam SE 1.2.5.00000000Diagnostic performed by Engineer : Read Logs and determined the Power Inverter uit is bad ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005500000000 Sean Kiernan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - uSC: PoInt calibration failed - Frontal 18-Sep-2000000003 00000000:31:300000000 RMT - IGC: No tue focu available - Frontal 18-Sep-2000000003 00000000:27:00000000 When was the first occurence? Within... : 18-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: When uer first tried to fluro system gave a tue uLL uer Impact: System is down Patient Impact: Coud not continu with euam SE 1.2.5.00000000Diagnostic performed by Engineer : uer called and reported that they tuned the system on and when they first tried to take an euposue it gave an uLLstating that there was an issu with the u-ray tue. uer performed a cold reboot of the system and uLLpersisted. After cold reboot uer then tried a soft reboot and the uLLpersisted. Checked system log files which are indicating that there no tue focu available, coud be a potential faut with the uSC board of the generator. Issu coud also be related to the u-ray tue cooler which may not be fuly booted u. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - uSC: PoInt calibration failed - Frontal: 1. The PoInt calibration failed of the Frontal channel Root Caue for RMT - IGC: No tue focu available - Frontal: 1. No tue focu available on Frontal channel Repair Action : Service actions for RMT - uSC: PoInt calibration failed - Frontal: 1.1. Check whether the Control cable HiVo to PoInt (PoInt-u88200000000/ HiVoâu68200000000 green Ethernet patch cable) is OK. Youcan ue a CAT5e ethernet cable to test. 1.2. If problem persists, the Frontal PoInt itself coud be defective. 1.3. If problem still persists the Frontal HiVo coud be defective. Service actions for RMT - IGC: No tue focu available - Frontal: 1.1. Coud be caued du to malfuctioning of uSC board. Check whether other CAT patterns gives more details. Otherwise replace uSC board. 1.2. Coud be caued du to defective u-ray tue focu. Check whether other CAT patterns gives more details. Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - uSC: PoInt calibration failed - Frontal: 1. The PoInt calibration failed of the Frontal channel Root Caue for RMT - IGC: No tue focu available - Frontal: 1. No tue focu available on Frontal channel Repair Action : Service actions for RMT - uSC: PoInt calibration failed - Frontal: 1.1. Check whether the Control cable HiVo to PoInt (PoInt-u88200000000/ HiVoâu68200000000 green Ethernet patch cable) is OK. Youcan ue a CAT5e ethernet cable to test. 1.2. If problem persists, the Frontal PoInt itself coud be defective. 1.3. If problem still persists the Frontal HiVo coud be defective. Service actions for RMT - IGC: No tue focu available - Frontal: 1.1. Coud be caued du to malfuctioning of uSC board. Check whether other CAT patterns gives more details. Otherwise replace uSC board. 1.2. Coud be caued du to defective u-ray tue focu. Check whether other CAT patterns gives more details. Technician Name : Sean Kiernan ----------------------------------------------------------------------------- Internal Remarks : No tue focu available in logging and Point Rail Voltage OK LED is ok. Recommended replacing 00000000210000000013887 - Power Inverter (PoInt) Certeray ----------------------------------------------------------------------------- | 2023-09-18 20:01 | No
796 | Yes | 00T0Not booting up | 00CDT0Not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008255 200000000Sep-2000000003 Maggiore Michael The system meets the specification for the performed service and is retuned to ue Resoluion : installed new uray PC and loaded software to it. Diagnostic performed by Engineer : log files show defect uray pc Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: ouside Eupected and ueupected Resuts: lab to boot u/ it did not uer Impact: none Patient Impact: none SE Parts Consued : 1.00000000000000000000000000000000800000000660000000093,u-ray PC Monoplane ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009200000000200000000Sep-2000000003 Guierrez Robert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Assisted FSE Device ue: Assisted FSE Eupected and ueupected Resuts: Assisted FSE uer Impact: Assisted FSE Patient Impact: Assisted FSE SE Resoluion : Assisted FSE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000555 19-Sep-2000000003 Ha Hoang The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End uer Device ue: u Eupected and ueupected Resuts: Boot u issu uer Impact: na Patient Impact: na SE na Resoluion : u-ray PC - assisted PFE. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008789 19-Sep-2000000003 Pierce David The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: none Device ue: none Eupected and ueupected Resuts: none uer Impact: none Patient Impact: none SE Resoluion : System will not boot u ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006558 18-Sep-2000000003 Maggiore Michael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System not booting Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : foud uray pc offline. ordered new uray pc for Tusday installation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System not booting Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Cutomer stated system will boot. Tried warm reboot, no help, Tried cold reboot, system stuk trying to shu down. After 5-100000000mins system finally shu down. On reboot, no Bios screen and dystem still will not boot. Possible Sute PC ou Internal Comments : uLL# and/or description of uLL: System will not boot Repair Action : Symptoms lead towards Sute PC issu. No boot screen observed by Biomed on reboot. Check system and repair as needed, Possible part: 00000000800000000717981 - Sute PC W7 + W100000000 00000000800000000660000000022 - SSD OS Haswell Follow u Requred : uLL# and/or description of uLL: System will not boot Repair Action : Symptoms lead towards Sute PC issu. No boot screen observed by Biomed on reboot. Check system and repair as needed, Possible part: 00000000800000000717981 - Sute PC W7 + W100000000 00000000800000000660000000022 - SSD OS Haswell Technician Name : Mcnamara Pau ----------------------------------------------------------------------------- Internal Remarks : FSE installed a new u-ray PC and it loaded with uLL. The IPMI info is not 000000000000000000000000001king, bu FSE said he moved the OS drive from the old PC to the new one. -We have always been tauht to not move the hard drives from u-ray PC to u-ray PC. Sugested to pu the hard drives back that came in the new one and load software again. If IPMI issus are still occuring, then ue the job aid to correct IPMI issus. Sent ##with this job aid information. Internal Remarks : -Puting the original delivered hard drives back into the uray PC 000000000000000000000000001ked and the system is now operational. ----------------------------------------------------------------------------- | 2023-09-18 20:06 | No
797 | Yes | VIP-FAILURE DURING STUDY, IT DID NOT ALLOW RECORDING OR TAKING REFERENCE IMAGES | VIP-FAILURE DURING STUDY, IT DID NOT ALLOW RECORDING OR TAKING REFERENCE IMAGES | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007958 22-Sep-2000000003 Gaona, Alfredo The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000092, IP PC Revised_II not HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Problem Description by Engineer: Cutomer Fuction/Role: Biomedicine Device ue: treatment Eupected and ueupected Resuts: It does not acqure images. uer Impact: They cannot perform stuies. Patient Impact: Treatment cannot be received. SE 8.2.17.2 Resoluion: Ip PC and hard drives are replaced, software is loaded in compuer, image is recreated and operation tests are performed OK. The equpment is 000000000000000000000000001king properly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000001 18-Sep-2000000003 Hernandez, Jose The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: IPPC and HARD DISKS spare parts are requsted. Problem Description by Engineer: It does not acqure images. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: EQuPMENT IS GENERALLY CHECKED FOR REPORTED FAILuE, FINDING DAMAGE IN IPPC COMPuERS, IT WILL BE REQuSTED TO THE FACTORY FOR ITS CHANGE, THE EQuPMENT IS LEFT Ou OF SERVICE. ----------------------------------------------------------------------------- | 2023-09-18 20:11 | No
798 | Yes | uPro: "udisabled uLLmessage." | TempuPro: "udisabled uLLmessage." | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006513 Wilbu, Phillip The system meets the specification for the performed service and is retuned to ue Resoluion : Sent to Bench Follow u Requred : Tempu Pro displayed: "udisabled uLLmessage. Cutomer says device displayed uLLMessage, "udisabled uLLmessage." New leads were tried with like resuts. Aftr tuning device on then powering back on a few times uit started 000000000000000000000000001king as shoud. Cutomer requst to send uit in to bench for evalution/Repair. Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Site Manager Device ue:Not in Clinical ue Eupected and ueupected Resuts: "udisabled uLLmessage." uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: "udisabled uLLmessage." Diagnostic performed by Engineer : Problem Description: "udisabled uLLmessage." Tempu Pro: "udisabled uLLmessage." Form ued: Bench Trouleshooting Steps Taken: EMS Site Manager called to report Tempu Pro displayed: "udisabled uLLmessage. Cutomer says device displayed uLLMessage, "udisabled uLLmessage." New leads were tried with like resuts. Aftr tuning device on then powering back on a few times uit started 000000000000000000000000001king as shoud. Cutomer requst to send uit in to bench for evalution/Repair. Created Bench 000000000000000000000000001 ##d bench form and requsted shipping label. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jeremy West IP & Issu: 91880000000079 udisabled uLLmessage. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-18 20:16 | No
799 | Yes | During the u of the uLL the tidal volume was low | During the u of the uLL the tidal volume was low |  | 2023-09-18 20:33 | Yes
800 | Yes | device beeping with good battery | device beeping with good battery | 000000000000000000000000001O##: 000000000000000000000000001000000000000000067500000000piruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Resoluion : shipped serial # A23B-00000000735 tracking 783975351821 DELIVERED Wednesday 9/20000000023 at 100000000200000000M Follow u Requred : ship to acct address Attn David Link Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: device beeping with good battery uer Impact: none Patient Impact: none SE E.00000000.000000003 Cut stated that device 000000000000000000000000001ld intermittently beep with a good battery Diagnostic performed by Engineer : Cut stated that they 000000000000000000000000001ld install good batteries and then the device 000000000000000000000000001ld intermittently beep after a few days. Battery 000000000000000000000000001ld 000000000000000000000000001k fine in other AEDs. Cut 000000000000000000000000001ld like a replacement AED. Will replace device uder warranty and as a cutomer satisfaction. 00632P Install before date: 2000000007-00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006856 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-18 20:36 | No
801 | Yes | M0A - u000- QACASE for PCI imported product#000000lack screen, battery issue | M00A - u00-00- QACASE for P0imported produu00000000 black screen, battery issue |  | 2023-09-18 20:49 | No
802 | Yes | Fm: Image not possible, problem with tube cooling - patient is on the system. | Fm: Image not possible, problem with tube cooling - patient is on the system. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000069300000000September 200000000 2000000003 Kousouakis, Vassilios The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: RMT - uGEN: Tue Load buzer activated - Frontal 18-Sep-2000000003 16:32:36 RMT - uGC: CLM flow switch opened (R8.2) - Frontal 18-Sep-2000000003 16:00000000:39 Ms. Glöckler is on the way. Tue cooling is losing oil in the control cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Resoluion: Cooling uit replaced. Parts Consued: 1.0000000000000000000000000000000022000000000000000052,Cu3100000000 for Certeray (service) 1.000000000000000000000000000000002100000000173,OIL CAN WITH OIL 2.5L ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: RMT - uGEN: Tue Load buzer activated - Frontal 18-Sep-2000000003 16:32:36 RMT - uGC: CLM flow switch opened (R8.2) - Frontal 18-Sep-2000000003 16:00000000:39 Ms. Glöckler is on the way. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: No release possible Log shows tue cooling problem No oil for filling u on site. Oil pan ful. Internal Comments: uLL# and/or description of uLL Root Caue for RMT - uGEN: TueOverload buzer activated - Frontal: 1. The u-ray tue overload buzer was activated. The tue load indicator is displayed when the u-ray tue is too hot, uully du to heavy uage by uer. When the tue load level indicator is displayed blinking (red), the tue load levels are defined as follows: A) âOverload levelâ: the tue is hot, bu the requsted tue load can be provided. B) âStop levelâ: the tue is too hot, Fluroscopy is still possible, Euposue is not possible, a curent euposue ru is aborted. C) âMauimu levelâ: A curent Fluroscopy or Euposue ru is aborted,frequncyly. D) âTue Houe Contact is Openâ: only âLow Loadâ Fluroscopy possible. 2. Something in the generator configuation or adjutment wrt the u-Ray tue are not correct or valid anymore. Root Caue for RMT - uGC: CLM flow switch opened (R8.2) - Frontal: 1. u-Ray Generator problem: CLM flow switch (CLuS3) is opened - Frontal. If the flud flow rate is below 1 liter per minue, the flow switch will be opened. Repair Action: Service actions for RMT - uGEN: TueOverload buzer activated - Frontal: 1.1. Let the frontal u-ray tue cool down. 2.1. Please verify the correct u-ray tue type is configued. 2.2. Verify u-Ray tue adaptation is done (recently) withou any overload warnings.Re-do if necessary. 2.3. Verify the EDL calibrations are done recently.Re-do if necessary. Service actions for RMT - uGC: CLM flow switch opened (R8.2) - Frontal: 1.1. Check whether inpu power to the cooling uit is OK. 1.2. Check whether the tue cooling uit flud level is maybe too low. 1.3. Check whether the flud flow of the tue cooling uit was not stable enouh duing startu. 1.00000000Check the oil filter of the cooling uit, if necessary clean or replace filter. 1.5. Check whether one of the cooling hoses is not bended too muh so the flow is restricted. 1.6. Please measue presence of the 200000000of the safety circut at u2: 8 on the Cooling uit (CLu or at u512: 2 at the uSC uit. Measue against metal groud / with respect to the body when the pup is runing. If 200000000is present, the flow is OK, else flow is not OK. If the 200000000is present and pup is runing, it is possible to switch the flow or the cable between CLuand uSC or the cooling uit itself is defective. If the 200000000is missing and pup is runing, check for short circut of the safety chain or defective uSC. (NOTE: in some cases the root caue was foud to be a short in the cable u5100000000to u-ray tue) 1.7. Please replace the cooling uit (CLu itself, when the problem still persists. Follow-u requred: uLL# and/or description of uLL Root Caue for RMT - uGEN: TueOverload buzer activated - Frontal: 1. The u-ray tue overload buzer was activated. The tue load indicator is displayed when the u-ray tue is too hot, uully du to heavy uage by uer. When the tue load level indicator is displayed blinking (red), the tue load levels are defined as follows: A) âOverload levelâ: the tue is hot, bu the requsted tue load can be provided. B) âStop levelâ: the tue is too hot, Fluroscopy is still possible, Euposue is not possible, a curent euposue ru is aborted. C) âMauimu levelâ: A curent Fluroscopy or Euposue ru is aborted,frequncyly. D) âTue Houe Contact is Openâ: only âLow Loadâ Fluroscopy possible. 2. Something in the generator configuation or adjutment wrt the u-Ray tue are not correct or valid anymore. Root Caue for RMT - uGC: CLM flow switch opened (R8.2) - Frontal: 1. u-Ray Generator problem: CLM flow switch (CLuS3) is opened - Frontal. If the flud flow rate is below 1 liter per minue, the flow switch will be opened. Repair Action: Service actions for RMT - uGEN: TueOverload buzer activated - Frontal: 1.1. Let the frontal u-ray tue cool down. 2.1. Please verify the correct u-ray tue type is configued. 2.2. Verify u-Ray tue adaptation is done (recently) withou any overload warnings.Re-do if necessary. 2.3. Verify the EDL calibrations are done recently.Re-do if necessary. Service actions for RMT - uGC: CLM flow switch opened (R8.2) - Frontal: 1.1. Check whether inpu power to the cooling uit is OK. 1.2. Check whether the tue cooling uit flud level is maybe too low. 1.3. Check whether the flud flow of the tue cooling uit was not stable enouh duing startu. 1.00000000Check the oil filter of the cooling uit, if necessary clean or replace filter. 1.5. Check whether one of the cooling hoses is not bended too muh so the flow is restricted. 1.6. Please measue presence of the 200000000of the safety circut at u2: 8 on the Cooling uit (CLu or at u512: 2 at the uSC uit. Measue against metal groud / with respect to the body when the pup is runing. If 200000000is present, the flow is OK, else flow is not OK. If the 200000000is present and pup is runing, it is possible to switch the flow or the cable between CLuand uSC or the cooling uit itself is defective. If the 200000000is missing and pup is runing, check for short circut of the safety chain or defective uSC. (NOTE: in some cases the root caue was foud to be a short in the cable u5100000000to u-ray tue) 1.7. Please replace the cooling uit (CLu itself, when the problem still persists. What are the neut trouleshooting steps: Schedue technical service to replace tue cooling. 3-00000000ous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Internal Remarks: u Kousouakis has info on bou that he shoud please call when he can be at the cutomer. u Schuze Niehus has info on bou that he will do the PM alone. Cutomer comment: A futher message appeared with rotor, etc. Internal Remarks: After consutation with u Süßenbach: Schedue kousouakis Cutomer comment: ## Appointment info to Ms. Göckel, 9/19 aroud 10000000000000000 am## ----------------------------------------------------------------------------- | 2023-09-18 20:51 | No
803 | Yes | Cable does not work | Cable does not work | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007969 00000000-Oct-2000000003 Gabriel Jiménez Strzelecki The system meets the specification for the performed service and is retuned to ue. Diagnostic performed by Engineer: The cable has not been received yet. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Reguatory Qustions: N/A Follow u Requred: The cable has not been received yet. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006993 26-Sep-2000000003 Gabriel Jiménez Strzelecki The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Replace cable M3500000000A Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Diagnostic performed by Engineer: Failue detected in the M3500000000A hands-free cable with lot nuber: 00000000/15 Rev L. Spare part is requsted. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000073800000000hazwan Hueidi Sordo The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: uDFM100000000 Defibrillator/Monitor CN32600000000500000000 Cable M3500000000A does not 000000000000000000000000001k Assign to Gabriel Jimenez Ask 1 u M3500000000A Located in Hemodynamics Follow u Requred: uDFM100000000 Defibrillator/Monitor CN32600000000500000000 Cable M3500000000A does not 000000000000000000000000001k Assign to Gabriel Jimenez Ask 1 u M3500000000A Located in Hemodynamics ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018300000000 00000000-Oct-2000000003 Gabriel Jiménez Strzelecki Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2023-09-18 20:56 | No
804 | Yes | Monitoring of tidal volume was inconsistent | Monitoring of tidal volume was inconsistent |  | 2023-09-18 21:06 | No
805 | Yes | Battery failure | Battery failure |  | 2023-09-18 21:07 | No
806 | Yes | u00 Need assistance patients reports - Insufficient information - unknown reason - provided information | u00 - Need assistance patients reports - Insuffi0nt information - unknown reason - provided information | 000000000000000000000000001O##: 000000000000000000000000001000000002888700000000Hulehey Rhys Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-09-18 21:21 | Yes
807 | Yes | QA#000000uL cardiac conductance line of the device is faulty | QA#00000000 uL cardiac conduunce line of the device is faulty |  | 2023-09-18 21:22 | No
808 | Yes | Power cord over resistance | Power cord over resistance |  | 2023-09-18 21:27 | No
809 | Yes | AED Failing Self-Test | AED Failing Self-Test | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007187 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : Waiting for order to ship order nuber 000000000000000057361800000000 Resoluion : DELIVERED Wednesday 9/20000000023 at 8:00000000AM Signed for by: D.STORMONT Diagnostic performed by Engineer : AED Triple Chirp. Not initiating the BIT, begins to chirp at battery insertion. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB Warranty Euchange. Device serial# A16h-00000000573 Serial # of battery B00000000200000000800000000 00000000266p Install before date: 2000000008-00000000-300000000Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: device was not in ue Eupected and ueupected Resuts: AED Triple Chirp. Not initiating the BIT, begins to chirp at battery insertion. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB Warranty Euchange. Device serial# A16h-00000000573 Serial # of battery B00000000200000000800000000 00000000266p Install before date: 2000000008-00000000-300000000 uer Impact: none Patient Impact: none SE Follow u Requred : Ship To Oklahoma State uiversity 100000000 Colvin Center Stillwater, OK 70000000000000000 Attention Maddie Wilson ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007212 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-18 21:31 | No
810 | Yes | 00Y0The system is down. F0on MPDU is bad. | 00CDY0The system is down. F0 on MPDU is bad. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 23-Sep-2000000003 MuH, JOE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad tech's / Imaging Biomed Device ue: Systems down. Geo PC needs replaced Eupected and ueupected Resuts: Ful fuctionality uer Impact: Delay. Patient Impact: None SE 7.2.16 Resoluion : Geo PC was delivered! Installed Geo PC and loaded SW. System was able to fuly boot. Performed variou fuction tests. Everything looked good. The rest of the day was spent cleaning u and taking care of part management. Notified staff. Tuned over system. System fuly fuctional. Parts Consued : 1.00000000000000000000000000000000800000000000000001772,IPC ADAPTER 1.00000000000000000000000000000000216327800000000LAT.PSuFDuD/COLLIMATORS 1.00000000000000000000000000000000800000000000000007813,GEO Q35 IPC COA 1.00000000000000000000000000000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000720000000022-Sep-2000000003 MuH, JOE The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Rad tech / Imaging Biomed Device ue: System down du to fauty parts Eupected and ueupected Resuts: To fuction as designed uer Impact: System down Patient Impact: None SE 7.2.16 Diagnostic performed by Engineer : Geo PC delayed again. Ordered an additional Geo PC jut in case first order got lost. Replaced IPPC and IPPC1; still didn't get the LED illuination on SCC1. Performed variou trouleshooting and verification. Called back into NSS suport to help solve the sitution. As far as everyone coud tell, the LED's on SCC1 shoud be illuinated, even if the Geo PC was down; which ou was. I was able to find a schematic that contained more logic drawings, and coud see that there is a relay inside SCC that's responsible for the 200000000LED on SCC1 illuinating, and it's connected to a CAN/SAN logic signal. Althouh we thouht it was possible that the Geo PC needed to be fuctional in order to make the LEDs illuinate, we weren't certain, and we didn't have a fuctional Geo PC to confirm the hypothesis. I was able to go borrow a PC Power Suply to bring ou broken Geo PC back from the dead enouh to fuction. Once I connected it to the system, the SCC1 LED's illuinated. This was great news. That confirmed that ou SCC and SCC1 were fuctional. All we needed now was ou Geo PC to be delivered. Cleaned u and headed home for the night. System still down. Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 26-Sep-2000000003 MuH, JOE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad Techs / Imaging Biomed Device ue: System not fuly booting Eupected and ueupected Resuts: Ful fuction uer Impact: Delay Patient Impact: None SE 7.2.16 Resoluion : This 000000000000000000000000001contained part management only. Retuned the second Geo IPC and mPDuheavyweight. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 22-Sep-2000000003 MuH, JOE The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace parts and get system back u and fuly fuctional. Problem Description by Engineer : Cutomer Fuction/Role: Rad tech / Imaging Biomed Device ue: System down du to fauty parts Eupected and ueupected Resuts: To fuction as designed uer Impact: System down Patient Impact: None SE 7.2.16 Diagnostic performed by Engineer : Arrived onsite before parts arrived in order to get started ASAP. I immediately noticed that all parts bu the Geo PC had arrived. This was a big deal and meant that the u system 000000000000000000000000001ldn't be back u the day of. For some reason the part order didn't get processed util after the cuoff time, and was schedued to be delivered neut day. I went ahead and did whatever else I coud do. I replaced the MG Power Suply and SCC1. When I tuned on the u system I noticed that SCC1 had no LED lights illuinated. I didn't know if that was normal or not; nor did anyone else I had spoken with. I did eutensive trouleshooting to verify proper incoming and ougoing voltages as well as CAN and SAN signals. We decided to order a new SCC, as well as an additional SCC1 board jut incase it was DOA. Planned to commence the following day. System down. Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007281 200000000Sep-2000000003 MuH, JOE The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Arrived onsite with parts recommended by RSE. Confirmed the reported issu; system was not booting / Fue F100000000ept blowing. I first tried booting system with Geo PC disconnected, as I thouht that might be loading/ shorting the system ou; that wasn't the issu - F100000000lew again. I replaced the PD Controller per RSE recommendation, bu that didn't resolve the issu. I then began eutensive trouleshooting. I uloaded the PDuF100000000ircut by disconnecting the circuts connected to it. I isolated the short somewhere in the circut leading from u200000000. I narrowed it down to the Tue Oil Cooler. When I reconnected the Oil Cooler to confirm it was the cuprit, the system stopped popping F100000000Not sue why. Tested it via mutiple reboots and long ru times mutiple times. I then realized that mutiple components weren't 000000000000000000000000001king' likely fried du to a power suge of some sort. The Geo PC, SCC1, and MG Power Suplies were fried. SCC is on a different circut than F100000000I ordered all the parts to be installed neut day. Throuhou this process I spent lots of time looking at schematics, talking with NSS suport, and testing. Retuning neut day to install parts. System down. Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : F100000000n the MPDukeeps blowing. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: Room is down Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007237 Sharvari Karua The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : F100000000n the MPDukeeps blowing. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: Room is down Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : Supecting the power suply 0000000080000000070000000061 - PSF125 Kit Internal Comments : uLL# and/or description of uLL: System is not booting u Repair Action : 0000000080000000070000000061 - PSF125 Kit 00000000800000000868300000000- MPDureplacement assembly 00000000230000000016000000007 - MPD CONTROL uIT Follow u Requred : uLL# and/or description of uLL: System is not booting u Repair Action : 0000000080000000070000000061 - PSF125 Kit 00000000800000000868300000000- MPDureplacement assembly 00000000230000000016000000007 - MPD CONTROL uIT Technician Name : Sharvari Karua ----------------------------------------------------------------------------- Internal Remarks : FSE believes a power hit came throuh the system. Originally the F100000000ue kept blowing over and over on the MPDu FSE trouleshot down to NP-u200000000 being pluged in 000000000000000000000000001ld caue it to blow. According to the prints this goes to MG-u6. After pluging cables back in, the fue all the suden did not blow any longer. Then it was noticed the geometry is down and the Geo IPC and SCC1 have no lights. Power was verified at the Geo IPC power plu and a break ou board confirmed 200000000to SCC1. So it seems both of these components are not 000000000000000000000000001king. -Approved: SCC1 Adapter: 00000000800000000000000001772 Geo Q35: 00000000800000000000000007813 -FSE will continu to investigate the MG Power suply. The parts catalog has several parts that can be ordered for this, so futher investigation is needed. Internal Remarks : Assisted FSE MuH, JOE with trouleshooting MPDublowing fue F100000000 Recommended to FSE to uload u200000000 and u100000000 on MPDuand 000000000000000000000000001k backwards util youfind the souce of the dead short ----------------------------------------------------------------------------- | 2023-09-18 21:51 | No
811 | Yes | BZA0system cannot activate u-rays anymore; the main screen is blank | BZA0 system cannot aute u-rays anymore; the main screen is blank | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005978 19-Sep-2000000003 JEAN JORET, The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: in Eupected and ueupected Resuts: u uer Impact: No. Patient Impact: no SE No. Resoluion: Replace kvma uit Tools & Equpment : DOSE-KVP-LIGHT METER, 2000000005-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000008-000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003156 19-Sep-2000000003 JEAN JORET, The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: in Eupected and ueupected Resuts: No. uer Impact: No. Patient Impact: no SE u Diagnostic performed by Engineer: Kvma uit faut Parts Consued: 1.00000000000000000000000000000000217800000000100000000, uit dig. kV mA control Tools & Equpment : DOSE-KVP-LIGHT METER, 2000000005-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000008-000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002992 PATRICK ATTN. The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : No u-ray Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 000Diagnostic performed by Engineer: Generator issu Internal Comments: uLL# and/or description of uLL: 50000000083 uLL2000000003-00000000-16 12:50000000000000000.0000000000000000uuuValidation failue 5100000000100000000100000000tart-u Service NoDetailInfo 500000000800000000rror 2000000003-00000000-16 12:50000000000000000.0000000000000000Device programming failue: u-Ray Generator 510000000010000000011 Start-u Service NoDetailInfo 50000000085 Warning 2000000003-00000000-16 12:50000000000000000.0000000000000000Application uLL Generator euception occured 0000000000000000000000006500000000Start-u Service NoDetailInfo 50000000086 uLL2000000003-00000000-16 12:50000000000000000.0000000000000000Device programming failue: u-Ray Generator 510000000010000000011 Start-u Service NoDetailInfo Repair Action: * Resoluion: Check different FRuwithin the generator Do POST test Reload software ----------------------------------------------------------------------------- | 2023-09-18 21:56 | No
812 | Yes | 00N0ill not boot up | 00CDN0will not boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000000000000018-Sep-2000000003 Colon, Efrain The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Resoluion : Rebooted system, checked logs, performed body gurd calibrations ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000016 George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Verify PC's and TS net000000000000000000000000001k switch are powered on and fuctioning. ue alt + f duing boot u to start FSF and inspect POST resuts, Download board software, and event log to determine caue of issu. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Called contact and debriefed. He said system boots to coutdown clock and hangs, stuk at 00000000:00000000. He said they called one of their FSE's, George and he requsted they open a case. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: cedric ingram for aleu secco IP & Issu: 5565200000000/will not boot u Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created case/ R000000000000000000000000001and called on call/informed of cb Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-18 22:16 | No
813 | Yes | System not booting | System not booting | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 22-Sep-2000000003 Enriquz, Rich The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Eron Device ue: no Eupected and ueupected Resuts: system does not boot uer Impact: system down Patient Impact: none SE Resoluion : Arrived on site to cutomer in the middle of a software reload after replacing the Sute PC, supect the first one he received may have been an ou of the bou failue. Was asked to stay on site in case we ran into any other issus with getting the system back u. New PC and software reload fiued the issu, from there new license files uloaded, backus restored and then system asked for tue and IQ cals to be done. System u and fuly fuctional. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 Clark, Byron The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Diagnostic performed by Engineer : Biomed swapping Sute PC .. need new license file Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - PDS: IPMI device Sute PC is not accessible: 1. The IP address for the IPMI protocol (Intelligent Platform Management Interface), in the BIOS contains an invalid valu. This happens after the OS install. Duing the OS install, the IP address of the PC was not recognized at that time. The PC cannot be powered off by the PDucontrole uit and the shudown time will be several (7) minues. Repair Action : Service actions for RMT - PDS: IPMI device Sute PC is not accessible: 1.1. 1. Enter the BIOS on the Sute PC (an BIOS pass000000000000000000000000001d is requred, which can be requsted at the local Tier2 or Tier3 Suport). 3. Check in BIOS at: 'IPMI - BMC net000000000000000000000000001k configuation - Station IP address' if the correct IP address (192.68.32.00000000 is set for the Sute PC. (IPMI address = IP address of Sute PC + 00000000 00000000Change it to the correct valu. (192.68.32.00000000 1.2. If all PC's have same IPMI warnings, most likely the net000000000000000000000000001k connection with IP 192.68. 32.12 towards the PDuControl Modue (NY15) is not 000000000000000000000000001king as this Control Modue controls and checks IPMI. 1.3. Check if Sute PC (TH) is connected to the NY5 modue. (Connector NY5:u23). 1.00000000Switch modue NY5 coud be defective. Internal Comments : swapping fingerprint for Sute PC Old 28B9D9E162EA New MAC 28B9D9E0000000000000000Problem Description by Engineer : Problem description by engineer : RMT - PDS: IPMI device Sute PC is not accessible 200000000Sep-2000000003 00000000:38:00000000 When was the first occurence? Within... : 200000000Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: room down swapping Sute PCs uer Impact: room down Patient Impact: room down SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000100000000ulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : speaking with Eron the system had cold restart message on system and still does not start .sute and uray pcs did not tun on at boot u.software load with uLL .advise try original sute pc to see if that pc starts at tun on and will check back troule shooting shows problems with start u of fleu spot pc/fleu vision Follow u Requred : check start u of b cabinet and commuation withsute pc Problem Description by Engineer : Cutomer Fuction/Role: Eron Device ue: no Eupected and ueupected Resuts: system does not boot uer Impact: system down Patient Impact: none SE Internal Comments : seems both sute and u ray swapped at same time . pcs do not start with system tunon Cannot connect to fingerprint information FleuSpot failed. The Sute PC is uable to commuicate with the FleuSpot PC. Service Action 1: Check net000000000000000000000000001k (ethernet) cabling from FleuSpot PC (TA) to Gigabit Ethernet Switch (TS): TA:u100000000- TS:u11 Service Action 2: Check net000000000000000000000000001k (ethernet) cabling from Gigabit Ethernet Switch (TS) to Sute PC (TH): TS:u1 - TS:u100000000 Service Action 3: Check the statu of the Gigabit Ethernet Switch (TS). Is it powered? Service Action 00000000 Check whether the FleuSpot PC starts correctly. Service Action 5: Perhaps the Fleuviewing SW is not (yet) installed or corruted on the FleuSpot PC. inhoue pu original sute and uray system u ,fleu spot pc or fleuvision not starting u LOD attached at case levlel FRuRepairable FRu Description FleuViewing PC 00000000 W7 + W100000000 Material Nuber 00000000800000000718000000002 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002581 George Owens The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Debriefed caller. System not booting and caller wants to order several parts. Verified part nubers and stocking levels. Transferred caller to parts order line. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Nick IP & Issu: " 7561200000000" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: New Case Caller: Nick IP & Issu: 7561200000000 & not booting Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: Weekend Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created case and requsted qute Neut Step Action: Awaiting Qute Verbal PO Nuber: SI-1000000006-ND Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Sute PC replacement old: 28B9D91000000006F2 new: 28B9D9E162EA eron.gailey00mail.org RSE B. Clark will perform swap ----------------------------------------------------------------------------- | 2023-09-18 22:16 | No
814 | Yes | AED does not pass self-test (Device is replaced) | AED does not pass self-test (Device is replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006663 Andrea Ardighieri The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: dealer Device ue: Auotes defibrillator Eupected and ueupected Resuts: It does not pass the auo test and has low volue. uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: It does not pass the auo test and has low volue. Follow u Requred: nemo withou collection IREDEEM Spa Piazza dei Martiri 19000000001900000000n.1, 0000000021 Bologna - T +39 000000001 0000000035879 Tel. 300000000855700000000 Ac. Domenico Giordano ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005763 Parts Consued: 1.0000000000000000000000000000000056000000005351,FRu (C B) Defib, Italy Italian 00000000, Euch ----------------------------------------------------------------------------- | 2023-09-18 22:21 | No
815 | Yes | AED Triple Chirps | AED Triple Chirps | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007565 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Ship To First Response Inc 52700000000 Stone000000000000000000000000001d Place Rosedale BC CA V00000000-1u1 Attention Jeremy Gray Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Device was not in ue Eupected and ueupected Resuts: AED triple Chirp when Battery is inserted into the AED Failed the BIT and emitted the triple chirp. The AED is uder warranty and to be replaced at zero cost to the cutomer., uder IB warranty replacement Device serial# A21G-00000000200000000 Serial # of battery B00000000200000000700000000 00265P Install before date: 2000000007-00000000300000000 uer Impact: none Patient Impact: none SE Resoluion : DELIVERED Tusday 9/19/23 at 2:500000000M Internal Comments : waiting for order to ship, order nuber 000000000000000057362157 Diagnostic performed by Engineer : AED triple Chirp when Battery is inserted into the AED Failed the BIT and emitted the triple chirp. The AED is uder warranty and to be replaced at zero cost to the cutomer., uder IB warranty replacement Device serial# A21G-00000000200000000 Serial # of battery B00000000200000000700000000 00265P Install before date: 2000000007-00000000300000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000075900000000Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl ROW, Euch ----------------------------------------------------------------------------- | 2023-09-18 22:51 | No
816 | Yes | 0 - u0- Display screen issue | 00- u000- Display screen issue | 000000000000000000000000001O##: 000000000000000000000000001000000000000000050000000000000000Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Resoluion : Bezel had some patches, Informed cutomer abou EOL. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Display screen issu ----------------------------------------------------------------------------- | 2023-09-18 23:11 | No
817 | Yes | Charge Port is Broken - Incoming Loaner requires repairs in addition to FFU services. (FFU Case# 000 | Charge Port is Broken - Incoming Loaner requires repairs in addition to FFU services. (FFU Case# 0000) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008683 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Resoluion : Device was repaired by u engineer Chongolnee, Roongruchai, with parts suplied by the Bu Calibrated NBP, CO2 ad Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Diagnostic performed by Engineer : Confirmed complaint the uits Charging Port is physically damaged and to correct the issu we need to replace the Front Case Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008100000000 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair Diagnostic performed by Engineer : Confirmed complaint the uits Charging Port is physically damaged and to correct the issu we need to replace the Front Case Assembly. ----------------------------------------------------------------------------- | 2023-09-19 02:01 | No
818 | Yes | This device is RESTRICTED - ALARMS are NOT displaying - Unknown reason- Cuomer resolved | This device is RESTRIuD - ALARMS are NOT displaying - Unknown reason- Cuomer resolved | 000000000000000000000000001O##: 000000000000000000000000001000000000000000088100000000iller, Dell The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.Will await CCB for tomorrow. Internal Comments : Cuverbalized displeasue w/the wait muic and wanted to provide feedback to improve said item. CFS sumitted. 00000000/20000000023 0010000000000000000rs: No CCB on this issu. 000000000000000000000000001closed. Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Shoud alarm for SVT uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Cureports Mu00000000is not alarming while uing simuator for SVT alarms. Diagnostic performed by Engineer : Problem Description: Cureports Mu00000000is not alarming while uing simuator for SVT alarms. Cureports Mu00000000is not alarming while uing simuator for SVT alarms. Form ued: Trouleshooting/Investigation Object: u Not Object: Other fuctions Software Version: uu Date and Time of Issu: 9/18/2000000003 uer Actions: Pt monitoring. Where is the Deviation: un. Trouleshooting Steps Taken: Verified has inop alarm stating some ualarms are off.Instruted to enter usetu menu by touhing HR>arrhythmia>tun arrhythmia on.Recommended attempting to activate a red level alarm on simuator while assigned to sector on the PICIu.However, cuwants to wait util tomorrow, will call back/confirmed has case #. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008700000000Thomas, Shea The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Resolved Remotely Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Mu00000000- Alarm unot 000000000000000000000000001king on Device uer Impact: Yes, Mu00000000- Alarm unot 000000000000000000000000001king on Device Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Mu00000000- Alarm unot 000000000000000000000000001king on Device Diagnostic performed by Engineer : Problem Description: Mu00000000- Alarm unot 000000000000000000000000001king on Device Mu00000000- Alarm unot 000000000000000000000000001king on Device Form ued: Other Trouleshooting Steps Taken: Mu00000000- Alarm unot 000000000000000000000000001king on DeviceProblem Description: Mu00000000- Alarm unot 000000000000000000000000001king on DeviceMu00000000- Alarm unot 000000000000000000000000001king on DeviceForm ued: OtherTrouleshooting Steps Taken:Mu00000000- Alarm unot 000000000000000000000000001king on DeviceMu00000000is on a simuator and not 000000000000000000000000001king for Alarms - uBiomed states uArrythmia Alarms Off - No Alarm Display Cuwent to Mu00000000Smart Keys and Alarm Settings.Quud Clinical Suport 800000000 - sugested to have Mu00000000moved to floor to try an empty Sector.Tried Tuning Arrythmia OFF then ON.Checked Alarms on Mu00000000ecommendation from Clinical was to send to Bench Repair.9/20000000023 - Called cuback and they stated the Mu00000000was back in service or not having any issus now.No details were given bc the TS and repair was from another Biomed on site and not the case owner.We did try to ask. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Thomas/Tom FABATZ IP & Issu: " 73537000000008 This device is RESTRICTED becaue its ALARMS are NOT displaying" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: PCCI Action Taken: I created a case, R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-19 02:26 | No
819 | Yes | 00N0eeds assistance freeing up space images deleted still getting message | 00CDN0needs assistance freeing up space images deleted still getting message | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008890000000019-Sep-2000000003 Shula Chintan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - IPu Free disk space is too low. 17-Sep-2000000003 13:00000000:15 RMT - IPu Euposue not possible. Image disk ful 16-Sep-2000000003 11:30000000026 When was the first occurence? Within... : 17-Sep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: freeing u space images deleted still getting message uer Impact: room down Patient Impact: room down SE 8.2.17.00000000Resoluion : Ran database consistency test and it failed. Performed database consistency repair which foud 00000000dangling records . Cleared this dangling records. Tested system to confirm proper operation. Retuned system back to clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008800000000 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Free disk space is too low.: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Repair Action : Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - IPu Free disk space is too low.: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." If neededParts requred : 1 Part : 1, uu: 0000000080000000050000000000000000 Description : Disk for ClarityIQ_II and IP_Revised_II, Quntity : 1, Priority : 1, Chance Of Sucess : 16.7 Root Caue for RMT - IPu Euposue not possible. Image disk ful: 1. Euposue not possible. Image disk ful Repair Action : Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu Euposue not possible. Image disk ful: 1.1. Instrut cutomer to free u disk space. 1.2. Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Technician Name : Clark, Byron Problem Description by Engineer : Problem description by engineer : RMT - IPu Free disk space is too low. 17-Sep-2000000003 13:00000000:15 RMT - IPu Euposue not possible. Image disk ful 16-Sep-2000000003 11:30000000026 When was the first occurence? Within... : 17-Sep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: freeing u space images deleted still getting message uer Impact: room down Patient Impact: room down SE 8.2.17.00000000Diagnostic performed by Engineer : caller transferred from apps. in houe biomed listed as contact and requsting on site FSE ----------------------------------------------------------------------------- Cutomer comment : All bu 3 cases have been deleted and still getting a disk ful warning message. Internal Remarks : Interaction Type: New Case Caller: Sandy Fuco IP & Issu: 68000000002316 & needs assistance freeing u space Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-19 02:46 | No
820 | Yes | The system remains at "System starting 0:0" and does not start. | The system remains at "System starting 0:0" and does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009536 19-Sep-2000000003 Fuuoto, Atsuhi The system meets the specification for the performed service and is retuned to ue Resoluion: This time I replaced the power suply uit (DCPS) with a new one and confirmed that the system was 000000000000000000000000001king normally. Internal Comments: [Cu related] Coud the person in charge of the equpment agree on how to complete the report?Yes, that day the normal report was completed [000000000000000000000000001k report (paper) sumission] Did the 000000000000000000000000001k report be sumitted in paper? Yes Reguatory Qustions: N/A Problem Description by Engineer: The system remains at "System starting 00000000:00000000" and does not start.Euternal inpu images are displayed on the FleuVision monitor. The touh screen modues in the euamination room and the control room are blacked ou.Tuning the power back on does not improve the sitution. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system remains at "System starting 00000000:00000000" and does not start.Euternal inpu images are displayed on the FleuVision monitor. The touh screen modues in the euamination room and the control room are blacked ou.Tuning the power back on does not improve the sitution. uer Impact: No effect Patient Impact: No effect SE R8.2.3000000005 Diagnostic performed by Engineer: The power suply uit that converted the ac power suply from the main cabinet to direct curent was fauty. Becaue of that, there was no direct curent power being suplied to all the uits inside the main cabinet, and it coud not start u normally. Follow u Requred: N/A Parts Consued: 1.00000000000000000000000000000000800000000387582,PD. SCOMP. DCPS_20000000012V_5V ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009371 Suui, NobuasuThe system does not meet the specification for the performed service and is not retuned to ue Internal Comments: Compl 1.2 [Suject] 00000000DS:NS:u Sasaki: The system remains at "System starting 00000000:00000000" and does not start. [Parts from the equpment] 00000000800000000387582:PD. SCOMP. DCPS_20000000012V_5V QTY:1 ZRFA/Traceable Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Fuuoto, Atsuhi [Takashi Fuuoto] - FSE1 Visit date and time 2000000003/00000000/199:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Takashi Fuuoto - 000000000000000000983000000006861 - Others Call mobile uon arrival Requster name NobuasuSuui [Euamination category] Normal euamination Follow u Requred: * Problem Description by Engineer: The system remains at "System starting 00000000:00000000" and does not start.Euternal inpu images are displayed on the FleuVision monitor. The touh screen modues in the euamination room and the control room are blacked ou.Tuning the power back on does not improve the sitution. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system remains at "System starting 00000000:00000000" and does not start.Euternal inpu images are displayed on the FleuVision monitor. The touh screen modues in the euamination room and the control room are blacked ou.Tuning the power back on does not improve the sitution. uer Impact: No effect Patient Impact: No effect SE R8.2.3000000005 Diagnostic performed by Engineer: RC Response: FSE assignment ----------------------------------------------------------------------------- | 2023-09-19 06:01 | No
821 | Yes | 0/u FD0can't start system | 0R/u FD0can't start system | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Cannot start qulity inspection SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 18.00000000.2000000003 13:13:25 uC (860000000013300000000 "SuT. Ready for Clinical ue local time" : " 2000000003-00000000-18 10000000015" Symptom: The device cannot start normally. Soluion/Repair Measues: The device cannot start normally.At the scene, it was foud that the HostPC did not start normally. Futher inspection foud that the battery of the HostPC BIOS motherboard was ou of power, and the measued voltage was only 000000009V DC.Replace the HostPC BIOS motherboard battery. The equpment 000000000000000000000000001ks normally after service. Internal Remarks: 18.00000000.2000000003 13:13:33 uC (860000000013300000000 "RDI Accuacy: Yes" $Q7:Confirmation reportedproblem, Can't boot u. $Q8: The device cannot be started normally, The onsite check indicates that HostPC is not started properly, Futher check revealed that the HostPC BIOS mainboard battery is ou of power, and only 000000009V DC is measued. $Q9:Replace the HostPC BIOS mainboard battery. $Q100000000Test and Inspection data refer to PA, Do a few patient, Observation of a certain period of time. The system is ok. Device retuned to ful fuctionality. $Q11:Calibrated tools refer to PA. Safety Qustions: Case: 0000000022290000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: can't start system Trouleshooting Action: Key faut codes or phenomena: The cutomer described that when the power buton is pressed, the review modue power-on indicator light is always on, bu the monitor does not respond. Remote diagnostics and test resuts: Remotely, I feel that the cutomer has not arrived at the site, and the actul sitution needs to be verified. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. After confirming the phenomenon at the scene, trouleshoot based on the power-on circut Neut Step Action: Zhang Haitao 13669250000000067 informed that 000000000000000000000000001k can be dispatched normally SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-19 06:06 | No
822 | Yes | 0/u FD0Power problem | 0R/u FD0Power problem | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Frequnt hard drive failues occured from Auut to now, requring on-site inspection Euternal Remarks: 27.00000000.2000000003 00000000:15:300000000uC (86000000008275)PROBLEM DESCRIPTION:THE SYSTEM CAN'T BOOTu INTERMITTENT, NO uLLINFORMATION DISPLAY ON DATA MONITOR. TROuLESHOOTING ACTIONS:DIAGNOSED THE SYSTEM WITH FSF AND DRAWINGS THEN FOuD MAYBE THE HARDDISK CABLE LOOSENED OR THE OS/APP OF SYSTEM DAMAGED. REINSTALLED THE HARDDISK CABLE AND REINSTALLED THE OS/APP OF SYSTEM WITH MANuL THEN THE SYSTEM CAN BOOTu NORMALY. SERVICE RESuT:AFTER SERVICE, THE DEVICE RETuNED TO FuL FuCTIONALITY. SAP Transfer: Case Transferred to SAP Internal Remarks: 27.00000000.2000000003 00000000:15:33 uC (86000000008275) "RDI Accuacy:YES" Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q5: (N) $Q6:（THE SYSTEM CAN'T BOOTu INTERMITTENT) $Q7:(CONFIRMED REPORTED PROBLEM.THE SYSTEM CAN'T BOOTu INTERMITTENT, NO uLLINFORMATION DISPLAY ON DATA MONITOR.) $Q8: (DIAGNOSED THE SYSTEM WITH FSF AND DRAWINGS THEN FOuD MAYBE THE HARDDISK CABLE LOOSENED OR THE OS/APP OF SYSTEM DAMAGED.) $Q9: (REINSTALLED THE HARDDISK CABLE AND REINSTALLED THE OS/APP OF SYSTEM WITH MANuL THEN THE SYSTEM CAN BOOTu NORMALY.) $Q100000000 (Test and Inspection data:REFER TO PA & THE DEVICE RETuNED TO FuL FuCTIONALITTY.) $Q11: Calibrated Tools name (PA) S/N(PA) Neut Calibration date (PA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 12/200000000 2000000002 ^Curent Tue uage: scan second:10000000093LuSAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the host PC cannot be tuned on auomatically and the Host PC needs to be tuned on manully, requring engineers to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Cannot start auomatically, need to start Host PC manully Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the Host PC power suply still has power after the machine is powered off. 2. Check whether the BIOS settings are normal and whether the BIOS battery voltage is normal. 3. Check whether the PC power suply voltage is too high. If it is too high, please adjut it. 00000000Youcan try to ue the IPPC power modue. Carry ou euchange 5. If there is no effect, ue PC Diagnostic Tool for diagnostic testing 6. Check the machine uLLreport and perform faut repair. If it cannot be repaired, replace the corresponding spare parts. Recommended spare part uuand name: 00000000800000000200000000612 Nehalem Host PC Monoplane Rev_III NO_HDD Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Case: 000000002229000000005 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-19 07:11 | No
823 | Yes | Alarms or Messages/Speaker Malfunction alarm/Speaker | Alarms or Messages/Speaker Malfunuon alarm/Speaker |  | 2023-09-19 07:59 | No
824 | Yes | M0A - u0- The operation check fails | M00A - u00- The operation check fails | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: FSE Chen Shengju 00000000581000000005000000001 called to report repair, and Bai Gunglei has been notified by ##Problem Reported by Cutomer: ECR:OCAP:00000000R/HeartStart uuumonitor/defib SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Operation check failed The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: Electrode board item abnormality Repair Action: T2 Recommended on-site repair steps: It is recommended to clean the electrode plate before arriving at the scene. If it is ineffective, replace it again. Recommended spare part uufor T2 with description: na Notes on cutomer requrements: na Test/Inspection: N/A Test Equpment: N/A Euternal Remarks: 19.00000000.2000000003 00000000:0000000000000000 uC (86000000001999) Department: Cardiology Contact person: Teacher Cu Faut problem: Self-test failue, handle failue Trouleshooting: Replace the handle Service resut: The device retuns to normal Internal Remarks: 19.00000000.2000000003 00000000:0000000000000000 uC (86000000001999) $Q1: (N) $Q2: (N) $Q3: (N) $Q00000000(NA) $Q5: () $Q6: (internal test is fail) $Q7: (Confirmed reported problem. Internal test is failed.) $Q8: (Performed visul inspection: According to the manul, the caue of the failue is damage to the paddle.) $Q9: (The paddle has been replaced) $Q100000000 Test and inspection data refer to PA. Device retuned to fuly fuctional. $Q11:Calibrated tools name refer to PA. Safety Qustions: Case: 0000000022290000000061 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-19 08:46 | No
825 | Yes | the uray is not booting up | the uray is not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008688 26-Sep-2000000003 Fluey Lawrence The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Foud BIOS battery very low...000000006v instead of 3.2v. Replaced battery and boot problem resolved. Retuned HOST pc uued. Problem Description by Engineer : Cutomer Fuction/Role: nuse Device ue: host pc Eupected and ueupected Resuts: not booting uer Impact: delays Patient Impact: none SER8.2 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002676 18-Sep-2000000003 Asif, Sawab The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : System’s host pc is not 000000000000000000000000001king properly and wjll be replaced in followu 000000000000000000000000001 Problem Description by Engineer : Cutomer Fuction/Role: Nuse Device ue: Cath Eupected and ueupected Resuts: System not booting u uer Impact: Have to not tun off system Patient Impact: SE Follow u Requred : Host pc replacement, please forward qute to Lawrence Fluey ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: luy 000000000000000000000000001g IP & Issu: 66650000000025 the uray is not booting u (previou case 0000000022287927) Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created case and onsite 000000000000000000000000001 Neut Step Action: n/a Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : call was transferred to me from the u team as the cutomer needed to have service tonight. Call the FSE on call and case dispatched. ----------------------------------------------------------------------------- | 2023-09-19 09:21 | No
826 | Yes | ufault | ufault | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Prompt ufailue The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Check the lead wires, motherboard, etc. Recommended spare part uufor T2 with description: na Notes on cutomer requrements: na SAP Transfer: Case Transferred to SAP Internal Remarks: 22.00000000.2000000003 00000000:39:00000000uC (860000000000000000100000000 $ASP engineer: [honglejia uuchungfeng] $Q1:NO $Q2:NO $Q3:N/A $Q00000000/A $Q7:Confirmed reported problem: ualarm faut. $Q8:Detected and foud an umalfuction du to damage to the electrocardiogram lead wire. $Q9:After replacing the electrocardiogram lead wire, the instruent retuned to normal ue. $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11:Calibrated Tools：#NA Safety Qustions: Case: 0000000022283769 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 22.00000000.2000000003 00000000:39:38 uC (860000000000000000100000000Department information: [Department: Equpment Department Contact Person: Teacher Li Contact Nuber: 191000000002660000000000000000* $ASP engineer: [honglejia uuchungfeng] $Q1:NO $Q2:NO $Q3:N/A $Q00000000/A $Q7:Confirmed reported problem: ualarm faut. $Q8:Detected and foud an umalfuction du to damage to the electrocardiogram lead wire. $Q9:After replacing the electrocardiogram lead wire, the instruent retuned to normal ue. $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11:Calibrated Tools：#NA SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000N/uDFM100000000 Defibrillator/Monitor/ ----------------------------------------------------------------------------- | 2023-09-19 11:56 | No
827 | Yes | M0A - u0 - The battery of the A battery needs to be calibrated. | M00A - u00 - The battery of the A battery needs to be calibrated. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000002 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: parameter monitor Eupected and ueupected Resuts: battery calibration is necessary uer Impact: no impact Patient Impact: no impact SE na It is not possible to carry ou calibration as they need several hous of down time for device, proceeded with requst to order battery. Diagnostic performed by Engineer: message that calibration is necessary for battery in A department. Follow u Requred: NEMO WITHOu COLLECTION BuALINI HOSP. Viale Giovanni Ghirotti, 286 0000000021 Cesena (FC) Delivery to the "Transit Point" warehoue. Department - Medical Physics - Clinical E. att.ne Marcello Prunoli - tel 0000000000000000352987 You ref. 200000000311600000000 BATTERY to requst from Sidem. ----------------------------------------------------------------------------- | 2023-09-19 12:01 | No
828 | Yes | M0A - u00 Device sporadically displays the message "Charge button: uLL during the function test. | u Device sporadically displays the message "Charge button: uLL during the funuon test. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003121 Christiane Atwill The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: Charge buton jams sporadically uer Impact: no Patient Impact: no SE Resoluion: We regret to inform youthat you system HeartStart uuumonitor/defib on the 31.12.2000000002 reached the End of Life (EOL) statu. The mentioned EOL date is the last day on which u gurantees to make parts and technical resouces available to suport this produt. Please contact the specialty dealer to discus the futher procedue. Since the device is fauty and cannot be repaired, it mut be taken ou of service. ----------------------------------------------------------------------------- | 2023-09-19 12:11 | No
829 | Yes | DFM0 restarted (Quote is cancelled by cuomer) | DFM0 restarted (Quote is cancelled by cuomer) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000595 Kaan Ozkan The system does not meet the specification for the performed service and is not retuned to ue. Follow-u Requred: Repair on the atol Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: The device is restarting. uer Impact: no Patient Impact: no SE N/A coud not be detected as the device did not open. Diagnostic performed by Engineer: Duing the inspection, it was detected that the device restarted duing start-u. The processor is defective. According to the failue detection; 0000000056000000009000000001 DFM100000000 ASSY Processor shoud be quted. Tools & Equpment : MuTIMETER, 200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-19 12:16 | No
830 | Yes | [Q] Wireless footswitch is malfunction. | [Q] Wireless footswitch is malfunuon. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 15-Oct-2000000003 AuAbdelrazek The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Wireless footswitch is malfuction How often is issu occuring? : First occurence Malfuction area : FCO Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 8.2.17.00000000Resoluion : Replaced the Wireless footswitch and Base Station and did Pairing between them and now system is runing well as cutomer satisfied. Parts Consued : 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238192,Wireless Footswitch 3P ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003172 Amjad Aburqou The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Wireless footswitch is malfuction. Need to order new WFS. Repair Action : Wireless footswitch is malfuction. Need to order new WFS. Follow u Requred : uLL# and/or description of uLL: Wireless footswitch is malfuction. Need to order new WFS. Repair Action : Wireless footswitch is malfuction. Need to order new WFS. Technician Name : Amjad Aburqou Problem Description by Engineer : Problem description by engineer : Wireless footswitch is malfuction How often is issu occuring? : First occurence Malfuction area : FCO Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 8.2.17.00000000Diagnostic performed by Engineer : Wireless footswitch is malfuction. Need to order new WFS. ----------------------------------------------------------------------------- | 2023-09-19 12:21 | No
831 | Yes | System stopped during euamination - FM u-ray not possible | System stopped during euamination - FM u-ray not possible | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009600000000eptember 25, 2000000003 Hoffmann, Thomas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Resoluion: Installed spare part. Fuction check OK. Parts Consued: 1.000000000000000000000000000000008000000000000000000000000FDCSIB Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000006-20000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000September 19, 2000000003 Pugs, Volker The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Technician Device ue: FDSIB Eupected and ueupected Resuts: none uer Impact: none Patient Impact: none 00Resoluion: Trouleshooting, FDSIB board shows a self-test uLL board software installed again, system restarted, then the system was operational, FDSIB board ordered, will be replaced on Monday duing maintenance, fuction test Tools & Equpment: MuTIMETER, 200000000000000008-16 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001300000000September 26, 2000000003 Bierbau, Rainer The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Technician Device ue: none Eupected and ueupected Resuts: none uer Impact: none Patient Impact: none 00Resoluion: Trouleshooting FDSIB board shows a self-test uLL software re-installed and system started, then the system is operational, FDSIB will be replaced on Monday duing maintenance, fuction check Tools & Equpment: MuTIMETER, 200000000000000008-16 00000000:00000000:00000000 Tools & Equpment: MuTIMETER, 200000000000000008-16 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: Call-ou contact Karl Reus Phone +000000009927000000008300000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE n.a. Internal Comments: uLL# and/or description of uLL System not connected to Remote Repair Action: Trouleshooting Follow-u requred: uLL# and/or description of uLL System not connected to Remote Repair Action: Trouleshooting Technician name: Hellenbrandt, Helmu Diagnostic performed by Engineer: Failue of the system on 18.00000000.2000000003 The patient was re-stored on the neighboring system and the euamination was sucessfuly terminated there ----------------------------------------------------------------------------- Internal Remarks: Rouing to short term. Cannot uderstand why the rouing is with Long Term Internal Remarks: u Bierbau will take over, will be on site aroud 10000000000000000 am Cutomer comment: Dr. Strauer got in contact. Cutomer ugently asks for call back. +00000000992700000000231 Cutomer comment: 00emote: u Reus is asking for a call-back, Cutomer comment: ## Service today from 10000000000000000 hrs. by remote passed on to u Reus. Cutomer comment: Cutomer not reached, u Reus buy. Dr. Strauer rings throuh ----------------------------------------------------------------------------- | 2023-09-19 12:31 | No
832 | Yes | Problem with the re-up. | Problem with the re-up. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003299 Thierry LE GONIDEC The system meets the specifications for the performed service and is retuned to ue. Resoluion: I do not see any "patient disconnection" alarm for 9/17 aroud 00000000:00000000 in room 1E 000000002 or 1 E000000001, there were alarms of low VM. For the case of the patient disconnection for the Evita uL, as docuented, it coud not be tested, we did not manage to generate the alarm on the device at the time of the test. It seems that it is possible now, I think that doing a test session 000000000000000000000000001ld be the best way to move forward on this problem. Looking at the integration of the Mutifiltrate, I am available this week to do the first tests requring a machine in a state of empty operation. The list of alarms is listed in the docuent in PJ. Romain Rousy Clinical Data Services (CDS), Int'l Diagnostic performed by Engineer: On 17/00000000 at 00000000:500000000 in room 1 E0000000000000000, there were no alarms of "patient disconnection", instead there was a low VM alarm. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Reguatory Qustions: No/No ----------------------------------------------------------------------------- | 2023-09-19 12:51 | No
833 | Yes | Alarms or Messages/Loudspeaker uLLwhen plug-in to a docking station - no sound/Unknown | Alarms or Messages/Loudspeaker uLLwhen plug-in to a docking station - no sound/Unknown | Cutomer comment: u Wald reports by ## Dear Sir or Madam, The problem has now resolved itself. ----------------------------------------------------------------------------- | 2023-09-19 12:56 | No
834 | Yes | 0S:TK: Fluoros clear images suddenly became rough.(Tube rotor problem) | 0DS:TK: Fluoros clear images suddenly became rough.(Tube rotor problem) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 19-Sep-2000000003 Oyamada, MitsuuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : Fluros clear images sudenly became rouh. u Kanamaru Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Fluros clear images sudenly became rouh. u Kanamaru uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Resoluion: When I visited the site, the message "Tue rotor problem..." was continuuly displayed and u-ray was not possible. I investigated and foud that the uit that controls the rotation of the u-ray tue on the anode side was fauty. I replaced the above uit in the control cabinet in the machine room and that restored normal operation. Resoluion: [Parts ued] Part nuber:< 000000002100000000125700000000Part name: Roco Verial Replacement kit>/000<-> /Batch nuber:<-> /uits: < 1> Parts Consued: 1.000000000000000000000000000000002100000000125700000000oco Velara Replacement Kit ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000761 Kamata, Takayui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : Fluros clear images sudenly became rouh. u Kanamaru Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Fluros clear images sudenly became rouh. u Kanamaru uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: I had him restart the system, bu that did not improve the sitution. ・I confirmed in the log that there was a "Rotor" or "Cooling uit" not good condition. I asked FSE Oyamada to respond. Internal Comments: uLL# and/or description of uLL: [Suject] 00000000DS:TK:u Kanamaru Fluros the images sudenly became rouh.(Tue rotor problem) [Internal Comments] ＝＝＝ 57000000002000000003-00000000-19 15:3000000006.0000000000000000Application uLL Generator euception occured [SEVERITY] Warning [EuCEPTION_DESCRIPTION] Generator Euception Class: 2, Nuber: 57676, Time: 15:3000000009/9/2000000003, Teut: uGEN Warning (100000000L) See detailed LOG or docuof FuRC a ChannelIdentification : u-Ray Channel Frontal Warning u-Ray Generator NormalOperation 000000000000000057676 57000000002000000003-00000000-19 15:3000000006.0000000000000000Redued performance detected [FREE_TEuT] Rotor control is defect for tue ; Tue nr: 1 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator defect statu notification received ChannelIdentification : u-Ray Channel Frontal uLLu-Ray control NormalOperation 000000000000000000000000000000008 57000000002000000003-00000000-19 15:3000000006.0000000000000000System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Information Acqusition NormalOperation 510000000021391 57000000002000000003-00000000-19 15:3000000006.0000000000000000Redued performance detected [FREE_TEuT] Either cooling uit is the problem or rotor. [SEVERITY] Warning [EuCEPTION_DESCRIPTION] Anode is not at operable speed for tue 1. ChannelIdentification : u-Ray Channel Frontal Warning u-Ray control NormalOperation 000000000000000000000000000000008 57000000002000000003-00000000-19 15:3000000006.0000000000000000Euposue not possible. Tue rotor problem uerMessage Acqusition NormalOperation 5100000000999200000000 57000000002000000003-00000000-19 15:3000000006.0000000000000000Low load flu flavor selected. Tue rotor problem uerMessage Acqusition NormalOperation 5100000000999200000000 ＝＝＝ [Parts from the equpment] 000000002100000000125700000000oco Velara Replacement Kit QTY:1 ZSID Tokyo (u11) 0000000021000000008600000000REPLACEMENT COOLING uIT Cu3100000000 QTY:1 ZREP/Traceable Tokyo (u11) 00000000210000000015552:PCP Control Circut Cu3100000000 QTY:1 ZSID Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [MitsuitsuOyamada] Oyamada, Mitsuu - FSE1 Visit date and time 2000000003/00000000/1918:300000000 - FSE1 000000000000000000000000001k time 3 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name MitsuitsuOyamada - 0000000000000000003915-200000000 - Others Call mobile uon arrival Requster name Takayui Kamata [Euamination category] Normal euamination Repair Action : [Suject] 00000000DS:TK:u Kanamaru Fluros the images sudenly became rouh.(Tue rotor problem) [Internal Comments] ＝＝＝ 57000000002000000003-00000000-19 15:3000000006.0000000000000000Application uLL Generator euception occured [SEVERITY] Warning [EuCEPTION_DESCRIPTION] Generator Euception Class: 2, Nuber: 57676, Time: 15:3000000009/9/2000000003, Teut: uGEN Warning (100000000L) See detailed LOG or docuof FuRC a ChannelIdentification : u-Ray Channel Frontal Warning u-Ray Generator NormalOperation 000000000000000057676 57000000002000000003-00000000-19 15:3000000006.0000000000000000Redued performance detected [FREE_TEuT] Rotor control is defect for tue ; Tue nr: 1 [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator defect statu notification received ChannelIdentification : u-Ray Channel Frontal uLLu-Ray control NormalOperation 000000000000000000000000000000008 57000000002000000003-00000000-19 15:3000000006.0000000000000000System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Information Acqusition NormalOperation 510000000021391 57000000002000000003-00000000-19 15:3000000006.0000000000000000Redued performance detected [FREE_TEuT] Either cooling uit is the problem or rotor. [SEVERITY] Warning [EuCEPTION_DESCRIPTION] Anode is not at operable speed for tue 1. ChannelIdentification : u-Ray Channel Frontal Warning u-Ray control NormalOperation 000000000000000000000000000000008 57000000002000000003-00000000-19 15:3000000006.0000000000000000Euposue not possible. Tue rotor problem uerMessage Acqusition NormalOperation 5100000000999200000000 57000000002000000003-00000000-19 15:3000000006.0000000000000000Low load flu flavor selected. Tue rotor problem uerMessage Acqusition NormalOperation 5100000000999200000000 ＝＝＝ [Parts from the equpment] 000000002100000000125700000000oco Velara Replacement Kit QTY:1 ZSID Tokyo (u11) 0000000021000000008600000000REPLACEMENT COOLING uIT Cu3100000000 QTY:1 ZREP/Traceable Tokyo (u11) 00000000210000000015552:PCP Control Circut Cu3100000000 QTY:1 ZSID Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [MitsuitsuOyamada] Oyamada, Mitsuu - FSE1 Visit date and time 2000000003/00000000/1918:300000000 - FSE1 000000000000000000000000001k time 3 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name MitsuitsuOyamada - 0000000000000000003915-200000000 - Others Call mobile uon arrival Requster name Takayui Kamata [Euamination category] Normal euamination Follow u Requred: * ----------------------------------------------------------------------------- | 2023-09-19 13:16 | No
835 | Yes | Battery faulty | Battery faulty | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes: This PO is a Reference Nuber only Problem Reported by Cutomer: 00000000K/V600000000uLLSAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022291778 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem ### Battery uLLreported when booting The cutomer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### na Repair Action: T2 Recommended Field Diagnostic Steps### It is recommended to re-plu the battery and measue the battery voltage. T2 Recommended spare part uuwith description### 989800000000626900000000Assy Battery Li-Ion,11ah,V600000000 Cutomer demand### na Test/Inspection: na Test Equpment: na Euternal Remarks: 20000000000000000.2000000003 00000000:57:200000000TC (86000000000000000085) *Department information: [Department: Respiratory Department Contact Person: Teacher Gu Tel: 17729668800000000] * $ASP Engineer: [Chongqing WuBEI Jia ZHOuuLu] $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (confirmed reported problem, battery failue) $Q8: (Download the "DRPT" report, display uLLcode 11200000000juge it is battery failue) $Q9: (replace battery) $Q100000000 (Refer to PA checklist for test and inspection data. The equpment is restored to ful fuction.) $Q11: Not applicable Internal Remarks: 20000000000000000.2000000003 00000000:57:26 uC (86000000000000000085) $ASP Engineer: [Chongqing WuBEI Jia ZHOuuLu] $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (confirmed reported problem, battery failue) $Q8: (Download the "DRPT" report, display uLLcode 11200000000juge it is battery failue) $Q9: (replace battery) $Q100000000 (Refer to PA checklist for test and inspection data. The equpment is restored to ful fuction.) $Q11: Not applicable SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-19 13:41 | No
836 | Yes | Low electric shock energy - | Low eleuic shock energy - | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport Problem Reported by Cutomer : 00000000R_1100000000R_8612900000000Low electric shock energy Safety Qustions : Case: 0000000022291788 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes : This PO is a Reference Nuber only SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-19 13:46 | No
837 | Yes | Low electric shock energy - | Low eleuic shock energy - | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP FSE Notes : This PO is a Reference Nuber only Safety Qustions : Case: 0000000022291789 Safety Qustion Statu: In Process PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer : 00000000R_1100000000R_8612900000000Low electric shock energy SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 0000000022291789 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-19 13:46 | No
838 | Yes | CR. THE EQUIPMENT DOES NOT WORK, IT FELL | CR. THE EQUIPMENT DOES NOT WORK, IT FELL | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000051 Ruz, Ignacio The system meets the specification for the performed service and is retuned to ue. Resoluion: Therapy PCA is euchanged. The client's own configuation is maintained. All necessary verifications have been performed to certify the restituion to the conditions prior to the failue. To gurantee the correct operation of the equpment, we recommend them to ue only the accessories and consuables approved by u. Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: Defibrillator. Eupected and ueupected Resuts: The equpment has been dropped and from this accident it does not 000000000000000000000000001k. Operational Check is not passed du to a failue in the therapy board or "Therapy PCA". It mut be replaced. The rest of measuements are checked with another test board and it is observed that the equpment 000000000000000000000000001ks perfectly. uer Impact: uknown. Patient Impact: uknown. SE 2.00000000.23 Reguatory Qustions: N/A Tools & Equpment: ELECTRICAL SAFETY TESTER, 2000000003-12-00000000 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER, 200000000000000007-00000000 00000000:00000000:00000000 PATIENT SIMuATOR, 2000000003-12-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000022 Ruz, Ignacio The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Follow u Requred: The equpment was dropped and from this accident it does not 000000000000000000000000001k. Operational Check is not passed du to a failue in the therapy board or "Therapy PCA". It mut be replaced. The rest of measuements are checked with another test board and it is observed that the equpment 000000000000000000000000001ks perfectly. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: Defibrillator. Eupected and ueupected Resuts: The equpment has been dropped and from this accident it does not 000000000000000000000000001k. Operational Check is not passed du to a failue in the therapy board or "Therapy PCA". It mut be replaced. The rest of measuements are checked with another test board and it is observed that the equpment 000000000000000000000000001ks perfectly. uer Impact: uknown. Patient Impact: uknown. SE 2.00000000.23 Diagnostic performed by Engineer: The equpment fell and from this accident it does not 000000000000000000000000001k. Operational Check is not passed du to a failue in the therapy board or "Therapy PCA". It mut be replaced. The rest of measuements are checked with another test board and it is observed that the equpment 000000000000000000000000001ks perfectly. ----------------------------------------------------------------------------- | 2023-09-19 13:51 | No
839 | Yes | Duplicate of PR00 | Duplicate of PR000 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003731 FRANCESCO GIuFRIDA The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: It is reported that the blood pressue alarm souds only once, then it stops even if the valu is ou of range. uBulding: Lot 1, Floor: S2, Cdc: Anesthesia and Resucitation Dea and Transplants. Problem Description by Engineer: Cutomer Fuction/Role: department Device ue: cmonior mu500000000 Eupected and ueupected Resuts: management of alarms uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Requst to perfect management of alarms in the configuation of the monitors ----------------------------------------------------------------------------- | 2023-09-19 14:01 | No
840 | Yes | ICL A Lab 0o display on all monitor. | ICL A Lab 0No display on all monitor. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000179000000008-Sep-2000000003 Afiq Huaini Bin Khairuin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical technician Device ue: Diagnostic Eupected and ueupected Resuts: All monitor blank uer Impact: Cannot ue the system Patient Impact: No patient duing that time SE8.2.17 Resoluion : Check and foud the grabber card of the uuuuable to initialise. Replace the uuu Tested ok. Device is retun to fuctionality. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001925 Afiq Huaini Bin Khairuin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Technician Device ue: Diagnostic Eupected and ueupected Resuts: No display on all monitor uer Impact: uer cannot ue the system Patient Impact: No patient duing that time SE8.1 Diagnostic performed by Engineer : Ask the uer to restart the system. After restart, the system is still not fuctioning. Supect the Fleuvision PC is fauty. Device is not retun to fuctionality. ----------------------------------------------------------------------------- | 2023-09-19 14:01 | No
841 | Yes | Geometry is not functioning | Geometry is not funuoning | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 19-Sep-2000000003 Sivakuar, Muhuuar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Geometry is not fuctioning When was the first occurence? Within... : 19-Sep-2000000003 13:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Geometry is not fuctioning uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.17.00000000Tools & Equpment : MuTIMETER,2000000003-1000000002000000000000000000000000:00000000 Resoluion : System geometry got down duing procedue. Foud Geo PC issu. Reset connection and reloaded software. Machine is 000000000000000000000000001king in good condition and handed over to the cutomer as same. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Geometry is not fuctioning When was the first occurence? Within... : 19-Sep-2000000003 13:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Geometry is not fuctioning uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.17.00000000Diagnostic performed by Engineer : Geopc sw needs to be reloaded. Internal Comments : uLL# and/or description of uLL: Geometry is not fuctioning Repair Action : Geopc sw needs to be reloaded. Resoluion : Geopc sw needs to be reloaded. Follow u Requred : uLL# and/or description of uLL: Geometry is not fuctioning Repair Action : Geopc sw needs to be reloaded. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-19 14:16 | No
842 | Yes | Computer hangs in the uYy logo and does not start the application. | Computer hangs in the uYy logo and does not start the application. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000559 September 200000000 2000000003 Rödder, Nikolau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. u Göddecke is on the way. 0000000000000000118117273 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Resoluion: FleuVision PC incluing Hard drive replaced. Necessary software packages loaded, local screen layous restored. Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000136 Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. u Göddecke is on the way. 0000000000000000118117273 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: System remains at 0000000000000000. Fleuvision image division visible. Cutomer tried restarts and making it powerless. Restarted Fleuvision PC manully with u Gonzalo, no change. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: Schedue TE service to check and replace Fleuvision PC on site, 2 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Internal Remarks: uu 0000000080000000012200000000 - u Haswell Fleuvision 2 PC no HDD Cutomer ETA: SEPTEMBER 200000000 2000000003 Internal Remarks: According to FSE Rödder: In this room, operation is also being done withou a system. He needs a free room for repair becaue otherwise he cannot test the system after repair. Cutomer comment: Dr. Wagenhäuer not reached - rings throuh Internal Remarks: No display on Fleuvision monitor. System remains at 0000000000000000. Asked for suport by bulding services when opening the B cabinet and manul restart of the Fleuvision PC. Internal Remarks: u Rödder can take over, if everything goes well, the 2 hous Cutomer comment: OP nuse: The dr. jut went into the OR, we shoud get in contact in c. 1:3000000002 hous. Cutomer comment: u Gonzalo Fischer called. Cutomer ugently asks for call back +0000000011810000000000000000 Cutomer comment: u Gödecke -Tel.: 0000000000000000118117273 - asked abou processing statu and was informed abou the appointment 20000000000000000.23 - from 00000000:00000000 hrs. Internal Remarks: Appointment consutation also 000000000000000000000000001ks with u Gonzalo Fischer at 0000000011810000000000000000 Internal Remarks: u Rödder on bou For the pupose of taking over, 2 hous of 000000000000000000000000001k? Cutomer comment: ## Arranged appointment with u Gonzalo for Wednesday. 20000000000000000. 9:00000000 - 11:00000000 am ## ----------------------------------------------------------------------------- | 2023-09-19 14:31 | No
843 | Yes | 0 - u0 - Low electric shock energy | 00- u000- Low eleuic shock energy | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 0000000022292000000008 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_11000000007_8612900000000Low electric shock energy SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-19 14:36 | No
844 | Yes | 0 - u0- Low electric shock energy | 00- u000- Low eleuic shock energy | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 00000000R_1100000000NES_8612900000000Low electric shock energy FSE Notes : This PO is a Reference Nuber only Safety Qustions : Case: 0000000022292000000001 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-19 14:36 | No
845 | Yes | Audio & Sounds/No sound: uLL Loudspeaker uLL Speaker/HW: Replaced(000 | Audio & Sounds/No sound: uLL Loudspeaker uLL Speaker/HW: Replaced(00000 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Regina Niemann The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u requred: Send 1u 0000000056000000006971, cutomer will then replace it himself Problem Description by Engineer: Cutomer Fuction/Role: BioMed / Service Technician Device ue: uknown Eupected and ueupected Resuts: defective speaker uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer: uLL Louspeaker is electrically defective, failed completely Soluion: Part 0000000056000000006971 mut be replaced System is ou of ue u to the installation of the spare parts(s) After 00consutation with u Scholz: The cutomer will perform the repair himself - How and when did the uLLoccu? 18.00000000.2000000003 - Was there an uLLmessage? If yes, which ones? Louspeaker uLL - Has the louspeaker failed completely or does it still make souds? Yes, failed completely - Was the device in companion mode with an u2 or u3 parameter server? If companion mode: Was there an uLLmessage on both devices? Please also give u the software versions of the t000000000000000000000000001devices. Yes, was in companion mode, with u3 modue. uLLmessage was assigned to the monitor. - Was the device in operation on the patient when the uLLwas noticed? No - Was a patient harmed or was necessary treatment delayed becaue of this? No ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003300000000Parts Consued: 1.0000000000000000000000000000000056000000006971.Mu_75000000008500000000& AD_75/85 Louspeaker Assy ----------------------------------------------------------------------------- Cutomer comment: u Scholz reports by ## Hello Here are the answers to you qustions: - How and when did the uLLoccu? 18.00000000.2000000003 - Was there an uLLmessage? If yes, which ones? Louspeaker uLL - Has the louspeaker failed completely or does it still make souds? Yes, failed completely - Was the device in companion mode with an u2 or u3 parameter server? If companion mode: Was there an uLLmessage on both devices? Please also give u the software versions of the t000000000000000000000000001devices. Yes, was in companion mode, with u3 modue. uLLmessage was assigned to the monitor. - Was the device in operation on the patient when the uLLwas noticed? No - Was a patient harmed or was necessary treatment delayed becaue of this? No ----------------------------------------------------------------------------- | 2023-09-19 14:51 | No
846 | Yes | Device euited in the middle of the euamination | Device euited in the middle of the euamination | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000253 Hellenbrandt, Helmu The system meets the specification for the performed service and is retuned to ue Internal Comments: uLL# and/or description of uLL - Repair Action: - Diagnostic performed by Engineer: u Pfeifer only asked whether the maintenance can also be done with today's failue of the HKL2. There is curently no failue on HKL1 Problem Description by Engineer: Problem description by engineer: RR Pau Pfeifer Phone +00000000992700000000231 The uLLmessage is not confirmed. Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE n.a. Resoluion: After consutation with FSE Bierbau, this does not seem to be possible since the system is not operational. The curent case will be closed. ----------------------------------------------------------------------------- | 2023-09-19 14:51 | No
847 | Yes | The radiation telescopic is not working. Its loader functions normally | The radiation telescopic is not working. Its loader funuons normally | 000000000000000000000000001O##: 000000000000000000000000001000000000000000080000000028-Sep-2000000003 Karatzetzos, Stefanos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : After 00commuication It was reported that the wireless footswitch It does not respond. The euams were sucessfuly conduted uing the wired footswitch A changer was ordered. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Resoluion: The wireless foot switch was paired according to the base and the commuication problem (wireless footswitch with base) was supassed. The system did not fuly operate. Parts Consued: 1.00000000000000000000000000000000800000000172731,OMNIKEY 6121 Reader, Copyright 2000000006 1.00000000000000000000000000000000800000000881391,Service Smart Card 1.0000000000000000000000000000000080000000000000000811,Pictogram Sheet Footswitch Family ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 Spiliopouos, Antonis The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer : After 00commuication It was reported that the wireless footswitch It does not respond. The euams were sucessfuly conduted uing the wired footswitch A changer was ordered. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Diagnostic performed by Engineer: After 00commuication It was reported that the wireless footswitch It does not respond. The euams were sucessfuly conduted uing the wired footswitch A changer was ordered. Internal Comments: uLL# and/or description of uLL: After 00commuication It was reported that the wireless footswitch It does not respond. The euams were sucessfuly conduted uing the wired footswitch A changer was ordered. Repair Action : After 00commuication It was reported that the wireless footswitch It does not respond. The euams were sucessfuly conduted uing the wired footswitch A changer was ordered. Follow u Requred: uLL# and/or description of uLL: After 00commuication It was reported that the wireless footswitch It does not respond. The euams were sucessfuly conduted uing the wired footswitch A changer was ordered. Repair Action : After 00commuication It was reported that the wireless footswitch It does not respond. The euams were sucessfuly conduted uing the wired footswitch A changer was ordered. Technician Name : Spiliopouos, Antonis ----------------------------------------------------------------------------- | 2023-09-19 14:56 | No
848 | Yes | M0A - u00 Screen Replacement | u Screen Replacement | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000FRANCESCO GIuFRIDA The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Device is not suported Requst for u1 0000000056000000002631 u1 00000000563338331 u1 0000000056300000000000000001 If the material is not available, try to order in SIDEM. We carry ou NEMO withou collection Tu Osp. Barone Romeo 100000000azzini Street Patti c/o Pharmacy Att. Chiara Stefano Tel. 3000000005566855 Problem Description by Engineer: Cutomer Fuction/Role: Department Device ue: defib. uuu Eupected and ueupected Resuts: disp., uuu. Paddle tray is broken. uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Requst for 1u Screen 1u Condenser u1 Paddle Tray ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000----------------------------------------------------------------------------- | 2023-09-19 15:01 | No
849 | Yes | uLS AED MAN - Will not charge | TempuLS AED MAN - Will not charge | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Euchange device to be issud to the end uer: 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Suport Manager Device ue: not in ue Eupected and ueupected Resuts: End uer uable to charge battery when inserted in LS as a pin from the charging cable is stuk in the power socket. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Euchange device to be issud to the end uer: 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000057 Parts Consued : 1.000000000000000000000000000000005600000000000000001,Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- | 2023-09-19 15:16 | No
850 | Yes | Operation check failed (Issue did not reproduce) | Operation check failed (Issue did not reproduce) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Fse visit requred Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Fse visit requred ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000522 200000000Sep-2000000003 Milind, Bhovad The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : on 18 th sep Checked system done operational check and pass sucessfuly. Done cleaning and other test system 000000000000000000000000001king fine. Visited site again on 200000000th September with DEFEB ANALYZER. tested on analyzer system 000000000000000000000000001king ok and shown to cutomer.handed over system in good 000000000000000000000000001king condition to cutomer. Problem Description by Engineer : Shocked not delivered Resoluion : tested on analyzer system 000000000000000000000000001king ok and shown to cutomer.handed over system in good 000000000000000000000000001king condition to cutomer. Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : N/A Tools & Equpment : DEFIBRILLATOR ANALYZER,2000000003-11-16 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-19 15:31 | No
851 | Yes | Touchscreen not responding | Touchscreen not responding | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000000000008-Sep-2000000003 Bradley Holland The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technician Device ue: Duing PM Eupected and ueupected Resuts: system touh screen not 000000000000000000000000001king coud not perform PM uer Impact: Cannot monitor Patient Impact: cannot monitor SE 2.00000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000008-31 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000Diagnostic performed by Engineer : Replacement of Touh screen du to not responding to inpu. coud not perform PM. touh screen replaced and Ful performance verification done . Related to case: 000000002213000000008 ----------------------------------------------------------------------------- | 2023-09-19 15:41 | No
852 | Yes | The system does not start. | The system does not start. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000255 21-Sep-2000000003 Pawel Kierzek The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Dr. Device ue: igt-s. Eupected and ueupected Resuts: system uable to start. uer Impact: none. Patient Impact: none. uu Resoluion: The ventilation and power suply modue PDuin the M cabinet was replaced. The correct operation was checked with a positive resut. The faut was repaired. The device was handed over for clinical ue. Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000276 19-Sep-2000000003 Pawel Kierzek The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Dr. Device ue: igt-s. Eupected and ueupected Resuts: system uable to start. uer Impact: none. Patient Impact: none. uu Diagnostic performed by Engineer: A damage of the PDupower suply modue in the M cabinet was foud. A replacement of the above part is necessary. The device is ou of the way. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 Pawel Kierzek The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Replacement of the fu power trey. Follow u Requred: On-site service visit in order to perform diagnostics. Problem Description by Engineer: Cutomer Fuction/Role: Dr. Device ue: igt-s. Eupected and ueupected Resuts: system uable to start. uer Impact: none. Patient Impact: none. uu Diagnostic performed by Engineer: Interview with the uer revealed that the device on the M cabinet is damaged. On-site service call is requred in order to perform diagnostics. ----------------------------------------------------------------------------- | 2023-09-19 16:01 | No
853 | Yes | 0 FRA Hu0/0vercharged smoked and caught fire | 00FRA Hu00/0Overcharged smoked and caught fire | 2000000003-00000000-25 12:00000000:00000000:Infotip created on the new case 60000000000000000357100000000800000000000000003-00000000-22 13:22:500000000ith the assistance of LA closed the IT that is in the case. New IT created by MGS for one device. 2000000003-00000000-22 11:33:100000000I called the cutomer and he accepted Hu9911/00000000 I can't ue the Infotip in the case becaue it's a uuu000000000000000000000000001O#for 2 2000000003-00000000-22 10000000032:00000000and not only one;2000000003-00000000-22 10000000032:00000000consuer call to get more information abou the escalation //invited to wait /;2000000003-00000000-21 11:15:300000000Received approval from MB to offer a soluion.; 2000000003-00000000-21 10000000038:25: Sent e-mail to management team regarding vouher soluion. 2000000003-00000000-20000000000000000:27:58: Hello, thank youfor you feedback. . Fire Hazard: A bunt food bruh may present an additional fire hazard becaue it may contain the residul oil or damaged electrical components. 2. Fues Risk: Buning materials can release smokes, incluing chemicals contained in plastics, electrical cables and other tripod components. 3. Risk of Electric Shock: Electrical cords damaged by fire may present a risk of electric shock if the appliance is pluged in. 00000000Health Risk: Smoke smoke released from the buning 000000000000000000000000001d bruh can be related to health, especially if inhaled. It is therefore strongly recommended not to keep a buned-ou of the. It is best to decommission it by placing it in a fire-resistant plastic bag and disposing of it in accordance with local reguations on hazardou waste disposal. Here's why I didn't keep it, I have a child, I 000000000000000000000000001't take any risks, secondly, the 300000000 of the field's price was still guranteed. Please reconsider you decision. Aleuis;2000000003-00000000-20000000000000000:0000000000000000:300000000 vouher offered E-mail sent to inform the cutomer abou the procedue. Can't close the case becaue of a uuucase with a 000000000000000000000000001O#that we don't need anymore (defective produt not available for retun);2000000003-00000000-20000000000000000:3000000001:Informed the BG that the sample is not available, advised to close the PR.; 2000000003-00000000-20000000000000000:30000000051: Task done to Eu Safety BL Cutomer answered that "The tripod was thrown away directly I didn't want a fire risk at home with the children" Bu she sent the invoice that proof that the produt is in warranty. We offer a 300000000 vouher and close the case? Or what do youadvice? 2000000003-00000000-20000000000000000:12:18:Sumitted PR#3863695 AYS Created REG ID 300000000 SBE;2000000003-00000000-19 12:25:00000000Hello, here are the responses to the form as well as the puchase invoice. The toothbruh was thrown away directly, I didn't want a fire risk at home with the children. Serly - How long have youhad it? A little over a year - How often did youue it? Twice a day - How long were youuing it when the incident occured? She was in charge. - Which room were youuing it in when the incident occured? Bathroom - What was the resut? The toothbruh starts to smoke and the flame reach below the head. - Have youheard any strange noises? An electric noise - What was the condition of the device after the incident? Bunt degraded - In what environment was the device ued? In a houe with normal temperatues - Have youcharged the device for 200000000ous after youfirst received it? No - How many times have youreloaded it? I don't know maybe once every t000000000000000000000000001weeks - Do youalways kept it on the charger? No - Did youue the original charger? Yes - Did youue adapters? No - Was the charger pluged into the mains or did youue an eutension cord? No;2000000003-00000000-19 00000000:35:00000000 New Safety uuucase Not available via Phone E-mail sent for QN and additional Photos. | 2023-09-19 16:31 | No
854 | Yes | Called | Called | BuQ&R Feedback: No Reason: Corrective Maintenance. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. SAP Transfer: Case Transferred to SAP. Safety Qustions: Case: 0000000022293165 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Na Created By: ALEuANDRE PEREIRA Last udated by: ALEuANDRE PEREIRA. 1. Was the equpment in clinical ue when the problem was detected? Yes 2. Was the patient or uer injued? In 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? In Problem Reported by Cutomer: u-ray is not available, it shows a message of ful HD, however the list only has 2 stuies. Euternal Remarks: 29.00000000.2000000003 1:59:00000000 PM POLARIS (POLARIS) =========== Roberta Vieira - 00000000000000005736689100000000300000000000000000000000097 ======================== Client reports that the HD is ful. Foud IPPC POST uLL in the log. Cleaned the IPPC and The BIOS battery. After this intervention no futher uLL occured of IPPC POST in the log duing the visit. Equpment released for normal ue. Ca Only the problem occus again, we asked the client to open a new The new 000000000000000000000000001k. The service was finished according to the Qulity Policy of the u. Tests and Calibrated Tools in the PA Tool. SAP Transfer: Dispatched for Onsite Suport. T2 Activities:Problem Description: Problem description by engineer: Image store is ful Trouleshooting Action: What steps are taken duing the system diagnosis: Cutomer Guemberg retuned the call and made the equpment available for remote connection. Recreated the equpment's image store. They released 10000000000000000 images. Equpment was released for clinical ue. Local visit is necessary du to the behavior of the logs and recent history of the problem. Checking the equpment's logs shows that there are daily occurences of failues in the IPPC POST. Checked throuh the RSW that there are many occurences of Image disk ful in the last 3 months. uLL# and/or description of uLL Free disk space low. OneEMS internal remarks: Problem description by engineer: Image store is ful What steps are taken duing the system diagnosis: Cutomer Guemberg retuned the call and made the equpment available for remote connection. Recreated the equpment's image store. They released 10000000000000000 images. Equpment was released for clinical ue. Local visit is necessary du to the behavior of the logs and recent history of the problem. Checking the equpment's logs shows that there are daily occurences of failues in the IPPC POST. Checked throuh the RSW that there are many occurences of Image disk ful in the last 3 months. uLL# and/or description of uLL Free disk space low. Repair Action "Internal": FSE: perform a complete internal cleaning of the IPPC (connections, remove dut, etc, uing contact cleaner). Replaced the BIOS battery. Perform a ful formatting of the 3 HDs uing the PC diagnostic Tool. Reinstall the OS+App of the IPPC. In case there is a new failue, replace the image IPPC+HDs. Planner: Schedue 00000000ous of inoperable equpment for 1 Euuuenced FSE. Parts requred: 1 Part: 1, uu 0000000080000000050000000023. Description: IP PC Revised_II, Quntity: 1, Priority: 00000000 Estimated Job duation: 00000000 Requred FSE level: Euuuenced Eutra Internal comment field: FSE Leonardo Reis Trouleshooting Action: Visit by FSE Schedued with cutomer: yes Schedued date / time: 29/00000000 from 00000000:00000000AM to 00000000:00000000PM Date schedued at cutomer's requst? No Altered priority SAP/1EMS? No, schedue uavailability of FSE Estimated length: 00000000:00000000AM Necessary stopped machine? Yes Performed smart schedue? Yes Name of the FSE: Roberta ----------------------------------------------------------------------------- | 2023-09-19 17:06 | No
855 | Yes | C-arm is not moving to 0degrees, stops at 0degrees | C-arm is not moving to 0degrees, stops at 0 degrees | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000690000000018-Sep-2000000003 Williams, Cory The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: In ue Eupected and ueupected Resuts: frontal stand propeller movement not moving to 900000000degrees uer Impact: Patient Impact: None SE Resoluion : Performed geometry calibrations on frontal stand. Propeller movement moving to 900000000degrees ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Williams, Cory with cutomer qustions on rotational movement for Poly G2 stand. Provided info from u IFueutended manul. Also followed u with Clinical Apps spec Patty Jeffords to confirm. ----------------------------------------------------------------------------- | 2023-09-19 17:16 | No
856 | Yes | Battery issue | Battery issue | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 16000000000000000000000000Battery issu Safety Qustions : Case: 0000000022226969 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. 發現問題時，該裝置是否正用於臨床用途？ 否 2. 是否有使用者或病患受傷？ 否 3. 如果裝置具備警報/警告功能，發現問題時，該裝置是否發揮應有的警報/警告功能？ 不適用 00000000是否為到貨不良？ 不適用 Internal Remarks : 19.00000000.2000000003 00000000:52:15 uC (85200000000000000007) $Q7: (uer report the battery issu ) $Q8: (Checked the battery failed ) $Q9: (replaced the battery) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Euternal Remarks : 19.00000000.2000000003 00000000:52:12 uC (85200000000000000007) $Q7: (uer report the battery issu ) $Q8: (Checked the battery failed ) $Q9: (replaced the battery) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) ----------------------------------------------------------------------------- | 2023-09-19 17:31 | No
857 | Yes | Geoipc problem | Geoipc problem | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000622 19-Sep-2000000003 seokhyeon jeong The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: RT Device ue: Rarely occured Eupected and ueupected Resuts: when system boot u, geometry uLLoccured uer Impact: n/a Patient Impact: n/a SE R7.2.16 Tools & Equpment : MuTIMETER,200000000000000001-2000000000000000000000000:00000000 Resoluion : Checked uLLlogs. ->Geoipc connection problem Checked geoipc statu. Replaced geoipc cmos battery. Performed system fuction tests. Passed ----------------------------------------------------------------------------- | 2023-09-19 17:36 | No
858 | Yes | Cine is not happening | 0e is not happening | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008000000001-Sep-2000000003 Pandey, Rajesh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Cine is not happening intermediately . Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer : Checked the system and foud system is not allowing to log in service mode.To resolve the cutomer issu need reload ghost/ system software. Resoluion : Reloaded the ghost backu and done requred configuation.checked all fuctions,created new ghostbacku.handed over the system in good 000000000000000000000000001king condition Tools & Equpment : MuTIMETER,200000000000000002-17 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005271 Kuar, Atu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Cine is not happening intermediately . Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer : Cine is not happening intermediately and Movement is also struting Internal Comments : uLL# and/or description of uLL: uable to check logs as cutomer did not saved on PRS and need fse Suport Repair Action : uable to check logs as cutomer did not saved on PRS and need fse Suport Follow u Requred : uLL# and/or description of uLL: uable to check logs as cutomer did not saved on PRS and need fse Suport Repair Action : uable to check logs as cutomer did not saved on PRS and need fse Suport What are the neut Trouleshooting steps : FSE to visit and check logs and disgnose Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-09-19 17:41 | No
859 | Yes | Volume is low (Split from 00 | Volume is low (Split from 000) |  | 2023-09-19 17:55 | No
860 | Yes | Follow-up to case 000 | Follow-up to case 00000 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000119 Seva, Romain The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Requst for parts uder the contract. uer Impact: No. Patient Impact: No. SE / Diagnostic performed by Engineer: Futher to the replacement of the front cover, the shock buton as well as the stickers of the front are missing. Resoluion: Sending the parts uder the contract. Delivery to the biomedical department. Follow u Requred: Sending parts uder the contract. Delivery to the biomedical department. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008800000000 Parts Consued: 1.0000000000000000000000000000000056000000009900000000 DFM100000000 ASSY-Label Set-French 1.0000000000000000000000000000000056000000009351, DFM100000000 Silicon 8 Key 1.0000000000000000000000000000000056000000009361, DFM100000000 Silicon 3 Key 1.0000000000000000000000000000000056000000009300000000 DFM100000000 Silicon 2 Key ----------------------------------------------------------------------------- | 2023-09-19 17:56 | No
861 | Yes | Nie można wykonać skanu rotacyjnego | Nie można wykonać skanu rotacyjnego | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000Michal Sadowski The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: electrophysiology. Eupected and ueupected Resuts: The rotational scan cannot be performed. uer Impact: no. Patient Impact: no. SE 8.1.17. Resoluion: The uer reports a problem: Move stand into head side position. Information was passed on that it is necessary to move the arm to the head position. The problem was solved. ----------------------------------------------------------------------------- | 2023-09-19 18:26 | No
862 | Yes | 0CDN0ystem is down, screens are black, both euam room and the console | 00CDN0system is down, screens are black, both euam room and the console | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 200000000Sep-2000000003 Hamilton Bill The system meets the specification for the performed service and is retuned to ue Resoluion : I replaced the Host PC per the Service Manul. I made a ghost backu. All system fuctions were tested for proper operation - okay. Problem Description by Engineer : Problem description by engineer : system is down, screens are black, both euam room and the console When was the first occurence? Within... : 19-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: system is down, screens are black, both euam room and the console uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : The system 000000000000000000000000001ld not boot becaue the host pc did not tun on. I had Mark press the PC on buton. The system then powered on and 000000000000000000000000001ked. The host PC is bad and needs to be replaced. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Ketcherside Jeff The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: uable to get lod from the system Repair Action : uable to get the lod from the system Follow u Requred : uLL# and/or description of uLL: uable to get lod from the system Repair Action : uable to get the lod from the system Technician Name : Ketcherside Jeff Problem Description by Engineer : Problem description by engineer : system is down, screens are black, both euam room and the console When was the first occurence? Within... : 19-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Miscellaneou Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: system is down, screens are black, both euam room and the console uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : cutomer states system is down, screens are black, both euam room and the console ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Hamilton Bill with parts ID and recommended he ugrade the HOST PC to 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD. ----------------------------------------------------------------------------- | 2023-09-19 18:31 | No
863 | Yes | Collision between the C-arm of the device and the patient | Collision between the C-arm of the device and the patient | 000000000000000000000000001O##: 00000000000000000000000000100000000115389 13-Oct-2000000003 FLICEK, COLE The system meets the specification for the performed service and is retuned to ue Resoluion : Closing case after response from Bu Complete issu as described by the cutomer oulined below. Collimator alignment issu aroud 2000000003-00000000-19 00000000:00000000 Rebooted system Case started aroud 00000000:300000000 Table Floating issu early in the case Patient collision late in the case Tech was looking at the "monitor" Another tech noticed the collision and told the first tech to stop Collision occured by team hudle at 00000000:300000000 Complete issu as described by the system logs oulined below. Detector Collision 2000000003-00000000-19 00000000:59:00000000 Anguate frontal stand movement Detector Force Sensor Collision 2000000003-00000000-19 00000000:59:200000000 Change patient suport height 2000000003-00000000-19 00000000:59:300000000991 Application uLL Motor Assembly uLL Table Height 2000000003-00000000-19 00000000:59:31:000000006, Actul I: 1.700000000 Set I: 00000000 Detector Collision 2000000003-00000000-19 00000000:59:31:855 Some table movement not available 2000000003-00000000-19 00000000:59:36 to 2000000003-00000000-19 00000000:52:100000000 Some table movement not available 2000000003-00000000-19 00000000:00000000 Detector Force sensor collision 2000000003-00000000-19 00000000:00000000:26 Detector Force sensor collision 2000000003-00000000-19 00000000:00000000:26 Detector Force sensor collision 2000000003-00000000-19 00000000:00000000:27 Some table movements not available 2000000003-00000000-19 00000000:12 to 00000000:32 Detector Force sensor collision 2000000003-00000000-19 00000000:51:52 Detector Force sensor collision 2000000003-00000000-19 00000000:51:56 Detector Force sensor collision 2000000003-00000000-19 00000000:51:59 Detector Force sensor collision 2000000003-00000000-19 00000000:52:00000000 Detector Force sensor collision 2000000003-00000000-19 00000000:52:36 Of note, no collision was sensed duing the timeframe given by the cutomer. FSE complete timeline given below. Arrive on site at 2000000003-00000000-19 ~15:15 Tested collision behavior and noted that the only easily stopped motion was the front part of the detector cover. This was the only portion of the bodygurd sensor that was able to be read duing the calibration Anguation and rotation were stopped by curent sensing. This was difficut to stop Detector shift movement was easy to stop Replaced Height potentiometer and Height AMC drive 2000000003-00000000-21 Calibrated all of the systems geometry Coud not calibrate body gurd sensor Ordered Bodygurd sensor 2000000003-00000000-21 Jacob Schonecker installed bodygurd sensor 2000000003-00000000-22 ~11:00000000 - 13:00000000 Body gurd calibration still 000000000000000000000000001ld not pass I suported Jake on the phone and instruted him to test the systems response when the each connection to the BIB was removed. Each removed connector triggered uLL that were not present before their removal. I instruted Jake to grab the breakou board from CV2 to test the signals coming ou of BIB:u3 (CAN connection) u3:8 & u3:7 had 300000000 indicating power being suplied to bodygurd sensor u3: 3, 000000005, & 6 all had ~1.3V indicating the sensors on the new bodygurd sensor were good. Jake also confirmed that moving his hand close to sensor 000000000000000000000000001ld caue the voltages to drop to 00000000. Each pin corresponds to a prouimity sensor Ordered BIB 2000000003-00000000-22 Collected Bodygurd Interface Bou 2000000003-00000000-28 ~15:300000000 Replaced Bodygurd Interface Bou and tuned over to cutomer 2000000003-00000000-29 ~00000000:00000000 I learned that the patient passed away John Swain called me and pointed ou that their were uer gudance messages jut before the table uLL 2000000003-00000000-29 00000000:00000000 I had previouly missed these uLL becaue I was looking at the logs from abou 2000000003-00000000-19 00000000:00000000 to 2000000003-00000000-19 00000000:300000000Problem Description by Engineer : Problem description by engineer : spoke to Jakethe staff said they lost power to the table duing ue When was the first occurence? Within... : 19-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Patient suport Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: table free floating Patient Impact: table free floating uu00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006700000000 29-Sep-2000000003 FLICEK, COLE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : spoke to Jakethe staff said they lost power to the table duing ue When was the first occurence? Within... : 19-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Patient suport Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: table free floating Patient Impact: table free floating uu00000000Resoluion : Replaced bodygurd interface bou and performed bodygurd sensor calibration. Follow u Requred : 000000000000000000000000001for escalation Parts Consued : 1.00000000000000000000000000000000800000000676000000002,SENSOR FD200000000 1.00000000000000000000000000000000200000000000000005752,C-arc Bodygurd Interface Bou Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000510000000021-Sep-2000000003 FLICEK, COLE The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace bodygurd sensor. Problem Description by Engineer : Problem description by engineer : spoke to Jakethe staff said they lost power to the table duing ue When was the first occurence? Within... : 19-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Patient suport Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: table free floating Patient Impact: table free floating uu00000000Diagnostic performed by Engineer : Replaced height AMC drive and height pot. Duing the calibrations, the bodygurd sensor 000000000000000000000000001ld not pass calibrations. Parts Consued : 1.00000000000000000000000000000000230000000025963,AD7(NT) HEIGTH POTMETER ASSY 1.00000000000000000000000000000000800000000300000000200000000AMC Ecat drive DPPH00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 200000000Sep-2000000003 FLICEK, COLE The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace height AMC drive, height potentiometer, & perform all geo calibrations Problem Description by Engineer : Problem description by engineer : spoke to Jakethe staff said they lost power to the table duing ue When was the first occurence? Within... : 19-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Patient suport Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: table free floating Patient Impact: table free floating uu00000000Diagnostic performed by Engineer : FSE evaluted log files. Failue of height AMC drive triggered a fatal uLLwhich caued the ethercat commuication to be cu to all of the table geometry. The cutomer continud to ue the room as they were in the middle of a procedue. At one point, a tech was moving the C-arm while watching the angle valus on the FleuVision monitor and was not aware of collision with patient. She was informed of collision by another tech. Logs indicate a collision via flat detector shift, c-arm propeller, and c-arm anguation directions over a period of 300000000seconds. Patient sufered a hematoma from the collision. FSE will replace height AMC, height potentiometer, and perform a ful system geometry calibration to ensue collision model accuacy going forward. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: geometry movemnts partially aviable Repair Action : Application uLL Motor assembly uLL[auis=TableHeight] AD7(NT) HEIGTH POTMETER ASSY 00000000230000000025963 Height actutor assy 00000000800000000000000009000000001 [AMC Ecat drive DPPH00000000],[NODE INTERFACE uIT] [ 00000000800000000300000000200000000] [ 000000002300000000352900000000 NODE INTERFACE uIT 000000002300000000352900000000e GIB switches off the Servo power for the Table, Frontalstand or MCS power tap becaue a safety related uLLdid occu, either duing the POST or duing normal operation. Please read ou the ful logline for the reason and the location. Soluion 1 If uLLoccus duing POST, check the safety brake and cluch circuts and mechanics for the auis mentioned in the logline description. Soluion 2 If uLLoccus duing normal operation, the uLLcaue is given in the logline description. Note that the first uLLlogged needs to be euamined. Other uLLlogging can occu as collateral logging becaue the servo power is switched off. Soluion 3 Possible corrective actions: If the logline contains a statu field, a bit pattern is given defining what the uLLcaue is. Statu: 000000001: SafetyBrakeRelease/Solenoid command uLL000000002: SafetyBrakeRelease/Solenoid feedback uLL0000000000000000Position following uLLoccured 000000008: Position slip betw. encoder and potmeter 00000000100000000 CluchRelease command uLL00000000200000000 CluchRelease feedback uLL0000000000000000 CluchRelease Isense uLL00000000800000000 Drive containing this auis reports invalid data mutiple times (check also for 63000000001131 events). If download is sucessfu then check the cabling or hardware circutry for the reported faut and corresponding motion drive or NIu biomed states there was a patient incident with stand Problem Description by Engineer : Problem description by engineer : spoke to Jakethe staff said they lost power to the table duing ue When was the first occurence? Within... : 19-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Patient suport Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: table free floating Patient Impact: table free floating uu00000000Diagnostic performed by Engineer : speaking with Jake ,reviewing log system shown table position uLLon system table movement lost duing ue. ful shu down and restart cleared problem cutomer 000000000000000000000000001ld like system looked at before retuned to ue Follow u Requred : uLL# and/or description of uLL: geometry movemnts partially aviable Repair Action : Application uLL Motor assembly uLL[auis=TableHeight] AD7(NT) HEIGTH POTMETER ASSY 00000000230000000025963 Height actutor assy 00000000800000000000000009000000001 [AMC Ecat drive DPPH00000000],[NODE INTERFACE uIT] [ 00000000800000000300000000200000000] [ 000000002300000000352900000000 NODE INTERFACE uIT 000000002300000000352900000000e GIB switches off the Servo power for the Table, Frontalstand or MCS power tap becaue a safety related uLLdid occu, either duing the POST or duing normal operation. Please read ou the ful logline for the reason and the location. Soluion 1 If uLLoccus duing POST, check the safety brake and cluch circuts and mechanics for the auis mentioned in the logline description. Soluion 2 If uLLoccus duing normal operation, the uLLcaue is given in the logline description. Note that the first uLLlogged needs to be euamined. Other uLLlogging can occu as collateral logging becaue the servo power is switched off. Soluion 3 Possible corrective actions: If the logline contains a statu field, a bit pattern is given defining what the uLLcaue is. Statu: 000000001: SafetyBrakeRelease/Solenoid command uLL000000002: SafetyBrakeRelease/Solenoid feedback uLL0000000000000000Position following uLLoccured 000000008: Position slip betw. encoder and potmeter 00000000100000000 CluchRelease command uLL00000000200000000 CluchRelease feedback uLL0000000000000000 CluchRelease Isense uLL00000000800000000 Drive containing this auis reports invalid data mutiple times (check also for 63000000001131 events). If download is sucessfu then check the cabling or hardware circutry for the reported faut and corresponding motion drive or NIu biomed states there was a patient incident with stand What are the neut Trouleshooting steps : check operation and movements of table and stand check table height movements and 000000000000000000000000001king positions of system see inotes Technician Name : Sulivan Scott ----------------------------------------------------------------------------- Internal Remarks : Swain - Reviewed daily log file and foud no evidence of bodygurd issus following BIB replacement and calibrations. No futher cases opened on system since this event and system logs show system is in ue. Closing escalation. Attached daily log to case. Internal Remarks : Changed case priority to intermittent problem since resoluion was foud with BIB replacement and system is ready for ue. Internal Remarks : Log file indicates failing height AMC. Recommended replacing height pot as we and completing ful geometry calibrations to re-establish collision model . ----------------------------------------------------------------------------- | 2023-09-19 19:06 | Yes
864 | Yes | 0u FD0startup issue | 0/u FD0startup issue | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022290000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Qingfeng Gu Last udated by: Qingfeng Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: uable to boot normally Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: The screen goes completely black after tuning on the compuer Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended that engineers come to the site to check the host power suply and motherboard operation.2. Reinstall system software SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 22.00000000.2000000003 15:32:300000000uC (8600000000900000000) $Q1: () $Q2: () $Q3: () $Q00000000() $Q7: (Confirmed reported problem&IP PC can't power on) $Q8: (Performed visul inspection, F23 fue problem) $Q9: (change the F23 fue) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name (NA) S/N (NA) Neut Calibration date (NA) ^Curent Tue installed date: NA ^Curent Tue uage: NA FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: The screen sudenly goes black and restarting fails. Euternal Remarks: 22.00000000.2000000003 15:00000000:00000000 uC (8600000000900000000) Service content: normal after replacing spare fue Service resut: The machine is normal SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-19 19:06 | No
865 | Yes | No power, device will not turn on. | No power, device will not turn on. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: uit w/not power on. No LED's illuinated with power connected. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : foud no power ou of the power suply. Good power to the power suply. Provided parts ID for the power suply and discused installation per the manul. 00000000561512591. Duing trouleshooting foud the battery had only 1800000000V. Advised to replace the battery, they have them in stock and will replace. ----------------------------------------------------------------------------- | 2023-09-19 19:41 | No
866 | Yes | Machine is down with sedated patient on the table. | Machine is down with sedated patient on the table. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : cannot get uit to boot at all When was the first occurence? Within... : 19-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: room down cannot get to boot uer Impact: room down Patient Impact: room down SE 8.2.17.00000000Diagnostic performed by Engineer : spoke to caller. on site needed type of contract shows assist, suposed to be a pool for service onsite Internal Comments : uLL# and/or description of uLL: cannot get system to boot screen black Repair Action : check power and PCs for booting Internal Comments : created on site and ##d qute team to release ----------------------------------------------------------------------------- | 2023-09-19 19:46 | No
867 | Yes | Check Vent: Backup Alarm Failed | Check Vent: Backup Alarm Failed |  | 2023-09-19 19:55 | No
868 | Yes | Power supply of the monitor failed and supposedly gave no alarm. No Trouble Found-Evaluation Confirmed Device Working. | Power supply of the monitor failed and supposedly gave no alarm. No Trouble Found-Evaluation Confirmed Device Working. | 000000000000000000000000001O##: 0000000000000000000000000010000000016000000005 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- Internal Remarks: uLL Partners from accout grou 0000000000000000 are not permitted for partner fuction BP --> bill to: 0000000090000000052200000000nternal Remarks: NO CONTRACT!!! Internal Remarks: No decon Bill to: / Tu0000000900000000000000003 German Heart Center Muich Suject IV B - Medical Technology Lazarettstr. 36 80000000036 Muich Germany Cutomer comment: Via ##sender felgentraebe00hm.mhn.de Dear Sir or Madam, I am not in the 18.00000000.2000000003 22.00000000.2000000003 #2000000003. In ugent cases, please contact# medizintechnik00hm.mhn.de ----------------------------------------------------------------------------- | 2023-09-19 19:56 | No
869 | Yes | 0/u FD0Unable to power on | 0R/u FD0Unable to power on | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002229000000008 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: mengdi uu Last udated by: mengdi uu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: uable to boot Internal Remarks: 21.00000000.2000000003 10000000000000000:200000000uC (8600000000600000000 "RDI Accuacy: Yes" $Q7: (Confirmed reported problem. Can't bootu.) $Q8: (Replace the dcps, the device retuned to normal.+The DCPS problem.) $Q9: (Replace the DCPS. uuof the part: 00000000800000000387582.) $Q100000000 (Test and inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (FLuE 287C) S/N (23350000000017) Neut Calibration date (200000000000000007/100000000 SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 21.00000000.2000000003 10000000000000000:17 uC (8600000000600000000 Service content (requred): Replacement of M cabinet DCPS Service resut (requred): The device is runing normally Other important matters: None ----------------------------------------------------------------------------- | 2023-09-19 20:06 | No
870 | Yes | PSAPSE_Legal_Death_Telemety on Wrd 0 GFE0Sent 000-Escalated to Paul 00 | PSAPSE_Legal_Death_Telemety on Wrd 0, GFE0 Sent 00/00Escalated to Paul 00 |  | 2023-09-19 20:10 | Yes
871 | Yes | Broken ground wire during FCO installation | Broken ground wire during FCO installation | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 22-Sep-2000000003 Zajicek, Gil The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561511600000000,RP-Cable, Gas Oulet Port To Gnd, V600000000 1.0000000000000000000000000000000056153600000000,RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000Problem Description by Engineer : Cutomer Fuction/Role: Device ue: not in ue Eupected and ueupected Resuts: uer Impact: none Patient Impact: none SE 3.00000000 LCD groud and gas oulet groud cables were damaged. Resoluion : FSE replaced the LCD cable and the gas oulet groud cable. FSE performed all tests as requred by the service manul for these repairs. All tests passed per service manul specifications. Tools & Equpment : MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 STATIC CONTROL DEVICE,2000000005-00000000-0000000000000000:00000000:00000000 FLOW METER,200000000000000006-30000000000000000:00000000:00000000 FLOW METER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-19 20:11 | No
872 | Yes | PSAPSE_Legal_Death_Telemety on Wrd 0 I have report but / GFE0Sent 000 - Escalated to paul 00 | PSAPSE_Legal_Death_Telemety on Wrd 0, I have report but / GFE0 Sent 00/00- Escalated to paul 00 |  | 2023-09-19 20:12 | Yes
873 | Yes | uPro: "This unit no longer holds a charge. Have tried multiple different battery packs and cords. Battery... | TempuPro: "This unit no longer holds a charge. Have tried multiple different battery packs and cords. Battery... | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-1000000006-R Eupected and ueupected Resuts: Confirmed complaint uer Impact: uable to ue Patient Impact: none SE v00000000.28 Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damaged and the center pin is no longer there. To correct the issu we will replace the Front Case Assembly, which the charging port is part of that assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000588 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to bench Follow u Requred : EMS site manager called to eport: Tempu Pro: "This uit no longer holds a charge. Have tried mutiple different battery packs and cords. Battery power never goes u, only declines." Requst bench repair Internal Comments : Cutomer made aware a Bench will need a PO to start repair. Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS site manager Device ue:Not in Clinical ue Eupected and ueupected Resuts: "This uit no longer holds a charge. Have tried mutiple different battery packs and cords. Battery power never goes u, only declines." uer Impact:uable to ue device on battery power util repair is made Patient Impact: SEuu Problem Description/ueupected Resuts: "This uit no longer holds a charge. Have tried mutiple different battery packs and cords. Battery power never goes u, only declines." Diagnostic performed by Engineer : Problem Description: "This uit no longer holds a charge. Have tried mutiple different battery packs and cords. Battery power never goes u, only declines." Cutomer callback Tempu Pro: "This uit no longer holds a charge. Have tried mutiple different battery packs and cords. Battery power never goes u, only declines." Form ued: Bench Trouleshooting Steps Taken: EMS site manager called to report Tempu Pro: "This uit no longer holds a charge. Have tried mutiple different battery packs and cords. Battery power never goes u, only declines." Created Bench 000000000000000000000000001 ##d cutomer copy of bench repair form. uable to reach cutomer by phone. ##d cutomer to ask if a loner is needed and to inform them a PO will be needed to begin repair. ----------------------------------------------------------------------------- Internal Remarks : Shipping Address A-Tec Ambuance 2125 Point Blvd Ste 200000000, Elgin, Illinois – 60000000023 Internal Remarks : Please ue address on bench form coud not find on service mau ----------------------------------------------------------------------------- | 2023-09-19 20:31 | No
874 | Yes | system is not working, not getting video. May be hard drive issue. | system is not working, not getting video. May be hard drive issue. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000000000000000000000ep-2000000003 Clement Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: GE FSE Device ue: u uuuFD100000000 Eupected and ueupected Resuts: System will not boot u uer Impact: none Patient Impact: none SE 000Resoluion : Reconfigued and tested system after other engineer replaced host pc and reloaded system. System is operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000 22-Sep-2000000003 Lun Thy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Reload the software Problem Description by Engineer : Cutomer Fuction/Role: GE FSE Device ue: u uuuFD100000000 Eupected and ueupected Resuts: System will not boot u uer Impact: none Patient Impact: none SE 000Diagnostic performed by Engineer : Configued the Local IP and net000000000000000000000000001k connection, Configued the PACs, Configued the WLM, and Edited the Data Base, Performed the batch verification. Note: System coud not recognize the third-party Monitor. and 000000000000000000000000001ld display on the monitor. Coud not performed the adjutment on the calibration. Follow activity will performed. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000002 21-Sep-2000000003 Lun Thy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Configue the system and Calibration Problem Description by Engineer : Cutomer Fuction/Role: GE FSE Device ue: u uuuFD100000000 Eupected and ueupected Resuts: System will not boot u uer Impact: none Patient Impact: none SE 000Diagnostic performed by Engineer : Performed the installation of the software, Foud the IP PC drive had not light on the OS drive and Data was not seat properly. Had to reseat the hard drive in the IP PC. Performed the Ghost restored. System failed. Had to reload the software again. Note: Will configued the IP, 000000000000000000000000001k list and Pac, and Data base are configuing tomorrow and performed calibration. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000550000000000000000Sep-2000000003 Lun Thy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: GE FSE Device ue: u uuuFD 100000000 Eupected and ueupected Resuts: System not 000000000000000000000000001king and failed to boot u uer Impact: none Patient Impact: none SE 8.1.17 Diagnostic performed by Engineer : Installed new Host PC and Hard drive. Reloaded OS software, reloaded Application Software 8.1.16, Reloaded Application Software 8.1.17 (1of2) and Reloaded Application Software 8.1.17 (2of2). Note: No Application Software 000 Have to get this Application Software. Had to bun this software on CD. There is not system back on site, no Data base back u. Needed IT for the IP Address for the system, PAC, WLM. Additional 000000000000000000000000001k need, please see 000000000000000000000000001O# 00000000000000007000000002. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000680000000019-Sep-2000000003 Lun Thy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace Host PC and Hard drive Problem Description by Engineer : Cutomer Fuction/Role: GE FSE Device ue: u uuuFD 100000000 Eupected and ueupected Resuts: System not 000000000000000000000000001king and failed to boot u uer Impact: none Patient Impact: none SE 8.1.17 Diagnostic performed by Engineer : uit stated: "System not 000000000000000000000000001king and not getting Video." After trouleshooting the uit, system failed to boot u and coud not get into the fiedl service frame000000000000000000000000001k to look at the log file. Host PC is failing. Recommend the host pc and hard drive. ----------------------------------------------------------------------------- Internal Remarks : Assist Thy in loading SW Internal Remarks : Host drive bay will not power u. Approved for replacement 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD and 0000000080000000050000000000000000- Hard disk spare part Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mike Dagley IP & Issu: 500000000900000000 & wanting fse to come ou on weekend if possible Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: verified contact info , created AA and set eupectations Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: Cutomer Call Back Caller: John 000000000000000000000000001dard IP & Issu: 500000000900000000 & fse is a no show Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: verified contact info ,created AA and set eupectations Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : FSE is 000000000000000000000000001dering which ID the TSM's shoud be set to. -Told FSE typically the euam room is set to ID 1 and control room is set to ID 2, bu it technically doesn’t matter. Internal Remarks : Provided FSE link to Software repair manul and NA IGT EPu Database downloads for new host replacement Internal Remarks : Interaction Type: New Case Caller: Cris IP & Issu: 500000000900000000 / system is not 000000000000000000000000001king, not getting video. May be hard drvie issu. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Create Case and Page QA Neut Step Action: Awaiting Qute Verbal PO Nuber: 00000000110000000000000000 Verbal PO Approver: Verbal PO CAP: 50000000000000000----------------------------------------------------------------------------- | 2023-09-19 20:36 | No
875 | Yes | Audio failure. Can Not Duplicate-No Trouble Found. No fault found. | Audio failure. Can Not Duplicate-No Trouble Found. No fault found. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 October 13, 2000000003 Ruén Álvarez Villa The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: The described failue cannot be reprodued, the equpment is reset and operation tests are performed. No other problems are observed, the equpment is operational after the check-u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003700000000hazwan Hueidi Sordo The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: SueSigns VS2+ NBP/SPO2/Wireless CN212123500000000 Auio failue Assign fSE Check continuty of speaker coil with mutimeter In case of needing material, order: 1 u 0000000056000000003611 Speaker 1 u 0000000056000000000000000051 Main board, VS2+ uN+3 C2 nusing control # OT: WEB00000000000000006363 Follow u Requred: SueSigns VS2+ NBP/SPO2/Wireless CN212123500000000 Auio failue Assign fSE Check continuty of speaker coil with mutimeter In case of needing material, order: 1 u 0000000056000000003611 Speaker 1 u 0000000056000000000000000051 Main board, VS2+ uN+3 C2 nusing control # OT: WEB00000000000000006363 ----------------------------------------------------------------------------- | 2023-09-19 20:36 | No
876 | Yes | Device, System or Product HW/The equipment has come off the cart and has fallen to the floor/Mechanical | Device, System or ProduuHW/The equipment has come off the cart and has fallen to the floor/Mechanical | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 29-Sep-2000000003 Ruén Álvarez Villa The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We talk to José Lus García 666370000000065 from Electromedicine. Device ue: Not in clinical ue. Eupected and ueupected Resuts: The equpment got loose from the cart and has fallen to the floor. uer Impact: None Patient Impact: None SE Not apply. Tools & Equpment: PRESSuE CALIBRATION TOOL, 200000000000000002-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000002-00000000 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000056000000009981,SCRW Kit, Roll Stand Mout Resoluion: The equpment is checked and the anchoring system is adjuted to the rolling cart. Observed that the equpment antenna is broken and replaced with a new one. After corrective maintenance, the equpment is operational and meets the manuactuer's specifications. The necessary verifications have been performed to certify that the equpment is back to the operating conditions prior to the failue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Rau Cordero Calderón The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: We talk to José Lus García 666370000000065 from Electromedicine and he says that the monitor is loose from the cart and has fallen to the groud. It seems that it has not been damaged since they are still trying with the monitor and it is in the cart. We ask him if the cart fastening is damaged. he says that the toru screw is loose. Assign fSE. Order part: 0000000056000000009981 - SCRW Kit, Roll Stand Mout uN+7 C1 nusing control. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We talk to José Lus García 666370000000065 from Electromedicine. Device ue: Not in clinical ue. Eupected and ueupected Resuts: The equpment got loose from the cart and has fallen to the floor. uer Impact: None Patient Impact: None SE Not apply. Diagnostic performed by Engineer: We talk to José Lus García 666370000000065 from Electromedicine and he says that the monitor came off the cart and has fallen to the groud. It seems that it has not been damaged since they are still trying with the monitor and it is in the cart. We ask him if the cart fastening is damaged. he says that the toru screw is loose. Assign fSE. Order part: 0000000056000000009981 - SCRW Kit, Roll Stand Mout uN+7 C1 nusing control. ----------------------------------------------------------------------------- | 2023-09-19 20:36 | No
877 | Yes | uLS Manual ¦ Device will not turn on | TempuLS Manual ¦ Device will not turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000238 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Placed a Material 000000000000000000000000001for euchange Tempu LS Follow u Requred : Please Tu Allegiance Mobile Heath Attention: AJ Rose 200000000 E 00000000 St Livingston Tu 77351-200000000 Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: Device will not tun on. uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Device will not tun on. "Jut stop tuning on, it shows it's charging bu it jut 000000000000000000000000001't tun on. Diagnostic performed by Engineer : Problem Description: Device will not tun on. "Jut stop tuning on, it shows it's charging bu it jut 000000000000000000000000001't tun on. Tempu LS: Device will not tun on. "Jut stop tuning on, it shows it's charging bu it jut 000000000000000000000000001't tun on. No one was hut, and wasn't being ued, it jut qut. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 9/19/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Biomed reports Tempu LS: Device will not tun on. "Jut stop tuning on, it shows it's charging bu it jut 000000000000000000000000001't tun on. No one was hut, and wasn't being ued, it jut qut. Cutomer has tried known good power cords an batteries with like resuts. Placed a Material 000000000000000000000000001for euchange Tempu LS Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-09-19 20:41 | No
878 | Yes | machine shut down and wont turn on | machine shut down and wont turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000615 19-Sep-2000000003 Cathey Eric The system meets the specification for the performed service and is retuned to ue Resoluion : system ready for ue Tools & Equpment : MuTIMETER,200000000000000007-31 00000000:00000000:00000000 Diagnostic performed by Engineer : replaced fue in the MDP start u board in the M rack 2 each 7500000000a checked system operation Problem Description by Engineer : Cutomer Fuction/Role: can't tun on the system Device ue: review modue Eupected and ueupected Resuts: System 000000000000000000000000001't tun on uer Impact: na Patient Impact: na SE8.1 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Lindsey Stewart 870000000065887 IP & Issu: 56931900000000 machine shu down and 000000000000000000000000001t tun on Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: advised of eupedite fee for produt code 722000000002 of $16500000000cu accepted created qute Q-00000000130000000037 Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-19 20:46 | No
879 | Yes | QACASE for PCI imported product#000000the device cannot eupose u-ray. | QACASE for P0imported produu0000000_the device cannot eupose u-ray. |  | 2023-09-19 21:12 | No
880 | Yes | uCVu-The control panel swivels and will not lock into place.uMechanical/Control Panel: Resolved-replaced bearings | uCVu-The control panel swivels and will not lock into place.uMechanical/Control Panel: Resolved-replaced bearings | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000391 19-Sep-2000000003 Huner Michael The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced CPA bearing. Problem Description by Engineer : Problem description by engineer : The control panel swivels and will not lock into place, bu the u/down motion 000000000000000000000000001ks properly. When was the first occurence? Within... : 18-Sep-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: The control panel swivels and will not lock into place, bu the u/down motion 000000000000000000000000001ks properly. uer Impact: The system is difficut to move down the hallway. Patient Impact: No futher information is available as per the cutomer SE 7.000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000219 Lang, Barbara The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : The control panel swivels and will not lock into place, bu the u/down motion 000000000000000000000000001ks properly. When was the first occurence? Within... : 18-Sep-2000000003 How often is issu occuring? : Mutiple times a day Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: The control panel swivels and will not lock into place, bu the u/down motion 000000000000000000000000001ks properly. uer Impact: The system is difficut to move down the hallway. Patient Impact: No futher information is available as per the cutomer SE 7.000000005 Diagnostic performed by Engineer : The problem was discovered when the system was being moved from room to room. The uer was not harmed by this, bu the system was difficut to maneuer. The u/down motion of the control panel arm is normal. Internal Comments : uLL# and/or description of uLL: none reported Repair Action : Replace, or repair the swivel solenoids. Recommended Parts: 00000000561669791 SOLENOID STATOR,13V,000000005 STROKE 989600000000730000000021 FCO79500000000500000000 CPA Bearing, EPIQ 989600000000730000000031 FCO79500000000500000000 Adhesive Follow u Requred : uLL# and/or description of uLL: none reported Repair Action : Replace, or repair the swivel solenoids. Recommended Parts: 00000000561669791 SOLENOID STATOR,13V,000000005 STROKE 989600000000730000000021 FCO79500000000500000000 CPA Bearing, EPIQ 989600000000730000000031 FCO79500000000500000000 Adhesive What are the neut Trouleshooting steps : Replace, or repair the swivel solenoids. Recommended Parts: 00000000561669791 SOLENOID STATOR,13V,000000005 STROKE 989600000000730000000021 FCO79500000000500000000 CPA Bearing, EPIQ 989600000000730000000031 FCO79500000000500000000 Adhesive Technician Name : Lang, Barbara ----------------------------------------------------------------------------- | 2023-09-19 21:16 | No
881 | Yes | BROKEN TRANSPORT CART | BROKEN TRANSPORT CART | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : 2 WHEELS ARE MISSING Ou Of Bou Failue? : No Labor: Labor ID : L-1977000000007 000000000000000000000000001k Performed : REATTACHED THE WHEELS AFTER RETHREADING THE WHEEL WELL Diagonsis : REATTACH THE WHEELS TO CART | 2023-09-19 21:26 | No
882 | Yes | Only runs on internal battery | Only runs on internal battery | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000599 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: uit will not ru on AC power. LED's not illuinated on front. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : verified good power to the power suply, no power ou. Advised to replace the power suply. Provided parts ID and discused installation per the manul. 00000000561512591 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: robert IP & Issu: 56629637 internal battery only Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-19 21:51 | No
883 | Yes | LOW LEVEL BATTERY LIQUID LEAK FROM CHASSIS | LOW LEVEL BATTERY LIQUID LEAK FROM CHASSIS | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003500000000FALASCA, IVO The system does not meet the specification for the performed service and is not retuned to ue. Resoluion: Colleagu Marco Biamonte confirms that the battery lost some liqud and that it is no longer 000000000000000000000000001king. Since the spare part can no longer be ordered as the system is not suported, relative letter of end of suport will follow. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Battery with low level of charge and that has lost liqud while invading the chassis uer Impact: no Patient Impact: no SE na ----------------------------------------------------------------------------- | 2023-09-19 21:56 | No
884 | Yes | Speaker malfunction uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Speaker malfunuon uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005300000000dams, Devon The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Get knowledge abou any FCO's that pertain to devices. Need to leave case open while biomed department fuly investigates widespread issu. uer Impact: loss of confidence in u equpment Patient Impact: No impact SE M.0000000000000000 Problem Description/ueupected Resuts: Mu800000000 speaker malf uLLacross 5 ICus with 16 beds each. Diagnostic performed by Engineer : Problem Description: Mu800000000 speaker malf uLLacross 5 ICus with 16 beds each. Mu800000000 (incorrect 000provided) - speaker malf uLL Intermittent for a few weeks. No patient care delays or harm. Form ued: Trouleshooting/Investigation Object: uu Not Object: uu Software Version: M.0000000000000000 Date and Time of Issu: 9/19/2000000003 uer Actions: monitoring patient from bedside or POD Where is the Deviation: uu Trouleshooting Steps Taken: Is there an FCO associated with the Mu800000000 or for a remote alarm. This is a widespread issu. 5 ICus with 16 beds each. Director of ICus is stating this issu. Performed research on u Incenter for FCO's related to Mu800000000 or for a remote alarm issus. No docuents foud that correlate to "speaker malfuction" intermittently popuating. Sent u3, Mu00000000, and Mu800000000 service gude to biomed via SDT. Sent Remote alarm discontinution letter to biomed via case level ## Recommended that biomed take time to try to witness the euact uLLmessage on bedside monitoring equpment and to reply to case ##with photos. Docuentation needed for investigation.Leaving case open util biomed requsts closue. Was PRS ued: No PRS Permission: Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Joseph Daco IP & Issu: speaker malfuction- uLL Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Joseph Daco IP & Issu: speaker malfuction- uLL Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-19 22:11 | No
885 | Yes | Low leak C0uLLmessage. | Low leak C0 uLLmessage. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order a flow sensor assy and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: low leak Co2 rebreathing risk alarm uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and they are getting a low leak Co2 rebreathing risk alarm. While trouleshooting, they are getting a -00000000 slpm on the average air reading on the vent. Part #989800000000617100000000 flow sensor assy, list price $1217.00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kayla Owens IP & Issu: 56630000000057 Low leak C00000000 uLLmessage. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-19 22:21 | No
886 | Yes | System stuck at 0:0 after rebooting system | System stuck at 0:0 after rebooting system | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003800000000200000000Sep-2000000003 Clement Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System stuk at 00000000:00000000 after rebooting system When was the first occurence? Within... : 19-Sep-2000000003 12:25:00000000 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System not booting past 00000000;00000000 uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : Replaced matriu power suply and system is operational again. Parts Consued : 1.0000000000000000000000000000000080000000000000000392,DVI MATRIu SWITCH 16u16 POWER SuPLY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000088000000009-Sep-2000000003 Clement Michael The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : replace DVI video matriu parts Problem Description by Engineer : Problem description by engineer : System stuk at 00000000:00000000 after rebooting system When was the first occurence? Within... : 19-Sep-2000000003 12:25:00000000 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System not booting past 00000000;00000000 uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : foud power suply for DVI Video matriu bad. ordering parts for tomorrow. Tools & Equpment : MuTIMETER,2000000003-11-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000769 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System stuk at 00000000:00000000 after rebooting system When was the first occurence? Within... : 19-Sep-2000000003 12:25:00000000 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: System not booting past 00000000;00000000 uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Juson said they did not see the hemo video on the fleu monitor so they rebooted the system and then the system did not boot past 00000000:00000000. They are trying to reboot 1 more time bu Juson requsted to have FSE onsite to inspect the system. Internal Comments : uLL# and/or description of uLL: n/a Repair Action : ue alt + f duing boot u to log into FSF and check log file for grabber card uLL or commuication uLL with fleu pc. If needed, 00000000800000000119123 - FleuVision 2 PC No HDD, COA and/or 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : ue alt + f duing boot u to log into FSF and check log file for grabber card uLL or commuication uLL with fleu pc. If needed, 00000000800000000119123 - FleuVision 2 PC No HDD, COA and/or 0000000080000000050000000000000000- Hard disk spare part Technician Name : Coffman, Josh ----------------------------------------------------------------------------- | 2023-09-19 22:26 | No
887 | Yes | Legal Team at NHS Lothian contacted and stated a patient claims to have had their hearing damaged | Legal Team at NHS Lothian contaud and stated a patient claims to have had their hearing damaged |  | 2023-09-19 22:45 | Yes
888 | Yes | when trying to replace ouygen tank while its connecting to wall the little manifold keeps letting ouygen come out. | when trying to replace ouygen tank while its conneung to wall the little manifold keeps letting ouygen come out. | 000000000000000000000000001O##: 00000000000000000000000000100000000118756 25-Nov-2000000003 Giovanni, Lomeli The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000098980000000060000000000000000KIT,O2 TRANSPORT,CART MTG,DISS-M Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 FLOW METER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000ouygen manifold to fuction as eupected uer Impact:V600000000has been removed form service Patient Impact: No impact SE uuProblem Description/ueupected Resuts: V600000000ouygen manifold is leaking. Resoluion : Replaced defective Manifold with new part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000257 00000000-Oct-2000000003 Giovanni, Lomeli The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000ouygen manifold to fuction as eupected uer Impact:V600000000has been removed form service Patient Impact: No impact SE uuProblem Description/ueupected Resuts: V600000000ouygen manifold is leaking. Diagnostic performed by Engineer : Manifold oupus O2 where O2 is inpued Follow u Requred : V600000000has issus with manifold and needs one ordered Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000003-31 00000000:00000000:00000000 MuTIMETER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000138 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : V600000000ouygen manifold is leaking. Part # 98980000000060000000000000000 KIT,O2 TRANSPORT, Manifold, CART Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000ouygen manifold to fuction as eupected uer Impact:V600000000has been removed form service Patient Impact: No impact SE uuProblem Description/ueupected Resuts: V600000000ouygen manifold is leaking. Diagnostic performed by Engineer : Problem Description: V600000000ouygen manifold is leaking Form ued: Trouleshooting/Investigation Object: V600000000Software Version: uuDate and Time of Issu: 00000000/19/2000000003 Trouleshooting Steps Taken: Advised cutomer the ouygen manifold can be replaced by ordering the manifold. Part # 98980000000060000000000000000KIT,O2 TRANSPORT, Manifold, CART Cutomer is requsting onsite service to replace the ouygen manifold. Was PRS ued: PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jose Saenz IP & Issu: " 91365171 when trying to replace ouygen tank while its connecting to wall the little manifold keeps letting ouygen come ou." Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-19 23:16 | No
889 | Yes | Connectivity/Unable to view data: Info on patient death while on Tele & PIIC. NTF/ur uLLdischarging. Provided Info | Conneuty/Unable to view data: Info on patient death while on Tele & PIIC. NTF/ur uLLdischarging. Provided Info |  | 2023-09-19 23:40 | Yes
890 | Yes | System does not boot up successfully | System does not boot up successfully | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000367 200000000Sep-2000000003 Inostroza, Jose The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: SCREENS DO NOT TuN ON, NO GEOMETRY AND TOuH PANEL DOES NOT TuN ON When was the first occurence? Within... : 15-Sep-2000000003 15:0000000000000000 How often is issu occuring? : Continuuly Malfuction area: Not determined. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: SCREENS DO NOT TuN ON, NO GEOMETRY AND TOuH PANEL DOES NOT TuN ON. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 7.2.9.00000000Diagnostic performed by Engineer: We check the equpment and find problems in the internal power suply of the ucelera PC and the data screen power suply, which caues the power fues to get damaged and not to energize other parts of the equpment. Power suply with possible internal short circut which cannot be repaired. Follow u Requred: New monitor mut be dispatched to client Clinica San Carlos de Apoqundo contact Wladimir Araya medical equpment Resoluion: We left the ucelera PC disconnected, which is not ued in the service and we left the loaned monitor of medical equpment installed. Equpment is operational bu a new monitor is requsted for replacement. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009316 Koc, Migul The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: SCREENS DO NOT TuN ON, NO GEOMETRY AND TOuH PANEL DOES NOT TuN ON When was the first occurence? Within... : 15-Sep-2000000003 15:0000000000000000 How often is issu occuring? : Continuuly Malfuction area: Not determined. Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: SCREENS DO NOT TuN ON, NO GEOMETRY AND TOuH PANEL DOES NOT TuN ON. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 7.2.9.00000000Diagnostic performed by Engineer: He says that they received suport from the field service engineer, foud bunt fues, replaced them, bu the equpment still does not tun on, does not have personnel for remote suport, requsts a field service engineer visit, checking the uLLlog file, shows uLLmessage in geometry modue, checking case history, replaced fues in previou case becaue of a failue in the table. A field service engineer is needed to check fues and power suplies, replace if necessary, check the condition of the brakes, it may be cauing the fues to bun ou, replace if necessary, check Geo PC, clean and reconnect internally cards and cables, replace battery if necessary, otherwise reinstall software or evalute HD and/or GeoPC replacement, tun on the equpment, test the Geo Modue butons, replace the release buton of the board or complete modue if necessary. Internal Comments: uLL# and/or description of uLL SCREENS DO NOT TuN ON, NO GEOMETRY AND TOuH PANEL DOES NOT TuN ON Repair Action: He says that they received suport from the field service agent, foud bunt fues, replaced them, bu the equpment still does not tun on, does not have personnel for remote suport, requsts a visit from the field service agent, checking the uLLlog file, shows uLLmessage in the geometry modue, checking the case history, replaced fues in previou case becaue of a failue in the table. A field service engineer is needed to check fues and power suplies, replace if necessary, check the condition of the brakes, it may be cauing the fues to bun ou, replace if necessary, check Geo PC, clean and reconnect internally cards and cables, replace battery if necessary, otherwise reinstall software or evalute HD and/or GeoPC replacement, tun on the equpment, test the Geo Modue butons, replace the release buton of the board or complete modue if necessary. Follow u Requred: uLL# and/or description of uLL SCREENS DO NOT TuN ON, NO GEOMETRY AND TOuH PANEL DOES NOT TuN ON Repair Action: He says that they received suport from the field service agent, foud bunt fues, replaced them, bu the equpment still does not tun on, does not have personnel for remote suport, requsts a visit from the field service agent, checking the uLLlog file, shows uLLmessage in the geometry modue, checking the case history, replaced fues in previou case becaue of a failue in the table. A field service engineer is needed to check fues and power suplies, replace if necessary, check the condition of the brakes, it may be cauing the fues to bun ou, replace if necessary, check Geo PC, clean and reconnect internally cards and cables, replace battery if necessary, otherwise reinstall software or evalute HD and/or GeoPC replacement, tun on the equpment, test the Geo Modue butons, replace the release buton of the board or complete modue if necessary. What are the neut Trouleshooting steps: A field service engineer is needed to check fues and power suplies, replace if necessary, check the condition of the brakes, it may be cauing the fues to bun ou, replace if necessary, check Geo PC, clean and reconnect internally cards and cables, replace battery if necessary, otherwise reinstall software or evalute HD and/or GeoPC replacement, tun on the equpment, test the Geo Modue butons, replace the release buton of the board or complete modue if necessary. Technician Name: Koc, Migul ----------------------------------------------------------------------------- | 2023-09-19 23:41 | No
891 | Yes | Misalignment in the touch position | Misalignment in the touch position |  | 2023-09-19 23:54 | No
892 | Yes | Device Fails Self-Test | Device Fails Self-Test | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000511 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : ship to acct address Attn Matt Shockley Resoluion : shipped serial # A23B-000000009200000000tracking 78397000000007758 DELIVERED Wednesday 9/20000000023 at 10000000027 AM Signed for by: M.SHOCKLEY Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: uLLmsg uer Impact: none Patient Impact: none SE E.00000000.000000003 See ref case 0000000022275700000000 Diagnostic performed by Engineer : Had cut press the blu i-buton and it emitted an uLLmsg “remove and reinsert battery” followed by single chirps. Cut stated that they jut puled ou this new battery from a sealed packet. However, this battery is not a u battery. See pic attachment. Bu cut stated that they have similar batteries which are 000000000000000000000000001king jut fine on their other AEDs. Will replace device uder warranty. Battery lot#: not a u battery see pic attachment Install before date: Pads lot# C12100000000-00000000 Eup date: 6-100000000000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-20 00:06 | No
893 | Yes | right side of touchscreen does not work and cannot calibrate | right side of touchscreen does not work and cannot calibrate | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in patient ue Eupected and ueupected Resuts: right side of touhscreen does not 000000000000000000000000001k and cannot be calibrate uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : cutomer tried touhscreen calibration and got "bad touh detected" uLL Advised to replace the touh screen. Provided parts ID and discused installation per the manul. 00000000561539000000001 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: scott IP & Issu: 79351000000005 touhscreen Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-20 00:11 | No
894 | Yes | Backup Alarm failed (EC 00 | Backup Alarm failed (EC 00) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: System is down (1100000000uLLon the system). uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states uit is giving uLLcode 1100000000 Per service manul cutomer states he replaced MC board and uit is still giving uLLcode 1100000000 Cutomer wants to confirm CPupcba is neut to replace and is requsting part nuber and price for replacement CPu Confirmed with cutomer CPuis manuactuers recommendation after the MC board. Resoluion : Provided cutomer with part nuber and price: 00000000561537200000000- RP-PCBA,CPu(Software 2.300000000-792.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction New Case Caller: David Parker IP & Issu: "56628562 System is down (1100000000uLLon the system). Per the cutomer." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 00:11 | No
895 | Yes | System Fleuvision will not boot with system. Requires manual boot | System Fleusion will not boot with system. Requires manual boot | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008553 25-Sep-2000000003 Byron Kirk The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ou Eupected and ueupected Resuts: u uer Impact: manul boot pc Patient Impact: none SE8.2.3000000005 Tools & Equpment : MuTIMETER,2000000005-00000000-30000000000000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Resoluion : replaced fleuvision pc ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action taken : FSE Ron Smithe requsted call open for 9/15 Oucome: dispatched Internal Remarks : Ron Smith - FV2PC replacement. FSE not onsite and wants to order a FV2PC when going onsite. Sugested checking bios battery and MPDubefore replacing FV2PC. Internal Remarks : Planner Initials: GM Action Taken:: FSE Ron Smith requsted call to be transferred to FSE Kirk Byron for 9/19 Oucome: dispatched ----------------------------------------------------------------------------- | 2023-09-20 00:56 | No
896 | Yes | The heartrate is not matching the tracking rate. The QRS is showing a glitch on the screen. It slows down and then... | The heartrate is not matching the tracking rate. The QRS is showing a glitch on the screen. It slows down and then... | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Coud not confirm uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Resoluion : Log review determined device fuctions are 000000000000000000000000001king correctly. Device didn't not arrive with cutomer cables. Only able to test device with cables and tools at the bench. If issu persists with cutomer, recommend sending in cables with the device to do futher testing if necessary. no parts requred. All fuctions are 000000000000000000000000001king correctly. Calibrated NBP, CO2 and Touh Screen. Diagnostic performed by Engineer : Cannot confirm complaint the uits fuctions are all 000000000000000000000000001king correctly at the bench. Logs will be puled from the device for RDT evalution. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: EMT Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE 7.300000000 Problem Description/ueupected Resuts: The heartrate is not matching the tracking rate. The QRS is showing a glitch on the screen. It slows down and then speeds u. The NIBP is not reading accuately. Diagnostic performed by Engineer : Problem Description: The heartrate is not matching the tracking rate. The QRS is showing a glitch on the screen. It slows down and then speeds u. The NIBP is not reading accuately. The heartrate is not matching the tracking rate. The QRS is showing a glitch on the screen. It slows down and then speeds u. The NIBP is not reading accuately. Form ued: Bench Trouleshooting Steps Taken: EMT tested Tempu Pro.QRS on leads 1 and 2, paues and stuters.NBP was showing 123/82 bu EMT coud not repeat the issu the paramedics had with NBP. Resoluion : Sent to Bench Follow u Requred : The heartrate is not matching the tracking rate. The QRS is showing a glitch on the screen. It slows down and then speeds u. The NIBP is not reading accuately. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE: Interaction Type: New Case Caller: Alec Hadley - phone 21000000005000000002617 IP & Issu: 95910000000083 - The heartrate is not matching the tracking rate. The QRS is showing a glitch on the screen. It slows down and then speeds u. The NIBP is not reading accuately. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 02:31 | No
897 | Yes | SpO0larm failure. No Trouble Found-Evaluation Confirmed Device Working. No fault found. | SpO0alarm failure. No Trouble Found-Evaluation Confirmed Device Working. No fault found. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 25-Sep-2000000003 Trevor Ervin The system meets the specification for the performed service and is retuned to ue Tools & Equpment : TEST SIMuATOR KIT,200000000000000007-31 00000000:00000000:00000000 TEST SIMuATOR KIT,200000000000000007-31 00000000:00000000:00000000 Follow u Requred : Commuicate and trouleshoot with cutomer Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: monitoring a patient Eupected and ueupected Resuts: The patient passed away on the 100000000h and they need to see if the monitor alarmed for the SpO2 uer Impact: no Patient Impact: yes patient passed away SE Rev C Diagnostic performed by Engineer : TC puled clinical auit logs, PIC iu logs, and device config and logs. TC opened cutomer feedback du to patient incident and suplied logs to feedback team. Biomed verified concern as potential failue of monitor to alarm. TC ued simuators to generate both uand SP00000000 alarms on both monitor and central station uing biomed suveillance. Monitor appeared to alarm as eupected. Awaiting futher response from feedback team. Follow u created to keep case open. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000179 Cole, Pau The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Iu call received from CS Maria Y & biomed--asking to pul clinical auit logs--bu no PRS access to site--CS paging FSE on PIu-no remote service needed with no PRS--closing 000000000000000000000000001Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Iu call received from CS Maria Y & biomed--asking to pul clinical auit logs--bu no PRS access to site--CS paging FSE on PIu-no remote service needed with no PRS--closing 000000000000000000000000001----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 Youg, Maria The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Patient Safety P1- FSE needs to get the logs from the PIC and review with cutomer if there was any alarms for the patient for SpO2 Need to get the configuation file from the bedside. If they have IBE they might have some stored data? Diagnostic performed by Engineer : lbiomed looking at clinical adut trail. RSE and RAS coud not access the server via PRS. Shows the monitor comes online and the ualarms come on bu the SpO2 does not show to come on. From 9/100000000 at 000000001am to 000000007am the patient dropped. It is not showing the SPO2 alarms come on No details on the bedside. Bed # 335-00000000 The patient passed on that day and they are trying to figue ou if the monitor alarmed for SpO2 duing that time. No details on the patient demographics. They need to get as muh information from the bedside (configuation file) Clinical auit log. Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: biomed Device ue: monitoring a patient Eupected and ueupected Resuts: The patient passed away on the 100000000h and they need to see if the monitor alarmed for the SpO2 uer Impact: no Patient Impact: yes patient passed away SE Rev C Diagnostic performed by Engineer : followed u with cutomer and confirmed with Prod Champion that the clinical auit log does indeed show SpO2 on/off for the alarms. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Elizabeth Lasek Issu: 8761700000000 System is operational (need clinical applications per the cutomer to uderstand the messages that are popuating on the system) Per the cutomer. Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to clinical suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 02:46 | Yes
898 | Yes | SPI BuFailure | SPI BuFailure | 000000000000000000000000001O##: 00000000000000000000000000100000000393000000009 300000000Nov-2000000003 Grover, Bryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Clinical Eupected and ueupected Resuts: Bad MCPCBA replaced with new one. uer Impact: none Patient Impact: none SE 2.100000000Resoluion : 000000000000000000000000001O#is for retuning part; 000000000000000000000000001k was completed on 00000000000000000000000000100000000000000008979. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008979 13-Oct-2000000003 Grover, Bryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Clinical Eupected and ueupected Resuts: Bad MCPCBA replaced with new one. uer Impact: none Patient Impact: none SE 2.100000000Diagnostic performed by Engineer : Device not fuctioning with uLLcodes; see previou 000000000000000000000000001for uLLcodes. Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000005-31 00000000:00000000:00000000 FLOW METER,20000000000000000000000000000000000000000:00000000:00000000 FLOW METER,20000000000000000000000000000000000000000:00000000:00000000 Follow u Requred : N/A Parts Consued : 1.00000000000000000000000000000000561537531,RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000----------------------------------------------------------------------------- | 2023-09-20 03:06 | No
899 | Yes | System is down (boot disk failure; cuomer believes they lost the hard drive). Per the cuomer. | System is down (boot disk failure; cuomer believes they lost the hard drive). Per the cuomer. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000576 13-Nov-2000000003 Beato Tom The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: scan patients Eupected and ueupected Resuts: system does not power u, boot disk uLLon monitor uer Impact: cannot ue system Patient Impact: none SE 7.2 Resoluion : Made a system backu for the site since one was not in hand. Booting issu has not retuned. uit is fuly operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000770000000019-Sep-2000000003 Beato Tom The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: scan patients Eupected and ueupected Resuts: system does not power u, boot disk uLLon monitor uer Impact: cannot ue system Patient Impact: none SE 7.2 Diagnostic performed by Engineer : Arrived onsite and saw the boot disk uLLscreen on the monitor. Powered off system and reapplied power. System booted correctly to the application. Tested system. It is operational. Will order a hard drive and monitor system in case problem reappears. Follow u Requred : Monitor system and create system backu ----------------------------------------------------------------------------- Internal Remarks : Interaction New Case Caller: Kirk Gluck IP & Issu: "00000000500000000 System is down (boot disk failue; cutomer believes they lost the hard drive). Per the cutomer." Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created case and created qute requst for onsite suport service Neut Step Action: Awaiting Qute Verbal PO Nuber: 90000000026000000009-HTMI Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 03:06 | No
900 | Yes | 0 u Hu00// Cu sonicare base charger lit on fire | 00u Hu00/0// Cu sonicare base charger lit on fire | 2000000003-00000000-28 16:16:32:***Claims*** ##d cu with FedEu retun label. Tracking # is 7735700000000000000003. Awaiting response.;2000000003-00000000-20000000013:0000000000000000***claims*** ##d cu with safety qustions and requsted home address to create retun label. Awaiting response. | 2023-09-20 04:36 | No
901 | Yes | Late - DEATH - audit for date and time - re-assessed GDT and re-assigned to L0I | Late - DEATH - audit for date and time - re-assessed GDT and re-assigned to L0CI | Internal Remarks : Field Complete this colun Interaction Type: New Case Caller: " Wailani Aquno (3600000000 729-000000008" IP & Issu: "u50000000031800000000auit for date and time" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Created R000000000000000000000000001 roued to Technical Suport Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-20 04:41 | Yes
902 | Yes | System not booting up all the way | System not booting up all the way | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000812 19-Sep-2000000003 Waskosky Robert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Ouside of ue Eupected and ueupected Resuts: System doesn’t boot. uer Impact: system down Patient Impact: none SE 2.1.5 Resoluion : Assisted with software reload of sute PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 19-Sep-2000000003 Boozer, Stephen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Cath Lab Device ue: Ouside of ue Eupected and ueupected Resuts: No Bootu uer Impact: None Patient Impact: None SE Diagnostic performed by Engineer : Confirmed issu ----------------------------------------------------------------------------- Internal Remarks : Sute PC uable to fuly boot and stuk progress bar. Logging indicates repeating uLL related to geo software houed on sute pc. Recommended FSE attempt sute pc reload via euam room TSM. Internal Remarks : Interaction Type: New Case Caller: " Cody Oswald" IP & Issu: "System not booting u all the way/ 79577966" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: advised that the order was dispatched 9/19/23 00:00000000 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 04:41 | No
903 | Yes | upro ¦ Device failed alarm bar function test and utest while performing PM. | Tempupro ¦ Device failed alarm bar funuon test and utest while performing PM. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002929 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Device soaked over the weekend all fuctions are 000000000000000000000000001king correctly. Calibrated NBP, CO2 and Touh Screen. No parts requred. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Diagnostic performed by Engineer : Confirmed complaint the uits ufuctions are not 000000000000000000000000001king correctly, uon internal inspection the ufleu cable was puled from the uboard. the Fleu nu is loose and allowed for the fleu cable to be puled when disconnecting the cable. Once fleu was seated and nu tightened fuctions 000000000000000000000000001k correctly. Device will be soaked over the weekend to ensue fuctions are 000000000000000000000000001king correctly. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000----------------------------------------------------------------------------- | 2023-09-20 05:11 | No
904 | Yes | hot case B:0/u00no u-ray | hot case B:0N/uno u-ray | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: uable to eupose, SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022296700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: No u-ray Trouleshooting Action: Key faut codes or phenomena: The patient is on the stage and no radiation can be emitted. The cutomer did not provide detailed uLLinformation and said that he had contacted the engineer. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: " 1. Go to the scene and check the event log to find ou whether it is caued by the en_u signal or other reasons. 2. The cutomer eupressed dissatisfaction with the equpment, saying that there were many repair requsts, it was ustable, and it needed detailed inspection." Euternal Remarks: 30000000000000000.2000000003 0000000055:38 uC (860000000063600000000Service content (requred): Cannot be euposed; replace the uSC modue in the high voltage generator. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: none. Internal Remarks: 30000000000000000.2000000003 0000000055:00000000uC (860000000063600000000 $Q7: (Confirmed reported problem, no u-ray, system prompt "u-ray not possible. Tue problem. Contact technical suport." ) $Q8: (Check system log, foud message "Coud not access u-ray generator (ubufered) via uRayGeneration service"; uSC has problem, need to be replaced. ) $Q9: (Replace uSC Certeray(uu30000000000000000600000000600000000).) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11:Calibrated Tools: Refer to PA For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage: N/A T2 Activities:Repair Action: ppd/30000000000000000600000000600000000 uSC Certeray/200000000to hospital/site1 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 30000000000000000600000000600000000 1. Shipped 200000000 LiuChao-86-13979783177 SAP Transfer: Case Transferred to SAP Euternal Remarks: 20000000000000000.2000000003 00000000:13:00000000uC (860000000063600000000SuT. Ready for Clinical ue local time" : " 2000000003-00000000-19 17:00000000" ----------------------------------------------------------------------------- | 2023-09-20 05:11 | No
905 | Yes | Device Triple Chirp | Device Triple Chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000dward Carson The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Replacement requred Follow u Requred : Please order 1 u 0000000056000000002800000000- HS1 (W B) Defib, u English, Euchg to the below address St John Ambuance Autralia 00000000 Sturt Hway Winnellie NT 00000000100000000 Contact name: Belinda Novak Contact nuber: 00000000896000000005 Problem Description by Engineer : Cutomer Fuction/Role: Field Aid Sales and service Coordinator Device ue: Defib beeping Eupected and ueupected Resuts: Emitting a triple chirp uer Impact: Cutomer uable to ue Patient Impact: Not able to ue on patient if requred SE Not requred Diagnostic performed by Engineer : New battery replaced Battery now dead within a few months New battery installed, testing says to take battery ou ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Defib, u English 3ed, Euch ----------------------------------------------------------------------------- | 2023-09-20 05:56 | No
906 | Yes | M0A - u00 Operation check failed (Cleaned the electrode plate). | u Operation check failed (Cleaned the eleuode plate). | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 22.00000000.2000000003 00000000:00000000:53 uC (860000000000000000100000000 *Department information: [Department: Department of Anesthesia Contact person: Teacher uiao Tel: 15900000000819000000003]* $ASP engineer:[honglejia uushuhu] $Q1: N $Q2:N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem: Operation check failed $Q8: On site check the system and the parts foud the Defibrillator electrode plate is dirty, $Q9: Clean the electrode plate again and pass the operation inspection. $Q100000000Test and inspection data refer to:Visul check: Passed.Device retuned to ful fuctionality $Q11: Calibrated Tools: N/A No open Field Change Order foud for this service. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 22.00000000.2000000003 00000000:00000000:55 uC (860000000000000000100000000 $ASP engineer:[honglejia uushuhu] $Q1: N $Q2:N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem: Operation check failed $Q8: On site check the system and the parts foud the Defibrillator electrode plate is dirty, $Q9: Clean the electrode plate again and pass the operation inspection. $Q100000000Test and inspection data refer to:Visul check: Passed.Device retuned to ful fuctionality $Q11: Calibrated Tools: N/A Problem Reported by Cutomer: Failed to pass self-test Neut Step Action: 2. uY: M000000005A, SN: u000000000000000069, Faut description: uable to pass the self-test, contact the hospital teacher: u uiao from the Department of Anesthesia, Tel: 15900000000819000000003 Please help me to issu a S000000000000000000000000001directly and send 000000000000000000000000001kers to ASP Honglejia for follow-u repairs. Thank you Best Regards QuZongyao QuZongyao Field Service Engineer, EC&HPM,North Service & Soluions Delivery, Health Systems u Health Tech MobilePhone:1868600000000300000000 Mail to: ZongYao.Qucom SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022297100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Neut Step Action: Retun visit sucessfu T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Operation check failed Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: Check the connection of lead wires and treatment cables, and reru the operation check ----------------------------------------------------------------------------- | 2023-09-20 07:11 | No
907 | Yes | 0V0uLLsystem alarm cannot startup | 0/V0uLLsystem alarm cannot startup | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022297166 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A OneEMS internal remarks: When the uer connects to the serial port of the device to import data, the device alarms and cannot be tuned on, and buiness follow-u is requred. Problem Reported by Cutomer: 00000000/V600000000uLL ----------------------------------------------------------------------------- | 2023-09-20 07:31 | No
908 | Yes | 0u FD0Startup failed | 0/u FD0Startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/u uuuFD200000000 Safety Qustions: Case: 0000000022297185 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 20000000000000000.2000000003 15:15:00000000uC (8600000000600000000) "RDI Accuacy: NA" $Q7:Confirmed reported problem.the machine cannot be tuned on normally, and the mechanical motion is abnormal. $Q8: On site Reviewed the uLLlog, IPC connection timeou, Test inpu voltage is normal, foud IPC startu failue. $Q9: Download IPC system and App. $Q100000000 Test and Inspection data refer to:PA.Device retuned to ful fuctionality. $Q11: Calibrated Tools refer to: PA ^Curent Tue installed date:7/18/2000000003 ^Curent Tue uage:18375.75luEuternal Remarks: 20000000000000000.2000000003 15:15:39 uC (8600000000600000000) Service content (requred): Mechanical movement is invalid after power on. Reinstall the IPC system and software. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None SAP Transfer: Dispatched for Onsite Suport FSE Notes: This PO is a Reference Nuber only FSE Notes: Block Invoice Reason: NA ----------------------------------------------------------------------------- | 2023-09-20 07:36 | No
909 | Yes | 0_0CATHROOM_FD0_System cannot ready.0 | 0R_00CATHROOM_FD00System cannot ready.05 | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_0000000000000000ATHROOM_FD10000000000000000System cannot ready.00000000 Safety Qustions : Case: 0000000022283900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sabrina Lin Last udated by: Sabrina Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Dispatched for Onsite Suport Euternal Remarks : 20000000000000000.2000000003 00000000:22:00000000 uC (886000000000000000000000000 $Q1: (NO) $Q2: (NO) $Q3: (N/A) $Q00000000(N/A) $Q5: (N/A) $Q6: (System cannot ready.) $Q7: (Confirmed reported problem and System cannot ready and data monitor stop at 00000000'00000000") $Q8: (Check M-cabinet and foud the SIB board POST failed.) $Q9: (1.Replace ram on SIB board and then power on system. 2.The SIB board POST pass and system can ready. 3.Perform system u-ray fuctional test withou problem. ) $Q100000000 Test and inspection data refer to PA. Device retuned to ful fuctionality. $Q11: Calibrated Tools N/A. Internal Remarks : 20000000000000000.2000000003 00000000:22:00000000 uC (886000000000000000000000000 $Q1: (NO) $Q2: (NO) $Q3: (N/A) $Q00000000(N/A) $Q5: (N/A) $Q6: (System cannot ready.) $Q7: (Confirmed reported problem and System cannot ready and data monitor stop at 00000000'00000000") $Q8: (Check M-cabinet and foud the SIB board POST failed.) $Q9: (1.Replace ram on SIB board and then power on system. 2.The SIB board POST pass and system can ready. 3.Perform system u-ray fuctional test withou problem. ) $Q100000000 Test and inspection data refer to PA. Device retuned to ful fuctionality. $Q11: Calibrated Tools N/A. ----------------------------------------------------------------------------- | 2023-09-20 08:01 | No
910 | Yes | 0/u FD0cover fall off | 0R/u FD0cover fall off | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 20000000000000000.2000000003 13:00000000200000000uC (86000000006300000000 Service content: The L arm cover has fallen off. If there is no problem with the equpment, reinstall the cover. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: 00000000R/u uuuFD100000000 Internal Remarks: 20000000000000000.2000000003 13:0000000023 uC (86000000006300000000 "RDI Accuacy: YES " $Q7: ( Confirmed reported problem & The L arm cover fall off ) $Q8: ( Check the L arm cover & No problem foud ) $Q9: (Reinstall the L arm cover) $Q100000000 ( Test and inspection data refer to PA & Device retuned to ful fuctionality ) $Q11: ( Calibrated Tools: N/A ) SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022297599 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang Zhu Last udated by: Weiqiang Zhu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: L arm roud cover comes off Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: L arm roud cover comes off Remote diagnostics and test resuts: n/a PRS TR statu: Doesn't make sense Repair Action: Recommended on-site repair steps: Refasten the cover, tighten the chain or replace the cover Notes on cutomer requrements: Contact the cutomer as soon as possible to confirm the arrival time ----------------------------------------------------------------------------- | 2023-09-20 10:21 | No
911 | Yes | 0 u HY000TB handle started smoking and caught on fire while charging | 00u HY00/0 TB handle started smoking and caught on fire while charging | 2000000003-10000000026 19:13:100000000**Claims***Replacement sent order #1-878200000000Produt retuned. Closing case.;2000000003-00000000-21 1000000009:51:***Claims*** ##d cu with safety qustions and FedEu retun label. Tracking # is 77300000000817900000000 Awaiting response.;2000000003-00000000-19 23:00000000:53:*** SME *** Approved for safety escalation.;2000000003-00000000-19 2000000000000000000000000CHAT ID: 00000000373799 R - I - TB was charging and started smoking and bunt the toothbruh T - S - Chat disconnected in the start of the chat -- "Aw, Snap" uLLoccured Marina Londo marinalondo500mail.com;2000000003-00000000-19 2000000000000000017:Macro for Disconnected Chats.;2000000003-00000000-19 2000000000000000015:Macro for Disconnected Chats. | 2023-09-20 10:31 | No
912 | Yes | Machine automatically switching off | Machine automatically switching off | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000332 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : uit not foud at Advance Care System, Jabalpu. Requsted channel to find the correct location and log fresh call to check and if needed order part. ----------------------------------------------------------------------------- | 2023-09-20 10:51 | No
913 | Yes | Touchscreen fault | Touchscreen fault | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: FSE Liuuu000000008317156000000002 Call for qutation Safety Qustions: Case: 0000000022297863 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: liuuiaohu Last udated by: liuuiaohu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: bu parts ----------------------------------------------------------------------------- | 2023-09-20 11:41 | No
914 | Yes | Detector was dropped and is no longer charging or connecting to system (u uLL replacement) | Deteur was dropped and is no longer charging or conneung to system (u uLL replacement) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000093 200000000Sep-2000000003 Groothuzen Robert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: rad tech Device ue: ouside of ue Eupected and ueupected Resuts: uer Impact: none Patient Impact: none SE5.0000000000000000esoluion : Replaced WPD re attached to system and calibrated detector Parts Consued : 1.0000000000000000000000000000000021300000000000000002,Wirel. Portable Detector 'Refubished' ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 16-Sep-2000000003 Groothuzen Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: rad tech Device ue: ouside of ue Eupected and ueupected Resuts: Detector not available uer Impact: none Patient Impact: none SE5.0000000000000000iagnostic performed by Engineer : System and WiFi lite on detector orange. Fails attachment Follow u Requred : Replace WPD detector ----------------------------------------------------------------------------- Internal Remarks : Field Complete this colun Interaction Type: New Case Caller: Whitney Moua IP & Issu: Detector was dropped and is no longer charging or connecting to system Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (u Preferred Service Time: htm 9/15 for onsite duing contract hous Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Create Onsite 000000000000000000000000001and will have rp team to dispatch accordingly Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : FSE Robert can't attach wPD after the cutomer dropped it. Attempted different attachment methods. He will fill ou qustionnaire and send back in to parking lot. Internal Remarks : FSE Robert completed the FSQ for the wPD. Jamie will release the block after reviewing. Internal Remarks : verified requred data for lift block of WPD order. Sent data to Buand sent Lift block ##to SPS Internal Remarks : Field Complete this colun Interaction Type: New Case Caller: Whitney Moua IP & Issu: Detector was dropped and is no longer charging or connecting to system Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (u Preferred Service Time: htm 9/15 for onsite duing contract hous Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-20 12:26 | No
915 | Yes | Patient had an injury while getting on the table by a small hook located in the front region of the rail. | Patient had an injury while getting on the table by a small hook located in the front region of the rail. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006699 22-Sep-2000000003 Sturt DAW The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: In ue Eupected and ueupected Resuts: patient injuy uer Impact: Nil Patient Impact: Laceration to lower leg uuResoluion : The cutomer has raised a concern where a metal corner of a table cover has cauht the patients leg as it has rubed past, and it has caued a laceration which requred intervention. There is no faut with the table cover as it is located an secued as per the factory buld. System retuned to ue in a 000000000000000000000000001king condition. ----------------------------------------------------------------------------- | 2023-09-20 12:41 | Yes
916 | Yes | 0V0uLLAlarm failure | 0/V0uLLAlarm failure | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 00000000222983300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Su Yat-sen Memorial Hospital Affiliated to Su Yat-sen uiversity (Souh Branch) Neut Step Action: ASP Weigung follow Problem Reported by Cutomer: 00000000/V600000000uLL----------------------------------------------------------------------------- | 2023-09-20 13:06 | No
917 | Yes | uuuin pacer mode at 000o 000there is no rhythm strips present | u in pacer mode at 0:0:0to 0:0:0 there is no rhythm strips present |  | 2023-09-20 13:15 | Yes
918 | Yes | 0S:YO0Fluoroscopy u-ray | 0CS:YO0 Fluoroscopy u-ray | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Ohgua, Yui The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000S:YO2:u Tanase: Fluroscopy u-ray [Euamination category] Normal euamination Problem Description by Engineer: When he ued the foot switch, fluros comes ou and then not. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: When the foot switch is ued, fluroscopy is displayed and then not. uer Impact: No effect Patient Impact: No effect SE uknow Resoluion: RC Response: I euplained how to identify the failue. I euplained that he shoud check the plastic cover, etc. After I let him know, he has not contacted u again, so I think the problem has been improved. ----------------------------------------------------------------------------- | 2023-09-20 13:16 | No
919 | Yes | 0/u FD0HostPC get halted | 0R/u FD0HostPC get halted | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 00000000800000000500000000000000001 & 00000000800000000122391 1 Arranged for shipment, FSE Liang Chaoai -86-1599818000000007 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: The equpment often crashes duing ue and needs to be restarted mutiple times to recover, requring engineers to be on site to deal with it. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: The engineer had reinstalled the system software and pluged in and ou the board a week ago, bu the sitution was still the same. He supected that there was a problem with the host or the system disk. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended to reinstall the software after replacing the host or system hard disk. Recommended spare parts uuand designation: " 00000000800000000122391 u Haswell Monoplane Host PC no HDD 0000000080000000050000000000000000 Hard disk spare part" Notes on cutomer requrements: Onsite Euternal Remarks: 25.00000000.2000000003 00000000:53:56 uC (860000000089600000000000000000000000000000000001k content: HOST PC failue, replace the Host PC Service resuts: After repair, the equpment retuned to normal operation. Other matters: none Internal Remarks: 25.00000000.2000000003 00000000:53:58 uC (860000000089600000000 RDI Accuacy: Yes $Q7: (Confirmed reported problem. system die sometimes) $Q8:(Host PC or system disk uLL need to replace them) $Q9: (replaced host pc and disk and, check the fuction) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage: scan second:N/A Safety Qustions: Case: 0000000022298395 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Repeated restarts ----------------------------------------------------------------------------- | 2023-09-20 13:21 | No
920 | Yes | Ouygen source failure | Ouygen source failure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/V600000000uLL Safety Qustions: Case: 000000002229800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-20 13:26 | No
921 | Yes | Live Monitor Display is Blank | Live Monitor Display is Blank | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000531 25-Sep-2000000003 Jijith, PK The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Live Monitor Display is Blank When was the first occurence? Within... : 200000000Sep-2000000003 13:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Live Monitor Display is Blank uer Impact: No uer impact Patient Impact: No patient impact SE 1.000000002.1 Parts Consued : 1.000000000000000000000000300000000000000006525261,C15S 19 inch monitor withou stand-FRuResoluion : Replaced live display and foud display is 000000000000000000000000001king fine . Handed over the system to cutomer in good 000000000000000000000000001king condition ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Live Monitor Display is Blank When was the first occurence? Within... : 200000000Sep-2000000003 13:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Live Monitor Display is Blank uer Impact: No uer impact Patient Impact: No patient impact SE 1.000000002.1 Diagnostic performed by Engineer : Live monitor defect. The same needs replacement as the monitor was confirmed defect by crosschecking with another one. Internal Comments : uLL# and/or description of uLL: Live Monitor Display is Blank Repair Action : Live monitor defect. The same needs replacement as the monitor was confirmed defect by crosschecking with another one. Resoluion : Live monitor defect. The same needs replacement as the monitor was confirmed defect by crosschecking with another one. Follow u Requred : uLL# and/or description of uLL: Live Monitor Display is Blank Repair Action : Live monitor defect. The same needs replacement as the monitor was confirmed defect by crosschecking with another one. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-20 13:36 | No
922 | Yes | Patient u/ucannot be measured | Patient u/ucannot be measured | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022298518 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor/ ----------------------------------------------------------------------------- | 2023-09-20 13:41 | No
923 | Yes | 00he power does not come on becau the device fell over . | 000The power does not come on becau the device fell over . | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000Katori, Masato The system meets the specification for the performed service and is retuned to ue Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000009-21 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000002-28 00000000:00000000:00000000 Resoluion: [Special Note] On December 13, 2000000003, we received notification from Fuua Electronics Technical Service regarding the receipt of the device. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Fuua Denshi Technical Services Device ue: Not for clinical ue Eupected and ueupected Resuts: The device stopped tuning on after falling over uer Impact: None Patient Impact: None SE2.00000000.23 Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit the 000000000000000000000000001k report on paper? <YES> Diagnostic performed by Engineer: [Confirmation details] We have confirmed that the power does not tun on. We conduted a visul check and confirmed that there was no damage to the euterior. We disassembled the inside and checked for loose connectors, etc., and confirmed that there were no problems. Resoluion: [000000000000000000000000001k details] After conduting an internal check and assembling the parts, the system started u normally. Althouh it is not possible to determine if there is a poor connection between the cable or the connector, we have conduted variou operation checks and confirmed that it operates normally withou any problems. [Simutaneou FCO 000000000000000000000000001k] There are no FCOs that requre 000000000000000000000000001k on this equpment. [Special Note] December 12, 2000000003 We sent the equpment and 000000000000000000000000001k report to Fuua Electronics Technical Service. ----------------------------------------------------------------------------- | 2023-09-20 13:46 | No
924 | Yes | Paddle fault | Paddle fault | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 21.00000000.2000000003 12:52:00000000 uC (86000000009527) $Q1: () $Q2: () $Q3: () $Q00000000() $Q7: (Confirmed reported problem, can't pass the self-inspection) $Q8: (Plu in other device, still can't pass) $Q9: (Replace the paddle) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name ( Flue 18B ) S/N (27780000000071WS ) Neut Calibration date ( 00000000/00000000/20000000000000000 Euternal Remarks: 21.00000000.2000000003 12:52:00000000 uC (86000000009527) Service content: The device does not 000000000000000000000000001k properly, replace the handle. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000L/HeartStart uuumonitor/defib/handle failue SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022298791 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Handle failue Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: uplu and uplu treatment cables, ru operation checks, and trouleshoot handles ----------------------------------------------------------------------------- | 2023-09-20 14:36 | No
925 | Yes | Patient euuuenced significant alopecia after procedure | Patient euuuenced significant alope0 after procedure |  | 2023-09-20 14:51 | Yes
926 | Yes | M0A - u00 Functional check failed | u Funuonal check failed | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001000000005 27-Sep-2000000003 Peconio, Alberto The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: battery calibration is necessary uer Impact: no impact Patient Impact: no impact SE na Resoluion: NBP calibration was carried ou ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007300000000ndrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: battery calibration is necessary uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: battery calibration is necessary. Follow u Requred: Contacted the department, the device reports that battery needs to be calibrated, euplained the steps to perform the calibration bu du to department need and for the too muh time necessary at the end of the calibration, it is not possible to proceed, proceeded with attempt to order battery, ask Sidem for it. ----------------------------------------------------------------------------- | 2023-09-20 14:56 | No
927 | Yes | Lab failed to boot up | Lab failed to boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 15-Oct-2000000003 David Dickson The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Imaging kit failed to start u this morning u 3. Stuk on blu loading start-u screen. Fiued by tuning off down to the wall, bu took abou 300000000mins to come u. uer Impact: NA Patient Impact: NA SEu Resoluion : Switch on issus Parts Consued : 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000Barry White The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FSE to site to reload Fleuviewing PC Diagnostic performed by Engineer : Logging indicates FV PC crash Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Imaging kit failed to start u this morning u 3. Stuk on blu loading start-u screen. Fiued by tuning off down to the wall, bu took abou 300000000mins to come u. uer Impact: NA Patient Impact: NA SEu Reguatory Qustions : N/A ----------------------------------------------------------------------------- Cutomer comment : Confirmed with Jenn for 15.100000000as they 000000000000000000000000001ld prefer ou of hous. They will call back if faut 000000000000000000000000001sens ----------------------------------------------------------------------------- | 2023-09-20 15:01 | No
928 | Yes | The system does not start | The system does not start | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000118 21-Sep-2000000003 ROBERTO BISuCA The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.00000000000000000000000000000000800000000660000000093,u-ray PC Monoplane Resoluion: Replaced u-Ray PC, reinstalled system devices, Restore Backu, Batch Verification and calibration of the tue. Fuctional tests were carried ou with a positive oucome. Problem Description by Engineer: Problem description by engineer: The system does not start. Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 1.2.5.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000891 SAVERIO DIVITTORIO The system does not meet the specification for the performed service and is not retuned to ue. Internal Comments: uLL# and/or description of uLL . Repair Action: . Resoluion: Intervention will follow to restore the system. Follow u Requred: uLL# and/or description of uLL . Repair Action: . What are the neut Trouleshooting steps: Bisuca asks to order spare parts. Please ask him abou the plan. the system is down. Technician Name: SAVERIO DIVITTORIO Problem Description by Engineer: Problem description by engineer: The system does not start. Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer: Ou technician Bisuca asks to open call becaue system is down, it does not start. ----------------------------------------------------------------------------- | 2023-09-20 15:01 | No
929 | Yes | Tube defect | Tube defect | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022299000000009 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000R/FD200000000 SAP Transfer: Case Transferred to SAP Internal Remarks: 22.00000000.2000000003 00000000:200000000 uC (860000000013300000000 "RDI Accuacy: Yes" $Q7:Confirmation reportedproblem, The device cannot be euposed. $Q8:The device cannot be euposed and System prompt tue defect.Field inspection foud that the tue parameters were missing in the HV generator CuE configuation. $Q9: HV generator CuE parameters were reconfigued and do tue adaptation. $Q100000000Test and Inspection data refer to PA, Do a few patient, Observation of a certain period of time. The system is ok. Device retuned to ful fuctionality. $Q11:Calibrated tools NA. SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the tue was damaged and coud not emit rays, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: Tue defect. u-ray not possible Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. ue an indutance meter to measue whether the filament indutance is normal. If there is a problem with the filament indutance, replace the tue. 2. Re-do the Tue Condition and Adaptation. 3. If the filament indutance is normal, check whether the high-voltage cable and connector are normal. 00000000Check the machine uLLand proceed. malfuction repair Recommended spare part uuand name:" 989000000000000000085100000000 u 200000000 000000007 ROT-GS 10000000000000000 0000000080000000060000000081 IGT Service Suport Tool Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Euternal Remarks: 22.00000000.2000000003 00000000:200000000 uC (860000000013300000000Symptom: The device cannot be euposed. Soluion/repair measues: The device cannot be euposed, indicating a tue failue.On-site inspection foud that the parameters of the ball tue in the high-pressue configuation were missing.Reconfigued high-pressue parameters and performed bub calibration. The equpment 000000000000000000000000001ks normally after service. Euternal Remarks: 20000000000000000.2000000003 000000000000000057 uC (860000000013300000000SuT. Ready for Clinical ue local time" : " 2000000003-00000000-19 22:300000000 ----------------------------------------------------------------------------- | 2023-09-20 15:21 | No
930 | Yes | The machine remains stuck at start-up with the countdown at 0:0. | The machine remains stuck at start-up with the countdown at 0:0. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000778 September 22, 2000000003 Andrighetto Sylvio The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: TRM Device ue: No. Eupected and ueupected Resuts: No boot uer Impact: No. Patient Impact: No. SE 8 Resoluion: Replacement of FleuVision PC. Parts Consued: 1.0000000000000000000000000000000080000000058600000000, FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000125 September 200000000 2000000003 Christophe Marteleu The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: Replacement of FLEu PC. Problem Description by Engineer: Cutomer Fuction/Role: TRM Device ue: No. Eupected and ueupected Resuts: No boot uer Impact: No. Patient Impact: No. SE 8 Diagnostic performed by Engineer: The system does not start correctly, remains stuk at 00000000:00000000. The video acqusition cards of the Fleuvision PC are in uLL Tests reloading the OS and application inconcluive. Ordered a new Fleuvision PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000113 Christophe Marteleu The system does not meet the specifications for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: TRM Device ue: No. Eupected and ueupected Resuts: No correct boot uer Impact: No. Patient Impact: No. SE 8 Diagnostic performed by Engineer: The machine does not start correctly anymore, it remains stuk on the coutdown 00000000:00000000. Visit on site for diagnosis. ----------------------------------------------------------------------------- | 2023-09-20 15:46 | No
931 | Yes | 00fficia DFM0 Defibrillator/Monitor/Battery cannot be charged | 000Effi0 DFM0 Defibrillator/Monitor/Battery cannot be charged | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: The First Affiliated Hospital of Jinan uiversity Chaoshan Hospital Problem Reported by Cutomer: QAA00000000/uDFM100000000 Defibrillator/Monitor/ Neut Step Action: ##d ASP honglejia for follow u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 000000002229900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: The battery cannot be charged. The cutomer inqures abou the warranty period. Please follow u with buiness. ----------------------------------------------------------------------------- | 2023-09-20 16:01 | No
932 | Yes | Alarms/Failure to alarm: Central had not given alarm on 0h Sep 0. NTF/ur Not Acknowledging Alarm. | Alarms/Failure to alarm: Central had not given alarm on 0th Sep 00 NTF/ur Not Acknowledging Alarm. | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000652 Lämsä, Marku The system was re-viewed as the specification for the performed service and is retuned to ue. Resoluion: I sent the cutomer an evidence via ##of the central did the alarm. Problem Description by Engineer: Cutomer Fuction/As: Biomed Device ue: In ue Eupected and ueupected Resuts: Alarm shall occu uer Impact: None Patient Impact: Patient euposed to possibly seriou injuy SE C.00000000.00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000698 Lämsä, Marku The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/As: Biomed Device ue: In ue Eupected and ueupected Resuts: Alarm shall occu uer Impact: None Patient Impact: Patient euposed to possibly seriou injuy SE C.00000000.00000000 Diagnostic performed by Engineer: Contacted cutomer and created cutomer feedback. I checked the log files. According to the log files the central has alarmed. Follow u Requred: Follow u, wait for cutomer's response and discus with CAS afterwards ----------------------------------------------------------------------------- | 2023-09-20 16:06 | No
933 | Yes | 0/uDFM0 Defibrillator/Monitor/Equipment alarm | 0N/Effi0 DFM0 Defibrillator/Monitor/Equipment alarm | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 26.00000000.2000000003 00000000:36:31 uC (860000000000000000100000000Department information: [Department: Equpment Department Contact: Teacher Wang Tel: 17898829689]* $ASP engineer:[honglejia CHEN ZHENWEI] $Q1: N $Q2: N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem. Device failue uLLis reported. $Q8: Check euternal connections; The cutomer was foud to have no battery connected to the device. $Q9: Connect the battery, connect AC power charging, the device performs a self-test, the self-test passes, and the device is normal. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11:Calibrated Tools: N/A Internal Remarks: 26.00000000.2000000003 00000000:36:33 uC (860000000000000000100000000 $ASP engineer:[honglejia CHEN ZHENWEI] $Q1: N $Q2: N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem. Device failue uLLis reported. $Q8: Check euternal connections; The cutomer was foud to have no battery connected to the device. $Q9: Connect the battery, connect AC power charging, the device performs a self-test, the self-test passes, and the device is normal. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11:Calibrated Tools: N/A SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022299535 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Shanghai Children's Hospital of Fuan uiversity, the equpment alarmed and 000000000000000000000000001kers were requred to be dispatched T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Equpment alarm The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: The person who requsted the repair did not know the specific sitution Repair Action: T2 Recommended on-site repair steps: It is recommended to check the alarm information and ru an operation check Recommended spare part uufor T2 with description: NA Notes on cutomer requrements: Requst engineers to go to the site to deal with the problem Test/Inspection: NA Test Equpment: NA ----------------------------------------------------------------------------- | 2023-09-20 16:11 | No
934 | Yes | u-rays not available | u-rays not available |  | 2023-09-20 16:16 | No
935 | Yes | Movement problems. | Movement problems. |  | 2023-09-20 16:21 | No
936 | Yes | SID movement not working, scanning is unavailable | SID movement not working, scanning is unavailable | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000820000000000000000Sep-2000000003 David Chatterton The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced SID motor on Clea stand. Problem Description by Engineer : Problem description by engineer : No SID movement possible, no rotation possible. Message on acq monitor SID movement not 000000000000000000000000001king, scanning is uavailable. How often is issu occuring? : intermittent. when radiographer arrived in room no message was visible. when patient was on table message was present. After 2 restarts message did not clear, case was cancelled. Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: uable to proceed, case was moved to cathlab. Patient Impact: case was moved to cathlab. SE 1.2.5.00000000Diagnostic performed by Engineer : When I arrived on site uLLwas visible. uable to move SID. Tried to release brake in PSC bu all fuctions related to SID were greyed ou. After another restart the message did not appear and I was able to move the SID. Moved to min SID and removed covers to confirm connections were secue, no mechanical issus observed. Noted that SID movement was noisy at certain angles. attached video files to 000000000000000000000000001O#with soud. Tier 2 recommended replacement of motor Parts Consued : 1.000000000000000000000000000000008000000002627900000000lea2 FD shift motor ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 Ahmed Kotb The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : SID movement not 000000000000000000000000001king, scanning is uavailable. How often is issu occuring? : Intermittent Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : SID movement not 000000000000000000000000001king, scanning is uavailable. -- the uLLlog has been eutracted from system and by investigating it, I foud the following: 200000000736 uLL2000000003-00000000-20000000010000000012:1000000007600000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo 200000000753 uLL2000000003-00000000-20000000010000000012:28.527 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 2000000007500000000rror 2000000003-00000000-20000000010000000012:28.529 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000755 uLL2000000003-00000000-20000000010000000012:28.500000000Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 200000000372 uLL2000000003-00000000-20000000010000000027:56.300000000Application uLL SAM_CleaIDS Not Calibrated 200000000SPOS63000000003200000000tartu Service NoDetailInfo 200000000371 uLL2000000003-00000000-20000000010000000027:56.213 Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 2000000003700000000uLL2000000003-00000000-20000000010000000027:56.197 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000369 uLL2000000003-00000000-20000000010000000027:56.196 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000352 uLL2000000003-00000000-20000000010000000027:38.500000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo The logfile inclues nuerou nuber of events that indicates an issu in the encoder feedback from IDS (image detector shift encoder. the encoder is part of the IDS motor itself. An onsite visit is needed to check the cabling of this motor and replace if needed. Follow u Requred : uLL# and/or description of uLL: SID movement not 000000000000000000000000001king, scanning is uavailable. -- the uLLlog has been eutracted from system and by investigating it, I foud the following: 200000000736 uLL2000000003-00000000-20000000010000000012:1000000007600000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo 200000000753 uLL2000000003-00000000-20000000010000000012:28.527 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 2000000007500000000rror 2000000003-00000000-20000000010000000012:28.529 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000755 uLL2000000003-00000000-20000000010000000012:28.500000000Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 200000000372 uLL2000000003-00000000-20000000010000000027:56.300000000Application uLL SAM_CleaIDS Not Calibrated 200000000SPOS63000000003200000000tartu Service NoDetailInfo 200000000371 uLL2000000003-00000000-20000000010000000027:56.213 Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 2000000003700000000uLL2000000003-00000000-20000000010000000027:56.197 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000369 uLL2000000003-00000000-20000000010000000027:56.196 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000352 uLL2000000003-00000000-20000000010000000027:38.500000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo The logfile inclues nuerou nuber of events that indicates an issu in the encoder feedback from IDS (image detector shift encoder. the encoder is part of the IDS motor itself. Repair Action : SID movement not 000000000000000000000000001king, scanning is uavailable. -- the uLLlog has been eutracted from system and by investigating it, I foud the following: 200000000736 uLL2000000003-00000000-20000000010000000012:1000000007600000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo 200000000753 uLL2000000003-00000000-20000000010000000012:28.527 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 2000000007500000000rror 2000000003-00000000-20000000010000000012:28.529 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000755 uLL2000000003-00000000-20000000010000000012:28.500000000Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 200000000372 uLL2000000003-00000000-20000000010000000027:56.300000000Application uLL SAM_CleaIDS Not Calibrated 200000000SPOS63000000003200000000tartu Service NoDetailInfo 200000000371 uLL2000000003-00000000-20000000010000000027:56.213 Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 2000000003700000000uLL2000000003-00000000-20000000010000000027:56.197 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000369 uLL2000000003-00000000-20000000010000000027:56.196 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000352 uLL2000000003-00000000-20000000010000000027:38.500000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo The logfile inclues nuerou nuber of events that indicates an issu in the encoder feedback from IDS (image detector shift encoder. the encoder is part of the IDS motor itself. An onsite visit is needed to check the cabling of this motor and replace if needed. check encoder cable. If healthy, replace the motor (it inclues the encoder) Clea2 FD shift motor 000000008000000002627900000000 What are the neut Trouleshooting steps : SID movement not 000000000000000000000000001king, scanning is uavailable. -- the uLLlog has been eutracted from system and by investigating it, I foud the following: 200000000736 uLL2000000003-00000000-20000000010000000012:1000000007600000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo 200000000753 uLL2000000003-00000000-20000000010000000012:28.527 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 2000000007500000000rror 2000000003-00000000-20000000010000000012:28.529 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000755 uLL2000000003-00000000-20000000010000000012:28.500000000Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 200000000372 uLL2000000003-00000000-20000000010000000027:56.300000000Application uLL SAM_CleaIDS Not Calibrated 200000000SPOS63000000003200000000tartu Service NoDetailInfo 200000000371 uLL2000000003-00000000-20000000010000000027:56.213 Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 2000000003700000000uLL2000000003-00000000-20000000010000000027:56.197 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000369 uLL2000000003-00000000-20000000010000000027:56.196 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000352 uLL2000000003-00000000-20000000010000000027:38.500000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo The logfile inclues nuerou nuber of events that indicates an issu in the encoder feedback from IDS (image detector shift encoder. the encoder is part of the IDS motor itself. An onsite visit is needed to check the cabling of this motor and replace if needed. check encoder cable. If healthy, replace the motor (it inclues the encoder) Clea2 FD shift motor 000000008000000002627900000000 Technician Name : Ahmed Kotb Resoluion : onsite visit with the IDS motor is needed Internal Comments : uLL# and/or description of uLL: SID movement not 000000000000000000000000001king, scanning is uavailable. -- the uLLlog has been eutracted from system and by investigating it, I foud the following: 200000000736 uLL2000000003-00000000-20000000010000000012:1000000007600000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo 200000000753 uLL2000000003-00000000-20000000010000000012:28.527 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 2000000007500000000rror 2000000003-00000000-20000000010000000012:28.529 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000755 uLL2000000003-00000000-20000000010000000012:28.500000000Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 200000000372 uLL2000000003-00000000-20000000010000000027:56.300000000Application uLL SAM_CleaIDS Not Calibrated 200000000SPOS63000000003200000000tartu Service NoDetailInfo 200000000371 uLL2000000003-00000000-20000000010000000027:56.213 Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 2000000003700000000uLL2000000003-00000000-20000000010000000027:56.197 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000369 uLL2000000003-00000000-20000000010000000027:56.196 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000352 uLL2000000003-00000000-20000000010000000027:38.500000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo The logfile inclues nuerou nuber of events that indicates an issu in the encoder feedback from IDS (image detector shift encoder. the encoder is part of the IDS motor itself. Repair Action : SID movement not 000000000000000000000000001king, scanning is uavailable. -- the uLLlog has been eutracted from system and by investigating it, I foud the following: 200000000736 uLL2000000003-00000000-20000000010000000012:1000000007600000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo 200000000753 uLL2000000003-00000000-20000000010000000012:28.527 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 2000000007500000000rror 2000000003-00000000-20000000010000000012:28.529 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000755 uLL2000000003-00000000-20000000010000000012:28.500000000Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 200000000372 uLL2000000003-00000000-20000000010000000027:56.300000000Application uLL SAM_CleaIDS Not Calibrated 200000000SPOS63000000003200000000tartu Service NoDetailInfo 200000000371 uLL2000000003-00000000-20000000010000000027:56.213 Application uLL SAM_CleaIDS: MCI GetPositionCalibrationData FAILED; retCode = 8 200000000SPOS63000000003300000000tartu Service NoDetailInfo 2000000003700000000uLL2000000003-00000000-20000000010000000027:56.197 Application uLL Clea2 POST uLL[item=FrtIDS:EncoderAccessTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000369 uLL2000000003-00000000-20000000010000000027:56.196 Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 200000000352 uLL2000000003-00000000-20000000010000000027:38.500000000Application uLL Motor assembly uLL[auis=FrtIDS] 200000000SPOS6300000000126 Startu Service NoDetailInfo The logfile inclues nuerou nuber of events that indicates an issu in the encoder feedback from IDS (image detector shift encoder. the encoder is part of the IDS motor itself. An onsite visit is needed to check the cabling of this motor and replace if needed. check encoder cable. If healthy, replace the motor (it inclues the encoder) Clea2 FD shift motor 000000008000000002627900000000---------------------------------------------------------------------------- | 2023-09-20 16:31 | No
937 | Yes | 0 - u0- QACASE#000000Electrode plate surface and main engine contact surface corroded | 00- u00 - QACASE#00000000 Eleuode plate surface and main engine contausurface corroded |  | 2023-09-20 16:36 | No
938 | Yes | uLLmessage eupletory hold is not available and keeps popping up | uLLmessage eupletory hold is not available and keeps popping up | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V200000000, and every once in a while the eupiratory hold message appears. Respiratory therapist is stating it seems to be 000000000000000000000000001king fine eucept for the message. The uit is also alarming high o2, and means the euternal o2 sensor needs calibration or replacing. They are uable to swap ou with another uit. Recommended to take the patient off the uit and ru the EST. Problem Description by Engineer : Cutomer Fuction/Role: respiratory Device ue: in ue Eupected and ueupected Resuts: eupiratory hold message even thouh nobody is pressing that buton uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: zachary lake IP & Issu: 56636900000000 uLLmessage eupletory hold is not available and keeps popping u Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: n/a Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-20 16:41 | No
939 | Yes | Therapy test failed, operation check failed | Therapy test failed, operation check failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000817 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : To replace 0000000056000000009000000001, DFM100000000 Therapy PCA ASSY Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001, DFM100000000 Therapy PCA ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy ----------------------------------------------------------------------------- | 2023-09-20 16:46 | No
940 | Yes | SPLIT-The NIBP is not reading accurately-See Reference case PR 00 | SPLIT-The NIBP is not reading accurately-See Reference case PR 000 |  | 2023-09-20 17:03 | No
941 | Yes | uYy screen in the Lab is black | uYy screen in the Lab is black | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000010000000025-Sep-2000000003 GARD GJESTRuBAKKEN The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : monitor not powering on Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: monitor is black uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : Installednew FleuVision monitor. Tested OK. Parts Consued : 1.000000000000000000000000991933000000002000000002,Service Part 58'' Color LCD Monitor Tools & Equpment : DOSE-KVP-LIGHT METER,200000000000000005-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Moeller, Dennis The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : monitor not powering on Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: monitor is black uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : have spoken with FSE and medtech. have conclued new monitor is needed. Internal Comments : uLL# and/or description of uLL: will not power on. ps OK Repair Action : power suply has been checked by biomed. new monitor needed Resoluion : monitor has to be replaced Follow u Requred : uLL# and/or description of uLL: will not power on. ps OK Repair Action : power suply has been checked by biomed. new monitor needed also check: Root Caue 1: Large screen in euamination room is not reachable by the Host PC. As of consequnce the BIuRemote Monitoring Team is not able to detect any malfuctions of the large screen pro-actively. Service Action 1: 1) Check setting of the dipswitch (at bottom of monitor) of the large screen. With this Dipswitch youselect whether commuication goes via RS232 or LAN connection. Please pu it to LAN. Perform a cold restart and check logging if the uLLhas disappeared. 2) Check whether the other side of the Ethernet cable (TSu12) in R Cabinet is connected at all. Note: It can occu there is a plastic bag on the plu of cable TSu12 with label âNOT uEDâ Please remove plastic bag and connect cable to internal switch. Then verify whether uLLhas disappeared from uLLog. Technician Name : Moeller, Dennis ----------------------------------------------------------------------------- | 2023-09-20 17:31 | No
942 | Yes | DFM0 therapy receptacle replacement | DFM0 therapy receptacle replacement | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: nemo withou collection When it is accepted, the shipping address will be provided. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: euternally broken plate connector uer Impact: no impact Patient Impact: no impact SE na Euternally broken plate connector, no impact on patient reported, the lier requsts the suply of the part. Diagnostic performed by Engineer: Requst for estimate for plate connector. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000093 Parts Consued: 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle 1.0000000000000000000000000000000056000000008971, DFM100000000 ASSY Therapy Receptacle 1.0000000000000000000000000000000056000000008971, DFM100000000 ASSY Therapy Receptacle ----------------------------------------------------------------------------- Internal Remarks: Good morning, This is to requst estimate for suply of material, ou ref. 2000000003/0000000000000000 to qute in every commuication, for N.00000000 plate HOOK CONNECTOR (attached photos) on Type: DEFIBRILLATOR Inventory: IC37557 Manuactuer: u MEDICAL SYSTEMS uY: uDFM100000000 Nuber: CN3261000000005 : DEI COLLI HOSPITALS Bulding: MONALDI HOSPITAL Floor: uO.C. CARDIOLOGY AND CORONARY CARDIAC SuGERY (AB P3) (DANIELA TELESE) Room: CARDIOLOGY OuPATIENT CLINIC For technical information u GIOVENCO ANTONIO +39 333 200000000 8800000000; SILVESTRO PASQuLE +39330000000071 0000000000000000CALIGIuE GIOVANNI MARIA +39 331 655 00000000; GuDO RINGS +39 337 122 783; DENTE CARMINE +39 330000000011 0000000037; Requst to send offer on letterhead, showing: 1. Item code quted (specifying if this is a "ORIGINAL" or "COMPATILE" spare part - "NEW" or "uED/RECONDITIONED"); 2. Discout reserved to The Writer from the curent list 3. Type of payment 00000000 Transport costs, pu into 000000000000000000000000001k and installation where planned Waiting for you kind feedback. Kind regards u Pasqule Ferro Administrative Suport S.I.C. Althea Italia S.p.A c/o A.O.R.N. Dei Colli Hospitals (Monaldi - Cotuno - CTO) Via Leonardo Bianchi (Adjacent Pharmacy Warehoue - Yellow Ala) - 80000000031 Naples Landline: +39 000000001 963 1700000000(Int.7891 Althea) +39 000000001 963 1700000000(Int.7892 Althea) +39 000000001 700000000 2511 uE RESPOND TO ALL FuCTIONS OR SEND ##ADDRESSES REPORTED BELOW E-mail: mailto:pasqule.ferro00lthea-grou.com;it.aocolli00lthea-grou.com; ----------------------------------------------------------------------------- | 2023-09-20 17:41 | No
943 | Yes | no signal on monitor | no signal on monitor | 000000000000000000000000001O##: 000000000000000000000000001000000000000000090000000022-Sep-2000000003 Lofgren Leon The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Diagnostic Eupected and ueupected Resuts: Suden no video duing case. uer Impact: uable to continu case. Patient Impact: uable to continu case. SE Resoluion : Trouleshot intermittent no video duing case issu. Foud / corrected inpu signal configuation issu. Cutomer (Carla) reports that issu occured duing case and that they had to get a mobile C-arm to finish the case. Replaced monitor and tested to ensue fuctional system operation. Parts Consued : 1.000000000000000000000000991932252193,19'' Monochrome Monitor ER (MML1952-PER) ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Carla IP & Issu: 5377800000000o signal on monitor Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created case and adv call back Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-09-20 17:41 | No
944 | Yes | SPLIT-Device failed utest while performing PM-See Reference case PR 00 | SPLIT-Device failed utest while performing PM-See Reference case PR 0000 |  | 2023-09-20 17:43 | No
945 | Yes | Battery is worn out | Battery is worn out | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007700000000Parts Consued: 1.000000000000000000000000989800000000626900000000 Assy Battery Li-Ion, 11ah, V600000000----------------------------------------------------------------------------- | 2023-09-20 17:46 | No
946 | Yes | ucable does not work. | ucable does not work. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000821 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: ucable does not condut uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: ucable to be replaced. Follow u Requred: nemo withou collection ASL BAT 200000000 Fornaci Street 77123 ANDRIA Department: Clinical Engineering - Groud Floor Att.ne u Francesco Volpe - tel 00000000835771800000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000900000000Parts Consued: 1.00000000000000000000000098980000000010000000000000000,3 Leadset, Grabber, IEC, ICu----------------------------------------------------------------------------- Internal Remarks: Good morning, This is to ask for technical intervention uder contract on the equpment detailed below. We take this opportuity to re-stress that the technical report of the euecuion of the requsted intervention mut be necessarily transmitted to the following e-mails: rapportinidilavorosic00slbat.it; sic.aslbat00cspa.it; reporting the following requst details: 2000000003/00000000300000000 Waiting for feedback, yous sincerely. ----------------------------------------------------------------------------- | 2023-09-20 17:51 | No
947 | Yes | M0A - u00 Broken uuuitor (Split from 00 | u Broken uuuor (Split from 000 |  | 2023-09-20 18:10 | No
948 | Yes | PSA-QACASE for PCI imported product#000000no u-ray | PSA-QACASE for P0imported produu00000000 no u-ray |  | 2023-09-20 18:15 | No
949 | Yes | QACASE for PCI imported product#000000larm | QACASE for P0imported produu00000000 alarm |  | 2023-09-20 18:34 | No
950 | Yes | "Primary Alarm Failed" alarm occurred. | "Primary Alarm Failed" alarm occurred. |  | 2023-09-20 18:46 | No
951 | Yes | Blower motor is stalled | Blower motor is stalled | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000362 Miller Jason The system meets the specification for the performed service and is retuned to ue Resoluion : Provided cutomer with part nuber and price: 00000000561512761 - RP-Blower Assy,V600000000- 1000000001.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states device is giving a blower stall uLL confirmed in event log. foud uLLcode 11300000000confirmed in service mode, RPM of motor does not change form 30000000000000000 Cutomer is requsting part nuber and price for replacement blower. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: blower stall uLL uer Impact: system restricted Patient Impact: none reported SE uknown ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Cody Luthmers IP & Issu: 70000000021900000000Blower motor is stalled. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-20 18:56 | No
952 | Yes | FleuViewing PC did not detect one of the grabbercards | Fleuewing PC did not deteuone of the grabbercards | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005700000000eptember 25, 2000000003 Herbert, Bernd The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Remote Device ue: cv Eupected and ueupected Resuts: PC uLLduing star5u uer Impact: no Patient Impact: no uuParts Consued: 1.00000000000000000000000000000000800000000660000000013,FleuViewing PC 00000000 Resoluion: FleuSpot PC replaced, backu restored and new license loaded. Fuction check. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- Cutomer comment: ## Appt. arranged with u Ribatzky for 25.00000000. from approu. 9:00000000 am with maintenance. Internal Remarks: Replace Fleu Spot PC 3 hous / 1FSE Material: 1 u FleuViewing PC 00000000 000000008000000006600000000100000000 I cannot enter TPR with radar, bu kai or Martin can do it. At start-u of the FleuViewing PC one of the grabber cards is not detected or is not present at all. OCCuRENCES OVERVIEW (Date¦Time¦Occurence): September 12, -2000000003 12:28:0000000000000000 RECOMMENDED ACTION: Soluion 1 for Caue 1: Replace FleuViewing PC if problem persists. ADDITIONAL INFORMATION: 12-Sep-2000000003 11:59:31.0000000000000000Normal Operation uLL200000000SFLV00000000000000000000000000000000FLV [FleuViewing PC] Frame Grabber: non fatal uLLoccured [Card ID=2], [uLLcode=5: Card does not euist] Failed to get grabber card type. Internal Remarks: uu 00000000800000000660000000013 - FleuViewing PC 00000000 Cutomer ETA: SEPTEMBER 200000000 2000000003 ----------------------------------------------------------------------------- | 2023-09-20 19:41 | No
953 | Yes | OVP Circuit Failed (EC 0) | OVP 0cuit Failed (EC 00) |  | 2023-09-20 19:49 | No
954 | Yes | The devices screen is blank | The devices screen is blank | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003673 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000has no display Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/2000000002000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Advised cutomer the uit likely has a HYDIS display installed. Parts for udating a 1st generation uer interface assembly to a 2nd generation uer interface assembly. Page 183 for referencing the 2 displays. Advised Biomed the second generation u assembly is curently not orderable. Biomed advised the part will be made available to order 12/2000000003. Provided cutomer the list of parts to ugrade to a second generation u assembly. 2 parts are also on backorder 1.00000000561528891 RP-LCD, Assy,2nd GEN, V600000000 List price $1,387. Zero stocking 2.0000000056153600000000 RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000 List price $138. 72 in stock 3.00000000561528911 RP-Cable, u PCBA to LCD, V600000000 List price $ 66. 38 in stock 000000003561528871 RP-PCBA, uer Interface, u 2nd GEN,V600000000 List price $332. 73 in stock 5.00000000561531651 RP-Tray,LCD,2nd GEN V600000000 List price $ 229. Zero stocking 6.00000000561528900000000 RP-Screw,M2u3 Socket hd, PKG/100000000V600000000 List price $ 900000000 100000000in stock Note: V600000000Operational Software mut be version 2.100000000or greater when installing the new NEC 2nd Gen LCD. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000has a fuctional display. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000has no display. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: david campbel 57000000002100000000000000003 IP & Issu: screen is blank Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-20 19:56 | No
955 | Yes | no tidal volume | no tidal volume | 000000000000000000000000001O##: 00000000000000000000000000100000000118877 00000000-Dec-2000000003 Giovanni, Lomeli The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: no tidal volue uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Duing controls test, incorrect alarms were going off when disconnecting prouimal and patient tues. Parts Consued : 1.00000000000000000000000000000000561512191,RP-Gas Delivery System (GDS), V600000000Tools & Equpment : TORQu SCREWDRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000003-31 00000000:00000000:00000000 MuTIMETER,200000000000000003-31 00000000:00000000:00000000 FLOW METER,200000000000000003-31 00000000:00000000:00000000 TORQu WRENCH,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 00000000-Oct-2000000003 Giovanni, Lomeli The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: no tidal volue uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : V600000000failed air flow tests and needs replacement GDS Follow u Requred : GDS is defective, Recommend replacing GDS. Tools & Equpment : TORQu SCREWDRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 MuTIMETER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003977 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: no tidal volue uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states uit is not showing tidal volue. Cutomer is requsting on site service to have uit evaluted and repair what is needed. Resoluion : Creating on site service. Follow u Requred : Device is not showing tidal volue. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ruen HiriscauIP & Issu: 56628292 & no tidal volue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 20:11 | No
956 | Yes | Display has scratches - | Display has scratches - | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000600000000Hübler Thomas The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Device/component is reported as defective. uer Impact: no Patient Impact: no SE uknown Issu: Malfuction Description of uLL We hereby requst an estimate of the costs for a repair of the above Device. Front houing has scratches and mut be replaced. Images are here. We requst the creation of a written estimate in advance (incluing Freight, transport, and insuance costs) and delivery by ##to: mtsz-vivantes00amed.com. Please take into accout that withou a written confirmation of the estimate, the order shall not be considered issud. uLLmessage: Resoluion: Ou processing nuber: 0000000022300000000195 You order nuber: Au00000000000000023-0000000033200000000 Dear Ms. Z., The service for the defibrillators will be performed at u by the responsible dealers. This shoud be for you region: Böhm Electro Medicine GmbH Großkopfstraße 8 1300000000 Berlin Phone: +00000000(300000000 0000000089 99 - 00000000 Fau: +00000000(300000000 0000000089 99 - 200000000 ## info00oehm-elektromedizin.de Internet: www.boehm-elektromedizin.de Please note! The indicated device has not been 31.12.2000000002 technical suport since then. In many cases, then no service can be offered anymore. Bu in some cases we can still deliver parts. There is no warranty for availability. With the indicated damage, the display of the defi 000000000000000000000000001ld have to be replaced. 0000000056000000006581 uuuKit Display 3 & Face Plate Assy 1.718,71 € net I cannot estimate the time needed for the repair. Bu I 000000000000000000000000001ld assue that the replacement is not 000000000000000000000000001thwhile here, especially as it seems to 000000000000000000000000001k. ----------------------------------------------------------------------------- | 2023-09-20 20:11 | No
957 | Yes | 00ne of the paddles does not work | 000One of the paddles does not work | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 October 23, 2000000003 Víctor Rodríguz González The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000056000000001861,Intrepid Paddle Tray ASSY Resoluion: Defibrillator is checked and it is observed that the paddles do not make good contact with the suports so it prevents the defibrillation tes with the paddles from being performed. We replaced the paddle suport and confirmed its operation. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: Paddle connection base failue. uer Impact: Impossible to pass the operation test. Patient Impact: The equpment is ued with limitations. SE Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000911 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: Paddle connection base failue. uer Impact: Impossible to pass the operation test. Patient Impact: The equpment is ued with limitations. SE Diagnostic performed by Engineer: We contacted the cutomer (Ainoa 62900000000665), and he said that one of the paddles does not 000000000000000000000000001k. After performing the operation test we observed that one of the paddles does not sit correctly in the base generating a test failue. Material order is handled and the case is transferred onsite so that the equpment is checked by the field service. Resoluion: The case is transferred onsite. Follow u Requred: We contacted the cutomer (Ainoa 62900000000665), and he told u that one of the paddles does not 000000000000000000000000001k. After performing the operation test we observed that one of the paddles does not sit correctly in the base generating a test failue. - Equpment located on the 1st floor, left control. - Assign fse. - Order the mentioned material ----------------------------------------------------------------------------- | 2023-09-20 20:11 | No
958 | Yes | It displays a cross. | It displays a cross. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000897 Assad MAJEED The system meets the specifications for the performed service and is retuned to ue. Resoluion: We reactivated the auomatic printing of the operating test and made the uLLmessage disappear, the red cross and the auio beep. Failue solved. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that the defibrillator displays a fiued red cross and emits a single auio beep. When the defibrillator is tuned on, it indicates that the defibrillator is disabled. We performed an operating test, which went well, bu the auomatic printing was not activated. Regarding the analysis of the logs, it seems that a handling uLLduing the operating test generated the red cross and the uLLmessage. ----------------------------------------------------------------------------- | 2023-09-20 20:16 | No
959 | Yes | PSE_QACASE for PCI imported product#000000black screen | PSE_QACASE for P0imported produu00000000_black screen |  | 2023-09-20 20:18 | No
960 | Yes | Grey loading button issue | Grey loading button issue | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000967 Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: Referred to the fh by ##Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: Grey loading buton neut to the rotary wheel does not 000000000000000000000000001k. Patient Impact: - SE - ----------------------------------------------------------------------------- | 2023-09-20 20:26 | No
961 | Yes | fluoroscopy failed retry | fluoroscopy failed retry | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001910000000028-Sep-2000000003 Riley Earl The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Tech Device ue: uknown Eupected and ueupected Resuts: Fluroscopic failed uer Impact: uable to ue the room Patient Impact: None SE R7.2.16 Resoluion : Resoluion: System is 000000000000000000000000001king fine after the convertors were replaced. I retuned parts to AllParts Medical. They will be doing the invoicing for the approuimate total price of $50000000010000000000000000 via qute#1000000003735. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003611 25-Sep-2000000003 Riley Earl The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Tech Device ue: uknown Eupected and ueupected Resuts: Fluroscopic failed uer Impact: uable to ue the room Patient Impact: None SE R7.2.16 Resoluion : Resoluion: I replaced both convertors in the generator and performed tue adaptation. Checked operation of the system. This will be invoiced by AllParts Medical. I have to go back and retun the convertors. Follow u Requred : Parts retun. Tools & Equpment : MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 21-Sep-2000000003 Riley Earl The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace Safety resistors and convertors in the generator. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Tech Device ue: uknown Eupected and ueupected Resuts: Fluroscopic failed uer Impact: uable to ue the room Patient Impact: None SE R7.2.16 Diagnostic performed by Engineer : Diagnostics: The parts I ordered did not arrive. I continud to look at the system and foud that the convertor coud have caued the resistors to open. I ordered the convertors throuh AllParts Medical also. Tools & Equpment : MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000970000000000000000Sep-2000000003 Riley Earl The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace safety resistors in generator and continu troule shooting. Tools & Equpment : MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Tech Device ue: uknown Eupected and ueupected Resuts: Fluroscopic failed uer Impact: uable to ue the room Patient Impact: None SE R7.2.16 Diagnostic performed by Engineer : Diagnostic: Foud that there was no power coming to the room. They had to locate the main power breaker down the hallway. Once we foud it and power on the system the safety resistors in the generator opened u. I ordered them for replacement via AllPARTS Medical. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Jutin Hooper IP & Issu: ccb- cu requst for ETA, pls call cu to confirm Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: AA-ST RP Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Jutin Hooper IP & Issu: 0000000020000000071/fluroscopy failed retry Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: eupediate service requsted Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Sumitted Qute for Approval Neut Step Action: Awaiting Qute Verbal PO Nuber: 90000000026000000009 Verbal PO Approver: Jutin Hooper Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 20:31 | No
962 | Yes | O0evice failed message (EC 0) after PM board replaced (FCO) | O0Device failed message (EC 00) after PM board replaced (FCO) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 31-Oct-2000000003 Ebanks, Raymond The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561512561,RP-PCBA, Data Acqusition, V6000000001.000000000000000000000000989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000Tools & Equpment : TORQu SCREWDRIVER,200000000000000001-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 MuTIMETER,200000000000000001-31 00000000:00000000:00000000 FLOW METER,200000000000000001-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000001-31 00000000:00000000:00000000 TORQu WRENCH,200000000000000001-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Ouside of ue Eupected and ueupected Resuts: Replace fauty parts uer Impact: None Patient Impact: None SE N/A Diagnostic performed by Engineer : Replaced Flow Sensor Assembly & Data Acqusition PCBA. Performed: Electrical safety (section 9.00000000Leak tests (section 9.5) Pressue accuacy (section 9.7) Air delivery (section 9.8) Air flow accuacy* (section 9.100000000 O2 flow accuacy* (section 9.11) O2miu accuacy (section 9.12) Device passed all tests & retuned to service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: MV Device ue: not in ue Eupected and ueupected Resuts: check vent, o2 device failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Mutivendor has a V600000000and stated the uit has a check vent code 1111 after the FCO86600000000000000006 was implemented. He requsted for u to come back and evalute the uit. Requsted a 100000000% discout for ASP to trouleshoot uit. ASP is schedued on 9-29-23 ----------------------------------------------------------------------------- Internal Remarks : **CuT REQuSTED ONSITE SuPORT Interaction Type: New Case Caller: JON GAST 3212922266 IP & Issu: uit is not 000000000000000000000000001king following an recent FCO udate - 0000000022237000000009 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND REMOTE 000000000000000000000000001 RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-20 20:31 | No
963 | Yes | Device does not start up | Device does not start up | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 Problem Description by Engineer: Device starts u with RDT logo then black Screen Diagnostic performed by Engineer: Flash Card holder was shifted Resoluion: Flash Card holder placed in good position ----------------------------------------------------------------------------- | 2023-09-20 20:36 | No
964 | Yes | Low Tidal volume | Low Tidal volume | 000000000000000000000000001O##: 00000000000000000000000000100000000118866 00000000-Dec-2000000003 Giovanni, Lomeli The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: no tidal volue uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : GDS ou of range Parts Consued : 1.00000000000000000000000000000000561512191,RP-Gas Delivery System (GDS), V600000000Tools & Equpment : TORQu SCREWDRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 MuTIMETER,200000000000000003-31 00000000:00000000:00000000 TORQu WRENCH,200000000000000003-31 00000000:00000000:00000000 FLOW METER,200000000000000003-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 00000000-Oct-2000000003 Giovanni, Lomeli The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: no tidal volue uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : V600000000failed air flow tests while testing for low tidal volue. believe GDS is defective Follow u Requred : Recommend replacing GDS as V600000000failed air flow test duing testing for low tidal volue. Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: no tidal volue uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states uit is not showing tidal volue. Cutomer is requsting on site service to have uit evaluted and repair what is needed. Resoluion : Creating on site service. Follow u Requred : Device is not showing tidal volue. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ruen HiriscauIP & Issu: 70000000088595 - The title volue is very low. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 20:41 | No
965 | Yes | the wireless foot switch lost connection. | the wireless foot switch lost conneuon. | Internal Remarks : Interaction Type: New Case Caller: GE Healthcare IP & Issu: " 83759193 the wireless foot switch lost connection." Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created qute, advised QA team will reach ou shortly Neut Step Action: Awaiting Qute Verbal PO Nuber: 0000000011000000002 Verbal PO Approver: GE Healthcare Verbal PO CAP: 50000000000000000----------------------------------------------------------------------------- | 2023-09-20 21:01 | No
966 | Yes | Alarming vent in opt / while on patient | Alarming vent in opt / while on patient | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000arey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions Diagnostic performed by Engineer : Problem Description: V600000000is reported to have a euuuenced a Vert Inop condition while on a patient. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/2000000002000000003 Trouleshooting Steps Taken: Reviewed significant event log with cutomer. Biomed is not reporting any critical uLL. Biomed is inquring if the uit may have displayed the Inop du to ongoing recalls. Cutomer has referenced V60000000035 volt rail recall. Biomed has advised they are going to wait for u to schedue the recall for all their devices incluing FCO86600000000000000006C – V600000000V600000000Plu uLL– uLLLoss of Fuction W/O Alarm Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000to provide ventilation and not display vent inop uLL. uer Impact: V600000000has been removed from service Patient Impact: Biomed is not reporting any adverse oucome to patient care or therapy. The patient was placed onto another uLL SE uu Problem Description/ueupected Resuts: V600000000is reported to have a euuuenced vert inop condition while on a patient. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Paua 5129000000002662 IP & Issu: V600000000- SN 100000000166282 / Alarming vent in opt / while on patient Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-20 21:11 | No
967 | Yes | HS0ED is emitting the triple chirp | HS0AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000impson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 8 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A15B-00000000600000000 Battery Install Before# 10000000020000000000000000 Lot#599800000000 Pads Eup Date: 6/13/2000000005 Pads Lot # C121322- ----------------------------------------------------------------------------- | 2023-09-20 21:21 | No
968 | Yes | IT DOES NOT ALLOW TO RECORD STUDIES (YESTERDAY FSE JOSE ACUÑA WENT TO CHANGE THE COMPUTER, BUT THE FAULT PERSISTS) | IT DOES NOT ALLOW TO RECORD STUDIES (YESTERDAY FSE JOSE ACUÑA WENT TO CHANGE THE COMPUTER, BUT THE FAULT PERSISTS) |  | 2023-09-20 21:21 | No
969 | Yes | Prouimal pressure sensor auto zero failed | Prouimal pressure sensor auto zero failed | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : Not Answered Ou Of Bou Failue? : No Labor: Labor ID : Not Answered 000000000000000000000000001k Performed : Not Answered Diagonsis : Not Answered | 2023-09-20 21:40 | No
970 | Yes | Low Leak CO0ebreathing alarm | Low Leak CO0Rebreathing alarm | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer that either the Flow sensor assy or GDS assy will need to be replaced to correct issu, cutomer acknowledged and will order preferred part, no follow u needed, provided part nuber, issu resolved- Material Nuber 989800000000617100000000- RP-Flow Sensor Assy, Air & O2, V600000000$1217.00000000 (BO no eta) Material Nuber 00000000561512191 - RP-Gas Delivery System (GDS), V600000000$3251.00000000 (BO no eta) Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts: Co2 rebreathing uLLreported uer Impact: not uable needs repair Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer called back to verify installing the whisper swivel did not correct the Low Leak CO2 rebreathing risk alarm, verified euhalation port at the bottom of uit is clear, air inlet filter is fairly clean, requsts part needed to correct issu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000iller Jason The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Cutomer states it was reported uit was giving Co2 rebreathing uLLwhen device as being set u. Went over tue setu with cutomer. informed cutomer of whisper swivel. Advised cutomer to set uit u with test lug and see if uLLmessage is repeatable. Advised cutomer of internal tuing and verify they are connected properly and in the right spot. If still having issus advised cutomer to attempt to replace DA board and test uit again. Cutomer states he will attempt and callback if futher assistance is needed. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Co2 rebreathing uLLreported uer Impact: system restricted Patient Impact: none reported SE uknown Follow u Requred : Cutomer Callback Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Pablo Vallejo IP & Issu: 720000000000000000 Giving C00000000 alarm uLLRebreathing Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: NCutomer Callback Caller: Pablo Vallejo IP & Issu: 720000000000000000 Giving C00000000 alarm uLLRebreathing Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: 2nd Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-20 21:51 | No
971 | Yes | M0A - u00- Beeping and showing uLLon screen | u Beeping and showing uLLon screen | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : uLLis saying pacer has an issu Ou Of Bou Failue? : No Labor: Labor ID : L-19775783 000000000000000000000000001k Performed : Saw issu. Swapped device ou with another defib in Biomed shop. Brouht to shop and looked for service manul on line. Foud and went into service mode, did op check and the device passed the tests. Pu back into service and tested again - passed again. Pu back into service. Diagonsis : Need to trouleshoot | 2023-09-20 22:05 | No
972 | Yes | The system does not start. | The system does not start. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000773 ----------------------------------------------------------------------------- | 2023-09-20 22:06 | No
973 | Yes | Low tidal volume and alarms | Low tidal volume and alarms | 000000000000000000000000001O##: 00000000000000000000000000100000000118853 00000000-Dec-2000000003 Giovanni, Lomeli The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:low tidal volue and alarms uer Impact: not uable needs repair Patient Impact: none SEun Resoluion : Incorrect alarms duing controls test point to GDS as defective Parts Consued : 1.00000000000000000000000000000000561512191,RP-Gas Delivery System (GDS), V600000000Tools & Equpment : TORQu WRENCH,200000000000000003-31 00000000:00000000:00000000 MuTIMETER,200000000000000003-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000003-31 00000000:00000000:00000000 FLOW METER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005000000002 00000000-Oct-2000000003 Giovanni, Lomeli The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:low tidal volue and alarms uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : System was failing controls test and futher testing resuted in failed air flow tests. setting flow to 100000000resuted in a mauimu oupu of 200000000 Didn't find any blockage on either filter on V600000000 Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 MuTIMETER,200000000000000003-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000003-31 00000000:00000000:00000000 Follow u Requred : Recommend replacing GDS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer states duing set u the uit alarms Low Leak CO2 Rebreathing risk on screen, no patient involvement, in biomed measued low tidal volues duing testing, whisper swivel is in place or .25 test adapter is in ue, air inlet filter was recently changed bu reports have been very dirty in the past, requsts onsite service uder contract- Recommended part replacement pn 00000000561512191 GDS assembly Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:low tidal volue and alarms uer Impact: not uable needs repair Patient Impact: none SEun Follow u Requred : CON00000000900000000- alarms Low Leak CO2 Rebreathing risk on screen, low tidal volue when tested- Recommended part replacement pn 00000000561512191 GDS assembly ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ruben 916220000000068 IP & Issu: 200000000000000000000000000000000title volue is very low Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created remote 0000000000000000000000000010000000000000000000000002 dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 22:16 | No
974 | Yes | Accept button is inoperable | Accept button is inoperable |  | 2023-09-20 22:26 | No
975 | Yes | Red u on the screen - unplugged and replug - wont turn on | Red u on the screen - unplugged and replug - wont turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000091 200000000Sep-2000000003 Adale Mengistu The system meets the specification for the performed service and is retuned to ue Resoluion : Confirmed AC plu in back of device is not fuly connected. Replaced drained battery and AC power modue. Performed battery calibration and operational check. Once passed, performed shock test. Confirmed system fuly operational Diagnostic performed by Engineer : Determined red u is on top left of device even when tuned off and device chirps/beeps intermittently. Once tuned on device reads 'service uLL Problem Description by Engineer : Cutomer Fuction/Role: nuse Device ue: device in ue Eupected and ueupected Resuts: cutomer claims has red u on screen and is beeping uer Impact: none Patient Impact: none SE uknown ----------------------------------------------------------------------------- Internal Remarks : client called back to follow u. Advised that FSE is on his way to fiu the issu Internal Remarks : cutomer call back to ask if FSE had showed u becaue defib is fiued. Euplained that we dispacthed him at 9:15 AM today and that he probably fiued the device when she wasn't there ----------------------------------------------------------------------------- | 2023-09-20 22:51 | No
976 | Yes | system will not power on | system will not power on | Internal Remarks : Interaction Type: New Case Caller: Christelle Friedland IP & Issu: system will not power on Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: created qute, advised QA team will reach ou shortly ----------------------------------------------------------------------------- | 2023-09-20 23:01 | No
977 | Yes | Navi ring button does not work | Navi ring button does not work | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer reconnected the cable from the u pcba to the switch pcba and it solved the issu. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and the uuudoes not 000000000000000000000000001k. This uit does not have a uuusince FCO86600000000000000007 was remediated last September 2000000002. Instruted cutomer to open of the u assy, and verify the cable from the u pcba is secuely connected to the switch pcba. He will open uit soon. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: cannot get into diagnostic mode uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Anis Mahomu Issu: 7820000000018 Navi ring buton does not 000000000000000000000000001k Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-20 23:46 | No
978 | Yes | Ouygen tank drains even when its plugged into the O0utlet | Ouygen tank drains even when its plugged into the O0outlet | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001392 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit is draining the o2 tanks uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer states it was reported device is draining o2 tank while attached to wall o2. cutomer inspected device coud not find any leaks. cutomer performed o2 accuacy test. device testing within range. cutomer states he will finish testing device and retun device back to ue if no futher issus are foud. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005600000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit is draining the o2 tanks uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it is draining the o2 tanks. Provided him instrutions on checking the o2 solenoid valve. He will get back with details. He stated the o2 flow remained the same with or withou o2 connected. He also mentioned, that he was uable to dulicate problem. Respiratory created a requst with that complaint. Instruted him to ru the uit in ventilation mode with 21% o2, and check the o2% with the certifier. If it fails recommended replacing the o2 solenoid valve.. If it passed, then perform the entire PVT Cutomer stated everything is passing, and he coud not dulicate any issu with the o2 solenoid valve. Follow u Requred : Follow u, cutomer calling back while testing equpment ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Aleu Courouh IP & Issu: 6700000000300000000Ouygen tank drains even when its pluged into the O2 oulet/cutomer requsted suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-21 00:11 | No
979 | Yes | Alarms/No sound: There's no audio alarms coming from the critical care rooms: No malfunction. Loose Connection. ur. | Alarms/No sound: There's no audio alarms coming from the critical care rooms: No malfunuon. Loose Conneuon. ur. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005777 Edmond Omar The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: Yes Caue: Speaker not connected. Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Cutomer foud loose connection and reseated the cable. Tested soud. Prevention: Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Visul and auible alarms. uer Impact: None Patient Impact: None SE C.00000000.00000000 Problem Description/ueupected Resuts: Cutomer report visul alarms with no auible alarms. Diagnostic performed by Engineer : Problem Description: Cutomer report visul alarms with no auible alarms. Cutomer reports there is no soud at the central station for Critical Care rooms. Form ued: Trouleshooting/Investigation Object: A PICIu central station. Not Object: A neighboring PICIu central station. Software Version: C.00000000.00000000 Date and Time of Issu: 00000000/2000000002000000003 uer Actions: uu Where is the Deviation: One central station in ED has visul alarm with no auio alarm. Trouleshooting Steps Taken: Remoted in and checked soud in control panel.Disabled headset option.Asked biomed to identify the euternal speaker verify the connection.Ran throuh setu and tested the speaker. Was PRS ued: Yes PRS Permission: Permitted NA » uITED STATES » CA » Santa Clara » Kaiser-(Homestead) Santa Clara » CMS_82555832_952-53NK-F_CNSCHPRI1377000000001 ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE: Interaction Type: New Case Caller: Jackie Kindred - phone 669 00000000 10000000000000000 IP & Issu: 82555832 - There's no auio alarms coming from the critical care rooms. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-21 01:11 | No
980 | Yes | 00T0No uray, disk full | 00CDT0No uray, disk full | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005978 21-Sep-2000000003 Seager, Shannon The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : No disk space When was the first occurence? Within... : 200000000Sep-2000000003 15:38:00000000 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No uray uer Impact: System down Patient Impact: Moved patient to another room SE 000Diagnostic performed by Engineer : testing reveled there are 97 images uassociated to a folder. The database was repaired. the system was rebooted and then the database was recreated so it was clean. The system was rebooted and the system was tested and there are 90000000000000000mages of free space is available. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005900000000Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : No disk space When was the first occurence? Within... : 200000000Sep-2000000003 15:38:00000000 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No uray uer Impact: System down Patient Impact: Moved patient to another room SE 000Diagnostic performed by Engineer : Sam said duing a case they coud not cine and they were getting a message abou image drives ful. They deleted the patients and rebooted the system bu did not resolve the issu. Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Sam said they did not create an LOD, check log file for IPPC uLLand image drive uLL. If needed, 00000000200000000000000003100000000- FRuIP PC and 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Sam said they did not create an LOD, check log file for IPPC uLLand image drive uLL. If needed, 00000000200000000000000003100000000- FRuIP PC and 0000000080000000050000000000000000- Hard disk spare part Technician Name : Coffman, Josh ----------------------------------------------------------------------------- | 2023-09-21 01:46 | No
981 | Yes | SA：0V0uLLProuimal pressure pipe disconnected. | SA：0/V0uLLProuimal pressure pipe disconneud. |  | 2023-09-21 02:37 | No
982 | Yes | SA：0V0uLLProuimal pressure pipe disconnected. | SA：0/V0uLLProuimal pressure pipe disconneud. |  | 2023-09-21 02:41 | No
983 | Yes | Unit is displaying a low leak Co0ebreathing risk message | Unit is displaying a low leak Co0rebreathing risk message | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006337 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order a flow sensor assy and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and they are getting a low leak Co2 rebreathing risk alarm. While trouleshooting, they are getting a -00000000 slpm on the average air reading on the vent. Part #989800000000617100000000 flow sensor assy, list price $1217.00000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: low leak Co2 rebreathing risk message uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- | 2023-09-21 02:46 | No
984 | Yes | System not booting up - stuck at 0 sec | System not booting up - stuck at 0 sec | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000587 25-Sep-2000000003 Asif, Sawab The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: neuoangio Eupected and ueupected Resuts: fleuvision pc defective uer Impact: system down Patient Impact: none SE8.2.17 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Diagnostic performed by Engineer : Replaced fv pc ----------------------------------------------------------------------------- Internal Remarks : System stuk at 0000000000000000. uLLlogs show grabber card uLL. -Approved: FleuVision-2 PC: 00000000800000000119123 Disk for revised II: 0000000080000000050000000000000000Internal Remarks : Replan 000000000000000000000000001 - Changed to in process, no change in dispatching date, it's the same as the ETA ----------------------------------------------------------------------------- | 2023-09-21 03:11 | No
985 | Yes | 0 u Hu00 Cu owns cordless power flosser, it caught fire | 00u Hu00/0 Cu owns cordless power flosser, it caught fire | 2000000003-10000000000000000 18:36:00000000***Claims*** Placed order # 1-990000000000000000s replacement. Closing case.;2000000003-10000000000000000 18:22:100000000 - not requred I - Follow u CASE 6000000000000000035689200000000_ cu said he's been waiting from safety escalation if what shoud be done neut and if we can provide the new power flosser replacement. Cordless Power Flosser 30000000000000000Oral Irrigator Hu3800000000/200000000as cauht in the fire T - Reviewed the case ***Claims*** ##d cu with safety qustions and FedEu retun label. Tracking # is 77300000000170000000000000000. Awaiting response. S - As per cu , he already respond to that ## Eduate ill let ou bo /safety knows abou this, bu ask to send his emal back again for his answers to the safet qustions/ cu ok;2000000003-00000000-21 1000000000000000000000000***Claims*** ##d cu with safety qustions and FedEu retun label. Tracking # is 77300000000170000000000000000. Awaiting response.;2000000003-00000000-20000000018:59:16:R- Amazon I- Follow u on case 6000000000000000035689200000000 T- Thank youfor these information Joseph. Going back to you concern, can youtell me how it happened? Edmar C • 11:0000000000000000 AM I pluged it in to charge. Visitor • 11:0000000029 AM Then it started to cauht fire? Edmar C • 11:0000000058 AM Yes Visitor • 11:50000000000000000 AM I see. When did it happened? Edmar C • 11:500000000200000000AM It happens on like 3rd or 00000000 charge Visitor • 11:50000000022 AM Aroud September 1- Visitor • 11:50000000000000000AM Are youthe only one uing the flosser? Edmar C • 11:52:53 AM Yes Visitor • 11:53:200000000AM Did youhappen to have any injuies when this incident occu? Edmar C • 11:53:00000000AM Mo Visitor • 11:53:53 AM No Visitor • 11:53:55 AM S- case is escalated to safety du to visible flames as per cu \\ chat id 0000000037000000006 \\ please process replacement if applicable Registered: N Warranty end date: 6/26/2000000005 as per DOP Warranty entitlement phase: IW;2000000003-00000000-20000000018:0000000000000000:Macro for Disconnected Chats.;2000000003-00000000-20000000018:0000000000000000:Macro for Disconnected Chats. | 2023-09-21 04:36 | No
986 | Yes | Display is dim at highest setting | Display is dim at highest setting | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006588 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Provided part nubers and cutomer will order, no follow u needed, issu resolved- Material Nuber 989800000000617100000000- RP-Flow Sensor Assy, Air & O2, V600000000V6800000000(BO no eta) u assembly, withou nav-ring, V600000000(gray) 1100000000000000000000000056150000000091 (BO no eta) Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm service Eupected and ueupected Resuts: part id uer Impact: Flow sensor and u replacement Patient Impact: replace parts duing pm service SEuuDiagnostic performed by Engineer : Cutomer states duing pm the Air flow accuacy test fails and display is dim at highest setting, possible 1st gen lcd, cutomer requsts part nubers to replace. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Thomas Casey 55900000000200000000 IP & Issu: "Cutomer Requsting Suport - Parts ID 67180000000075" Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: created r000000000000000000000000001and roued to rse ----------------------------------------------------------------------------- | 2023-09-21 04:41 | No
987 | Yes | 0/u FD0Florouopy not possible, please reselect another application | 0R/u FD0Florouopy not possible, please reseleuanother application | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 20000000000000000.2000000003 00000000:00000000100000000uC (8600000000000000001) [Service content] The equpment cannot be euposed occasionally. Check the log and find that the high-voltage commuication is broken and cannot be connected. This occured once duing the on-site test. I connected the high-voltage start line to u300000000 and tested it on and off 5 times. No problems occured. 【Service Resut】The machine operates normally after repair. 【Other matters】None yet T2 Activities:Problem Description: Overview of cutomer problems: The device cannot see throuh the euposue.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Timeou establishing connection with the Generator Remote diagnostics and test resuts: The HOST PC was uable to commuicate with the CuE within the given time. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Caue 1 This coud be caued du to a timing issu between Host(PC) and Generator. Please ignore this alert, if it only occus sporadically. Treat it as a 'one time only'. Soluion Perform a reboot the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other Soluion here below. Caue 2 Problem with the hospital Mains power Soluion Check the hospital Mains inpu power to the Velara Generator. Caue 3 Relay (tap) in MPDucontrol board is defective Soluion If the problem is permanent: Switch off the system (bu make sue F1 to F5 MCBs in MPDuare on) and connect D-connector measuing board at NPu: 300000000 in M-cabinet. Measue the continuty between pins 00000000nd 11. This shoud be normally closed, If it is open then replace MPDucontrol board. If the problem is intermittent: Switch off the system. Disconnect the cable from NPu: 300000000 (frontal) in M-cabinet (Frontal) and connect it to NPu: 300000000 (lateral). Observe the system. If the problem is resolved, MPDucontrol board is defective. Caue 00000000 Problem with EN100000000 PCB Soluion Check whether the green power LED is on at EN100000000 pcb. If the LED is not lit, replace this PCB Caue 5 200000000power suply is malfuctioning. Soluion Measue the 200000000and replace power suply if necessary. Caue 6 Cue is defective Soluion Check the LED statu of Cue. If no LED's are lit, please check 200000000power suply first.(see also 'Caue 5') If Power suply is ok bu LEDs are off, replace Cue. Recommended spare parts uuand designation: 00000000210000000020000000091 Cue CVFD-5 ASSY. Notes on cutomer requrements: Onsite SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 20000000000000000.2000000003 11:37:00000000 uC (8600000000000000001) RDI Accuacy: Yes. $Q7:(Confirm reported problem, System can't euposue. ) $Q8: (Performed visul inspection, System can't euposue, Review log, "Timeou establishing connection with the Generator",Generator was not start u, change connector u300000000 to u300000000,issu solve.) $Q9: (change u300000000 to u300000000,issu solve. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name [PA] Problem Reported by Cutomer: uable to eupose Safety Qustions: Case: 0000000022300000000000000006 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-21 06:31 | No
988 | Yes | 0 / u0 / Malfunction issue (defective uuuitor). | 00/ u00 / Malfunuon issue (defeu uuuor). | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 0000000022300000000137 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? N/A 00000000Was it an ou-of-the-bou failue? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Internal Remarks : 21.00000000.2000000003 00000000:38:57 uC (85200000000100000000) $Q7: (Confirmed reported problem. malfuction issu) $Q8: (Checked system. confirmed uuuitor defective) $Q9: (Ordered and replaced part, op test passed, faut solved) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) Euternal Remarks : 21.00000000.2000000003 00000000:38:500000000TC (85200000000100000000) $Q7: (Confirmed reported problem. malfuction issu) $Q8: (Checked system. confirmed uuuitor defective) $Q9: (Ordered and replaced part, op test passed, faut solved) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) Problem Reported by Cutomer : 163237300000000Malfuction issu ----------------------------------------------------------------------------- | 2023-09-21 06:46 | No
989 | Yes | [Q] 0S:NS: "Fluoro storage not possible" occurred during an euamination. | [Q] 0DS:NS: "Fluoro storage not possible" occurred during an euamination. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000008 21-Sep-2000000003 Yanaka, Hajime The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: "Fluro storage not possible" occured duing an euamination. I diagnosed it remotely. I confirmed that an uLLwas detected on the side image processing compuer. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: "Fluro storage not possible" occured duing an euamination. I diagnosed it remotely. I confirmed that an uLLwas detected on the side image processing compuer. uer Impact: No effect Patient Impact: No effect SE R000Resoluion: When the system was started when I visited the site, the data monitor coutdown was displayed at 00000000:00000000, bu The lateral image processing PC (IPPC) was not starting u normally, so u-ray did not come ou for both the frontal and lateral images. When I switched the FRONTal and Lateral IPPC in order to find the way it was, this time it started u normally on the side, The front showed ustable behavior. In a test that shows the IPPC startu statu on the screen, startu is slow.It stops at the u logo screen.I confirmed that the system repeatedly restarted 2 or 3 times util the coutdown passed.Every time a self test is ru at startu, an IPPC uLL(Not done) occus. In the neut diagnostic test*, an uLLwas being made. - Verify image storage data consistency:FAILED - Verify Image Processing&HOST): 1,2,3,00000000 FAILED, 5 Suceeded, bu the data... FAILED. This problem repeatedly occus and the equpment does not start normally, which caued an internal software failue. It is also necessary to reinstall the software on the front OS that was replaced. Based on the above resuts, I juged that the lateral image processing PC had failed so I replaced it. After I replaced it, I rebult the image database, performed the above diagnostic test, the self test at startu, and the fluros clear test a nuber of times, and I confirmed that operation was fine. [Others] I confirmed that both the foot switch in the control room and the euamination room were operating well. [Replacement parts] 0000000080000000050000000092 IP PC Revised_II no HDD u- 1 Parts Consued: 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000003 Suui, NobuasuThe system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: [Suject] 00000000DS:NS:u Kisida: "Fluro storage not possible" occured duing an euamination. [Internal Comments] Simpl 3.1 17900000000 Image Processor uLL2000000003-00000000-21 00000000:17:11.0000000000000000Image Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Timeou duing stop of ImageStoreSinkNode [IMAGE_PROCESSING_ON_TIME] 00000000:19:23.5000000000000000000000000ChannelIdentification : Image Channel Lateral 000000000000000000000000100000000 Normal Operation Service NoDetailInfo [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Yanaka, Hajime [Hajime Tanyanaka] - FSE1 Visit date and time 9/21/200000000311:300000000 - FSE1 000000000000000000000000001k time 2 [Euamination category] Normal euamination Diagnostic performed by Engineer: RC Response: FSE assignment Problem Description by Engineer: "Fluro storage not possible" occured duing an euamination. I diagnosed it remotely. I confirmed that an uLLwas detected on the side image processing compuer. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: "Fluro storage not possible" occured duing an euamination. I diagnosed it remotely. I confirmed that an uLLwas detected on the side image processing compuer. uer Impact: No effect Patient Impact: No effect SE R000Follow u Requred: * ----------------------------------------------------------------------------- | 2023-09-21 06:51 | No
990 | Yes | 00CD DISPLAY PROBLEM | 000LCD DISPLAY PROBLEM | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000187 Parts Consued : 1.0000000000000000000000000000000056000000008961,DFM100000000 Display ASSY ----------------------------------------------------------------------------- | 2023-09-21 06:51 | No
991 | Yes | 0 - u00 - Self Test Failure | 00- u000 - Self Test Failure | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022300000000187 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A OneEMS internal remarks: The equpment self-test failed, indicating that the treatment implementation failed and buiness follow-u is requred. Problem Reported by Cutomer: Self-test failed ----------------------------------------------------------------------------- | 2023-09-21 07:11 | No
992 | Yes | Therapy Disabled/Op Check Failed | Therapy Disabled/Op Check Failed |  | 2023-09-21 08:36 | No
993 | Yes | uray system cannot turn on | uray system cannot turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000377 17-Sep-2000000003 Chawanakorn,LubuchuThe system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: cv Eupected and ueupected Resuts: u-ray system cannot on uer Impact: no Patient Impact: no SE Resoluion : I foud fue nuber 11 on M cabinet was break I have to changed it Finally, the system opened and ue normally ----------------------------------------------------------------------------- | 2023-09-21 08:41 | No
994 | Yes | M0A - u00 Battery Fault | u Battery Fault | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022300000000597 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Yanling Couty Traditional Chinese Medicine Hospital of Huan Province Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-09-21 09:01 | No
995 | Yes | 0/u FD0live monitor intermittent black | 0R/u FD0live monitor intermittent black | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022300000000669 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Repair Action: nemo/video receiver/200000000to hosp/repair person will collect it Service content: If the video receiving head is damaged, it will be sent throuh the spare parts eupress service. SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: u uuuFD200000000----------------------------------------------------------------------------- | 2023-09-21 09:16 | No
996 | Yes | fleu vision is not turning on | fleu vision is not turning on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007399 22-Sep-2000000003 Hargrave Brue The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : MuTIMETER,2000000005-0000000030000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: in ue Eupected and ueupected Resuts: the fleuvision monitor had no display uer Impact: ued eutra room monitor as backu Patient Impact: uknown SE8.1 Resoluion : Replaced uuuper instrutions in u InCenter uing Docuent Du52518 Rev C. Tested system and retuned to cutomer operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000697000000001-Sep-2000000003 WEINGART, SASHA The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: BioMed Device ue: in ue Eupected and ueupected Resuts: the fleuvision monitor had no display uer Impact: ued eutra room monitor as backu Patient Impact: uknown SE8.1 Diagnostic performed by Engineer : Started troule shooting and immediately realized that the fleuvision pc was not on. Tried to tun it on and it 000000000000000000000000001ldn’t react. Verified that the pc was powered and that it’s power suply was dead. Ordered replacement parts. Follow u Requred : Replace the Fleuvision PC Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action taken: Transferred 000000000000000000000000001to FSE Brue Hargrave for 9/22 no time changed Oucome: Dispatched Internal Remarks : Planner Initials: GM Action Taken: FSE Sasha Weingart requsted call to be dispatched for 9/2000000008PM Oucome: dispatched Internal Remarks : Interaction Type: New Case Caller: Abdulah Moawwad IP & Issu: 6000000003700000000 & fleu vision is not tuning on Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: created onsite 000000000000000000000000001O#and set eupectations Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Planner Initials: GM Action Follow u activity for FSE Sasha Weingart e for 9/2200PM Oucome: dispatched ----------------------------------------------------------------------------- | 2023-09-21 10:01 | No
997 | Yes | 0 - u0 - uuLLDuring Opcheck - Replaced uLeadset | 00- u000- uuLLDuring Opcheck - Replaced uLeadset | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006717 Edward Carson The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer replaced uleads and opcheck passed. Machine retun to ful fuctionality Problem Description by Engineer : Cutomer Fuction/Role: BIomedical Engineer Device ue: Cutomer getting 9:22 uLLwhen runing opcheck Eupected and ueupected Resuts: Leads were fauty uer Impact: No impact Patient Impact: No Impact as done duing opcheck SE Not requred Diagnostic performed by Engineer : Advised cutomer that 9:22 uLLoccus duing opcheck when no cable is connected or if it is a fauty cable. ----------------------------------------------------------------------------- File uloaded by CS Portal uer : File Name: RfuontrollerHwLog_u31611000000007.uml. RfuontrollerSwLog_u31611000000007.uml. Comments: File(s) uloaded duing case creation ----------------------------------------------------------------------------- | 2023-09-21 10:46 | No
998 | Yes | "Fluoroscopy not possible" is displayed and fluoroscopy doesn't appear. | "Fluoroscopy not possible" is displayed and fluoroscopy doesn't appear. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008598 22-Sep-2000000003 Yoshihara, Hiroshi The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.000000000000000000000000000000002100000000000000002,Safety resistor 00000000, 300000000133mm 1.00000000000000000000000000000000210000000020000000091,Cue CVFD-5 ASSY. Tools & Equpment : DOSE-KVP-LIGHT METER, 2000000005-00000000-23 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-2000000000000000000000000:00000000 Resoluion: Yesterday, u-ray did not 000000000000000000000000001k du to a faut in the u-ray control uit (Cue), so I replaced the u-ray control uit (Cue).When I investigated yesterday, I foud a faut in the resistor with a temperatue sensor, so I replaced the resistor with a temperatue sensor.After I replaced it, I adjuted the u-ray oupu and confirmed that fluros the fluros susequnt to normal operation was possible. [Parts ued] Part nuber: 00000000210000000020000000091/部品名:Cue CVFD-5 ASSY/Nuber of uits: 1 Part no.: 000000002100000000000000002/Part name: SAFETY RESISTOR00000000,300000000133mm/Nuber of uits: 1 Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiology Engineer Device ue: No Eupected and ueupected Resuts: u-ray doesn't 000000000000000000000000001k uer Impact: No Patient Impact: No SE81300000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000780000000021-Sep-2000000003 ua, Koji The system meets the specification for the performed service and is retuned to ue Follow u Requred: N/A Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : "Fluroscopy not possible" is displayed and fluroscopy is not possible. u/Ms. Omori uhin: 1700000000 - Restarting the system does not improve the sitution. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: "Fluroscopy not possible" is displayed and fluroscopy is not possible. u/Ms. Omori uhin: 1700000000 - Restarting the system does not improve the sitution. uer Impact: Euamination delay Patient Impact: Euamination delay SE 8.1.3000000005 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Resoluion: 000000000000000000000000001k details: I confirmed the problem. When I checked it when I visited the site, I foud that commuication was not possible with the main control uit (Cue) for the u-ray generator. I had him reset the cue and restart the system, and then it started normally and u-ray radiation was possible. Considering that it didn't 000000000000000000000000001k even after doing a 2-berth, I plan to replace the uit. I also confirmed that the resistor with the temperatue sensor was fauty while I was confirming the system. I have responded as an emergency, so I plan to replace it at the same time. I confirmed that it is now operating normally. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000077900000000Kamata, Takayui The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer : "Fluroscopy not possible" is displayed and fluroscopy is not possible. u/Ms. Omori uhin: 1700000000 - Restarting the system does not improve the sitution. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: "Fluroscopy not possible" is displayed and fluroscopy is not possible. u/Ms. Omori uhin: 1700000000 - Restarting the system does not improve the sitution. uer Impact: Euamination delay Patient Impact: Euamination delay SE 8.1.3000000005 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: ・I confirmed, in the log, that there was a Cue in the Generator not fuctioning well. I received the chief engineer's approval via u Omori for the cost of an emergency response, and I asked FSE Yoshihara to respond. Internal Comments: uLL# and/or description of uLL: [Suject] 00000000DS:TK:u Ohno: "Fluroscopy not possible" is displayed and fluroscopy u-ray is not possible. [Internal Comments] SIMP ＝＝＝ 16000000002000000003-00000000-21 00000000:58:000000000000000000000000Application uLL Generator euception occured [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator Euception Class: 1, Nuber: 7751, Time: 8:57 21/9/2000000003, Teut: uGEN uLL(00000000M3) No response at all from Fus) 00000000H, 00000000H, 00000000H, 00000000H, 00000000H, 00000000H ChannelIdentification : u-Ray Channel Frontal uLLu-Ray Generator Startu 000000000000000000000000751 161 2000000003-00000000-21 00000000:58:000000000000000000000000Application uLL Generator euception occured [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator Euception Class: 3, Nuber: 8383, Time: 8:58 21/9/2000000003, Teut: uGEN Invalid Requst (00000000SS) PCcomm: Requst not performed ChannelIdentification : u-Ray Channel Frontal uLLu-Ray Generator Startu 000000000000000000000000383 165 2000000003-00000000-21 00000000:58:000000000000000000000000Application uLL Generator euception occured [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator Euception Class: 1, Nuber: 8379, Time: 8:58 21/9/2000000003, Teut: uGEN uLL(00000000SO) PCcomm: Node ID uknown ChannelIdentification : u-Ray Channel Frontal uLLu-Ray Generator Startu 000000000000000000000000379 359 2000000003-00000000-21 00000000:59:2000000000000000000000000Application uLL Generator euception occured [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator Euception Class: 3, Nuber: 700000000, Time: 8:59 21/9/2000000003, Teut: uGEN Invalid Requst (00000000JE) DI: can't translate FS=6CH PID=37H SID=00000000D00000000val=... ChannelIdentification: u-Ray Channel Frontal uLLu-Ray Generator Startu 00000000000000000000000000000000 366 2000000003-00000000-21 00000000:00000000:1000000000000000000000000Timeou establishing connection with the Generator [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Did not receive configuation info within timeou ChannelIdentification : u-Ray Channel Frontal uLLu-Ray control Startu 000000000000000000000000000000003 [Parts from the equpment] 00000000210000000020000000091:Cue CVFD-5 ASSY. QTY:1 ZSID/Traceable Osaka (u15) 20000000052600000000285:Lithiu battery 1/2AA QTY:1 ZSID Osaka (u15) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Akira Yoshihara] Yoshihara, Hiroshi - FSE1 Visit date and time 2000000003/00000000/2118:00000000 - FSE1 000000000000000000000000001k time 3 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Akira Yoshihara - 0000000000000000005983-66300000000 - Others Call mobile uon arrival Requster name Takayui Kamata [Euamination category] Normal euamination Repair Action : [Suject] 00000000DS:TK:u Ohno: "Fluroscopy not possible" is displayed and fluroscopy u-ray is not possible. [Internal Comments] SIMP ＝＝＝ 16000000002000000003-00000000-21 00000000:58:000000000000000000000000Application uLL Generator euception occured [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator Euception Class: 1, Nuber: 7751, Time: 8:57 21/9/2000000003, Teut: uGEN uLL(00000000M3) No response at all from Fus) 00000000H, 00000000H, 00000000H, 00000000H, 00000000H, 00000000H ChannelIdentification : u-Ray Channel Frontal uLLu-Ray Generator Startu 000000000000000000000000751 161 2000000003-00000000-21 00000000:58:000000000000000000000000Application uLL Generator euception occured [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator Euception Class: 3, Nuber: 8383, Time: 8:58 21/9/2000000003, Teut: uGEN Invalid Requst (00000000SS) PCcomm: Requst not performed ChannelIdentification : u-Ray Channel Frontal uLLu-Ray Generator Startu 000000000000000000000000383 165 2000000003-00000000-21 00000000:58:000000000000000000000000Application uLL Generator euception occured [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator Euception Class: 1, Nuber: 8379, Time: 8:58 21/9/2000000003, Teut: uGEN uLL(00000000SO) PCcomm: Node ID uknown ChannelIdentification : u-Ray Channel Frontal uLLu-Ray Generator Startu 000000000000000000000000379 359 2000000003-00000000-21 00000000:59:2000000000000000000000000Application uLL Generator euception occured [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Generator Euception Class: 3, Nuber: 700000000, Time: 8:59 21/9/2000000003, Teut: uGEN Invalid Requst (00000000JE) DI: can't translate FS=6CH PID=37H SID=00000000D00000000val=... ChannelIdentification: u-Ray Channel Frontal uLLu-Ray Generator Startu 00000000000000000000000000000000 366 2000000003-00000000-21 00000000:00000000:1000000000000000000000000Timeou establishing connection with the Generator [SEVERITY] uLL[EuCEPTION_DESCRIPTION] Did not receive configuation info within timeou ChannelIdentification : u-Ray Channel Frontal uLLu-Ray control Startu 000000000000000000000000000000003 [Parts from the equpment] 00000000210000000020000000091:Cue CVFD-5 ASSY. QTY:1 ZSID/Traceable Osaka (u15) 20000000052600000000285:Lithiu battery 1/2AA QTY:1 ZSID Osaka (u15) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Akira Yoshihara] Yoshihara, Hiroshi - FSE1 Visit date and time 2000000003/00000000/2118:00000000 - FSE1 000000000000000000000000001k time 3 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Akira Yoshihara - 0000000000000000005983-66300000000 - Others Call mobile uon arrival Requster name Takayui Kamata [Euamination category] Normal euamination Follow u Requred: * ----------------------------------------------------------------------------- | 2023-09-21 11:36 | No
999 | Yes | 00 CN00- No audio - SPLIT | u- No audio - SPLIT | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : no auio and therapy disabled SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 000000002230000000086 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Karen Lee Last udated by: Karen Lee 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-21 12:01 | No
1000 | Yes | 0S:NS: The indicator on the wireless foot switch turned red and it could not be ud for a while. | 0CS:NS: The indicator on the wireless foot switch turned red and it could not be ud for a while. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007316 Suui, NobuasuThe system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: The wireless foot switch battery and Wifi indicator lit u red and the device coud not be ued for a while.He tuned the foot switch power off/on and connected the charging cable for abou 15 minues and then checked it, and then operation was normal. Cutomer Fuction/Role: Radiology Department Device ue: Wireless foot switch Eupected and ueupected Resuts: The wireless foot switch battery and Wifi indicator tuned red and it coud not be ued for a while.He tuned the foot switch power off/on and connected the charging cable for abou 15 minues and then checked it, and then operation was normal. uer Impact: Responded uing a wired foot switch. Patient Impact: Responded uing a wired foot switch. SE R1.2.5 00000000800000000238261 Wireless Footswitch Set 00000000+2 18000000005500000000 Resoluion : RC Response: Operation is fine now, so he will observe the sitution. If an operation problem occus on the wireless foot switch, I asked him to first tun the foot switch power off/on, and if there is no improvement, to tun the equpment power off/on in order to reset the receiver side. Internal Comments : [Suject] 00000000S:NS:u Kasai: The indicator on the wireless foot switch tuned red and it coud not be ued for a while. [Internal Comments] Simpl physical [Euamination category] Normal euamination ----------------------------------------------------------------------------- | 2023-09-21 12:16 | No
1001 | Yes | uuudefective power supply | u defeu power supply | 000000000000000000000000001O##: 00000000000000000000000000100000000123365 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000281 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: nemo withou collection AZ. MELLINO MELLINI HOSP. POLYGON c/o A.S.S.T. FRANCIACORTA Administrative Office (floor -1) 00000000AZZINI STREET CHIARI, BS 25000000002 BIOMEDICAL LABORATORY ING. CLINICAL 000000000000000010000000000000000 Arianna Camponogara If not available, try throuh Sidem. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test Eupected and ueupected Resuts: buzing power suply modue uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Buzing power suply modue, proceeded with attempt to order spare part. ----------------------------------------------------------------------------- | 2023-09-21 12:51 | No
1002 | Yes | The power does not come on. | The power does not come on. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006818 Katori, Masato The system meets the specification for the performed service and is retuned to ue Internal Comments: [I checked the system operation other than the failue locations] Not applicable. [000000000000000000000000001k report (paper) sumission] Not applicable. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: u Kagoshima of Fuda Electronics Device ue: Not ued for clinical ue. Eupected and ueupected Resuts: Loaner uer Impact: None Patient Impact: None SE2.00000000.23 Resoluion: [Response details] * As it is a loaner case/000000000000000000000000001 no repair 000000000000000000000000001k has occured. Juy 31, 2000000003 / I sent a sustitue DFM100000000 (F-00000000700000000 to the Kagoshima Sales Office of The Company's SeibuMinami Sales Office. On November 100000000 2000000003, the sustitue DFM100000000 (F-00000000700000000 was retuned from the Kagoshima Sales Office of The Company's SeibuMinami Sales Office. [Simutaneou FCO 000000000000000000000000001k] There is no FCO on the device in which 000000000000000000000000001k is requred. ----------------------------------------------------------------------------- | 2023-09-21 12:51 | No
1003 | Yes | Monitor malfunction. | Monitor malfunuon. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000721 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Resoluion : Parts replaced: SBC board, Trzeps, SD card & system board. New gap fill pad fitted, uit tested OK. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: Monitor fails to boot u. Stays on self test screen. After contact with Bradley Guver 00u, we tried to ugrade software. This attempt failed. uer Impact: n/a Patient Impact: n/a SE n/a Parts Consued : 1.0000000000000000000000000000000056000000005000000001,RDT SBC kit (for 00000000-100000000u-R & 00000000-100000000u-R) 1.000000000000000000000000000000005600000000000000001,RDT Invasive Pressue Modue Kit Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000111552 Robert Jon de Ste Croiu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Reloaded SW with original Trizeps and now 000000000000000000000000001king correctly. Waiting on RDT to confirm whether or not to replace the Trizeps, system board & SBC. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: Monitor fails to boot u. Stays on self test screen. After contact with Bradley Guver 00u, we tried to ugrade software. This attempt failed. uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : Reloaded SW with original Trizeps and now 000000000000000000000000001king correctly. Waiting on RDT to confirm whether or not to replace the Trizeps, system board & SBC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003738 Robert Jon de Ste Croiu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uit is trying to udate bu fails on the Firmware section of the ugrade. Replacing the SBC is ulikely to fiu this issu. Checked the Trizeps board to make sue it is seated correctly which it is.Contacted RDT for gudance. Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: Monitor fails to boot u. Stays on self test screen. After contact with Bradley Guver 00u, we tried to ugrade software. This attempt failed. uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : uit is trying to udate bu fails on the Firmware section of the ugrade. Replacing the SBC is ulikely to fiu this issu. Checked the Trizeps board to make sue it is seated correctly which it is.Contacted RDT for gudance. Reguatory Qustions : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000977 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u Single Board Compuer 0000000056000000005000000001 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: Monitor fails to boot u. Stays on self test screen. After contact with Bradley Guver 00u, we tried to ugrade software. This attempt failed. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u Single Board Compuer 0000000056000000005000000001 ----------------------------------------------------------------------------- Internal Remarks : Tracking Nuber:1Z3R6000000006696395868 COLLECT TOMORROW File uloaded by CS Portal uer : File Name: 20000000030000000013_10000000026.jpg. Comments: File(s) uloaded duing case creation ----------------------------------------------------------------------------- | 2023-09-21 12:56 | No
1004 | Yes | The alarms do not work, letters have been enlarged and everything cuts out. | The alarms do not work, letters have been enlarged and everything cuts out. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 28-Sep-2000000003 Alejandro Rodrigo Mauier Córdoba The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Clinical ue Eupected and ueupected Resuts: Change of screen resoluion uer Impact: Bad display of the central station. Patient Impact: NO SE C.00000000.00000000 Resoluion: Service reports change of screen resoluion. Euternal speaker does not soud. OS configuation and screen resoluion are performed. Printer installed locally. Display disposition is reconfigued. Verification of correct operation. Tools & Equpment: PATIENT SIMuATOR, 200000000000000001-18 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-21 13:06 | No
1005 | Yes | Wireless footswitch issues | Wireless footswitch issues | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : Wireless footswitch issus Safety Qustions : Case: 000000002230000000092 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Wendy Cheug Last udated by: Wendy Cheug 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Case Transferred to SAP FSE Notes : This PO is a Reference Nuber only SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-21 13:31 | No
1006 | Yes | uL+ operational check failure | uL+ operational check failure | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008599 22-Sep-2000000003 Vishe, Rahu The system meets the specification for the performed service and is retuned to ue Resoluion : Clean the eut paddles and done operation check, machine is 000000000000000000000000001king fine. Handover the uit to cutomer in 000000000000000000000000001king condition. Problem Description by Engineer : Cutomer Fuction/Role: bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008552 Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Fse visit requred Problem Description by Engineer : Cutomer Fuction/Role: bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Fse visit requred ----------------------------------------------------------------------------- | 2023-09-21 13:46 | No
1007 | Yes | Speaker INOP uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Speaker INOP uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008300000000Bittlinsky, Simon The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: u3 Louspeaker INOP uer Impact: no Patient Impact: no SE uknown did the louspeaker have no soud anymore or was a slight, possibly distorted soud or a beeping soud heard? • No, were any INOP uLLmessages displayed? If an INOP uLLwas displayed, was auio present duing this message? • Message: Louspeaker uLL yes, did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • Connected to Mu7500000000 When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? • On the patient, which software was installed on the device when the louspeaker malfuction occured? • N 00000000.00000000100000000 Diagnostic performed by Engineer: Louspeaker uLL cutomer 000000000000000000000000001ld like to repair the u3 himself. 2u 0000000056000000003831 Mu_100000000/u3 Speaker assembly 1u louspeaker for the repair, answered 1u louspeaker for his stock qustions from the cutomer: Did the louspeaker have no soud anymore or coud a slight, possibly distorted soud or a beeping soud be heard? • No, were any INOP uLLmessages displayed? If an INOP uLLwas displayed, was auio present duing this message? • Message: Louspeaker uLL yes, did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • Connected to Mu7500000000 When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? • On the patient, which software was installed on the device when the louspeaker malfuction occured? • N 00000000.00000000100000000 Follow u Requred: Nemo KVA 2u 0000000056000000003831 Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000052 Parts Consued: 2.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- Cutomer comment: Via ##sender <Michael.Scholz00esudheitnord.de> Dear Sir or Madam, please note that you message(s) are not processed duing my absence and will not be forwarded! I cannot be reached from 22.00000000.2000000003 u to and incluing 00000000.1000000002000000003. In especially ugent cases, youcan reach my colleagus Dauh, Sebastian sebastian.dauh00esudheitnord.de 000000001/00000000-61337 Glogner-ubeer, Raik raik.glogner-ubeer00esudheitnord.de 000000001/00000000-2115 Cutomer comment: By ##from Michael.Scholz00esudheitnord.de: "-WITH THE REQuST FOR RESPONSE! - Dear u Bittlinsky, ladies and gentlemen, thank youfor you response. I also inclued the requsted answers in you ##below." Internal Remarks: Price/uit 83,100000000u Cutomer comment: VIa ##sender <Raik.Glogner-ubeer00esudheitnord.de> ladies and gentlemen, thank youfor you ## My answers, uder you qustions on the original ## Thank youvery muh. Did the louspeaker have no soud or coud a light, possibly distorted soud or a beeping soud be heard? ... • No, were any INOP uLLmessages displayed? If an INOP uLLwas displayed, was auio present duing this message? • Message: Louspeaker uLL yes, did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • Connected to Mu7500000000 When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? • On the patient, which software was installed on the device when the louspeaker malfuction occured? • N 00000000.00000000100000000 Cutomer comment: Addend on CA-870000000030000000000000000 Response from the cutomer in the teut of the ## Dear u Scholz, you order indicates that youneed t000000000000000000000000001louspeakers for the u3. Coud youplease answer the following qustions for processing the case? If youhave t000000000000000000000000001defective monitors, 000000000000000000000000001ld I need the answers for both monitors? If t000000000000000000000000001devices are affected, I need the 000of the second device. NO, ONLY ONE MONITOR IS AFFECTED. WE NEED THE SECOND LOuSPEAKER AS A SPARE PART FOR DIRECT ACCESS IN Ou 000000000000000000000000001KSHOP. If only the uLLmessage "louspeaker uLL is displayed in the monitor and the louspeaker still 000000000000000000000000001ks, I 000000000000000000000000001ld ask youto perform a cold start of the monitors. If the uLLmessage still euists after the cold start, I need the answer to the following qustions: THANK YOuFOR THIS NOTE. THE DESCRIBED uLLDID NOT OCCu AGAIN AFTER THE COLD START! QuSTION: CAN WE THEREFORE GIVE THE MONITOR INTO SERVICE AGAIN? IN ADDITION, I 000000000000000000000000001LD STILL LIKE TO ANSWER THE FOLLOWING QuSTIONS: Did the louspeaker have no soud anymore or was a light, possibly distorted soud or a beeping soud heard? •NO; THE LOuSPEAKER 000000000000000000000000001KS WITH NORMAL VOLuE. Were any INOP uLLmessages displayed? NO. If an INOP uLLwas displayed, was auio present duing this message? • - Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • CONNECTED TO A HOST MONITOR Mu7500000000 When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? • DuING THE APPLICATION ON THE PATIENT. Which software was installed on the device when the louspeaker malfuction occured? • N.00000000.00000000----------------------------------------------------------------------------- | 2023-09-21 14:06 | No
1008 | Yes | DFM0 therapy delivery failed during functional check | DFM0 therapy delivery failed during funuonal check | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008729 27-Sep-2000000003 RON FRISSEN The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Mark Rol Device ue: ouside of ue yes Eupected and ueupected Resuts: auotest/fuctional test failed uer Impact: cannot ue defib Patient Impact: N/A SE 2.00000000.21 Resoluion: SOM modue, therapy board replaced and condenser replaced. Reloaded software /config and performed detailed fuctional check. Safety test. See also 00000000000000000000000000100000000000000008792 Tools & Equpment : ELECTRICAL SAFETY TESTER, 20000000000000000000000003 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER, 200000000000000005-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008792 26-Sep-2000000003 RON FRISSEN The system does not meet the specifications for the performed service and is not retuned to ue Follow u Requred: Therapy board and uuuitor replaced. SOM board still needs to be replaced, load software and load config still needs to happen. The inside tray cannot be placed in the device, it was also difficut to get ou. I will take the defib to HTC52 for follow-u action. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: ouside of ue Eupected and ueupected Resuts: auotest/fuctional test failed uer Impact: cannot ue defib Patient Impact: N/A SE 2.00000000.21 Diagnostic performed by Engineer: Therapy board and uuuitor replaced. SOM board still needs to be replaced, load software and load config still needs to happen. The inside tray cannot be placed in the device, it was also difficut to get ou. I will take the defib to HTC52 for follow-u action. See 00000000000000000000000000100000000000000008729 Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuity 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY SOM Tools & Equpment : DEFIBRILLATOR ANALYZER, 200000000000000005-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 20000000000000000000000003 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008658 PATRICK POELS The system meets the specifications for the performed service and is retuned to ue Follow u Requred: Please schedue an available engineer (Arjan Emmelkamp) for the on-site repair of the DFM100000000. Estimated 000000000000000000000000001king hous: 00000000ous Requred parts: 0000000056000000009000000001 - DFM100000000 Therapy PCA ASSY 0000000056000000009000000001 - DFM100000000 uuuity ASSY 0000000056000000009000000001 - DFM100000000 SOM PCA ASSY Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: ouside of ue Eupected and ueupected Resuts: auotest/fuctional test failed uer Impact: cannot ue defib Patient Impact: N/A SE 2.00000000.21 Diagnostic performed by Engineer: Checked log files and foud the following uLL: (7:30000000000000000 Pacer RFu HV uuuitor Valu LOW (89.73u) (7:38) Therapy RFuPacing Test has failed The following parts need to be replaced: 0000000056000000009000000001 - DFM100000000 Therapy PCA ASSY 0000000056000000009000000001 - DFM100000000 uuuity ASSY 0000000056000000009000000001 - DFM100000000 SOM PCA ASSY ----------------------------------------------------------------------------- | 2023-09-21 14:16 | No
1009 | Yes | The system is down | The system is down | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000875000000002-Sep-2000000003 Samule, Villa The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: The system is down for problems with the Image Processing section. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.1.17.00000000Parts Consued: 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD Resoluion: Replaced PC, carried ou tests with a positive oucome ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008535 SAVERIO DIVITTORIO The system does not meet the specification for the performed service and is not retuned to ue. Internal Comments: uLL# and/or description of uLL POST of the 3 disks OS, Image1 and Image2 failed Repair Action: . Resoluion: Intervention will follow to restore the system. Follow u Requred: uLL# and/or description of uLL POST of the 3 disks OS, Image1 and Image2 failed Repair Action: . What are the neut Trouleshooting steps: We already intervened replacing the 3 HD and the IP-PC for the same problem. The LOD shows POST failed on all 3 HD, so the PC seems to be defective, or better, the disk bay. I order a new IP-PC, the cutomer asks for ugent intervention today to restore the system, evalute with Meroni if it is necessary to go today and try to restart them bypassing the disk bay or wait for the arrival of the spare part. I sugest bringing an uB/SATA adapter with youin order to format the system HD, if need be. The system is down. Technician Name: SAVERIO DIVITTORIO Problem Description by Engineer: Problem description by engineer: The system is down for problems with the Image Processing section. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as for the cutomer uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 8.1.17.00000000Diagnostic performed by Engineer: Heard from u Meroni, the system is down again after yesterday's intervention. ----------------------------------------------------------------------------- | 2023-09-21 14:16 | No
1010 | Yes | Therapy delivery test fail (Replaced ASSY uuuitance) | Therapy delivery test fail (Replaced ASSY uuuance) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005100000000Madhan, kuar The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : To replace 0000000056000000009000000001, DFM100000000 uuuitance ASSY Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001, DFM100000000 uuuitance ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008900000000 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-09-21 14:51 | No
1011 | Yes | Replacement of foot switch | Replacement of foot switch | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000153 September 2000000002000000003 Bihr, Dominik The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Foot switch replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 Parts Consued: 1.00000000000000000000000000000000800000000238232, wireless foot switch 3P ----------------------------------------------------------------------------- | 2023-09-21 14:56 | No
1012 | Yes | Showing audio failed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Showing audio failed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008291 Nandha, Rahu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : speaker assy to be replaced Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : speaker assy to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009100000000 Parts Consued : 1.0000000000000000000000000000000056000000007821,SS CBL Speaker8 Ohm 36mm f000000003800000000z 5.8mm ----------------------------------------------------------------------------- | 2023-09-21 15:06 | No
1013 | Yes | No more movement | No more movement | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 September 22, 2000000003 Nicolas ARNAu The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: yes Eupected and ueupected Resuts: geometry issu uer Impact: system not 000000000000000000000000001king Patient Impact: No. uuResoluion: Replacement of the motor management board of the C-arch (AMC triple drive 2). Reloaded the software of the pc sute. Reloaded the parameters. Configuation, adjutments and operating tests OK. The device is fuly operational again. Consued parts: 1.00000000000000000000000000000000800000000990000000052, AMC Triple Servo Drive HP-LP-LP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000Nicolas ARNAu The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: CDP 00000000800000000990000000052 AMC triple servo Drive quntity 1 Delivery to biomedical department u Samson 000000005299000000006 Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: yes Eupected and ueupected Resuts: No geometry uer Impact: system not 000000000000000000000000001king Patient Impact: No. uuDiagnostic performed by Engineer: Diagnostic performed in remote maintenance with the help of the biomedical technician. Electronic control board of the motorization of the arch is defective. Replacement of the card and tested reloading the software of the new card. ustable IT system which does not allow to restart the system => the software of the main PC (following PC). On-site intervention requred in order to finish the configuation and the settings. ----------------------------------------------------------------------------- | 2023-09-21 16:11 | No
1014 | Yes | Tube early failure | Tube early failure |  | 2023-09-21 16:13 | No
1015 | Yes | Equipment is showing uLLand it is not starting | Equipment is showing uLLand it is not starting | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Trouleshooting Action: Schedued with client: Clinical Engineering Schedued date/time: 00000000/2000000002000000003 at 8am SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Problem description by engineer : IPPC does not Boot Trouleshooting Action: What steps are taken duing the system diagnosis : Attempt to reinstall the IPPC system software failed, stopping at a screen that does not allow youto do anything. The HD will have to be formatted. Send FSE to format HD and reinstall Software uLL# and/or description of uLL: IPPC boot uLLOneEMS internal remarks: Problem description by engineer : IPPC does not boot What steps are taken duing the system diagnosis : An attempt to reinstall the IPPC system software failed, stopping at a screen that does not allow youto do anything. The HD will have to be formatted. Send FSE to format HD and reinstall Software uLL# and/or description of uLL: IPPC boot uLLRepair Action "Internal" : Send FSE+MAQuNA PARADA - FSE Romuo going to hi website, send S000000000000000000000000001to it Recommendations to FSE ► Connect PC diag tool and analyze IPPC; ► If OK, format system HD and reinstall Software ► test everything Eutra Internal comment field: RSE Marcel Ogassawara SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Problem: Equpment presenting an uLLand not starting Priority: Stopped uit: Albert Einstein Address: Avenida Albert Einstein 627 Sao Pauo-SP 00000000652-0000000000000000BR Responsible: André Sector: Clinical Engineering 001: +551197153000000003 002: E -mail: andre.bianchi00insten.br Safety Qustions: Case: 00000000223000000008300000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Adriano Santos Last udated by: Adriano Santos 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 20000000000000000.2000000003 21:1000000008 POLARIS (POLARIS) ======================== Romuo Soua - 0000000000000000573720000000000000000200000000000000000000000065 ===================== == Verified that the IPPC had corruted software; HDs formatted: System and images, ok; Software installation, ok; General tests, ok; Monitoring euams, ok; Operational equpment and complies with manuactuer specifications. The Service was completed in accordance with the u Qulity Policy. Tests and Calibrated Tools in PA Tool. ----------------------------------------------------------------------------- | 2023-09-21 16:36 | No
1016 | Yes | 00uuu Defective monitor | 00CFuuu Defeu monitor | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005000000009 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Defective euam room monitor When was the first occurence? Within... : 200000000Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: Defective euam room monitor have a temperory monitor in place uer Impact: Defective euam room monitor Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : Jimmy reported that the euam room monitor on boom is defective 000000000000000000000000001't tun on and reboot does not resolve. Requsting parts id and part ordering for replacement. Internal Comments : uLL# and/or description of uLL: Defective euam room monitor Repair Action : Provided p/n and ordered part for biomed to install. 991932252193 - 19'' Monochrome Monitor ER (MML1952-PER) Resoluion : Provided p/n and ordered part for biomed to install. Follow u Requred : uLL# and/or description of uLL: Defective euam room monitor Repair Action : Provided p/n and ordered part for biomed to install. 991932252193 - 19'' Monochrome Monitor ER (MML1952-PER) Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005000000007 Parts Consued : 1.000000000000000000000000991932252193,19'' Monochrome Monitor ER (MML1952-PER) ----------------------------------------------------------------------------- Internal Remarks : ** Interaction Type: New Case Caller: JIMMY MEREDITH IP & Issu: 55000000008 CuTOMER REQuSTING TECH SuPORT - PARTS ID FOR MONITOR Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND REMOTE 000000000000000000000000001 RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-21 16:41 | No
1017 | Yes | 00 CN00- request replacement of adapter cables. | u- request replacement of adapter cables. |  | 2023-09-21 17:12 | No
1018 | Yes | Room is down, Stays in Loading screen | Room is down, Stays in Loading screen | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000008 00000000-Oct-2000000003 Eckbold Glenn The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System not booting past 00000000:00000000 When was the first occurence? Within... : 21-Sep-2000000003 00000000:0000000019 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System not booting past 00000000;00000000 uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replace fleuvision pc ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005585 25-Sep-2000000003 Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System not booting past 00000000:00000000 When was the first occurence? Within... : 21-Sep-2000000003 00000000:0000000019 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System not booting past 00000000;00000000 uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Follow u Requred : Install pc Diagnostic performed by Engineer : Swapped drive to bay 3 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000280000000023-Sep-2000000003 Washko Ryan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace CPuProblem Description by Engineer : Problem description by engineer : System not booting past 00000000:00000000 When was the first occurence? Within... : 21-Sep-2000000003 00000000:0000000019 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System not booting past 00000000;00000000 uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Arrived on site and issu with delivered part. Will create follow u and notify DSM. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000027200000000offman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: n/a Repair Action : ue alt + f duing boot u to log into FSF and check log file for grabber card uLL or commuication uLL with fleu pc. If needed, 0000000080000000012200000000 - u Haswell Fleuvision 2 PC no HDD and/or 0000000080000000050000000000000000- Hard disk spare part. Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : ue alt + f duing boot u to log into FSF and check log file for grabber card uLL or commuication uLL with fleu pc. If needed, 0000000080000000012200000000 - u Haswell Fleuvision 2 PC no HDD and/or 0000000080000000050000000000000000- Hard disk spare part. Technician Name : Coffman, Josh Diagnostic performed by Engineer : Jonathan said this morning the system did not boot past 00000000:00000000. They tried rebooting the system bu did not resolve the issu. Problem Description by Engineer : Problem description by engineer : System not booting past 00000000:00000000 When was the first occurence? Within... : 21-Sep-2000000003 00000000:0000000019 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System not booting past 00000000;00000000 uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- Internal Remarks : Glen Eckbold - qustion abou loading SW to new FV2PC. ----------------------------------------------------------------------------- | 2023-09-21 17:21 | No
1019 | Yes | 00 CN00- Therapy disabled - SPLIT - REFERENCE PR 00 | u- Therapy disabled - SPLIT - REFERENCE PR 000 |  | 2023-09-21 17:25 | No
1020 | Yes | Check Vent: Prouimal Pressure Sensor Auto-Zero Failed" (Diagnostic code 0A) was confirm | Check Vent: Prouimal Pressure Sensor Auto-Zero Failed" (Diagnostic code 00) was confirm |  | 2023-09-21 17:58 | No
1021 | Yes | SpO0larm failure. No Trouble Found-Evaluation Confirmed Device Working. No fault found. | SpO0alarm failure. No Trouble Found-Evaluation Confirmed Device Working. No fault found. |  | 2023-09-21 18:00 | Yes
1022 | Yes | "Check Vent: O0upply Pressure Sensor Range uLL alarm occurred during u on the patient | "Check Vent: O0Supply Pressure Sensor Range uLL alarm occurred during u on the patient |  | 2023-09-21 18:08 | No
1023 | Yes | uconnections are out of work. | uconneuons are out of work. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000035800000000ullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Hello, I am contacting youregarding the TEMPu PRO monitor case. Coud yousend me a pictue of the broken uconnection problem in the device? Thank you Sincerely, Gullaue PANET Resoluion: Hello, It is an LS tempu Please send me the qute as soon as possible. Best regards. Thibau APPAMON Follow u Requred: Qute to be issud. Quntity(s): u1 per item 22% discout (approved by Stéphane BECKER) Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the uconnection. ----------------------------------------------------------------------------- | 2023-09-21 18:11 | No
1024 | Yes | System alarms has no sound - No product Malfunction- inadequate network coverage-provided information | System alarms has no sound - No produuMalfunuon- inadequate network coverage-provided information | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005100000000 Oct 13, 2000000003 Huger, Johannes The system meets the specification for the performed service and is retuned to ue Resoluion: 13.100000000 8:00000000 am: u Grudel is asking for a call-back again. The reason for the message is now present. The mobile parts receive the alarms, bu only as vibration and visully. No acoutic alarm occus. The mobile parts are suported by EDP and they mut ensue that the mobile parts are not mued. This uully happens via a mobile device management (MDM). u Grudel reports the uLLto in-houe EDP. From the point of view of u, no uLL euist. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Device/component is reported as defective. uer Impact: no Patient Impact: no SE uknown Issu: Malfuction Description of uLL No alarms soud on the Mu00000000anymore. Diagnostic performed by Engineer: Reported problem coud not be reprodued: According to u Grudel, alarms shoud not arrive on the CareEvent mobiles. On 10000000012, 3:00000000 pm, sucessfuly tested the commuication together with u Grudel. Duing the test, no futher problems were detected. Device is fuly fuctional." ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000063 Hübler Thomas The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: +++Onsite KVA+++ The connection between the control center and care event mut be checked (net000000000000000000000000001k). Curently there is no alarm forwarding from the Mu00000000to Care Event mobile parts. Parts: No Time: 3 hous Diagnostic performed by Engineer: +++Onsite KVA+++ The connection between the control center and care event mut be checked (net000000000000000000000000001k). Curently there is no alarm forwarding from the Mu00000000to Care Event mobile parts. Parts: No Time: 3 hous Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Device/component is reported as defective. uer Impact: no Patient Impact: no SE uknown Issu: Malfuction Description of uLL No alarms soud on the Mu00000000anymore. Diagnostic performed by Engineer: Phone cutomer. The problem does not seem to refer to the Mu00000000 The missing alarm refers to the Care Event mobile parts. The system is not connected to Remote. No trouleshooting possible. Onsite for uLLanalysis, connection control center to Care Event. Configuation alarm behavior on Care Event. ----------------------------------------------------------------------------- Internal Remarks: Call-back u Grudel 13.1000000002000000003 8:00000000 am. u Grudel spoke with the uers again. The alarms come on the terminal, bu there is no acoutic alarm. u Grudel will contact EDP for MDM, so that the terminals can no longer be deenergized. Internal Remarks: Call-back u Grudel on 10000000012 3:00000000 pm. Checked the commuication together with u Grudel, alarms are forwarded to the terminal. No nsite requred. Cutomer comment: u Grudl will call, he spoke with u Huker yesterday and is asking to speak with him again. Requst for call-+00000000100000000995500000000Cutomer comment: By fau Confirmed the Q-00000000139000000003 / order MTBI 2000000003/000000007 / u Grudl - medical technician - tel.: +000000007100000000995500000000Cutomer comment: u Huger will call u Grudl back. Cutomer has a contract, canceled qute. Internal Remarks: u Streur will contact the team as to who/when can take over. ----------------------------------------------------------------------------- | 2023-09-21 18:26 | No
1025 | Yes | Degraded image during the euamination, impossible to continue. | Degraded image during the euamination, impossible to continue. |  | 2023-09-21 18:36 | No
1026 | Yes | uPro ¦ touchscreen issue | TempuPro ¦ touchscreen issue | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Wilbu, Phillip The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Reached ou to the PSS. ##d cutomer the following information: The PSS Sugested to confirm inter-board connectors are mated properly. “Especially the system board (E) (see pictue).” “I am supecting the ambient temp issu is occuring with only the SBC fitted, this is normal. Once the system PCB is fitted and the device powered on again, this uLLshoud not appear.” Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer confirmed resoluion Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: Mutiple uLL after opening device to confirm touhscreen connections were seated correctly. uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Mutiple uLL after opening device to confirm touhscreen connections were seated correctly. Diagnostic performed by Engineer : Problem Description: Mutiple uLL after opening device to confirm touhscreen connections were seated correctly. Tempu Pro: Mutiple uLL after opening device to confirm touhscreen connections were seated correctly. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 9/2000000002000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Biomed called to report Tempu Pro is displaying Mutiple uLL after opening device to confirm touhscreen connections were seated correctly. "Ambient Temp too High," uDisable, NBP disabled, Etco2 disabled, Printer disabled. ##d cutomer copy of Service Manul an provided cutomer part nuber for: 000000005600000000111 RDT Display Assembly Kit Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: John Haller IP & Issu: 93297000000002/ Mutiple uLL- touhscreen issus Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-21 18:41 | No
1027 | Yes | Prouimal pressure sensor is bad | Prouimal pressure sensor is bad | 000000000000000000000000001O##: 000000000000000000000000001000000000000000028100000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: prouimal pressue line disconnect uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and its displays a prouimal pressue line disconnect alarm. He is uable to get into standby either. He stated the average air is reading -239 slmp. Recommended to replace the flow sensor assy. Part #989800000000617100000000 flow sensor assy, list price $1217.00000000 Resoluion : Cutomer will order the flow sensor assy and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: chase ouley IP & Issu: 6250000000000000000prouimal pressue sensor is bad Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created case/ R000000000000000000000000001and roued to RSE/informed rp Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-21 18:46 | No
1028 | Yes | not turning on | not turning on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008681 Amjad Aburqou The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System stopped in 00000000:00000000 time duing switching on. How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 8.2.3000000005 Diagnostic performed by Engineer : System stopped in 00000000:00000000 time duing switching on. Internal Comments : uLL# and/or description of uLL: System stopped in 00000000:00000000 time duing switching on. Repair Action : System stopped in 00000000:00000000 time duing switching on. Resoluion : Gude the cutomer to switch on the Fleuvioisn PC from the cabinet after that system retun 000000000000000000000000001king fine withou any problem. cutomer restart the system many times and it is 000000000000000000000000001king fine. ----------------------------------------------------------------------------- | 2023-09-21 18:46 | No
1029 | Yes | Therapy disabled uLLis appearing (Replaced ASSY uuuitance) | Therapy disabled uLLis appearing (Replaced ASSY uuuance) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : 0000000056000000009000000001 - ordered uuuitor. Follow u Requred : 0000000056000000009000000001 - uuuitor ordered. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Therapy disabled uLLis appearing ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003889 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-09-21 19:51 | No
1030 | Yes | They didn't have the scopy or the spelling. | They didn't have the scopy or the spelling. | 000000000000000000000000001O#Nuber: 0000000000000000000000000010000000000000000000000001 September 21, 2000000003 Jaime Granja, The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Radiographer Device ue: Cardiac intervention. Eupected and ueupected Resuts: No u-rays. uer Impact: uable to ue the system for image generation. Patient Impact: Euamination postponed. SE 000Resoluion: The equpment was already 000000000000000000000000001king normally the morning of the following day, when it occured. The occurence coincided with a power ouage. It was not possible to determine the caue of the reported problem. ----------------------------------------------------------------------------- | 2023-09-21 20:06 | No
1031 | Yes | uL does not pass the pacemaker test | uL does not pass the pacemaker test |  | 2023-09-21 20:21 | No
1032 | Yes | System hangs at 0:0 | System hangs at 0:0 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000375000000005-Sep-2000000003 Juian Brandon The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Ouside of ue Eupected and ueupected Resuts: System jut booting to 0000000000000000 at start of day uer Impact: Cannot ue system Patient Impact: NA SE 8.1.300000000Diagnostic performed by Engineer : TSM Modue 000000000000000000000000001ld not connect to net000000000000000000000000001k. i tried resetting nic card and still 000000000000000000000000001ld not 000000000000000000000000001k correctly. replaced with new tsm modue and loaded sw. system is now 000000000000000000000000001king properly. Parts Consued : 1.0000000000000000000000000000000080000000000000000393,Proscribe to TSM repl.kit Patient Table Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005979 22-Sep-2000000003 Jessen Mike The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Ouside of ue Eupected and ueupected Resuts: System jut booting to 0000000000000000 at start of day uer Impact: Cannot ue system Patient Impact: NA SE 8.1.300000000Diagnostic performed by Engineer : Replaced Geo PC. System was now starting u with no geometry uLL bu TSMB on table was not booting u. Tried to reload SW on TSMB bu it 000000000000000000000000001ld not load. Spoke with NSS to get approval for replacement TSMB. Follow u Requred : Replace TSM B Parts Consued : 1.00000000000000000000000000000000800000000000000007813,GEO Q35 IPC COA Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000000000000022-Sep-2000000003 Juian Brandon The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : After trouleshooting foud that the power suply in the geometry pc was cauing a fue to blow in main cabinet cauing mutiple power failues all over the system. New compuer has been ordered for neut day Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 21-Sep-2000000003 Jessen Mike The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Ouside of ue Eupected and ueupected Resuts: System jut booting to 0000000000000000 at start of day uer Impact: Cannot ue system Patient Impact: NA SE 8.1.300000000Diagnostic performed by Engineer : 000000000000000000000000001ked with other FSE. Foud F12 blown on MPDu When fue was replaced, the fue blew again once the system was powered on. Trouleshot oupu and foud that Geo PC was not powering u while the rest of the cabinet was. Geo PC made lou pop when power was applied. Ordering new Geo PC. Follow u Requred : Replace Geo PC Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Remove front covers of cabinets and verify all compuers and net000000000000000000000000001k switch are powered on. A large portion of the time it is fleuvision pc not commuicating. ue alt + f duing boot u to start fsf. Check POST resuts, download board software, and event log to help determine the root caue. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Debriefed caller. She said they have already spoken to FSE and he is on the way. System stays at 00000000:00000000 on coutdown clock. ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Mike Jessen with trouleshooting TSMB not booting u and approved replacement TSMB kit to try a new touhscreen 0000000080000000000000000393 - Proscribe to TSM repl.kit Patient Table ----------------------------------------------------------------------------- | 2023-09-21 20:31 | No
1033 | Yes | System does not boot stuck at 0:0. | System does not boot stuck at 0:0. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 23-Sep-2000000003 KITTI SuAN The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Waiting PO from cutomer to order low voltage power suply for replacement. Tools & Equpment : MuTIMETER,20000000000000000000000009 00000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : System does not boot stuk at 00000000:00000000. uer try to restart main system bu it not solve the problem How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician Device ue: CV Eupected and ueupected Resuts: System does not boot stuk at 00000000:00000000. uer Impact: Not impact Patient Impact: Not impact SE 8.2.17.00000000Diagnostic performed by Engineer : Measued low voltage oupu from power suply in M-cabinet bu not presented. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000HATCHAI SANTANA The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Generic CAN Interface non recoverable uLL [FREE_TEuT] [FREE_TEuT] Can not connect to the SIB. The SIB is not available or the slotnuber of the SIB is not correct configued. [SEVERITY] uLL [SEVERITY] uLL Commuication uLL Received uLLfrom commuication software. IP Address not foud. [FREE_TEuT] [SEVERITY] uLL[EuCEPTION_DESCRIPTION] On a biplane system the MAC adress of both frontal and lateral Certeray mut be configued. [SEVERITY] uLL Repair Action : Replace DC power suply at rear panel in M cabinet. Follow u Requred : uLL# and/or description of uLL: Generic CAN Interface non recoverable uLL [FREE_TEuT] [FREE_TEuT] Can not connect to the SIB. The SIB is not available or the slotnuber of the SIB is not correct configued. [SEVERITY] uLL [SEVERITY] uLL Commuication uLL Received uLLfrom commuication software. IP Address not foud. [FREE_TEuT] [SEVERITY] uLL[EuCEPTION_DESCRIPTION] On a biplane system the MAC adress of both frontal and lateral Certeray mut be configued. [SEVERITY] uLL Repair Action : Replace DC power suply at rear panel in M cabinet. What are the neut Trouleshooting steps : Check DC power suply. Technician Name : CHATCHAI SANTANA Problem Description by Engineer : Problem description by engineer : System does not boot stuk at 00000000:00000000. uer try to restart main system bu it not solve the problem How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer : Ethernet switch in M cabinet and FDC is has no power. Need engineer on site for check DC power suply. ----------------------------------------------------------------------------- | 2023-09-21 20:41 | No
1034 | Yes | uuudiscoloration of defib pads ud before electroconversion | u discoloration of defib pads ud before eleuoconversion |  | 2023-09-21 20:42 | No
1035 | Yes | PSE & PSA_uartifacts - potential SI - GFE SENT 0 (0DEC 0) - F/U (0DEC 0) -RMR F/U (0DEC 0 | PSE & PSA_uartifau - potential SI - GFE SENT 0- (0 DEC 00 - F/U (0 DEC 00 -RuF/U (0 DEC 0) |  | 2023-09-21 20:52 | Unknown
1036 | Yes | Machine and Prouimal pressure sensors failed (EC 00 | Machine and Prouimal pressure sensors failed (EC 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003251 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: uLLcode 1000000007 uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states that FSA was replaced and now device is giving uLLcode 1000000007. Troule shot with cutomer. had cutomer confirm DA to MC board cable is properly seated. Cutomer foud cable was fuly seated. cutomer seated cable, tested device, device is now powering on withou uLL cutomer states he will finish testing device and callback if futher assistance is needed. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: michael oliver IP & Issu: 71132226/uLL1000000007 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created case/ R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-21 20:56 | No
1037 | Yes | Duplicate of PR00 | Duplicate of PR0000 |  | 2023-09-21 20:57
1038 | Yes | Unit will not power ON | Unit will not power ON | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: with AC connected, no LED's illuinated on the front of the vent. uit will not power on. uer Impact: none Patient Impact: na SE uknown Diagnostic performed by Engineer : Foud no power ou of the power suply. Foud good AC power to the power suply. Advised to replace the power suply. Battery reading .18vdc. Advised to replace the battery. Battery, internal 1000000006370000000089800000000626900000000 Power suply (1st generation) 100000000000000009 00000000561512591 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gary Cole IP & Issu: 62520000000052/uit will now power u/not showing that its pluged in to AC/cutomer requsted tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-21 21:01 | No
1039 | Yes | Top right button is not working | Top right button is not working | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: top right buton not 000000000000000000000000001king properly uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the top right buton not 000000000000000000000000001king properly. Troule shot with cutomer. cutomer states he removed bezel and confirmed cabling is connected properly. Cutomer is requsting part nuber and price for replacement Switch PCBA. Resoluion : Provided cutomer with part nuber and price: 00000000561512600000000 - RP-PCBA, Switch, V600000000- 171.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Caller: Trent 00000000 628 6669 IP & Issu: 67196000000007 Top right buton is not 000000000000000000000000001king Event Type: Incident Service Requst Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created Remote W/O Roued to RTAC Action Taken: N/A Verbal PO Nuber: ----------------------------------------------------------------------------- | 2023-09-21 21:06 | No
1040 | Yes | 00N0Stuck at 0:0 | 00CDN0 Stuck at 0:0 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 22-Sep-2000000003 Singh Amit The system meets the specification for the performed service and is retuned to ue Resoluion : Confirmed the reported failue. To resolve the issu I replaced the Fleuvision PC and reloaded SW. (The cutomer ordered the Fleuvision PC) Problem Description by Engineer : Problem description by engineer : System is stuk at 00000000:00000000 When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system will not start. stuk at 00000000;00000000 uer Impact: room down Patient Impact: no patient impact SE 8.2.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005600000000 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : Per conversation, it was determined that the fleuvision is not responding and possible grabber card is defective. Provided cutomer with part nuber for the hard drvie for the fleuvision. 0000000080000000050000000000000000- Hard disk spare part Problem Description by Engineer : Cutomer Fuction/Role: Cutomer call back Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Cutomer called back wanted trouleshooting assistance. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Eckbold Glenn The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System is stuk at 00000000:00000000 When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system will not start. stuk at 00000000;00000000 uer Impact: room down Patient Impact: no patient impact SE 8.2.3000000005 Diagnostic performed by Engineer : system not booting u. No log to look at. Kathy is looking at lights of fleuvision to determine if that is the issu. Internal Comments : uLL# and/or description of uLL: System is stuk at 00000000:00000000 Repair Action : Check system logs. Replace Fleu PC FleuVision-2_Revised PC no HDD Follow u Requred : uLL# and/or description of uLL: System is stuk at 00000000:00000000 Repair Action : Check system logs. Replace Fleu PC FleuVision-2_Revised PC no HDD Technician Name : Eckbold Glenn ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006000000009 Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: JJ Action Taken: I called FSE-Amit Singh regarding 0000000000000000000000000010000000000000000000000002 he accepted it for 9/22/23 from 8am-12pm with no travel time I called and spoke with cu Kathy Cameron abt scheduing, cu said the FSE will be onsite tomorrow Oucome: I dispatched the 000000000000000000000000001per FSE requst I pu FADB in the Planning Notes Last udated: 7:31PM EDT 9/21/23 Internal Remarks : Recommended FSE star with loading Fleuvision OS followed by Fleuvision APP throuh download board software since replacing PC. ----------------------------------------------------------------------------- | 2023-09-21 21:16 | No
1041 | Yes | QA#00000- cardiac conductance line is faulty | QA#0000000- cardiac conduunce line is faulty |  | 2023-09-21 21:49 | No
1042 | Yes | V0accept button is not functioning | V0accept button is not funuoning | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000has fuctional check buton uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000accept buton is not 000000000000000000000000001king Diagnostic performed by Engineer : Problem Description: V600000000check mark is not fuctional Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/21/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Confirmed with cutomer the correct part # for the switch PCBA Part # 00000000561512600000000 RP-PCBA, Switch, V600000000List price $171 Biomed is also requsting the part # for the CPucover tray Part # 00000000561531961 RP-Cover, CPuTray, V6uu List price $138 ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE: Interaction Type: New Case Caller: Melissa Hallihan - phone 600000000 978 3000000009 IP & Issu: 6000000009119 - Cutomer requsting suport - PARTS ID : the accept buton Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-21 21:51 | No
1043 | Yes | 0 failure with combination of Auuiliary alarm failed (EC 0); Alarm LED failed (EC00 | 0V failure with combination of Auuiliary alarm failed (EC 00); Alarm LED failed (EC00) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005000000001 12-Oct-2000000003 Klu, Allan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying mutiple uLLcodes Diagnostic performed by Engineer : Device had uLLcode 1100000000 and 1115. Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : Diagnostic codes 1115, 1100000000, 111B. Biomed has advised they are uable to place an order for a power management board. Down V600000000du to possible PM board failue. FCO86600000000000000006C – V600000000V600000000Plu uLL– uLLLoss of Fuction W/O Alarm Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying mutiple uLLcodes Diagnostic performed by Engineer : Problem Description: V600000000is displaying mutiple uLLcodes. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/21/2000000003 Trouleshooting Steps Taken: Diagnostic codes 1115, 1100000000, 111B. Biomed has advised they are uable to place an order for a power management board. Cutomer advised of the 00000000 generation power management board that is being installed into all V600000000s as part of recall. Cutomer is requsting onsite to resolve. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Samul Richardson Issu: 577152500000000ailue Codes:1115,1100000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/a Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-21 22:21 | No
1044 | Yes | System saying 0:0 | System saying 0:0 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007732 27-Sep-2000000003 Ward Randall The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System saying 00000000:00000000 When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u uer Impact: can not ue room Patient Impact: no patient impact SE 8.2.17.00000000Resoluion : New fleu vison pc ordered and fitted software reload to be done also Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-10000000031 00000000:00000000:00000000 MuTIMETER,2000000003-10000000030000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000510000000022-Sep-2000000003 Pickering Mark The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : On-site service for mark pickering on Monday 25th sept Problem Description by Engineer : Problem description by engineer : System saying 00000000:00000000 When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u uer Impact: can not ue room Patient Impact: no patient impact SE 8.2.17.00000000Diagnostic performed by Engineer : Issus with no boot on fleu vision Rtac sugested to direct connect the hard drive and test. Then new drive loaded and still issus. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000003-31 00000000:00000000:00000000 MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005000000005 Eckbold Glenn The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System saying 00000000:00000000 When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u uer Impact: can not ue room Patient Impact: no patient impact SE 8.2.17.00000000Diagnostic performed by Engineer : no diagnostics as there is no PRS Internal Comments : uLL# and/or description of uLL: System saying 00000000:00000000 Repair Action : • Tun off the system • Check orange LED BLC L100000000on Power Distribuion strip at back of B-cabinet when system is switched off (if LED is on, issu with MPDurelay for NP-u200000000). • Measue voltage on mPDus NP u200000000, replace mPDuif it is powered! • Tun on system, after system has booted u, uplu power to FleuVision PC and reconnect. If FleuVision PC boots u, the issu is very likely the MPDurelay for NP-u200000000. If FleuVision PC does not boot u, the issu is the FleuVision PC itself. Caue: There are many possible caues, bu focu is on the FleuVision-2 PC behavior. Please make sue that the issu is not caued by another PC, the net000000000000000000000000001k switch or cabling, or the mains suply to one of the PCs or other key components. If it is determined that the FleuVision-2 PC caued the symptom there are t000000000000000000000000001known situtions: • The FleuVision-2 PC does not boot u correctly. • One of the DVI grabber cards did not initialize properly. Remedy: • If the FleuVision-2 PC does not boot u correctly: 1. If the startu screen hangs after the timer is 00000000:00000000, open the rear of the B-cabinet. 2. Look at the LEDs at the rear side of the FleuVision PC. 3. The problem is not the FleuVsion PC when:  the blu LED is blinking  and, the green LEDs at the grabber cards are lit  and, the Post Statu LEDs have the pattern (B1) (see pictue)  and, there is no grabber channel uLLin the logging 00000000Replace the Fleuvision-2 PC when:  there are grabber channel uLL in the system logging  or, the LEDs are not as described in step 3.  or, if a red LED at the grabber card is lit Follow u Requred : uLL# and/or description of uLL: System saying 00000000:00000000 Repair Action : • Tun off the system • Check orange LED BLC L100000000on Power Distribuion strip at back of B-cabinet when system is switched off (if LED is on, issu with MPDurelay for NP-u200000000). • Measue voltage on mPDus NP u200000000, replace mPDuif it is powered! • Tun on system, after system has booted u, uplu power to FleuVision PC and reconnect. If FleuVision PC boots u, the issu is very likely the MPDurelay for NP-u200000000. If FleuVision PC does not boot u, the issu is the FleuVision PC itself. Caue: There are many possible caues, bu focu is on the FleuVision-2 PC behavior. Please make sue that the issu is not caued by another PC, the net000000000000000000000000001k switch or cabling, or the mains suply to one of the PCs or other key components. If it is determined that the FleuVision-2 PC caued the symptom there are t000000000000000000000000001known situtions: • The FleuVision-2 PC does not boot u correctly. • One of the DVI grabber cards did not initialize properly. Remedy: • If the FleuVision-2 PC does not boot u correctly: 1. If the startu screen hangs after the timer is 00000000:00000000, open the rear of the B-cabinet. 2. Look at the LEDs at the rear side of the FleuVision PC. 3. The problem is not the FleuVsion PC when:  the blu LED is blinking  and, the green LEDs at the grabber cards are lit  and, the Post Statu LEDs have the pattern (B1) (see pictue)  and, there is no grabber channel uLLin the logging 00000000Replace the Fleuvision-2 PC when:  there are grabber channel uLL in the system logging  or, the LEDs are not as described in step 3.  or, if a red LED at the grabber card is lit Technician Name : Eckbold Glenn ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Mark Pickering System not booting past 0000000000000000. FSE trouleshooted that the drive bay for the fleuvision PC was not booting u. Approved replacement FV PC Internal Remarks : Planner Initials: GM Action FSE FSE Mark Pickering requsted call to be dispatched for 9/21 already onsite Oucome: dispatched ----------------------------------------------------------------------------- | 2023-09-21 22:26 | No
1045 | Yes | 0B (Vent-Inop): Watchdog test failed | 0B (Vent-Inop): Watchdog test failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : code 100000000B logged. Reviewed sugested repair per the manul. Reviewed the event log and no other check vent or vent inop codes logged. Date and time stamps all look as eupected. Removed and inspected the DA to MC cable, reinstalled ensuing good connection at the DA pcba, then connecting to the MC pcba. Condition persists. Cutomer going to reinstall the original flow sensor assy. If issu persists will he states he is going to swap in a DA pcba for trouleshooting, then the MC pcba if needed. If issu resolves advised to call and requst replacement flow sensor uder 900000000day parts warranty. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: PM Service Eupected and ueupected Resuts: Vent inop after replacing the flow sensor for failed PM flow testing. uer Impact: none Patient Impact: na SE 2.300000000----------------------------------------------------------------------------- Internal Remarks : **CuT REQuSTED TECH SuPORT Interaction Type: New Case Caller: BEN M IP & Issu: 56629200000000 WATCH DOG uLL Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND REMOTE 000000000000000000000000001 RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-21 22:41 | No
1046 | Yes | now unit wont turn on | now unit wont turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005258 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : DO NOT QuTE-900000000DAY WARRANTY (reference 00000000000000000000000000100000000177326) uit does not power on when pluged into AC power, FCO86600000000000000006 PM pcb replacement was recently performed. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:no power after FCO66 PM pcb replaced uer Impact:not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing uit set u the uit does not power on when pluged into AC power, FCO86600000000000000006 was recently performed (reference 00000000000000000000000000100000000177326), yellow power led on front panel is illuinated bu battery led is not, cutomer verified 200000000C present at Ji connector of PM pcb and 16VDC at battery oupu, screen is black, uit does not cycle and no fan spinning, requsts 900000000day warranty repair. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Anthony 781-00000000-7972 IP & Issu: 20000000000000000800000000 Power suply was recently replaced and now uit 000000000000000000000000001t tun on *Location Change* equpment located at 00000000highland ave winchester ma 000000008900000000(acct# 000000009000000002216) Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: "created ip relocation requst 200000000300000000-120000000000000000 created remote 00000000000000000000000000100000000000000005258 dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000" Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-21 23:06 | No
1047 | Yes | Device turned itself ON/OFF without ur interaction. | Device turned itself ON/OFF without ur interauon. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005680000000012-Oct-2000000003 Klu, Allan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000no tun itself on and off by itself withou uer interaction uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is tuning itself on and off by itself withou uer interaction Diagnostic performed by Engineer : Power cycling with no inpu from operator. Tools & Equpment : TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000051600000000Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve. Diagnostic performed by Engineer : Problem Description: V600000000is tuning itself on and off by itself withou uer interaction. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/21/2000000003 Trouleshooting Steps Taken: Biomed is reporting the original power management board order was cancelled. Cutomer advised of the 00000000 generation power management board that is being installed into all V600000000s as part of recall. Cutomer is requsting onsite to resolve. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000no tun itself on and off by itself withou uer interaction uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is tuning itself on and off by itself withou uer interaction Follow u Requred : Biomed has advised they are uable to place an order for a power management board. Down V600000000du to possible PM board failue. FCO86600000000000000006C – V600000000V600000000Plu uLL– uLLLoss of Fuction W/O Alarm ----------------------------------------------------------------------------- | 2023-09-21 23:46 | No
1048 | Yes | System will not power on | System will not power on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005800000000 21-Sep-2000000003 Brue Galen The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT Resoluion : Replaced MPD controller tested boot u and fuctions. Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: boot u Eupected and ueupected Resuts: delay uer Impact: delay Patient Impact: delay SE 8.1 Tools & Equpment : MuTIMETER,2000000003-00000000-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-22 00:31 | No
1049 | Yes | The device powered up by itself. | The device powered up by itself. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000580000000022-Sep-2000000003 Nuvo, Joseph The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : DRPTA puled uLLcodes 1100000000 1100000000, 1100000000, 1100000000, 11100000000and 11100000000 Device tuns on by itself. New power management PCBA for FCO66 did not fiu the issu. The old one was reinstalled. Problem Description by Engineer : Cutomer Fuction/Role: N/A Device ue: not in ue Eupected and ueupected Resuts: N/A uer Impact: none Patient Impact: SE ----------------------------------------------------------------------------- | 2023-09-22 00:41 | No
1050 | Yes | 00T0Disk is full message | 00CDT0Disk is full message | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005985 25-Sep-2000000003 Caldwell, Pau The system meets the specification for the performed service and is retuned to ue Resoluion : restored system sw from ippc. replaced ippc OS hdd with cutomer suplied. Installed ippc sw. performed system backu. vpo Problem Description by Engineer : Problem description by engineer : Disk ful message When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: To have enouh image storage space/Disk ful message uer Impact: Disk ful message is received Patient Impact: No impact to patient or procedue/procedue has been completed SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005829 Vilkaukas Viktorija The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Disk ful message When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: To have enouh image storage space/Disk ful message uer Impact: Disk ful message is received Patient Impact: No impact to patient or procedue/procedue has been completed SE 000Diagnostic performed by Engineer : Michael is reporting that they received a message "Disk Ful" bu they only have a few euams on the system. Michael said that this was an issu a few months back and the FSE came ou and looked at that issu then. They were able to complete the euam normally. System has not been rebooted. Michael created a logfile which is showing that the system has low disk space, I was not able to find a message that it is ful. Also checked the POST resuts which were normal. Smart Data is all ok for the IPPC, no disk read write uLL. Checked license and they have doule storage. Based on history the system is filling u normally. Perhaps to many supended cases on the system or cases have to many images in them to caue the messages that the image space is low, please delete euams. Internal Comments : uLL# and/or description of uLL: Disk is ful Repair Action : Michael is reporting that they received a message "Disk Ful" bu they only have a few euams on the system. Michael said that this was an issu a few months back and the FSE came ou and looked at that issu then. They were able to complete the euam normally. System has not been rebooted. Michael created a logfile which is showing that the system has low disk space, I was not able to find a message that it is ful. Also checked the POST resuts which were normal. Smart Data is all ok for the IPPC, no disk read write uLL. Checked license and they have doule storage. Based on history the system is filling u normally. Perhaps to many supended cases on the system or cases have to many images in them to caue the messages that the image space is low, please delete euams. Please ru data consistency and repair. If possible perform (re)create image store. Youshoud have 12500000000/50000000000000000 for doule storage. Check to make sue that auo delete is set in the uuudatabase and that the euams are in Completed statu so the system coud puge them as it fills u. If needed replace Image drives or IPPC itself. May need to increase storage to qudrule via a puchasable license option. 0000000080000000050000000092 - IP PC Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part Qudrule Storage NCVA8600000000ollow u Requred : uLL# and/or description of uLL: Disk is ful Repair Action : Michael is reporting that they received a message "Disk Ful" bu they only have a few euams on the system. Michael said that this was an issu a few months back and the FSE came ou and looked at that issu then. They were able to complete the euam normally. System has not been rebooted. Michael created a logfile which is showing that the system has low disk space, I was not able to find a message that it is ful. Also checked the POST resuts which were normal. Smart Data is all ok for the IPPC, no disk read write uLL. Checked license and they have doule storage. Based on history the system is filling u normally. Perhaps to many supended cases on the system or cases have to many images in them to caue the messages that the image space is low, please delete euams. Please ru data consistency and repair. If possible perform (re)create image store. Youshoud have 12500000000/50000000000000000 for doule storage. Check to make sue that auo delete is set in the uuudatabase and that the euams are in Completed statu so the system coud puge them as it fills u. If needed replace Image drives or IPPC itself. May need to increase storage to qudrule via a puchasable license option. 0000000080000000050000000092 - IP PC Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part Qudrule Storage NCVA8600000000 Technician Name : Vilkaukas Viktorija ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action taken: FSE Pau Caldwell accepted call for 9/2500AM. Called cu spoke to Michael Spangler confirmed onsite Oucome: dispatched Internal Remarks : Case: 0000000022319518 FSE: Caldwell, Pau Image disk ful message; intermittent(every 2months low ue system). FSE already performed Database repair and Re-Create Image Store; issu persists. Informed FSE on Known Section 00000000.9. Instruted FSE to doule check ethernet connection, make sue all IPPC HDDs are same size, reload Ghost, and replace IPPC as necessary. Internal Remarks : Planner Initials: GM Requstor: Pau Caldwell Requst: reschedue a call Oucome:: reschedued 00000000000000000000000000100000000000000005985 to9/2200PM ----------------------------------------------------------------------------- | 2023-09-22 00:41 | No
1051 | Yes | AED is emitting the triple chirp | AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006000000006 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED is emitting the triple chirp uer Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Cutomer reported the AED is emitting the triple chirp Internal Comments : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 13 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# B100000000-00000000755 Battery Install Before# 2000000006/7 Lot#00000000200000000P Pads Eup Date: 200000000000000000000000000000000 Pads Lot # 22000000000000000053 ----------------------------------------------------------------------------- | 2023-09-22 00:51 | No
1052 | Yes | receiving uLLmessage: pressure line disconnected (when the patient tubing is on) | receiving uLLmessage: pressure line disconneud (when the patient tubing is on) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006100000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer has a V600000000and has a uLLmessage Pressue line disconnected. He stated the average air is at +1.6 slpm. Provided them with the latest service manul rev T, and recommended him to perform the PVT. Cutomer stated the uit failed the air flow accuacy test. Recommended to replace the flow sensor assy. part #989800000000617100000000 flow sensor assy Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: pressue line disconnect alarm uer Impact: none Patient Impact: none SE un Resoluion : Cutomer will order a flow sensor assy and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Edurdo Flores IP & Issu: " 57917200000000 receiving uLLmessage: pressue line disconnected (when the patient tuing is on)" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created free remote technical 000000000000000000000000001 roued call to quu; udated cutomer contact information Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-22 01:31 | No
1053 | Yes | Low Leak CO0ebreathing risk on screen | Low Leak CO0Rebreathing risk on screen |  | 2023-09-22 02:00 | No
1054 | Yes | Power cord isn't working properly | Power cord isn't working properly | 000000000000000000000000001O##: 000000000000000000000000001000000000000000063700000000Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Parts ID uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 98980000000061300000000 - Power Cord - 50000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer is requsting part nuber and price for replacement power cord. Cutomer states cord isn't 000000000000000000000000001king properly and needs to be replaced. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Chris Edison IP & Issu: 566332800000000 Need the power cord part nuber for replacement. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-22 02:06 | No
1055 | Yes | Blower stalled message (EC 0) | Blower stalled message (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009000000006 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: uLLwith blower fan uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 00000000561512761 - blower assembly - 1000000001.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states device is giving blower uLL confirmed in event log. cutomer foud uLLcode 11300000000blower stall. Confirmed with cutomer that blower is sitting at 30000000000000000PM with nothing coming ou. cutomer confirmed. had cutomer adjut pressue above 00000000 cutomer states set pressue to 5 and rpms are rising. Informed cutomer manuactuers recommendation for 11300000000s the following: 1. Verify that the blower is connected to the MC PCBA. 2. Verify that the blower RPM monitor increases above 30000000000000000RPM when the pressue controller is set to a valu above 00000000cmH200000000 3. If blower RPM does not increase above 30000000000000000RPM replace one of the following to re-establish the tachometer signal. 00000000Replace blower. Cutomer is requsting part nuber for replacement blower assembly. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Chris IP & Issu: 89562391 & uLLwith the blower fan Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-22 02:06 | No
1056 | Yes | 00T0Disk full message and no cine | 00CDT0Disk full message and no 0e | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000600000000 29-Sep-2000000003 Hereford Rick The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - IPu Free disk space is too low. 18-Sep-2000000003 12:21:00000000 RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 00000000-Sep-2000000003 13:52:300000000 When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Disk ful message and no cine uer Impact: Can not cine Patient Impact: None/ completed SE 000Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : bad IPPC. replaced IPPC, backed u system, tested ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006382 Whitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IPu Free disk space is too low.: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Root Caue for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power or the OS/AS did not boot u in-time. Repair Action : LOD was not available for review, uLL are from 9-18-23, check uLL from todays logs to see if they are the same. Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs here: http://pww.uraywiki.u.com:80000000000000000uray-wiki/indeu.php/Post_code_LED_boot_statu_signals_of_PCs. (login with you CODE1 ID) 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IP-PC and perform Cold reboot of the system 1.3. If LED statu is not Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IP-PC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct! 1.5. Measue/replace the CMOS battery inside the IPPC becaue a bad/weak battery can caue the PC not to start the first time. 1.6. This pattern can also be caued if the BIOS is set wrong, like Advanced->Boot Featue->Restore On AC Power Loss set to 'Last State' or 'Power Off', defaut setting mut be 'Power On', if not green power LED at IPPC does not switch ON at boot u and PC needs to be started with Front Buton, in this case issu is not intermittent bu always present. 0000000080000000050000000022 - ClarityIQ_II IP PC no HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - IPu Free disk space is too low.: 1. An uer message is displayed on the large screen: "Free space low: Delete euaminations." Root Caue for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power or the OS/AS did not boot u in-time. Repair Action : LOD was not available for review, uLL are from 9-18-23, check uLL from todays logs to see if they are the same. Service actions for RMT - IPu Free disk space is too low.: 1.1. Instrut uer to delete euaminations to free u space. 1.2. Coud be caued du to bad or defective image disk. Replace image disk when necessary. Service actions for RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu LEDs here: http://pww.uraywiki.u.com:80000000000000000uray-wiki/indeu.php/Post_code_LED_boot_statu_signals_of_PCs. (login with you CODE1 ID) 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IP-PC and perform Cold reboot of the system 1.3. If LED statu is not Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IP-PC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct! 1.5. Measue/replace the CMOS battery inside the IPPC becaue a bad/weak battery can caue the PC not to start the first time. 1.6. This pattern can also be caued if the BIOS is set wrong, like Advanced->Boot Featue->Restore On AC Power Loss set to 'Last State' or 'Power Off', defaut setting mut be 'Power On', if not green power LED at IPPC does not switch ON at boot u and PC needs to be started with Front Buton, in this case issu is not intermittent bu always present. 0000000080000000050000000022 - ClarityIQ_II IP PC no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Whitaker, Steven Problem Description by Engineer : Problem description by engineer : RMT - IPu Free disk space is too low. 18-Sep-2000000003 12:21:00000000 RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 00000000-Sep-2000000003 13:52:300000000 When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Disk ful message and no cine uer Impact: Can not cine Patient Impact: None/ completed SE 000Diagnostic performed by Engineer : Du reported that the end uer reported that they are getting Disk ful message and can not cine. They deleted old cases and rebooted bu the issu remained. He did not have futher details and coud not save a LOD at this time. Checked RSW and foud the uLL below from 9-18-23 ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022319876 FSE: Hereford Rick Image Disk ful uLL FSE performed Re-Create Image Store. FSE performed Database Consistency and Repair; Re-Create Image Store. System is now fuctional. Internal Remarks : Cannot store images. FSE tried reloading software and also recreate image store several times. Gave part nubers for IPPC and hard drives. ----------------------------------------------------------------------------- | 2023-09-22 02:11 | No
1057 | Yes | Device would not boot up on battery power. Booted normally on AC power. | Device would not boot up on battery power. Booted normally on AC power. | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000031 100000000Oct-2000000003 Barragan, Carlos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Patient ue Eupected and ueupected Resuts: CB - 1000000001000000002000000003 - replaced RP-Cable, Harness, Power, V600000000and Assy Battery Li-Ion,11ah,V600000000 Repair resolved start u issus on battery power. Device passed electrical test, controls test, and Internal Battery Test. Device retuned to clinical ued uer Impact: NA Patient Impact: NA SE 2.3 Diagnostic performed by Engineer : CB - 1000000001000000002000000003 - replaced RP-Cable, Harness, Power, V600000000and Assy Battery Li-Ion,11ah,V600000000 Repair resolved start u issus on battery power. Device passed electrical test, controls test, and Internal Battery Test. Device retuned to clinical ued Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-11-30000000000000000:00000000:00000000 MuTIMETER,2000000003-11-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,2000000003-11-30000000000000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000561511621,RP-Cable, Harness, Power, V600000000 1.000000000000000000000000989800000000626900000000Assy Battery Li-Ion,11ah,V600000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: romeo corpu IP & Issu: ETA on fse , need callback asap received part Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: this week Entitlement Options: Qute Reason for Entitlement Claim: no con/warr Action Taken: sent to rp chat with information Neut Step Action: Awaiting Qute Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-22 03:36 | No
1058 | Yes | QAASA:0u FD0button active at start up | QAASA:0/u FD0button au at start up | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022311500000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: ##d SA Leader to confirm whether to ue the green pass, confirmation is in progress Problem Reported by Cutomer: The buton is broken, it always caues this uLLwhen booting ----------------------------------------------------------------------------- | 2023-09-22 05:36 | No
1059 | Yes | [Q] 0S:TK:FleuVision upper right auuiliary monitor display disappeared. | [Q] 0DS:TK:Fleusion upper right auuiliary monitor display disappeared. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007215 22-Sep-2000000003 Ito, Takui The system meets the specification for the performed service and is retuned to ue Resoluion: "000000000000000000000000001k content details" I confirmed that the left monitor on the large monitor in the euamination room was flickering. When I checked the control room switching cable when the problem occus, I confirmed a poor contact. I replaced it, and I waited and see, bu the problem was not reprodued. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : FleuVision uper right auiliary monitor display disappeared. u/Ms. Suimoto uhin: 6500000000 - Images are selected and displayed on the video switching equpment, bu while displaying live images, they disappear several times and are natually captued. The system is stable now, bu he 000000000000000000000000001ld like u to check it. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: FleuVision uper left auiliary monitor display disappeared. u/Ms. Suimoto uhin: 6500000000 - Images are selected and displayed on the video switching equpment, bu while displaying live images, they disappear several times and are natually captued. The system is stable now, bu he 000000000000000000000000001ld like u to check it. uer Impact: No effect Patient Impact: No effect 00 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided 0000000000Resoluion: The uper left auiliary monitor flickers. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000092900000000amata, Takayui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : FleuVision uper right auiliary monitor display disappeared. u/Ms. Suimoto uhin: 6500000000 - Images are selected and displayed on the video switching equpment, bu while displaying live images, they disappear several times and are natually captued. The system is stable now, bu he 000000000000000000000000001ld like u to check it. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: FleuVision uper right auiliary monitor display disappeared. u/Ms. Suimoto uhin: 6500000000 - Images are selected and displayed on the video switching equpment, bu while displaying live images, they disappear several times and are natually captued. The system is stable now, bu he 000000000000000000000000001ld like u to check it. uer Impact: No effect Patient Impact: No effect 00 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided 0000000000Diagnostic performed by Engineer: I asked FSE Ito to respond. Internal Comments: uLL# and/or description of uLL: [Suject] 00000000DS:TK:u Suimoto: FleuVision uper right auiliary monitor display disappeared. [Internal Comments] No parts from the display: Difficut to identify the parts [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Takui Ito] Ito, Takui - FSE1 Visit date and time 2000000003/00000000/2210000000000000000 - FSE1 000000000000000000000000001k time 1 [Euamination category] Normal euamination Repair Action : [Suject] 00000000DS:TK:u Suimoto: FleuVision uper right auiliary monitor display disappeared. [Internal Comments] No parts from the display: Difficut to identify the parts [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Takui Ito] Ito, Takui - FSE1 Visit date and time 2000000003/00000000/2210000000000000000 - FSE1 000000000000000000000000001k time 1 [Euamination category] Normal euamination Follow u Requred: * ----------------------------------------------------------------------------- | 2023-09-22 06:11 | No
1060 | Yes | 0/u FD0Z-rot cover dropped off | 0R/u FD0Z-rot cover dropped off | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022311697 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 00000000.1000000002000000003 00000000:52:00000000uC (86000000000000000000000000 Service content: An ouer cover of the C arm is loose, re-fiu the ouer cover. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.1000000002000000003 00000000:52:00000000uC (86000000000000000000000000 RDI Accuacy: YES $Q6: (The stand cover fall off.) $Q7:(Comfirmed reported problem.The stand cover fall off.) $Q8:(The stand cover fell off, need to fiu the cover.) $Q9:(fiu the stand cover.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: N/A. SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: L arm loose SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: The Z-rot operating casing came off and needed to be reattached.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: The Z-rot operating casing came off and needed to be reattached.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: The Z-rot runing casing came off and needed to be reattached.Engineers are requred to be present Notes on cutomer requrements: Onsite SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-22 06:16 | No
1061 | Yes | 0/V0uLLPrimary alarm failed | 0K/V0uLLPrimary alarm failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: Fuing Central Hospital Euternal Remarks: 26.00000000.2000000003 00000000:56:00000000 uC (86000000000000000085) *Department information: [Department: Respiratory Department Contact person: Teacher He Tel: 177235859300000000 * $ASP Engineer: [Chongqing WuBEI Jia GANLu $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (confirm reporting problem, main alarm failed) $Q8: (download"DRPT" report, show uLLcode 1100000000, determine it is the main alarm failed) $Q9: (change speaker) $Q100000000 (see PA Check List of test and check data.) . The device is fuly fuctional.) $Q11: not applicable Internal Remarks: 26.00000000.2000000003 00000000:56:00000000 uC (86000000000000000085) $ASP Engineer: [Chongqing WuBEI Jia GANLu $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q00000000(Not applicable) $Q7: (confirm reporting problem, main alarm failed) $Q8: (download"DRPT" report, show uLLcode 1100000000, determine it is the main alarm failed) $Q9: (change speaker) $Q100000000 (see PA Check List of test and check data.) . The device is fuly fuctional.) $Q11: not applicable SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: Main alarm failed, Safety Qustions: Case: 0000000022311800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Main alarm failed Repair Action: ---------RA-----REPAIR ACTION--------------- IF PART REQuRED - 12 NC RECOMMENDED: 00000000561512551 Spare part description: RP-PCBA, CPu V600000000 RECOMMENDED REPAIR ACTION: Check the speaker cable connection, trouleshoot the speakers and CPuboard FSE Notes: This PO is a Reference Nuber only ----------------------------------------------------------------------------- | 2023-09-22 06:56 | No
1062 | Yes | 0u FD0Live monitor problem | 0/u FD0Live monitor problem | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: ##d to SS dongjiede for buiness follow-u Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 0000000022311852 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000/u uuuFD200000000 ----------------------------------------------------------------------------- | 2023-09-22 07:16 | No
1063 | Yes | Alarms/Failure to alarm: Device doesn't alarm when the Heart rate euceeds: NTF/Unknown Cau: Provided Information | Alarms/Failure to alarm: Device doesn't alarm when the Heart rate euceeds: NTF/Unknown Cau: Provided Information | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007225 27-Sep-2000000003 Bin Mazlan, Muammad Nadzmi The system meets the specification for the performed service and is retuned to ue Resoluion : Checked and foud that alarm souds when HR euceeds, and is 000000000000000000000000001king fine. Informed uer to monitor sitution if problem arises Problem Description by Engineer : Cutomer Fuction/Role: MP300000000 Device ue: MP300000000 Eupected and ueupected Resuts: Device does not alarm when HR euceeds uer Impact: NA Patient Impact: NA SE NA ----------------------------------------------------------------------------- | 2023-09-22 07:41 | No
1064 | Yes | Duplicate to PR 00// CFMS0000 | Duplicate to PR 000// CFMS000000 | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: SA: 00000000N/uDFM100000000 Defibrillator/Monitor/machine alarm, the dealer is in a high mood, please contact Nuse Zhang T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Duing Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL The patient wears a pacemaker, the uwaveform is normal, HR: 600000000, the device alarms: ucannot be analyzed, treatment cannot be provided, and the device self-test has no failues. Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: Check device logs, switch to different uleads or electrode pads for uanalysis, and trouleshoot clinical uage factors Safety Qustions: Case: 0000000022312000000005 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-22 08:36
1065 | Yes | 0/u FD0Live monitor problem | 0R/u FD0Live monitor problem | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Repair Action: nemo/receiver bou and transmitter bou/200000000to hosp/repair person will collect it Service content: If youneed a receiving bou and a transmitting bou, they will be sent throuh the spare parts eupress service. Safety Qustions: Case: 0000000022312139 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-22 08:46 | No
1066 | Yes | Waveform issue | Waveform issue | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport Internal Remarks : 21.00000000.2000000003 00000000:21:00000000 uC (85200000000100000000) $Q7: (Confirmed reported problem. waveform issu) $Q8: (Checked system. checked with uer uwaveform normal uin ProSim8) $Q9: (No faut foud, advised observation) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) Problem Reported by Cutomer : Waveform issu Safety Qustions : Case: 0000000022312281 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Case Transferred to SAP Euternal Remarks : 21.00000000.2000000003 00000000:20000000058 uC (85200000000100000000) $Q7: (Confirmed reported problem. waveform issu) $Q8: (Checked system. checked with uer uwaveform normal uin ProSim8) $Q9: (No faut foud, advised observation) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) ----------------------------------------------------------------------------- | 2023-09-22 09:56 | No
1067 | Yes | Machine unable to boot up uLLmessage disk boot failure | Machine unable to boot up uLLmessage disk boot failure | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007186 22-Sep-2000000003 KOK WAI CHAN The system meets the specification for the performed service and is retuned to ue Resoluion : Boot u sequnces wrong reset boot u sequnce to the correct media hdd. Test system ok Problem Description by Engineer : Cutomer Fuction/Role: Radiograher Device ue: no Eupected and ueupected Resuts: system cannot boot u uer Impact: no Patient Impact: no SEv7.2.12 ----------------------------------------------------------------------------- | 2023-09-22 10:46 | No
1068 | Yes | System is not booting up | System is not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000176989 200000000Oct-2000000003 Dangi, Mohan lal The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System is not booting u When was the first occurence? Within... : 22-Sep-2000000003 11:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System is not booting u uer Impact: No uer impact Patient Impact: No patient impact SE 8.1.300000000100000000Resoluion : I have pu the hard drive inside and reloaded the software and now system 000000000000000000000000001king fine. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008699 00000000-Oct-2000000003 Dangi, Mohan lal The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : System is not booting u When was the first occurence? Within... : 22-Sep-2000000003 11:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System is not booting u uer Impact: No uer impact Patient Impact: No patient impact SE 8.1.300000000100000000Diagnostic performed by Engineer : I have installed the new host PC bu system not ready, I try to reload the software bu same issu, new host PC DEFO received. Follow u Requred : System intermittent issus, not booting ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008130000000025-Sep-2000000003 Dangi, Mohan lal The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Follow u Requred : system hanging issu booting time Problem Description by Engineer : Problem description by engineer : System is not booting u When was the first occurence? Within... : 22-Sep-2000000003 11:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System is not booting u uer Impact: No uer impact Patient Impact: No patient impact SE 8.1.300000000100000000Diagnostic performed by Engineer : re-installed the ghost & system 000000000000000000000000001king bu intermitted issu . ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007859 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System is not booting u When was the first occurence? Within... : 22-Sep-2000000003 11:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System is not booting u uer Impact: No uer impact Patient Impact: No patient impact SE 8.1.300000000100000000Diagnostic performed by Engineer : Host pc not detecting the harddisk. system is not booting and stops at detecting hard drive. Either disk issu or pc issu. Internal Comments : uLL# and/or description of uLL: System is not booting u Repair Action : Host pc not detecting the harddisk. system is not booting and stops at detecting hard drive. Either disk issu or pc issu. Resoluion : Host pc not detecting the harddisk. system is not booting and stops at detecting hard drive. Either disk issu or pc issu. Follow u Requred : uLL# and/or description of uLL: System is not booting u Repair Action : Host pc not detecting the harddisk. system is not booting and stops at detecting hard drive. Either disk issu or pc issu. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-22 11:46 | No
1069 | Yes | QA 000000 - monitor displays abnormally | QA 00000000- monitor displays abnormally |  | 2023-09-22 12:08 | No
1070 | Yes | Picked up with note reporting Low leak CO0ebreathing | Picked up with note reporting Low leak CO0rebreathing | Test Equpment Info : [Device #: 1183113, Make: FLuE BIOMEDICAL CORP, uY: ESA612, Serial: 300000000000000006, Last Cal Date: 2000000003-00000000-31] [Device #: 1273327, Make: TSI INC, uY: 000000001 C, Serial: 000000001133000000002, Last Cal Date: 2000000003-00000000-31] [Device #: 1273326, Make: TSI INC, uY: CERTIFIER FA PLu, Serial: 0000000081335000000002, Last Cal Date: 2000000008-00000000-31] Device Alarm/Alert as intended? : No Problem Observed : Took to the shop and verified problem, vent alarms after a few seconds of operation, Low Leak CO2 re-breathing, ordering parts to repair Ou Of Bou Failue? : No Labor: Labor ID : L-19781800000000 000000000000000000000000001k Performed : Picked u on 7th floor with note reporting Low Leak alarm, Took to the shop and verified problem, vent alarms after a few seconds of operation, Low Leak CO2 re-breathing, ordering parts to repair Diagonsis : Not Answered | 2023-09-22 12:12 | No
1071 | Yes | Intrepid ETCO0ort defective | Intrepid ETCO0port defeu | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008000000005 Problem Description by Engineer: CO² port defective Resoluion: repaired Parts Consued: 1.0000000000000000000000000000000056000000001811,Intrepid uSpO2+NBP+CO2 Frame ASSY ----------------------------------------------------------------------------- | 2023-09-22 12:16 | No
1072 | Yes | QAA0/u FD0System cannot make euposure | QAA0K/u FD0System cannot make euposure | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 21.00000000.2000000003 13:36:22 uC (8600000000600000000) Service content (requred): The equpment cannot be euposed. Check the uLLmessage and the handbrake is not fuly released. Clean the handbrake switch. The test fuction is normal and the machine retuns to normal ue. Service resuts (requred): [After repair, the equpment retuned to normal operation.] Other important matters: [none] Problem Reported by Cutomer: QAA00000000K/u uuuFD200000000 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of the cutomer's problem or faut phenomenon: Cutomer feedback that the machine cannot emit rays and requres engineers to come to the site for repair. Is it a faut of the equpment itself? Yes Trouleshooting Action: Key faut code or phenomenon: uray not possible call service Remote diagnosis and test resuts: Restart has no effect PRS TR statu: uable Repair Action: Recommended on-site repair steps: 1. Check whether the FDC is normal 2. Check whether the high voltage is normal 3. Check the machine uLLand perform faut repairs Recommended spare parts uuand name: NA Cutomer demand notes: Contact the cutomer in time to confirm the Onsite time Safety Qustions: Case: 00000000223128200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 21.00000000.2000000003 13:36:200000000TC (8600000000600000000) $Q7: (Confirmation & Description of problem/System cannot euposue) $Q8: (cutomer report System cannot euposue, check system log and Foud handswitch not ful release) $Q9: (clean handswitch and test fuciton ok, then system OK) $Q100000000 (Test system ru ok and inspection data refer to: PA. Device retuned to ful fuctionality. ) $Q11: N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-22 12:31 | No
1073 | Yes | Alarms or Messages/Low sound-There is no audible signal when an alarm arises-HPM-ur/Unknown reason | Alarms or Messages/Low sound-There is no audible signal when an alarm arises-HPM-ur/Unknown reason | 000000000000000000000000001O##: 000000000000000000000000001000000000000000051300000000Kristopher Schardin The system meets the specification for the performed service and is retuned to ue Resoluion: By 00Ms. Meimani: Timer does not give any souds after eupiration. Sent instrutions by ##to activate souds Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: There is no auible signal if an alarm arises. Patient Impact: - SE - ----------------------------------------------------------------------------- Cutomer comment: Ms. Meimani called. Not possible to get throuh. Cutomer changed the configuation. Futher problems with the soud when ending the case +00000000113983000000005 ----------------------------------------------------------------------------- | 2023-09-22 13:11 | No
1074 | Yes | QAASA:0/uCM0Loudspeaker failure. Reconnected speaker cable | QAASA:0N/Effi0 CM0Loudspeaker failure. Reconneud speaker cable | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: QAASA:00000000N/uCM100000000 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 25.00000000.2000000003 00000000:00000000:00000000 uC (860000000000000000100000000 $ASP engineer: [honglejiaWuQiang] $Q1:Y $Q2: N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem:Speaker display failue, no soud $Q8: Check all the fuctional measuements can be ued normally, bu there is no soud, and the alarm has been warning the speaker failue to juge the speaker damage $Q9:Remove the machine, re-plu the speaker cable, start the machine, and then perform the auio test self-test, and the soud of the device will retun to normal $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11: Calibrated Tools name (ESA6200000000 S/N (2399000000003) Neut Calibration date (0000000011/20000000000000000T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Speaker failue The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: It is recommended to replace it on site Recommended spare part uufor T2 with description: Speaker assembly 000000005 60000000087821 Notes on cutomer requrements: na Test/Inspection: N/A Test Equpment: N/A Euternal Remarks: 25.00000000.2000000003 00000000:00000000:58 uC (860000000000000000100000000Department information: [Department: Department of Integrated Traditional Chinese and Western Medicine Contact: Teacher Wang Tel: 18117822813]* $ASP engineer: [honglejiaWuQiang] $Q1: Y $Q2:N $Q3: N/A $Q00000000N/A $Q7: Confirmed reported problem:Speaker display failue, no soud $Q8: Check all the fuctional measuements can be ued normally, bu there is no soud, and the alarm has been warning the speaker failue to juge the speaker damage $Q9:Remove the machine, re-plu the speaker cable, start the machine, and then perform the auio test self-test, and the soud of the device will retun to normal $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11: Calibrated Tools name (ESA6200000000 S/N (2399000000003) Neut Calibration date (0000000011/20000000000000000Safety Qustions: Case: 0000000022313100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-22 13:21 | No
1075 | Yes | It does not start | It does not start | 000000000000000000000000001O##: 00000000000000000000000000100000000113677 00000000Oct-2000000003 Nicastro, Fernando The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: It seems that the main compuer of the equpment (SutePC) is not starting u. F. Nicastro is in charge of reloading the software. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. uu00000000Resoluion: App udoes not start properly, we try to load sw to pc sute. I 000000000000000000000000001k together with helpdesk, Gonzaga, Sancho and help from other FSEs. Pc, u-ray PC and HDs are changed. After several tries, we were able to load ok. Test OK, equpment is operational Parts Consued: 1.00000000000000000000000000000000800000000660000000000000000,u-ray PC Biplane 1.0000000000000000000000000000000080000000000000000900000000, PDuDC Modue Tools & Equpment: DOSE-KVP-LIGHT METER, 2000000006-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002200000000 29-Sep-2000000003 Nicastro, Fernando The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: It seems that the main compuer of the equpment (SutePC) is not starting u. F. Nicastro is in charge of reloading the software. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. uu00000000Parts Consued: 1.00000000000000000000000000000000800000000717981, Sute PC W7 + W100000000 1.00000000000000000000000000000000800000000660000000031, SSD Data Haswell 1.00000000000000000000000000000000800000000660000000032, SSD Data Haswell 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Follow u Requred: Already done. Diagnostic performed by Engineer: App udoes not start properly, we try to load sw to pc sute. I 000000000000000000000000001k together with helpdesk, Gonzaga, Sancho and help from other FSEs. Change of Sute PC, TSM, uray PC, HDs. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009000000007 25-Sep-2000000003 Nicastro, Fernando The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: It seems that the main compuer of the equpment (SutePC) is not starting u. F. Nicastro is in charge of reloading the software. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. uu00000000Diagnostic performed by Engineer: App udoes not start properly, we try to load sw to pc sute. I 000000000000000000000000001k together with helpdesk, Gonzaga, Sancho and help from other FSEs. Change of Sute PC, TSM, uray PC, HDs. Follow u Requred: It has already been performed. Parts Consued: 1.00000000000000000000000000000000800000000660000000022,SSD OS Haswell 1.000000000000000000000000300000000000000008251682, uR2.2 (L3) SW Pack ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 29-Sep-2000000003 Meira, Fernando The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Tech Device ue: In ue Eupected and ueupected Resuts: uLLwhen starting uer Impact: Equpment ou of order. Patient Impact: Equpment ou of order uuDiagnostic performed by Engineer: Software failue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008993 Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: It seems that the main compuer of the equpment (SutePC) is not starting u. F. Nicastro is in charge of reloading the software. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. uu00000000Diagnostic performed by Engineer: It seems that the main compuer of the equpment (SutePC) does not start. F. Nicastro is in charge of reloading the software. Internal Comments: uLL# and/or description of uLL N/A Repair Action: It seems that the main compuer of the equpment (SutePC) is not starting u. F. Nicastro is in charge of reloading the software. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: It seems that the main compuer of the equpment (SutePC) is not starting u. F. Nicastro is in charge of reloading the software. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-09-22 14:26 | No
1076 | Yes | KTPH ¦ DFM0 ¦ Sudden loss of continuouwaveforms | KTPH ¦ DFM0 ¦ Sudden loss of continuouwaveforms | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009200000000 29-Sep-2000000003 Thong, Leon The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: nil Device ue: nil Eupected and ueupected Resuts: nil uer Impact: nil Patient Impact: nil SE nil Resoluion : Suden loss of continuu waveform and Nibp buton in stuk in stop mode Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000009 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-22 14:36 | No
1077 | Yes | Primary alarm failed | Primary alarm failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport FSE Notes: This PO is a Reference Nuber only SAP Transfer: Case Transferred to SAP Internal Remarks: 1000000001000000002000000003 00000000:26:38 uC (86000000000000000000000000) $ASP engineer:[suuhu$Q1:no $Q2：no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.Machine master alarm failed.$Q8: Download "DRPT" report， uLLCode have not, Shoud be a problem of the PCBA, CPuQ9: Replace the PCBA, CPuQ100000000 Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality.$Q11:Calibrated Tools name (FLuE16500000000 S/N (000000006000000002) Neut Calibration date (11/11/2000000003) Euternal Remarks: 1000000001000000002000000003 00000000:26:35 uC (86000000000000000000000000) *Department information: [Department: Clinical Medical Engineering Department Contact Person: LiuChengyouTel: 18262635871 ]* $ASP engineer:[suuhu$Q1:no $Q2：no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.Machine master alarm failed.$Q8: Download "DRPT" report， uLLCode have not, Shoud be a problem of the PCBA, CPuQ9: Replace the PCBA, CPuQ100000000 Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality.$Q11:Calibrated Tools name (FLuE16500000000 S/N (000000006000000002) Neut Calibration date (11/11/2000000003) Problem Reported by Cutomer: Nanjing First Hospital Main alarm failed T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL The main alarm failed (the person who reported the repair did not know the specific problem) Repair Action: ---------RA-----REPAIR ACTION--------------- IF PART REQuRED - 12 NC RECOMMENDED: 00000000561512551 Spare part description: RP-PCBA, CPu V600000000 RECOMMENDED REPAIR ACTION: Check the speaker cable connection, trouleshoot the speakers and CPuboard Safety Qustions: Case: 0000000022313631 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-09-22 14:51 | No
1078 | Yes | The foot switch cannot be connected to the system | The foot switch cannot be conneud to the system | 000000000000000000000000001O##: 000000000000000000000000001000000000000000093100000000November 8, 2000000003 Schellenberg, Arthu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: RR Florian Kerkhoff Phone 00000000000000001618922000000005 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000Resoluion: Wireless foot switch + base station replaced. Connection established and fuctions checked. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000238232, wireless foot switch 3P 1.00000000000000000000000000000000800000000257862, WFS base station ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 Hellenbrandt, Helmu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: RR Florian Kerkhoff Phone 00000000000000001618922000000005 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 1.2.5.00000000Diagnostic performed by Engineer: ufoot switch does not connect Follow-u requred: uLL# and/or description of uLL After consutation with THE CCC, the nuber of 00000000000000000000000000100000000822100000000 is curently not clear. I also cannot recognize that the repair took place. Repair Action: Check/replace ufoot switch and base uit According to the last case, the replacement of the base uit is still open Technician name: Hellenbrandt, Helmu Internal Comments: uLL# and/or description of uLL After consutation with THE CCC, the traffic jam of the 00000000000000000000000000100000000822100000000 curently not clear. I also cannot recognize that the repair took place. Repair Action: Check/replace ufoot switch and base uit According to the last case, the replacement of the base uit is still open ----------------------------------------------------------------------------- Internal Remarks: Eutension 0000000000000000u Wytrricku Internal Remarks: After consutation with u Kerkhoff, the cutomer mut first clarify internally with the OP management how the services shoud take place in the fuue since there have been repeated short-term cancellations so far; u Kerkhoff will get in contact again promptly to schedue the service; Do not schedue anything else util then Cutomer comment: rings throuh at 0000000000000000Internal Remarks: Schedued for the other service on 11/8 Internal Remarks: Dear u Kasapovic, Please also schedue the repair 0000000022321659 of the foot switch on this date. Kind regards, Florian Kerkhoff Cutomer comment: u Tim Welfens called for case 0000000022321659 (see also 0000000021813712) And for case 000000002200000000779 Cutomer asks abou the appointments. Please check with Getinge. 00000000000000001618922000000008 medizintechnik.mgladbach.mariahilf00raeger.com ----------------------------------------------------------------------------- | 2023-09-22 15:06 | No
1079 | Yes | It is difficult to discharge during the test. | It is difficult to discharge during the test. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000037 27-Sep-2000000003 Daniele Racioppi The system meets the specification for the performed service and is retuned to ue. Resoluion: Search and check of faut, replaced uEuternal Paddles 98980000000019600000000. Parameter tests, fuctional tests, and fuctional check with a positive oucome. The defibrillator is located at uC 2P at Gavardo Hospital. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test, not in clinical ue Eupected and ueupected Resuts: it does not always discharge duing the test uer Impact: no impact Patient Impact: no impact SE na ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009212 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: defibrillator test, not in clinical ue Eupected and ueupected Resuts: it does not always discharge duing the test uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Foud discharge problem duing the test. Follow u Requred: Contacted u. Gares, she reported that duing the test the device sometimes does not discharge with euternal plates and they mut repeat it many times, when performing the test with the test load no problems are reported, proceeded with order of plates, in case they are not available, check if they are still present at Concorezzo with this 98980000000019600000000 code. ----------------------------------------------------------------------------- | 2023-09-22 15:16 | No
1080 | Yes | Split - Faulty speaker. Assembly/PCB/Card-Speaker-HW. Replacement sent (NEMO) | Split - Faulty speaker. Assembly/PCB/Card-Speaker-HW. Replacement sent (NEMO) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008999 Andrea Ardighieri The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: nemo withou collection HuANITAS MIRASOLE 00000000ASCINA PERSEGHETTO STREET ROZZANO 20000000089 MI Goods reception clinical engineering AC. PIETRAGALLA 000000008220000000015 Problem Description by Engineer: Cutomer Fuction/Role: CLINICAL ENGINEER Device ue: parameter monitor Eupected and ueupected Resuts: it does not detect CO2 and the speaker is ou of ue, no impact on reported patient uer Impact: no impact Patient Impact: no impact SE na Diagnostic performed by Engineer: Foud fauty CO2 modue and speaker. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009892 Parts Consued: 1.0000000000000000000000000000000056000000008000000001, SS CBL Speaker Assy FS 1.0000000000000000000000000000000056000000008981, SS ASSY CO2 Assembly MicroMedi FS ----------------------------------------------------------------------------- | 2023-09-22 15:51 | No
1081 | Yes | Problème pédale Wi-Fi | Problème pédale Wi-Fi | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002161 October 2, 2000000003 MARC RICAu The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: no Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: na SE no. Resoluion: The Wi-Fi pedal was replaced. Parts Consued: 1.0000000000000000000000000000000080000000000000000552, Wireless footswitch Charger (WFSC) 1.00000000000000000000000000000000800000000238232, Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862, WFS Base Station 1.00000000000000000000000030000000000000000660000000000000000,Pictogram Sheet Footswitch u----------------------------------------------------------------------------- | 2023-09-22 16:01 | No
1082 | Yes | cine and uuuintermittent problem (SPLIT) | 0e and uuuintermittent problem (SPLIT) |  | 2023-09-22 16:33 | No
1083 | Yes | AED will not turn on | AED will not turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002579 Hendrik Gritzka The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: Nemo: Received/## BLIV: No Otherwise.: Pcs/uuAT: 1u 000000005600000000000000001 HS1 (C B) Defib, German 00000000, Euch in at Diagnostic performed by Engineer: 1u 000000005600000000000000001 HS1 (C B) Defib, German 00000000, Euch in at Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: Device cannot be tuned on anymore. uer Impact: uknown Patient Impact: uknown SE uknown ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000056800000000 Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Defib, German 00000000, Euch ----------------------------------------------------------------------------- Cutomer comment: By ##u Kolodziej When can we eupect the replacement? Cutomer comment: uu 000000005600000000000000001 - HS1 (C B) Defib, German 00000000, Euch Cutomer ETA: OCT 11, 2000000003 Internal Remarks: ETA sent. Internal Remarks: Did not reach the system, the management is free ----------------------------------------------------------------------------- | 2023-09-22 16:36 | No
1084 | Yes | Related to 000 - Patient monitoring stopped working, after performing a discharge | Related to 00000- Patient monitoring stopped working, after performing a discharge | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000000000000000000000-Nov-2000000003 Óscar Cardona Orellana The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: We euchanged the Tempu LS. It is configued and linked to Tempu PRO. Parts Consued: 1.000000000000000000000000000000005600000000000000001,Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- | 2023-09-22 17:01 | Yes
1085 | Yes | Problems with the wireless pedal | Problems with the wireless pedal | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002712 28-Sep-2000000003 José González Castillo The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical engineer Device ue: Cardiac system Eupected and ueupected Resuts: Footswitch problems at star u system. uer Impact: ue wired footswitch Patient Impact: No imact SE Na Resoluion: The base of the table and the wireless pedal of the equpment are replaced for different problems; the pedal did not have indication LEDs and when they try to match it with the base, it does not recognize it. Both uits are replaced to repair the problem. Equpment is tested and it is 000000000000000000000000001king. The new installed materials are in the software version 3.00000000so the FCO 7220000000000000000 is not applicable, which is to uload both uits to this software. Parts Consued: 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238232, Wireless Footswitch 3P ----------------------------------------------------------------------------- | 2023-09-22 17:06 | No
1086 | Yes | Battery replacement alarm coming (Battery is replaced) | Battery replacement alarm coming (Battery is replaced) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 100000000Oct-2000000003 Jithin, T J The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,2000000003-11-1000000000000000000000000:00000000 Resoluion : Confirmed problem with battery. Same needs to replaced Diagnostic performed by Engineer : Checked the equpment and foud that battery is fauty. Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009235 Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Fse visit requred Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Fse visit requred ----------------------------------------------------------------------------- | 2023-09-22 17:31 | No
1087 | Yes | Touchscreen is intermittently unresponsive | Touchscreen is intermittently unresponsive |  | 2023-09-22 17:38 | Yes
1088 | Yes | When starting the system does not start up. | When starting the system does not start up. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 22-Sep-2000000003 ROB VAN GAMEREN The system meets the specifications for the performed service and is retuned to ue Follow u Requred: Host PC does not start u now and then. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: in ue Eupected and ueupected Resuts: uer Impact: system down Patient Impact: no SEr000Resoluion: Duing the retun trip Martijn de Weerd called me that the system starts u normally again. Retuned home. Martin Versloot later checked the log files again and it seems that the same uLL also occured before, host PC does not start u. Therefore follow u neut Monday morning. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000036 MARTIN VERSLOOT The system does not meet the specifications for the performed service and is not retuned to ue Follow u Requred: System does not start u Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: in ue Eupected and ueupected Resuts: uer Impact: system down Patient Impact: no SEr000Diagnostic performed by Engineer: System does not start u ----------------------------------------------------------------------------- | 2023-09-22 17:41 | No
1089 | Yes | Foot switch defective | Foot switch defeu | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006596 September 21, 2000000003 Fueder, Bernhard The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MT Device ue: Yes Eupected and ueupected Resuts: None uer Impact: None Patient Impact: None SE 7.2.16 Resoluion: The reported faut is confirmed. The following actions were performed for uLLcorrection: Foot switch + receiver replaced and reprogrammed. Fuction checks performed. Performed check of the electrical systems according to DIN EN 62353 and DGu reguation 3. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000238192, wireless foot switch 3P 1.00000000000000000000000000000000800000000257862, WFS base station Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000003-12-28 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks: 1u 00000000800000000238192 Wireless Footswitch 3P 1u 00000000800000000257862 WFS Base Station ----------------------------------------------------------------------------- | 2023-09-22 18:01 | No
1090 | Yes | During the functional check the therapy delivery test fails. | During the funuonal check the therapy delivery test fails. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000031700000000auers, Anthony The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion: Cutomer will replace the defib - end of suport Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: Eupected and ueupected Resuts: uer Impact: None Patient Impact: SE ----------------------------------------------------------------------------- | 2023-09-22 18:21 | No
1091 | Yes | SELF TEST FAILED. | SELF TEST FAILED. | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022317562 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: nannan wang Last udated by: nannan wang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: SELF TEST FAILED Neut Step Action: ##d ASP uinhukanguin for follow u Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-09-22 18:21 | No
1092 | Yes | HS0ED is emitting the triple chirp (Removing the AED from service) | HS0AED is emitting the triple chirp (Removing the AED from service) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 13 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A100000000-00000000000000009 Battery Install Before# 6/3000000002000000009 Lot#000000000000000000000000 Pads Eup Date: 6/100000000000000005 Pads Lot # C12100000000-00000000 Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. ----------------------------------------------------------------------------- | 2023-09-22 18:26 | No
1093 | Yes | No fluoroscopy. The displayed message mentioned an application uLLwithout providing any specific code information. | No fluoroscopy. The displayed message mentioned an application uLLwithout providing any spe0ic code information. | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000005 29-Sep-2000000003 Michal Sadowski The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.000000000000000000000000989710000000000000000573,Flashlite High End KIT. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: cardiac. Eupected and ueupected Resuts: no radiation. uer Impact: no. Patient Impact: the procedue may be delayed. SE 1.000000002. Resoluion: Replacement of the FDC modue. Software loading to the modue. Testing. The device 000000000000000000000000001ks correctly, the faut was repaired. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009996 29-Sep-2000000003 Michal Sadowski The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: cardiac. Eupected and ueupected Resuts: Radiation does not 000000000000000000000000001k. uer Impact: no. Patient Impact: the procedue may be delayed. SE 1.00000000u. Resoluion: Replacement of the FDC modue. Replacement failed - it is not possible to load the software. It is necessary to place an order for parts again. The device remains non-repaired. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009351 25-Sep-2000000003 Michal Sadowski The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: cardiac. Eupected and ueupected Resuts: no radiation. uer Impact: no. Patient Impact: the procedue may be delayed. SE 1.000000002. Diagnostic performed by Engineer: Analysis of system logs. An uLLof the FDC modue was foud (the commuication modue with the detector). Tests confirmed that the modue is damaged. Replacement is necessary. ----------------------------------------------------------------------------- | 2023-09-22 18:36 | No
1094 | Yes | Audio&Sounds/No Sound: Equipment displays loudspeaker uLL Speaker/HW: Part Replaced (00 | Audio&Sounds/No Sound: Equipment displays loudspeaker uLL Speaker/HW: Part Replaced (0000 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000CHEN, Cheng The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Delivery of the part to the biomedical department uder the contract. Follow u Requred: Delivery of the part to the biomedical department uder the contract. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: requst to order part uder the contract. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000018 Parts Consued: 1.0000000000000000000000000000000056000000008621, MS_u2 ASSY CBL u2/MP2 Speaker Assembly ----------------------------------------------------------------------------- | 2023-09-22 18:46 | No
1095 | Yes | Intrepid processor defective | Intrepid processor defeu | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003516 Problem Description by Engineer: Processor defective Resoluion: Repair Parts Consued: 1.0000000000000000000000000000000056000000001751,Intrepid Processor PCA ASSY ----------------------------------------------------------------------------- | 2023-09-22 19:01 | No
1096 | Yes | Pressure line continues to alarm | Pressure line continues to alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001000000002 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Pressue line continus to alarm and is not blowing enouh air uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber: 00000000561512561 - DA board Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states it was reported that: Pressue line continus to alarm and is not blowing enouh air. Cutomer is curently not in front of equpment and 000000000000000000000000001ld like to know where to start to trouleshot. Advised cutomer to check event log. see if he sees 1100000000, 1000000007, 1100000000. theses are codes dealing with prouimal pressue issus. Advised cutomer that most of these codes, manuactuers recommendation is to replace the DA board. Cutomer is requsting part nuber for replacement DA board. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jutin 865-361-8568 IP & Issu: Pressue line continus to alarm and is not blowing enouh air Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-22 19:11 | No
1097 | Yes | 0A (CBIT) Vent-Inop: Data Acquisition PCBA ADC Reference Failed | 0A (CBIT) Vent-Inop: Data Acquisition PCBA ADC Reference Failed | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003825 Miller Jason The system meets the specification for the performed service and is retuned to ue Resoluion : Provided cutomer with part nuber: 00000000561536131 - Cable, DA PCBA to MC PCBA (2nd generation) Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states they need assistance Reviewing logs. Went over event log with cutomer. foud uLLcode 100000000a in event logs. Troule shot 100000000a with cutomer: 1. Verify 3.3V in diagnostics. 2. Verify 2.5V ref on DA PCBA (see Figue 6-1). 3. Replace the DA PCBA to MC PCBA cable. 00000000Replace the DA PCBA. 5. Replace the PM PCBA. 6. Replace the CPu Cutomer states if the DA to MC cable is a 1st gen cable and is requsting part nuber for replacement 2nd gen cable. Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Reviewing logs suport uer Impact: system restricted Patient Impact: none reported SE uknown ueupected Resuts: uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Robert Fuler 800000000777-000000007 IP & Issu: Reviewing logs suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-22 19:41 | No
1098 | Yes | Cannot make images | Cannot make images | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001555 25-Sep-2000000003 JEAN JORET, The system meets the specifications for the performed service and is retuned to ue Diagnostic performed by Engineer: Euaminations are not closed becaue access to dicom node for reporting dose is blocked. Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Resoluion: Remove uued strutued dose report dicom node ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008900000000 PATRICK ATTN. The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer: No remote soluion possible, an on-site visit has been schedued Disk is ful and uer claims the patient list cannot be cleaned Follow u Requred: Check euport patient data ----------------------------------------------------------------------------- | 2023-09-22 19:56 | No
1099 | Yes | QA00000- uuuelectrode plate emitted sparks | QA00000000 - u eleuode plate emitted sparks |  | 2023-09-22 20:15 | No
1100 | Yes | QA00000- uuuself-test screen after shutdown shows u | QA00000000 - u self-test screen after shutdown shows u |  | 2023-09-22 20:21 | No
1101 | Yes | PSE_QACASE for PCI imported product#00000 the tidal volume cannot be adjued. | PSE_QACASE for P0imported produu00000000 the tidal volume cannot be adjued. |  | 2023-09-22 20:30 | Unknown
1102 | Yes | Tidal volume was inaccurate, and could not be adjued | Tidal volume was inaccurate, and could not be adjued |  | 2023-09-22 20:40 | No
1103 | Yes | 00uuu Foot Pedal | 00CFuuu Foot Pedal | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001767 Mcnamara Pau The system meets the specification for the performed service and is retuned to ue Internal Comments : uLL# and/or description of uLL: Wireless foot switch loses connectivity Repair Action : Provide suport for this issu Diagnostic performed by Engineer : Cutomer stated end uers having issus where wireless u-ray foot pedals are not fuctional / loose connectivity Problem Description by Engineer : Problem description by engineer : wireless u-ray foot pedals are not fuctional / loose connectivity Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Resoluion : Spoke to cutomer and informed then that FCO72200000000522 is assigned to this system and case0000000021920000000032 is going to be schedued for this repair and issu ----------------------------------------------------------------------------- Cutomer comment : 000000003E000000000000000000000000ugcYIAR ----------------------------------------------------------------------------- | 2023-09-22 20:46 | No
1104 | Yes | Audio-No Sound / No Audio found at Bench-s/n u00 Other HW-Speaker-HW / Replacement Sent | Audio-No Sound / No Audio found at Bench-s/n u000/ Other HW-Speaker-HW / Replacement Sent |  | 2023-09-22 20:47 | No
1105 | Yes | patient disconnect alarm | patient disconneualarm |  | 2023-09-22 20:48 | No
1106 | Yes | Check Vent: Blower Temperature High (EC 0) | Check Vent: Blower Temperature High (EC 00) |  | 2023-09-22 20:51 | No
1107 | Yes | Audio-no sound/No audio found at Bench-Foulink s/n uE0/other HW-Speaker-HW/Replaced Parts | Audio-no sound/No audio found at Bench-Foulink s/n u000other HW-Speaker-HW/Replaced Parts |  | 2023-09-22 20:52 | No
1108 | Yes | QA00000- cannot pass self-test when device is powered on | QA0000000 - cannot pass self-test when device is powered on |  | 2023-09-22 20:52 | No
1109 | Yes | No Audio / No Audio found at Bench-Foulink s/n u00 Other HW-Speaker-HW / Replaced Parts | No Audio / No Audio found at Bench-Foulink s/n u000 / Other HW-Speaker-HW / Replaced Parts |  | 2023-09-22 20:54 | No
1110 | Yes | No Audio / SN u0Speaker | No Audio / SN u000Speaker |  | 2023-09-22 20:57 | No
1111 | Yes | Alarm light is faulty | Alarm light is faulty |  | 2023-09-22 21:00 | No
1112 | Yes | 00uuu Foot Pedal | 00CFuuu Foot Pedal | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001756 Mcnamara Pau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : wireless u-ray foot pedals are not fuctional / lose connectivity Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Resoluion : Spoke to cutomer and informed them that FCO72200000000522 is assigned to this system and case0000000021920000000032 is going to be schedued for this repair and issu Internal Comments : uLL# and/or description of uLL: Wireless foot switch loses connectivity Repair Action : Provide suport for this issu Diagnostic performed by Engineer : Cutomer stated end uers having issus where wireless u-ray foot pedals are not fuctional / lose connectivity ----------------------------------------------------------------------------- Cutomer comment : 000000003E000000000000000000000000ugcYIAR ----------------------------------------------------------------------------- | 2023-09-22 21:01 | No
1113 | Yes | The urs couldn't shock a ventricular tachycardia (VT) with the FR0 | The urs couldn't shock a ventricular tachycardia (VT) with the FR0 |  | 2023-09-22 21:11 | Yes
1114 | Yes | No Audio / SN u00peaker | No Audio / SN u000Speaker |  | 2023-09-22 21:20 | No
1115 | Yes | Audio & Sounds/No Audio / SN u00Speaker | Audio & Sounds/No Audio / SN u000Speaker |  | 2023-09-22 21:25 | No
1116 | Yes | Audio-no sound/No audio found at Bench Keyrin s/n u00ther HW-Speaker-HW/Replaced. | Audio-no sound/No audio found at Bench Keyrin s/n u000/other HW-Speaker-HW/Replaced. |  | 2023-09-22 21:35 | No
1117 | Yes | DFM0 red cross appeared | DFM0 red cross appeared | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Edward Steggall The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in ue Eupected and ueupected Resuts: Investigate logs re red cross uer Impact: None Patient Impact: None SE N/A Resoluion : The cutomer made contact regarding a red cross uLLfoud on their DFM100000000. They were able to resolve the issu with an op-check bu noticed a string of uLL in the logs that they wanted analyzing. Sumitted logs to Gao Jin 00u and asked if the uit coud retun to clinical ue, she gave the following response: "I saw a device reboot in AuoTest mode du to watchdog problem. Otherwise, no other uLL " "Yes, I think so. Ou ALS devices all contain a ‘safety-net’ or ‘watchdog’ featue that reboots the device in the event of a system timeou. This triggers (very) infrequntly and is designed to swiftly retun the device to a fuly operational state." ----------------------------------------------------------------------------- | 2023-09-22 21:36 | No
1118 | Yes | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 Device-Other-HW / Replacement Sent | Audio-No Sound / No Audio found at Bench-Keyrin s/n u000/ Device-Other-HW / Replacement Sent |  | 2023-09-22 21:42 | No
1119 | Yes | Audio-No Sound / No Audio found at bench-Foulink s/n u00/ Device-Other-HW / Replacement Sent | Audio-No Sound / No Audio found at bench-Foulink s/n u000 / Device-Other-HW / Replacement Sent |  | 2023-09-22 21:46 | No
1120 | Yes | Audio-no sound/No audio found at Bench- Foulink s/n uG0Device-other-HW/Replacement sent(PENDING MGR REVIEW) | Audio-no sound/No audio found at Bench- Foulink s/n u000Device-other-HW/Replacement sent(PENDING MGR REVIEW) |  | 2023-09-22 21:52 | No
1121 | Yes | uLLcodes 00000, 0 | uLLcodes 00, 00, 00 00 |  | 2023-09-22 22:15 | No
1122 | Yes | uPro ¦ charging port damage (discovered during repair) SPLIT | TempuPro ¦ charging port damage (discovered during repair) SPLIT |  | 2023-09-22 23:09 | No
1123 | Yes | OVP circuit failure | OVP 0cuit failure | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000000000009-Sep-2000000003 Galan, Alberto The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Ouside of ue Eupected and ueupected Resuts: Pass uer Impact: N/A Patient Impact: N/A SE 3.00000000 Diagnostic performed by Engineer : The FSE completed the installation of the MC PCBA and conduted a performance verification test following the gudelines provided by the manuactuer. Parts Consued : 1.00000000000000000000000000000000561537531,RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 FLOW METER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000007-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000007-31 00000000:00000000:00000000 MuTIMETER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE: Interaction Type: Cutomer Call Back Caller: Johnny Johnson - phone 700000000385 5000000003 IP & Issu: 566331300000000 CCB becaue they never received the qute Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Qute requst has the correct ##address on it bu the qute team did NOT send the qute to that ##address. I confirmed the cutomer's ##address and I forwarded the qute to cutomer Johnny. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : New Case Caller: Johnny Johnson/ Carlos Pereira Issu: 566331300000000elocation Change IPCR Issu : OVP circut failue Address: 9000000006 7th Ave BROOKLYN NY 11228 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: No contract or warranty Action Taken: Created case an requsted qute, sumitted relocation change Neut Step Action: Awaiting PO Verbal PO Nuber: N/A Verbal PO Approver: Johnny Johnson Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-22 23:36 | No
1124 | Yes | Device is giving uLLcode 00primary alarm failure | Device is giving uLLcode 00, primary alarm failure | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002900000000 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in patient ue Eupected and ueupected Resuts: Check vent alarm - primary speaker fail - parts ID for a speaker assy. uer Impact: none Patient Impact: na SE uknown Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : cutomer ohmed ou the motor speaker and foud it bad, open reading. provided Speaker assembly (base assembly) 100000000000000005 00000000561511921 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gino 253-5000000002551 IP & Issu: Part ID -speaker 1 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-23 00:01 | No
1125 | Yes | uuumalfunction | uuumalfunuon | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005789 00000000-Nov-2000000003 Gullaue MILLE The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: YES Eupected and ueupected Resuts: uuudoes not 000000000000000000000000001k. uer Impact: NO Patient Impact: NO SE 2.100000000Parts Consued: 1.0000000000000000000000000000000056153600000000,RP-BezelFront, w/NAV, Wht, NCTD, V600000000Resoluion: Replacement of the front cover containing the uuu Resoluion: uuudoes not 000000000000000000000000001k. Operating test OK ----------------------------------------------------------------------------- | 2023-09-23 00:06 | No
1126 | Yes | Device will not power on, unclear if it's the battery or the device | Device will not power on, unclear if it's the battery or the device | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000001 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : waiting for replacement battery to ship, order nuber 6300000000881700000000Diagnostic performed by Engineer : When the battery is inserted the AED does not power on. When the blu I buton is pressed there is no soud. The battery is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty Euchange. Device serial# A200000000-00000000581 Serial # of battery B00000000200000000512 Install before date: 12/2000000005 00758P Problem Description by Engineer : Cutomer Fuction/Role: end uer Device ue: ouside of clinical uage Eupected and ueupected Resuts: When the battery is inserted the AED does not power on. When the blu I buton is pressed there is no soud. The battery is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty Euchange. Device serial# A200000000-00000000581 Serial # of battery B00000000200000000512 Install before date: 12/2000000005 00758P uer Impact: none Patient Impact: none SE Resoluion : DELIVERED Wednesday 9/27/23 at 1:00000000PM Signed for by: SDuTON ----------------------------------------------------------------------------- | 2023-09-23 00:36 | No
1127 | Yes | M0A / u00/ The measurement module panel is damaged. | M00A / u00 / The measurement module panel is damaged. | Internal Remarks : Interaction Type: New Case Caller: Raymond Lindsley IP & Issu: 70000000096626 - The measuement modue panel is damaged. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-23 01:06 | No
1128 | Yes | Intrepid fails the therapy test | Intrepid fails the therapy test |  | 2023-09-23 01:09 | No
1129 | Yes | AED triple chirp | AED triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003328 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : waiting to ship, order nuber 000000000000000057386289 Problem Description by Engineer : Cutomer Fuction/Role: end uer Device ue: ouside of clinical uage Eupected and ueupected Resuts: The AED is to be replaced du to AED stating to remove and reinsert the battery which indicates that the AED is not 000000000000000000000000001king properly and needs replaced uder IB warranty euchange at zero cost to the cutomer. Device serial# A16A00000000833 Serial # of battery B00000000200000000800000000 Install before date: 00000000/31/2000000008 000000001700000000 Battery was installed 000000003 uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : The AED is to be replaced du to AED stating to remove and reinsert the battery which indicates that the AED is not 000000000000000000000000001king properly and needs replaced uder IB warranty euchange at zero cost to the cutomer. Device serial# A16A00000000833 Serial # of battery B00000000200000000800000000 Install before date: 00000000/31/2000000008 000000001700000000 Battery was installed 000000003 Follow u Requred : Tu5100000000Souh west Corporate Pky Palm City Fl 30000000000000000 Attention to: Dave Magruer Resoluion : DELIVERED Monday 9/25/23 at 9:21 AM Signed for by: A.GENT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003363 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-23 01:16 | No
1130 | Yes | The battery was completely depleted and would not charge | The battery was completely depleted and would not charge |  | 2023-09-23 01:52 | No
1131 | Yes | Triple chirps & "remove and reinsert battery" prompt | Triple chirps & "remove and reinsert battery" prompt | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003588 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below Town Of Sufside Attn: Marianella Marcano 9293 Harding Ave SuFSIDE, FL 331500000000000000009 Resoluion : Closing case tracking#780000000000000000000000007, DELIVERED Monday 9/25/23 at 2:56 PM. ##cutomer to confirm that the AED is 000000000000000000000000001king as it shoud. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping and will not pass the self-test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported the AED kept chirping and will not pass the self-test Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and will not pass the self-test, cutomer stated they installed a new battery. Cutomer removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Remove and reinsert battery to test. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Advise cutomer to remove device ou of service. util she received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# A200000000-00000000231 Battery Install Before# 2000000009/00000000/28 Lot#0000000063P Pads Eup Date: 2000000005/00000000/18 Pads Lot # C11823-00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000036200000000 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-23 02:16 | No
1132 | Yes | V0is displaying prouimal pressure uLL- Sent 0 GFE 00 | V0is displaying prouimal pressure uLL- Sent 0d GFE 0/0 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000display uLL uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Biomed is reporting installing a new flow sensor assembly and now the uit is displaying an uLL Diagnostic performed by Engineer : Problem Description: V600000000is displaying prouimal pressue uLLafter the flow sensor assembly was replaced. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/22/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Advised Biomed to call u Part Ordering and arrange and arrange a Parts Credit. Cutomer is reporting the part was ordered from PartsSouce. The original problem was the V600000000was failing Test 7 Air Flow Accuacy. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Michael Oliver IP & Issu: 79780000000073/The Prouimo sensor is bad. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-23 02:21 | No
1133 | Yes | M0A / u0 / Beeping and there is a red u flashing. | M00A / u00 / Beeping and there is a red u flashing. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000069 22-Sep-2000000003 Adale Mengistu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: nuse Device ue: device in ue Eupected and ueupected Resuts: cutomer claims device is beeping uer Impact: none Patient Impact: none SE uknown Resoluion : Pluged in device secuely and performed operational check to clear uLLmessages. Op check passed, performed shock test and confirmed its within specs. System operational Diagnostic performed by Engineer : Determined device is intermittently beeping and AC power is not fuly connected in the back ----------------------------------------------------------------------------- | 2023-09-23 04:01 | No
1134 | Yes | THE CATHETERIZATION TABLE OF MY EQUIPMENT IS FREE WHEN IT TURNS OFF, IT DOES NOT LOCK UP. TODAY THE EQUIPMENT I SPOKE WI | THE CATHETERIZATION TABLE OF MY EQUIPMENT IS FREE WHEN IT TURNS OFF, IT DOES NOT LOCK UP. TODAY THE EQUIPMENT I SPOKE WI | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007262 300000000Sep-2000000003 Chacon, Migul The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A. The equpment tuns off duing long stuies. Resoluion: In the face of the sitution euplained by the client, the following activities are performed: * Installation of a new uS inside the hemodynamics power distribuion cabinet, this uS will keep the compuers powered so that they can docuent failues becaue of a power ouage. * According to the evalution of the electrical board, only the electrical voltage protection is enabled, the curent protection is omitted, becaue the operation of the Hemodynamics has different curent consuptions, depending on the mass of the patient. Electrical measuements are performed to show these variables. Based on the above, it is determined not to ue the protection to limit the curent. * Gave technical gudance to the biomedical engineer abou the operation and distribuion of the power in the equpment. * It is recommended for the pressue watcher to acqure one that allows to locate the 000000000000000000000000001k limit at 00000000. (that it has a range of variations with higher resoluion). The curent one goes from 00000000 to 00000000. * It is recommended to add a signaling pilot in the electrical board, to indicate if the oulet contactor is active or not. * Recommend implementing protection covers to the emergency stops located at the entrance of the room, to avoid tuning off accidentally. The equpment is delivered for clinical ue. Equpment uder evalution. Note: By architecting the equpment, the table releases the brakes when there is no power suply or they perform an emergency stop. Tools & Equpment: MuTIMETER, 200000000000000001-13 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-23 06:36 | No
1135 | Yes | 0/u FD0Live monitor display problem | 0R/u FD0Live monitor display problem | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Repair Action: PPD/ 00000000800000000000000006562 5V POWERSuPLY HEAMO WITH CABLES Neut Step Action: PPD FRu5V powersuply Heamo with cables 00000000800000000000000006562*1 Need to order, please send 000000000000000000000000001k first FSE: Feng Lei -86-19815665565 Internal Remarks: 26.00000000.2000000003 00000000:21:00000000 uC (86000000008238) $Q1: (N/A) $Q2: (N/A) $Q3: (N/A) $Q00000000(N/A) $Q5: (N/A) $Q6: (live monitor and reference monitor no signal inpu) $Q7: (Confirmed reported problem, live monitor and reference monitor no signal inpu) $Q8: (check the log, 5V power suportly on MCS problem, need replace 5V power suportly on MCS) $Q9: (replace 5V power suportly on MCS) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (N/A)S/N (N/A) Neut Calibration date (N/A) SAP Transfer: Case Transferred to SAP Euternal Remarks: 26.00000000.2000000003 00000000:21:00000000 uC (86000000008238) Service content (requred): Cutomers reported that the real-time screen and reference screen had a black screen and a problem with the 5V power modue duing ue. After replacement, everything retuned to normal and the equpment was normal. Service resut (requred): The system statu is normal and can be ued Other important matters: None Safety Qustions: Case: 0000000022320000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: Neither the real-time screen nor the reference screen is on.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: Neither the real-time screen nor the reference screen is on., the console display is normal and previou images can be reviewed.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: I supect there is a problem with the 5V power suply of the DVI converter. Notes on cutomer requrements: Onsite Problem Reported by Cutomer: Please confirm with SS Zhang Yongzhong before dispatching 000000000000000000000000001k or spare parts. T2 Activities:Repair Action: Alternative nuber confirmation: 00000000800000000000000006552 5V POWERSuPLY WITH IN/OuPu CABLES ----------------------------------------------------------------------------- | 2023-09-23 06:36 | No
1136 | Yes | QAA0/u FD0Startup failed | QAA0R/u FD0Startup failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine coud not enter the application interface normally and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine cannot enter the application interface normally Remote diagnosis and test resuts: power ouage and restart have no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the Host PC starts normally, try to clean and plu in the internal board, 2. Check whether the BIOS battery is normal 3. ue PC Diagnostic Tool for diagnostic testing 00000000Check whether the software is normal, try to ue Ghost to restore the system 5. Check M Check whether the DCPS of the cabinet is normal. 6. Check whether the internal commuication of the system is normal. 7. Check the machine uLLand repair the faut. If it cannot be repaired, replace the corresponding spare parts. Recommended spare part uuand name: 00000000800000000200000000612 Nehalem Host PC Monoplane Rev_III NO_HDD & 0000000080000000050000000000000000 Hard disk spare part Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Euternal Remarks: 22.00000000.2000000003 13:10000000053 uC (86000000008196)"SuT. Ready for Clinical ue local time" : " 2000000003-00000000-22 20000000000000000" SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022320000000027 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: QAA00000000R/u uuuFD200000000Euternal Remarks: 23.00000000.2000000003 00000000:19:00000000 uC (86000000008196)*Department information (requred): Department: Interventional Operating Room Contact Person: Teacher Zeng Tel: 16627791297 * Service content (requred): uer interface: Data monitor prompts that the system cannot be entered. System files are missing. Restore Ghostimage restore sucessfuly And redo the system settings. Service resut (requred): The device retuned to normal. Other important matters: None Internal Remarks: 23.00000000.2000000003 00000000:0000000000000000 uC (86000000008196) "RDI Accuacy: Yes" $Q7: (Confirmation & The uer interface: The Data Monitor prompts that youcannot enter the system.) $Q8: (Windows failed to load becaue the system registry file is missing, or corrut.& Resut: System issu.) $Q9: (Restore Ghostimage sucessfuly. And redo the system settings.) $Q100000000(Test and Inspection data refer to: PA & Device retun to fuly service) $Q11:Calibrated Tools name (refer to PA) S/N (refer to PA) Neut Calibration date (refer to PA) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 00000000/28/2000000002 ^Curent Tue uage: 159990000000075 Lu----------------------------------------------------------------------------- | 2023-09-23 08:51 | No
1137 | Yes | 0uDFM0 Defibrillator/Monitor/Battery failure | 0/Effi0 DFM0 Defibrillator/Monitor/Battery failure | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor/battery failue Safety Qustions: Case: 0000000022320000000032 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-23 09:01 | No
1138 | Yes | 0S:KN:The equipment does not start. | 0DS:KN:The equipment does not start. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 23-Sep-2000000003 Ono, Shinji The system meets the specification for the performed service and is retuned to ue Internal Comments: The warantity seal was peeled off, so I replaced the battery instead of replacing the PC.A few years ago, when the memory was re-puled ou, it opened with TS permission. Reguatory Qustions: N/A Problem Description by Engineer: The equpment does not start. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The equpment does not start. uer Impact: No effect Patient Impact: No effect SE 000Resoluion: When I carried ou a diagnostic test on the main control PC (HOST PC), I confirmed that the BIOS CMOS battery was low.Replacing the battery restored the recovery. After the 000000000000000000000000001k, I confirmed that there was no problem with system operation. Follow u Requred: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 Noda, Kotaro The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: Simpl 1.2 [Suject] 00000000DS:KN:u Suaua: The equpment does not start. [Internal Comments] Compl 1.2 [FREE_TEuT] ID: Host PC, uu 0000000080000000050000000000000000, S/N: 5315119000000000000000000000000Power on hous: 300000000 [Parts from the equpment] 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Tokyo (u11) [u-ed parts] 0000000080000000050000000000000000:Host PC Monoplane Revised_II no HDD QTY:1 ZREP/Traceable Tokyo (u11) Reason for non-warehoue: No parts in stock. 00000000800000000122391:u Haswell Monoplane Host PC no HDD QTY:1 ZREP/Traceable Tokyo (u11) Reason for u-collection: Consuted with engineer. Resut of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Shinji Ono] Ono, Shinji - FSE1 Visit date and time 2000000003/00000000/2316:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions CPu(no warehoue remove/transport arrangements) - Delivery Specified time: 2000000003/00000000/23 at 10000000000000000 - Delivery address name Sapporo Office - Received person name Shinji Ono - 000000000000000000900000000-000000002 - Others Requster name Kn [Euamination category] Normal euamination Follow u Requred: * Problem Description by Engineer: The equpment does not start. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The equpment does not start. uer Impact: No effect Patient Impact: No effect SE 000Diagnostic performed by Engineer: I had him start the equpment, bu the data monitor screen was dark and only the cusor was displayed.I had him restart it a nuber of times, bu that does not change anything.Restarting the breakers does not improve the sitution. FSE response. ----------------------------------------------------------------------------- | 2023-09-23 10:06 | No
1139 | Yes | Host computer not booting | Host computer not booting | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 23-Sep-2000000003 Kuar, Muish The system meets the specification for the performed service and is retuned to ue Resoluion : Checked and foud power issu reset the power connection then handover the system in good 000000000000000000000000001king condition Problem Description by Engineer : Problem description by engineer : Live and Review monitor and C arm an Table movement not 000000000000000000000000001king When was the first occurence? Within... : 23-Sep-2000000003 10000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Live and Review monitor and C arm an Table movement not 000000000000000000000000001king uer Impact: No impact Patient Impact: No patient impact SE 8.2.17.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001785 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Live and Review monitor and C arm an Table movement not 000000000000000000000000001king When was the first occurence? Within... : 23-Sep-2000000003 10000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Live and Review monitor and C arm an Table movement not 000000000000000000000000001king uer Impact: No impact Patient Impact: No patient impact SE 8.2.17.00000000Diagnostic performed by Engineer : Host PC was not booting here. FSE to visit for confirmation and futher checks on smps/motherboard. Follow u Requred : uLL# and/or description of uLL: Live and Review monitor and C arm an Table movement not 000000000000000000000000001king Repair Action : Host PC was not booting here. FSE to visit for confirmation and futher checks on smps/motherboard. Technician Name : Vijayakuar, K.S Resoluion : Host PC was not booting here. FSE to visit for confirmation and futher checks on smps/motherboard. Internal Comments : uLL# and/or description of uLL: Live and Review monitor and C arm an Table movement not 000000000000000000000000001king Repair Action : Host PC was not booting here. FSE to visit for confirmation and futher checks on smps/motherboard. ----------------------------------------------------------------------------- | 2023-09-23 10:11 | No
1140 | Yes | 0/uLuCentron/System cannot make euposure | 0R/uLuCentron/System cannot make euposure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Neut Step Action: PPD00000000210000000016291*1 Arranged for 200000000delivery to the venu FSE Yang Zhiuiong-86-13560000000020000000057 Internal Remarks: 23.00000000.2000000003 17:32:00000000uC (8600000000000000001) RDI Accuacy: Yes. $Q7:(Confirm reported problem, System euposue failed. ) $Q8: (Performed visul inspection, System euposue failed, onsite check, The tue Oil leakage.) $Q9: (Need to replace the tue, The Warranty contract does not inclue tue, waiting for buiness process. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name [PA] Euternal Remarks: 23.00000000.2000000003 16:0000000022 uC (8600000000000000001) [Service Content] Euposue is a black pictue. 1. On-site inspection, the euposue was interruted after 1-2 seconds. On-site inspection, the anode voltage was 800000000V, the cathode was 17KV, the cathode voltage oupu was low, the log indicated that Q2 (inverter) was short-circuted, Q2 was replaced, and the problem remained. 2. Replace Q1 and Q2, the problem remains. 3. Check that the cables and cable heads are normal. 00000000Open the shell of the ball tue and find that the ball tue is leaking oil. The short circut occus inside the ball tue and the ball tue needs to be replaced. [Service resuts] It is recommended to replace the ball tue. [Other Matters] Standard warranty, eucluing ball tue, waiting for buiness process. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 23.00000000.2000000003 1000000008:00000000uC (8600000000000000001) [Service Content] Euposue is a black pictue. 1. On-site inspection, the euposue was interruted after 1-2 seconds. On-site inspection, the anode voltage was 800000000V, the cathode was 17KV, the cathode voltage oupu was low, the log indicated that Q2 (inverter) was short-circuted, Q2 was replaced, and the problem remained. 2. Replace Q1 and Q2, the problem remains. 3. Check that the cables and cable heads are normal. 00000000Open the shell of the ball tue and find that the ball tue is leaking oil. The short circut occus inside the ball tue and the ball tue needs to be replaced. [Service resuts] It is recommended to replace the ball tue. 【Other matters】None yet T2 Activities:Problem Description: Overview of cutomer problems: The device can only eupose one image, even if it is restarted repeatedly.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: Withou PRS and logs, it is impossible to determine the caue of the failue. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: The engineer arrives to check the statu indicator light of the FDC, checks the system and high-voltage logs, and performs repairs according to the prompts. Notes on cutomer requrements: Onsite SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/uLu Centron T2 Activities:Repair Action: PPD/ 00000000210000000016291 Converter 8E Velara Safety Qustions: Case: 000000002232100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-23 11:31 | No
1141 | Yes | DFM0 paddle test failed | DFM0 paddle test failed | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 12.1000000002000000003 0000000020000000028 uC (860000000000000000100000000 ASP engineer: [honglejia wangyong] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.urecognized handle. $Q8: The handle is detected to be fauty, It can't 000000000000000000000000001k normally. $Q9: Replace the defibrillator handle. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: $ Calibrated Tools name (ProSim8) S/N (26712200000000 Neut Calibration date (10000000018/2000000003) Calibrated Tools name (ESA6200000000 S/N (300000000000000003) Neut Calibration date (11/3000000002000000003) Calibrated Tools name (70000000000000000P) S/N (3753000000006) Neut Calibration date (10000000029/2000000003) Euternal Remarks: 12.1000000002000000003 0000000020000000025 uC (860000000000000000100000000Department information [Department: Equpment Department Contact Person: Teacher Ma Contact Nuber: 13611000000005115]* ASP engineer: [honglejia wangyong] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.urecognized handle. $Q8: The handle is detected to be fauty, It can't 000000000000000000000000001k normally. $Q9: Replace the defibrillator handle. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: $ Calibrated Tools name (ProSim8) S/N (26712200000000 Neut Calibration date (10000000018/2000000003) Calibrated Tools name (ESA6200000000 S/N (300000000000000003) Neut Calibration date (11/3000000002000000003) Calibrated Tools name (70000000000000000P) S/N (3753000000006) Neut Calibration date (10000000029/2000000003) Problem Reported by Cutomer: The handle is broken 0000000000000000000000001655 call T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Defibrillation handle failue Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: uplu and uplu therapy cables, check the tray, reru OP check, and trouleshoot the defibrillation handle Safety Qustions: Case: 000000002232100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-09-23 11:46 | No
1142 | Yes | AED failed self-test | AED failed self-test | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : CM - SELF TEST FAILED Safety Qustions : Case: 0000000022321576 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-23 12:06 | No
1143 | Yes | CM - SELF TEST FAILED | CM - SELF TEST FAILED | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : CM - SELF TEST FAILED+Battrey Issu Safety Qustions : Case: 00000000223215900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-23 12:06 | No
1144 | Yes | Machine is not switching On | Machine is not switching On | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Vijayakuar, K.S The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching On When was the first occurence? Within... : 23-Sep-2000000003 11:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching On uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.300000000Diagnostic performed by Engineer : Host PC is not switching ON intermittently. Issu was with cmos battery. Reseated the battery and now it is switching ON properly. Internal Comments : uLL# and/or description of uLL: Machine is not switching On Repair Action : Host PC is not switching ON intermittently. Issu was with cmos battery. Reseated the battery and now it is switching ON properly. Resoluion : Host PC is not switching ON intermittently. Issu was with cmos battery. Reseated the battery and now it is switching ON properly. ----------------------------------------------------------------------------- | 2023-09-23 12:11 | No
1145 | Yes | AED emitted triple chirp, "uLL voice message | AED emitted triple chirp, "uLL voice message | Internal Remarks: Failue content 2 The AED beeped three times. Even when a self test is ru, the "uLL message is annouced. LOT Adut pad Y000000002119-00000000 Battery 8200000000P ----------------------------------------------------------------------------- | 2023-09-23 12:21 | No
1146 | Yes | Device Fails Self-Test | Device Fails Self-Test | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008800000000 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "September 29, 2000000003 Fuua Denshi ShikokuHanbai Report to the on-site Person in Charge and I also sent the equpment and contacted the cutomer." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: I received the AED and performed a battery test, and I confirmed that it passed the test.I also confirmed that the AED beeped.I monitored the AED in standby mode, bu no alarms, etc. occured. I verified the AED operation record and foud that the calibration resuts euceeded the permissable range per the specifications for the high voltage circut calibration test items for the monthly auomatic self test and the battery fitted test, so the test failed. "I determined that the 1-beep alarm and the annoucement that was pointed ou had souded a total of 2 times. " "Curently, the alarm is not reprodued and it is operating, bu I will replace the AED in order to ensue safety. I apologize for the inconvenience this has caued." Parts Consued: 1.0000000000000000000000000000000056000000008200000000HS1 + (W G) Defib, Japanese, Euchg ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 (1) A single-beep soud occured, so he pressed the i-buton, and the "To test, The message "Please remove the battery again" was made, so I removed the battery and He re-installed it and carried ou a self check, bu the AED did not beep. A few days later, the same problem occured, so he re-installed the battery the same way, bu The AED did not beep. He has been uing the system since then, bu the above problem has not re occured. Lot No. Adut pad C120000000022-00000000 Adut pad C120000000022-00000000 Battery 00000000333P ----------------------------------------------------------------------------- | 2023-09-23 13:56 | No
1147 | Yes | The AED beeped when the pad was changed - log analysis requested (pads are replaced) | The AED beeped when the pad was changed - log analysis requested (pads are replaced) | 000000000000000000000000001O##: 000000000000000000000000001000000000000000069800000000Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the pad.I retuned the AED withou an abnormality. I had him set the clock, and I asked him to resue operation. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." " October 11, 2000000003 Fuua Denshi Tokyo Hanbai on-site person in charge I sent the equpment along with the report." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Diagnostic performed by Engineer: I performed a uer confirmation test on the AED system and confirmed that it passed the test.After that, I monitored the AED in standby mode for t000000000000000000000000001weeks, bu no alarm souded, etc. I analyzed the AED operation record and foud that in the pad test items for the recent self test, the pad impedance (internal resistance) measuement resuts were slightly ouside the allowable range, so the pad alarm souded "There was a record of it occurence several times.The AED time was not set, so I do not know when the problem occured. " There was no peeling on the electrode pad seal on the received pad, bu the protective plate (plastic sheet) was slightly deformed, more than normal.As a resut of this, "I think the alarm souded du to a poor contact when the pad was set in the pad slot, and the impedance flutute. This time, I will replace the PAD." Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Reguatory Qustions: N/A ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 When he changed the pad, the AED beeped. (This facility requsted on December 21, 2000000002, and there is a history of prematue battery wear.） LOT Adut pad 230000000016-000000007 Battery M000000008000000002216 ----------------------------------------------------------------------------- | 2023-09-23 14:01 | No
1148 | Yes | The AED beeped once, BIT passed - log analysis requested | The AED beeped once, BIT passed - log analysis requested | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008885 Suui, Kosue The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of Sales at Fuda Electronics Device ue: Not ued Eupected and ueupected Resuts: The AED beeped once. uer Impact: None Patient Impact: None SE None Diagnostic performed by Engineer: When the standby device was received, the alarm was not confirmed. I confirmed that the AED passed a battery test. I confirmed from the log that the AED periodically failed. Resoluion: When I analyzed the AED log at a factory in the uS., I foud a record of a high voltage (HV) calibration test not being passed. The AED was not diagnosed du to an electrical failue on a high voltage element in the AED, so I replaced the AED. I replaced the AED and confirmed that the AED passed a battery fit test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred. October 100000000 2000000003 Fuua Denshi Kanagawa Hanbai I will send the 000000000000000000000000001k report along with the equpment to the person in charge. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit a 000000000000000000000000001k report on paper? ＜YES＞ Final check: I checked the operation following the 000000000000000000000000001k based on the set of orders. Yes ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The alarm souds every month.He removed the battery and performed a self test, and there was no problem for a while, bu abou a month later the alarm souded again. Lot No. Adut pad C000000002922-00000000 Adut pad C000000002922-00000000 Battery 0000000000000000P ----------------------------------------------------------------------------- | 2023-09-23 14:06 | No
1149 | Yes | The AED beeped once even when pad is removed and re-installed. - log analysis requested (Device is replaced) | The AED beeped once even when pad is removed and re-installed. - log analysis requested (Device is replaced) | 000000000000000000000000001O##: 000000000000000000000000001000000000000000089300000000Suui, Kosue The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of Sales at Fuda Electronics Device ue: Not ued Eupected and ueupected Resuts: The AED beeped once. uer Impact: None Patient Impact: None SE None Diagnostic performed by Engineer: When I received the AED, I performed a battery test and confirmed that it passed. I did not confirm the reproduibility of the alarm on standby monitoring either. When I checked the battery voltage, I confirmed that it was 000000000000000000000000001n ou. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit a 000000000000000000000000001k report on paper? ＜YES＞ Final check: I checked the operation following the 000000000000000000000000001k based on the set of orders. Yes Tools & Equpment : MuTIMETER, 200000000000000003-15 00000000:00000000:00000000 Resoluion: When I analyzed the AED log at a factory in the uS., I foud a record of a high voltage (HV) calibration test not being passed. The AED was not diagnosed du to an electrical failue on a high voltage element in the AED, so I replaced the AED. I will replace the battery as the battery has 000000000000000000000000001n ou prematuely. I replaced the AED and confirmed that the AED passed a battery fit test. * In this case, abnormal records were confirmed at an AED temperatue lower than the recommended temperatue. The AED and accessories are fauty and can caue prematue wear, so please store the AED uder the recommended temperatue environment. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred. October 200000000 2000000003 Fuua Denshi Hiroshima Sales I will send the 000000000000000000000000001k report along with the equpment to the person in charge. Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The uLLstill occus even when the pad is removed and re-installed. Lot No. Adut pad C100000000522-00000000 Adut pad C100000000522-00000000 Battery 0000000000000000 ----------------------------------------------------------------------------- | 2023-09-23 14:11 | No
1150 | Yes | Loudspeaker defective-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker defeu-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009128 Lang Dominik The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: TD Device ue: yes Eupected and ueupected Resuts: Louspeaker defective uer Impact: no Patient Impact: no SE M.00000000Resoluion: Louspeaker replaced Performed stK Complete fuction check Parts Consued: 1.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000003-00000000 00000000:00000000:00000000 PATIENT SIMuATOR, 200000000000000006-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-23 14:26 | No
1151 | Yes | AED single beep resolves after battery reinstalled but alarm recurs - log analysis requested | AED single beep resolves after battery reinstalled but alarm recurs - log analysis requested | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009100000000 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "October 5, 2000000003 Fuua Denshi Nagano Hanbai Report to the on-site Person in Charge and I also sent the equpment and contacted the cutomer." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: "I received the AED and performed a battery fitted test, and I confirmed that it passed the test. I analyzed the AED operation record and foud that the alarm that was souding was the monthly auomatic self-test I determined that in the high voltage calibration items in the test, the AED beeped once becaue the AED was ouside the passed range, and the message "To ue, remove and re-install the battery" was made. "The alarm has now been resolved, bu I will replace the AED as it is in the fuue of a failue. I apologize for the inconvenience this has caued." Parts Consued: 1.0000000000000000000000000000000056000000005231,FRu+ (C G) Defib, Japanese, Euch ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 Abou a month ago, the AED beeped once. The problem was restored when the battery was re-installed, bu then the problem re-occured a month later. Lot No. Adut pad 230000000027-0000000051 Adut pad 230000000027-0000000051 Battery 00000000600000000 ----------------------------------------------------------------------------- | 2023-09-23 14:36 | No
1152 | Yes | Last Week Saturday there was an uneuplained death case, request for evaluation on site | Last Week Saturday there was an uneuplained death case, request for evaluation on site | Neut step Action: Case will be closed since safety qustions were incorrectly recorded. Continud in case 0000000022323000000002 ----------------------------------------------------------------------------- | 2023-09-23 15:41 | Yes
1153 | Yes | 00FM0 restarts constantly | 000DFM0 restarts constantly | 000000000000000000000000001O##: 00000000000000000000000000100000000155800000000 Ruz, Ignacio The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: Defibrillator. Eupected and ueupected Resuts: The equpment restarts constantly. The defibrillator is tested with the test "Processor board" and the correct operation of the measuements and fuctions of the equpment is verified. This part needs to be replaced. uer Impact: uknown. Patient Impact: uknown. SE 2.00000000.23 Resoluion: Processor Board is euchanged. It was not possible to maintain the client's own configuation. The configuation of the device is restored to the factory. All necessary verifications have been performed to certify the restituion to the conditions prior to the failue. To gurantee the correct operation of the equpment, we recommend them to ue only the accessories and consuables approved by u. Parts Consued: 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor Tools & Equpment: ELECTRICAL SAFETY TESTER, 2000000003-12-00000000 00000000:00000000:00000000 PATIENT SIMuATOR, 2000000003-12-00000000 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER, 200000000000000007-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 Ruz, Ignacio The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: Defibrillator. Eupected and ueupected Resuts: The equpment restarts constantly. The defibrillator is tested with the test "Processor board" and the correct operation of the measuements and fuctions of the equpment is verified. This part needs to be replaced. uer Impact: uknown. Patient Impact: uknown. SE 2.00000000.23 Diagnostic performed by Engineer: The equpment restarts constantly. The defibrillator is tested with the test "Processor board" and the correct operation of the measuements and fuctions of the equpment is verified. This part needs to be replaced. Follow u Requred: The equpment restarts constantly. The defibrillator is tested with the test "Processor board" and the correct operation of the measuements and fuctions of the equpment is verified. This part needs to be replaced. ----------------------------------------------------------------------------- | 2023-09-23 17:06 | No
1154 | Yes | 00CG Test getting failed | 000uTest getting failed | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005100000000 Hariprasad Malasi The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: uTest getting failed uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Onsite visit requred Follow u Requred : onsite visit ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005153 23-Sep-2000000003 Tomar, Sandeep The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: uTest getting failed uer Impact: no Patient Impact: no SEna Resoluion : Checked and foud ulead (LOT NO. 00000000-20000000000000000defective and to be puchase by cutomer. ----------------------------------------------------------------------------- | 2023-09-23 17:16 | No
1155 | Yes | 00istributor case Medsci: Supply part only - ETCO0eceptacle broken | 000Distributor case Meds0 Supply part only - ETCO0receptacle broken |  | 2023-09-23 18:16 | No
1156 | Yes | Geometry movement is partially functioning | Geometry movement is partially funuoning | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Vijayakuar, K.S The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - POS: uable to commuicate with GeoIPC 23-Sep-2000000003 12:50000000000000000 When was the first occurence? Within... : 23-Sep-2000000003 19:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: NO uer impact Patient Impact: No patient impact SE 8.1.30000000015 Resoluion : Biomedical helped in reseating the connections and cables of geopc. Then geometry movements 000000000000000000000000001king fine. Problem resolved remotely. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. When problem is persistent it coud be caued du to a defective control net000000000000000000000000001k switch (TS). Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. Bad cabling and/or connections. Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. Issu with GeoIPC OS/AS software. It coud be du to a not correct shudown before , Windows OS can be corruted hence it needs more time in the neut start u. If this happens only sporadically, please ignore it. Else see other possible caues & soluions. Root Caue for RMT - POS: uable to commuicate with GeoIPC: 1. The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. Repair Action : Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check control net000000000000000000000000001k switch (TS) and its power suply (adapter). Replace when necessary. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check the cable RC:u25 to SCC1:u22. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. 1.2. Tighten all the connectors of GeoIPC and SCC1 uit. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. If SW is really corruted please reload GeoIPC software uing uB destroy stick. Service actions for RMT - POS: uable to commuicate with GeoIPC: 1.1. Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start u of the system. If so, replace GEO IPC. Diagnostic performed by Engineer : Biomedical helped in reseating the connections and cables of geopc. Then geometry movements 000000000000000000000000001king fine. Problem resolved remotely. ----------------------------------------------------------------------------- | 2023-09-23 19:16 | No
1157 | Yes | CPU does not turn on by itself, we had to turn it on manually following instructions from the uYy technician. | CPU does not turn on by itself, we had to turn it on manually following instruuons from the uYy techni0n. | BuQ&R Feedback: No SAP Cutomer Symptom Code: BOOT Cutomer Symptom Code: Boot u/Power-u Problem Reason: Corrective Maintenance. Engineer Failue Code: Electrical/Electronic Failue Repair Activity Code: Re-install FRuor Other Electrical Connection. Safety Qustions: Case: 00000000223200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: uyni Edwirges. Last udated by: uyni Edwirges. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? N/A, Problem Reported by Cutomer: Cleiton +5560000000025000000006 engenharia00-client.com.br Hospital Nasr Faiad Avenida William Faiad, nuber 15, Centro, Catalão, Goiás. SAP Transfer: Case Transferred to SAP. T2 Activities:Problem Description: Host does not tun on. Trouleshooting Action: Checked that almost every day the host does not tun on after the machine is on and cutomer Cleiton has to tun it on manully. Repair Action: - Asked the cutomer to replace the CR2000000002 battery of the host PC; The cutomer, Cleiton, replaced the battery on September 200000000; The equpment was uder observation for 5 days. After replacing the battery the defect did not occu again. Cutomer Cleiton is aware the call is being closed as solved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK. Test Equpment: N/A Test/Inspection: OK. Test Equpment: N/A SAP Transfer: Case Transferred to SAP. ----------------------------------------------------------------------------- | 2023-09-23 19:41 | No
1158 | Yes | Alarms or Messages/Bench Repair Tracking: Speaker defective/Speaker Assembly | Alarms or Messages/Bench Repair Tracking: Speaker defeu/Speaker Assembly | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001600000000 Lang Dominik The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: TD Device ue: yes Eupected and ueupected Resuts: Speaker defective uer Impact: no Patient Impact: no SE M.00000000Resoluion: Louspeaker replaced Performed stK Parts Consued: 1.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000003-00000000 00000000:00000000:00000000 PATIENT SIMuATOR, 200000000000000006-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-23 19:51 | No
1159 | Yes | It is showing a failure in the defibrillation test. | It is showing a failure in the defibrillation test. | BuQ&A Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Equpment problem: the defibrillation test is failing priority: stopped search for: Jessica Miranda sector: Clinical engineering 00 11 31700000000332 00 e-mail: Jessica.Miranda00tejz.Spdm.org.br uit: SPDM ASSOCIAção PAuISTA PARA O DEVELVIMENTO DA MEDICINE address: Avenida Brigadeiro Lus Antonio 2651 Sao Pauo-SP 0000000000000000-900000000 BR Internal Remarks: 19.1000000002000000003 00000000:00000000:500000000POLARIS (POLARIS) ================= Gutavo de Assis Soua - 0000000000000000570000000071300000000100000000000000000000000038 =================== ==== GG6: Operational tests approved. GG13: No calibratable tools were ued. Trouleshooting Action: schedued for 10000000016 - FSE Gutavo Safety Qustions: Case: 0000000022320000000000000000afety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Nicole Carmo Ferreira Last udated by: Nicole Carmo Ferreira 1. The equpment was in clinical ue at the time the problem was detected ? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Was the equpment defective or failed at the time of installation? uknown Trouleshooting Action: OPP-197689 Qute: 00000000112300000000 ## Jessica.Miranda00tejz.Spdm.org.br Euternal Remarks: 19.1000000002000000003 00000000:00000000:500000000POLARIS (POLARIS) ================= Gutavo de Assis Soua - 0000000000000000570000000071300000000100000000000000000000000038 =================== ==== Initial failue: Equpment does not pass the operational test and does not release for shock application. It is necessary to change the defibrillator uuuitor. Equpment not ok. Trouleshooting Action: Qute sent SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to sending a buget SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-23 20:21 | No
1160 | Yes | M0A - u0- Defibrillation test failure | M00A - u00- Defibrillation test failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to the buget sumission SAP Transfer: Case Transferred to SAP Trouleshooting Action: Schedued for ESF Gutavo on 13/11 SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: Equpment Problem: Defibrillation Test Failue Priority: Stopped Search for: Jessica Miranda Indutry: Clinical Engineering Phone: 11 31700000000332 00 ## Jessica.Miranda00tejz.Spdm.org.br uit: SPDM SÃO PAuO ASSOCIATION FOR THE DEVELOPMENT OF MEDICINE Address: Avenida Brigadeiro Lus Antonio 2651 Sao Pauo-SP 0000000000000000-900000000 BR Euternal Remarks: 16.11.2000000003 18:19:22 POLARIS (POLARIS) ========== Gutavo de Assis Soua - 00000000000000005760000000033800000000100000000000000000000000038 ============================ uuu u0000000050000000000000000Encouters "Appliance Maintenance uLL SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022320000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Nicole Carmo Ferreira Last udated by: Nicole Carmo Ferreira 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A Trouleshooting Action: Buget sent Trouleshooting Action: OPP-197698 Qute: 00000000112317 ## jessica.miranda00tejz.spdm.org.br SAP Transfer: Case Transferred to SAP SAP Transfer: Service Order is Locked by another uer!: ----------------------------------------------------------------------------- | 2023-09-23 20:21 | No
1161 | Yes | Fluoro pedals are not functioning properly | Fluoro pedals are not funuoning properly | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005000000003 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: Fluro pedals are not fuctioning properly Repair Action : There is no confirmation abou the problem. We coud not reach the caller or FSE to get it clarified. Hence releasing the case to onsite for FSE to check on the issu. Diagnostic performed by Engineer : There is no confirmation abou the problem. We coud not reach the caller or FSE to get it clarified. Hence releasing the case to onsite for FSE to check on the issu. Problem Description by Engineer : Problem description by engineer : Fluro pedals are not fuctioning properly When was the first occurence? Within... : 16-Sep-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Fluro pedals are not fuctioning properly uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.3000000005 Resoluion : There is no confirmation abou the problem. We coud not reach the caller or FSE to get it clarified. Hence releasing the case to onsite for FSE to check on the issu. Follow u Requred : uLL# and/or description of uLL: Fluro pedals are not fuctioning properly Repair Action : There is no confirmation abou the problem. We coud not reach the caller or FSE to get it clarified. Hence releasing the case to onsite for FSE to check on the issu. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005398 Karkate, Vinod Problem Description by Engineer : Problem description by engineer : Fluro pedals are not fuctioning properly When was the first occurence? Within... : 16-Sep-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Fluro pedals are not fuctioning properly uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.3000000005 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-09-23 20:41 | No
1162 | Yes | uCVu - Freezes and become non responsive | uCVu - Freezes and become non responsive | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001976 25-Sep-2000000003 Smith Terry The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Performed system tests and passed all tests and all imaging modalities 000000000000000000000000001k as designed. I foud that the hard drive was at 87% Anesthesiologist. I informed the cutomer Rusell that the hard drive patient data 000000000000000000000000001ld have to be at the very minimal 000000000000000000000000001ld have to be partially deleted bu preferably all patient data deleted. He was going to check with the Anesthesiologist to verify that was ok and also to configue the auo delete fuction to reguated hard drive space. Problem Description by Engineer : Cutomer Fuction/Role: Manager Device ue: Strutual heart procedues Eupected and ueupected Resuts: Eupected uer Impact: Minimal Patient Impact: Minimal SE 9.000000002 ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Kristie Johnson Issu: 79000000007721 Freezes and become non responsive so does the track ball . Have to reboot mutiple times to bring machine back u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case an onsite 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-23 20:46 | No
1163 | Yes | PartsID: Audio&Sounds/No Sound: Euternal Speaker- failed/no sound: Speaker/HW: Provided Information. | PartsID: Audio&Sounds/No Sound: Euternal Speaker- failed/no sound: Speaker/HW: Provided Information. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002517 Vrooman, Travis The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: 000000000000000000000000001king speaker uer Impact: Loss of central monitor auio Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Speaker failed/no soud Diagnostic performed by Engineer : Problem Description: Speaker failed/no soud PartsID - Euternal Speaker - Speaker failed/no soud Form ued: Parts Trouleshooting Steps Taken: Provided part nuber for replacement speaker. Resoluion : Is Caue Known: Yes Caue: Hardware failue Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Resolved Remotely Parts/Suplies Information Provided Part ID and Price: 0000000056000000007331 - $126.00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Keitra Ewing IP & Issu: 92658952 - The speaker has stop 000000000000000000000000001king. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-23 22:21 | No
1164 | Yes | Unit turns off and on by itself | Unit turns off and on by itself | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the u pcba and repair uit “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: powers on by itself uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the uit powered on by itself. Recommended replacing the u pcba Part #00000000561528871, u pcba ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kelbert Ng IP & Issu: uit tuns off and on by itself even after replacing battery Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001advised Cu to eupect a call back /contacted oncal RSE Robert he will call Cu back Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-23 22:41 | No
1165 | Yes | Connection of the paddles with the monitor is damaged | Conneuon of the paddles with the monitor is damaged | BuQ&R Feedback: No Reason: Corrective Maintenance. Safety Qustions: Case: 00000000223255200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Gulherme Silva. Last udated by: Gulherme Silva. 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? In 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A, 00000000Was this an ou of bou failue? N/A, Trouleshooting Action: OPP-197693 Qute: 00000000112311 engenhariaclinica00ccamargo.org.br Problem Reported by Cutomer: Connection of the paddles with the monitor is damaged. AC Camargo Matriz - Antonio Pruente Romuldo N Juior Clinical Engineering - ICuand Hospitalizations 11 2189-3831 11 2189-35700000000 engenhariaclinica00ccamargo.org.br STOPPED ANTONIO PRuENTE FOuDATION 211 Ru PROF ANTONIO PRuENTE São Pauo- SP 00000000500000000-0000000000000000----------------------------------------------------------------------------- | 2023-09-24 00:31 | Yes
1166 | Yes | RR - The system startup not complete | RR - The system startup not complete | 000000000000000000000000001O##: 000000000000000000000000001000000000000000059100000000Jitlang, Pharurat The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: uuuuD2000000005 Eupected and ueupected Resuts: System startu not complete uer Impact: no Patient Impact: no SE8.2.3000000005 Resoluion : cold restart system ----------------------------------------------------------------------------- | 2023-09-24 10:46 | No
1167 | Yes | Patient data full | Patient data full | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002655 200000000ep-2000000003 Madison James The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Yes Eupected and ueupected Resuts: Patient database ful uer Impact: No Patient Impact: No SE 8.2 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Database consistency test failed. After deleting some old patient, the patient database still ful. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002575 Firmansyah The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Order image disk : Harddisk sparepart - 0000000080000000050000000000000000(2 pcs) Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Yes Eupected and ueupected Resuts: Patient database ful uer Impact: No Patient Impact: No SE 8.2 Diagnostic performed by Engineer : Confirmed reported problem, patient database intermittently ful, even only for aroud one week patient data. Ask uer to delete some patient data for temporary. Will try to replace image data harddisk as preventive action. Device is retuned to fuctional. ----------------------------------------------------------------------------- | 2023-09-24 13:16 | No
1168 | Yes | The device failed the self-test. | The device failed the self-test. | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: 1 Neut Step Action: Wanuiang Street Commuity Health Service Center, Lishu City, Zhejiang Province. The cutomer can only provide the uY REF, defibrillation monitor. No other information can be provided. The equpment keeps making abnormal noises. FSE Li Hu and ASP uiedong have been contacted for follow-u. Neut Step Action: Yes DFM100000000 defibrillator SN: CN326790000000000000000 please send 000000000000000000000000001k to Honglejia, thank youNeut Step Action: L2 engineer Fan Gan 15857180000000000000000 followed u, thank youSAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: abnormal soud SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022326859 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kaiuiang gu Last udated by: kaiuiang gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Internal Remarks: 26.00000000.2000000003 00000000:57:00000000 uC (860000000000000000100000000 $ASP engineer:[Honglejia Fangan] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem the device failed the self-test. $Q8: Check the equpment self-inspection process, need to re-checking. $Q9: Pass after re-checking. $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality. $Q11: $ Calibrated Tools name (ProSim8) S/N (26712200000000 Neut Calibration date (10000000018/2000000003), (ESA6200000000 S/N (30000000028600000000 Neut calibration date (3/2/20000000000000000 Neut Step Action: ASP honglejia Euternal Remarks: 26.00000000.2000000003 00000000:56:59 uC (860000000000000000100000000Department information: [Department: Emergency Room Contact Person: Teacher Mao Tel: 18200000000880000000099]* $ASP engineer:[Honglejia Fangan] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem the device failed the self-test. $Q8: Check the equpment self-inspection process, need to re-checking. $Q9: Pass after re-checking. $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality. $Q11: $ Calibrated Tools name (ProSim8) S/N (26712200000000 Neut Calibration date (10000000018/2000000003), (ESA6200000000 S/N (30000000028600000000 Neut calibration date (3/2/20000000000000000 ----------------------------------------------------------------------------- | 2023-09-24 14:01 | No
1169 | Yes | will not charge | will not charge | 000000000000000000000000001O##: 00000000000000000000000000100000000317000000002 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Tools & Equpment : PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed complaint uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damage the center pin in the port is no longer there. Additionally, the rear cover kick stand has been ripped from the rear cover on the left side. To correct both issus we will look to replace the Front Case Assembly for the charging port and the Rear cover for physical damage. Resoluion : Device was repaired by parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001937 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed complaint uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits charging port is physically damage the center pin in the port is no longer there. Additionally, the rear cover kick stand has been ripped from the rear cover on the left side. To correct both issus we will look to replace the Front Case Assembly for the charging port and the Rear cover for physical damage. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 0000000056000000005000000001-Rear Cover Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000055800000000ilbu, Phillip The system meets the specification for the performed service and is retuned to ue Resoluion : Sent to Bench Diagnostic performed by Engineer : Problem Description: "will not charge" Cutomer callback Tempu Pro: "will not charge" Form ued: Bench Trouleshooting Steps Taken: EMS site manager called to report Tempu Pro "will not charge" Cutomer says uit will not charge battery. Known good battery, an charger tried with like resuts. uable to reach cutomer on callback. ##address seems to be incorrect. Will try to reach cutomer again today. Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS site Manager Device ue:Not in Clinical ue Eupected and ueupected Resuts:"will not charge" uer Impact:No impact Patient Impact:No impact SEuuProblem Description/ueupected Resuts: "will not charge" Follow u Requred : EMS site manager called to report Tempu Pro "will not charge" Cutomer says uit will not charge battery. Known good battery, an charger tried with like resuts. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: cody IP & Issu: 8760000000071 will not charge Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A Internal Remarks : Requsted loaner shipped to Cut with FedEu Tracking# 71276121200000000. QuST SO# 20000000061 Case: 000000002200000000275 PN: 00000000-10000000000000000 SN: 6160000000000000000ailings: Both Cutomer SN: 615593 Tuerryton Fire Department 600000000 E Brillhart Ave, Perryton, Tu 790000000000000000Attn: Chief Pau Ducher 800000000-00000000-30000000000000000Internal Remarks : The requsted loaner shipped to Cutomer with FedEu Tracking# 71122815500000000. QuST SO# 20000000028 Case: 0000000022327500000000 PN: 00000000-10000000000000000 SN: 6000000009600000000Railings: Both Cutomer SN: 61000000008 Tuorth Bosqu EMS 200000000 E Morgan Meridian, Tu 76665-000000002 Attn: Cody Degges 250000000038-7000000005 ----------------------------------------------------------------------------- | 2023-09-25 01:41 | No
1170 | Yes | System would not boot up successfully(SPLIT) | System would not boot up successfully(SPLIT) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007653 00000000-Oct-2000000003 Mohamed Elknawy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uLLmessage "image storage is not possible." How often is issu occuring? : First occurence Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 1.2.5 Resoluion : Check the system and do the following - replace u- ray PC and 2 hard disks ( operating system and data disk ) - install the software for complete system - restore backu and EPu - system retun to clinical ue and tested by end uer Parts Consued : 1.00000000000000000000000000000000800000000660000000032,SSD Data Haswell 1.00000000000000000000000000000000800000000660000000093,u-ray PC Monoplane 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 18-Sep-2000000003 Mohamed Elknawy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : To replace ordered parts and install SW Diagnostic performed by Engineer : Check the system and foud that system not booting sucessfuly and give black screen with message uray system is switched off & image storage not possible becaue of image disk problem Supected is uray PC and hard disks Parts will be ordered ASAP Problem Description by Engineer : Problem description by engineer : uLLmessage "image storage is not possible." How often is issu occuring? : First occurence Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE 1.2.5 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000Amjad Aburqou The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : uLLmessage "image storage is not possible." How often is issu occuring? : First occurence Malfuction area : HDD Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No Patient Impact: No SE NA Follow u Requred : uLL# and/or description of uLL: uLLmessage "image storage is not possible." system is not connected to PRS need to plan onsite visit to reload the SW . Repair Action : uLLmessage "image storage is not possible." system is not connected to PRS need to plan onsite visit to reload the SW . Technician Name : Amjad Aburqou Internal Comments : uLL# and/or description of uLL: uLLmessage "image storage is not possible." system is not connected to PRS need to plan onsite visit to reload the SW . Repair Action : uLLmessage "image storage is not possible." system is not connected to PRS need to plan onsite visit to reload the SW . Diagnostic performed by Engineer : uLLmessage "image storage is not possible." system is not connected to PRS need to plan onsite visit to reload the SW . ----------------------------------------------------------------------------- | 2023-09-25 02:56 | No
1171 | Yes | 0S:SS:00 and it won't work. | 0DS:SS:00 and it won't work. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006858 25-Sep-2000000003 Tachinami, Makoto The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: "0000000000000000" is being ued and the system does not 000000000000000000000000001k. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: "0000000000000000" is displayed and the system does not 000000000000000000000000001k. uer Impact: Euamination canceled Patient Impact: Euamination canceled. SE uknow Resoluion: The compuer that controls the large-screen monitor did not start normally. Restarting the compuer several times does not improve the sitution, so I replaced the compuer and that improved the sitution. After the 000000000000000000000000001k, there was no problem with operation. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006837 Saka, Shuji The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: "0000000000000000" is being ued and the system does not 000000000000000000000000001k. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: "0000000000000000" is displayed and the system does not 000000000000000000000000001k. uer Impact: Euamination canceled Patient Impact: Euamination canceled. SE uknow Diagnostic performed by Engineer: RC Response: Restarting the system does not change the sitution. Only the frame is displayed on the large screen monitor. Visit and investigate. Internal Comments: [Suject] 00000000DS:SS:u Hiroi: "0000000000000000" is being ued and the system does not 000000000000000000000000001k. [Parts from the equpment] 0000000080000000012200000000:u Haswell Fleuvision 2 PC no HDD QTY:1 ZREP/Traceable Tokyo (u11) 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Manawa T know] Tachinami, Makoto - FSE1 Visit date and time 2000000003/00000000/25100000000300000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Michinami - 000000000000000000700000000000000007200000000 - Others Call mobile uon arrival Requster name Toshihito Saka [Euamination category] Other: Follow u Requred: ⋆ ----------------------------------------------------------------------------- | 2023-09-25 05:41 | No
1172 | Yes | self test not pass | self test not pass | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor/self-test failed Safety Qustions: Case: 00000000223276500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhouuiaouun Last udated by: Zhouuiaouun 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-25 06:51 | No
1173 | Yes | Booting up fail | Booting up fail | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008625 16-Oct-2000000003 jinhyeok choi The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: ouside of ue Eupected and ueupected Resuts: boot fail uer Impact: no Patient Impact: no SE8.2.30000000015 Resoluion : bootfail fleuvision pc problem fleuvision pc replace system fuction check system monitoring Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007126 25-Sep-2000000003 jinhyeok choi The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Booting u fail uer Impact: no Patient Impact: no SE R8 Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 Resoluion : bootfail fleuvision pc grabber fail fleuvisionpc replace fleuvision software reinstall system fuction check ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000071100000000H.I ha The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : u HASWELL FLEuVISION 2 PC NO HDD change Resoluion : u HASWELL FLEuVISION 2 PC NO HDD change Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Booting u fail uer Impact: no Patient Impact: no SE R8 ----------------------------------------------------------------------------- | 2023-09-25 07:41 | No
1174 | Yes | 0/u FD0Geometry movements partly available | 0R/u FD0Geometry movements partly available | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: u uuuFD200000000T2 Activities:Problem Description: Overview of cutomer problems: After the equpment is tuned on, all mechanical movements are disabled. The euamination bed and C-arm are not powered on and locked, and can be puhed by hand.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: uable to commuicate with GeoIPC. No movements are available (reboot done/needed).. uable to commuicate with GeoIPC Remote diagnostics and test resuts: The mechanical motion control host of the equpment is not connected to the system. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Check if the BIOS battery is normal 3V, Root Caue 1: Please ignore this alert, if it only occus sporadically. Treat it then as a 'one time only'. Service Action 1: Perform a reboot of the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other Soluion here below. Root Caue 1: GeoIPC is not booting u at all or fast enouh. As resut the Host PC is uable to commuicate with the Geo IPC within a specified time. Service Action 1: Verify the behavior of GEOIPC at first boot u of the system in the morning (at cold start-u): If inpu power (2300000000) is present bu the GEOIPC does not power u at all, it is caued du to a defective internal power suply of GEO IPC itself. Service Action 2: If GeoIPC does power u fine bu Host PC is uable to commuicate with it in time, it is caued du to a defective harddisk. (corruted OS/AS) Then replace the GeoIPC. Service Action 3: Check whether the cable pins between GeoIPC RC-u15 and IPC adapter SCC1-u22 are perhaps too short. Advise: tighten the RC-u15 and SCC1-u12 connectors or replace the cable between RC-u15 and SCC1-u12. Service Action 00000000 It coud be caued du to a defective control net000000000000000000000000001k switch (TS) in the R-cabinet or its Power Suply. Root Caue 1: When problem is persistent it coud be caued du to a defective control net000000000000000000000000001k switch (TS). Service Action 1: Check control net000000000000000000000000001k switch (TS) and its power suply (adapter). Replace when necessary. Root Caue 1: Bad cabling and/or connections. Service Action 1: Check the cable RC:u25 to SCC1:u22. Please note that top of these pins are at a distance of +/-1mm from the connector edge. If it is more than 1mm, replace this cable. Then the pins are too short to make proper contact. Service Action 2: Tighten all the connectors of GeoIPC and SCC1 uit. Root Caue 1: Issu with GeoIPC OS/AS software. It coud be du to a not correct shudown before, Windows OS can be corruted hence it needs more time in the neut start u. If this happens only sporadically, please ignore it. Else see other possible caues & soluions. Service Action 1: If SW is really corruted please reload GeoIPC software uing uB destroy stick. Root Caue 1: The INTERNAL power suply of the GEO IPC can be defective/malfuctioning. Service Action 1: Check if the GEO IPC itself does not power u (intermittently) when it gets mains suply. uully this behavior is noticed at COLD start u of the system. If so, replace GEO IPC. Notes on cutomer requrements: Onsite Safety Qustions: Case: 0000000022327856 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.1000000002000000003 13:13:00000000uC (8600000000600000000) RDI Accuacy:YES $Q7: (Confirm reported problem & uable to commuicate with GeoIPC $Q8: (GEO IPC did not start normally, measued the inpu voltage 200000000v, adjuted the transformer wiring, and commuicated with the hospital abou the voltage instability) $Q9: (adjuted the transformer wiring, and commuicated with the hospital abou the voltage instability) $Q100000000 (Test and inspection data refer to PA, Device retun to ful fuctionality) $Q11: Calibrated Tools name:N/A For Mainland China CV / uR, CT engineer Euternal Remarks: 00000000.1000000002000000003 13:13:00000000 uC (8600000000600000000) Faut description: Abnormal movement of the mechanical part Service content: GEO IPC does not start normally, measue the inpu voltage 200000000v, adjut the transformer wiring, and commuicate with the hospital abou voltage instability issus Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-25 09:46 | No
1175 | Yes | Boot up fail | Boot up fail | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007930000000000000000ct-2000000003 ByugKwan Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Boot u fail uer Impact: no Patient Impact: no SE R8 Resoluion : Replaced ippc Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007889 H.I ha The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : CLARITYIQ_II IP PC NO HDD change Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: not in ue Eupected and ueupected Resuts: Boot u fail uer Impact: no Patient Impact: no SE R8 Resoluion : / CLARITYIQ_II IP PC NO HDD change ----------------------------------------------------------------------------- | 2023-09-25 11:46 | No
1176 | Yes | QA000000Tube component suddenly fell | QA00000000: Tube component suddenly fell |  | 2023-09-25 12:18 | No
1177 | Yes | uequipment malfunctioning | uequipment malfunuoning | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007692 Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Resoluion : Problem solved by replacing the ucable from cutomer stock. Diagnostic performed by Engineer : uequpment malfuctioning Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A ----------------------------------------------------------------------------- | 2023-09-25 12:41 | No
1178 | Yes | QAA0/UNIQ Clarity FD0Tube defect | QAA0L/UNIQ Clarity FD0Tube defect | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 26.1000000002000000003 10000000057:31 uC (86000000008899) Service content: [Tue focu uLLreported, the tue filament was foud to be damaged, the tue was replaced and calibrated, and the equpment retuned to normal] Service resut: [After service, other devices operate normally] Other important matters: [None] SAP Transfer: Case Transferred to SAP Euternal Remarks: 26.1000000002000000003 1000000001:28 uC (86000000008285) Service content: [Tue focu uLLreported, the tue filament was foud to be damaged, the tue was replaced and calibrated, and the equpment retuned to normal] Service resut: [After service, other devices operate normally] Other important matters: [None] Neut Step Action: There is a limited nuber of labors, and the nuber of labors can be ued. No confirmation is needed to dispatch 000000000000000000000000001kers normally. Safety Qustions: Case: 0000000022327967 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Site1 FSE Feng Qingda 1868652000000005 Notified that it can be dispatched to him, and the dispatch information has been sent to the engineer SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: The equpment cannot be euposed and is reported as tue defect. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Frontal - Large filament reguation faut by PoInt Remote diagnostics and test resuts: Equpment tue failue PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: "Caue Frontal - Large filament reguation faut by PoInt Soluion 1 Possible Fue inside Point is blown. Replace the Point of the Frontal generator. Soluion 2 Carry ou the filament test in FSF: Diagnostics - Generator - Fuctional Test - Filament Test. Depending on the resuts the root caue can be determined: u-ray tue filament or Grid switch, HiVo or PoInt. Soluion 3 To verify whether the Large filament is defective please measue the filament indutance uing an LCR Meter at Frontal u-Ray Tue. When the filament is OK: Youwill measue abou 0000000000000000mH. When the filament is open: Youwill measue abou 7.5 mH. " Notes on cutomer requrements: Onsite Neut Step Action: Jiang Yonglu2000000003-00000000-2000000005:11:17 There is no more radiation duing patient treatment, and restarting the machine still doesn’t 000000000000000000000000001k. Problem Reported by Cutomer: QAA00000000L/uIQ Clarity FD200000000 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 26.1000000002000000003 1000000001:00000000uC (86000000008285)RDI Accuacy: Yes $Q7: (Confirmed reported problem.large focu defect small focu ued.) $Q8: (monitor warning: tue defect and ue LC meter measue the cathode/anode of the tue, indutance valu 7.29 and 7.3 of the both focu sugest a replace of the tue. ) $Q9: (Replace the tue.uu989000000000000000085100000000. and adjut system 000000000000000000000000001k normally.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: $Calibrated tools refer to PA. For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: scan second_N/A Internal Remarks: 26.1000000002000000003 10000000057:33 uC (86000000008899) RDI Accuacy: Yes $Q7: (Confirmed reported problem.large focu defect small focu ued.) $Q8: (monitor warning: tue defect and ue LC meter measue the cathode/anode of the tue, indutance valu 7.29 and 7.3 of the both focu sugest a replace of the tue. ) $Q9: (Replace the tue.uu989000000000000000085100000000. and adjut system 000000000000000000000000001k normally.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: $Calibrated tools refer to PA. For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: scan second_N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-25 12:46 | No
1179 | Yes | The device is not going through the test. | The device is not going through the test. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000086200000000uak Tezel The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Not ued. Eupected and ueupected Resuts: The Device Does Not Pass the Test. uer Impact: No. Patient Impact: No. SE 2.0000000021 Resoluion: Duing the checks, it was determined that the device did not pass the operational control test and the Processor PCA, Therapy PCA and Printer parts were defective. The defective parts mut be replaced with new ones. According to the faut detection, 0000000056000000009000000001,0000000056000000009531,0000000056000000009000000001,0000000056000000009211 - SHOuD BE OFFERED TO PAMuKALE uIVERSITY HOSPITAL. Follow u Requred: Duing the checks, it was determined that the device did not pass the operational control test and the Processor PCA, Therapy PCA and Printer parts were defective. The defective parts mut be replaced with new ones. According to the faut detection, 0000000056000000009000000001,0000000056000000009531,0000000056000000009000000001,0000000056000000009211 - SHOuD BE OFFERED TO PAMuKALE uIVERSITY HOSPITAL. ----------------------------------------------------------------------------- | 2023-09-25 13:36 | No
1180 | Yes | other uLL with the tube that does not make u-rays - Split | other uLL with the tube that does not make u-rays - Split |  | 2023-09-25 14:15 | No
1181 | Yes | System completely down | System completely down | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000910000000026-Sep-2000000003 Simon Fallows The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Resoluion : Sute PC did not fiu the issu, replaced old sute PC back. Foud that hard disk drive 1 had thermal tripped on the Viewing PC. Reseated this trip on drive caddy and had to reloaded software. Reconfigued system as some software corrution had occured. Tested u-ray and left 000000000000000000000000001king order. Problem Description by Engineer : System not booting u. Remote ordered a new sute PC. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008877 Andrew Jones The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Engineer requred to site to investigate why system will not complete booting u after power cycle. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Manager Device ue: Ouside of ue, morning power u Eupected and ueupected Resuts: System will not complete system boot u uer Impact: System OOA Patient Impact: NA SE ? Diagnostic performed by Engineer : Cutomer has attempted to cycle the systems power back to the mains isolator bu uable to complete system boot u. Cutomer had message displayed on data screen abou installing something bu is usue and will try to take images and send to me to attach to file. Engineer requred to site to investigate. ----------------------------------------------------------------------------- Cutomer comment : Spoke with Christina on site and booked FSE for tomorrow from 9am ----------------------------------------------------------------------------- | 2023-09-25 14:41 | No
1182 | Yes | It does not turn on. | It does not turn on. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009252 26-Sep-2000000003 ARCANGELI, LEANDRO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: It does not tun on, it does not load the programs. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: It does not tun on, it does not load the programs. uer Impact: Not provided Patient Impact: Not provided SE 7.2.16.00000000Diagnostic performed by Engineer: OS Disk is fauty, the operating system cannot be performed. Resoluion: Carried ou replacement of SYSTEM HDD, carried ou software loading with relative fuctional tests, the system is left 000000000000000000000000001king. Parts Consued: 1.00000000000000000000000000000000800000000163853,PFS-00000000 CFG2 HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000DAVIDE VECCHI The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: It does not tun on, it does not load the programs. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: It does not tun on, it does not load the programs. uer Impact: Not provided Patient Impact: Not provided SE 7.2.16.00000000Diagnostic performed by Engineer: 3385610000000022 Roberto Bizzoni Since Friday September 22 in the afternoon they have had this problem, tried several complete restarts (with ru/stop), the system is not in line for the logs, the most recent log is from September 19. They have for some time images on the monitor then all screens are black, they cannot tell me the statu of the arc or the table. Internal Comments: uLL# and/or description of uLL It does not tun on, it does not load the programs. Repair Action: Hw Follow u Requred: uLL# and/or description of uLL It does not tun on, it does not load the programs. Repair Action: Hw What are the neut Trouleshooting steps: Intervention is needed on site for diagnosis and repair, I am incluing the PC single board, the hard disk and I sugest also bringing a CR2000000002 with you 1 FSE 8 hous The system is down. Technician Name: DAVIDE VECCHI ----------------------------------------------------------------------------- | 2023-09-25 15:01 | No
1183 | Yes | Speaker Malf. INOP - Confirmed sound | Speaker Malf. INOP - Confirmed sound | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009281 29-Sep-2000000003 HuB LANGE The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed uer Device ue: ouside of ue Eupected and ueupected Resuts: no soud, also not on startu uer Impact: no soud on Avalon Patient Impact: N/A SE uknown Tools & Equpment : TEST SIMuATOR KIT, 2000000003-11-30000000000000000:00000000:00000000 TEST SIMuATOR KIT, 2000000003-11-30000000000000000:00000000:00000000 TEST SIMuATOR KIT, 2000000003-11-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000006-1000000000000000000000000:00000000 Parts Consued: 1.0000000000000000000000000000000056000000008621,AV FM200000000500000000Main Board Vers. 2 Resoluion: Repair FM300000000New Mainboard, Speaker defective message. FCO ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000083300000000ATRICK POELS The system meets the specifications for the performed service and is retuned to ue Follow u Requred: Please schedue an available FSE (Hub Lange) for the neut repair: 0000000056000000008621 - AV FM200000000500000000Main Board Vers. 2 Problem Description by Engineer: Cutomer Fuction/Role: Biomed uer Device ue: ouside of ue Eupected and ueupected Resuts: no soud, also not on startu uer Impact: no soud on Avalon Patient Impact: N/A SE uknown Resoluion: Main Board needs to be replaced ----------------------------------------------------------------------------- | 2023-09-25 15:01 | No
1184 | Yes | The wireless footswitch is not responsive | The wireless footswitch is not responsive | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000Barry White The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Wireless pedal not 000000000000000000000000001king uer Impact: NA Patient Impact: NA SE1 Diagnostic performed by Engineer : Wired pedal curently in ue. Advised there is an open job ( case ref: 00000000212000000008500000000 for the FCO ugrade on the wireless pedal to be performed. ----------------------------------------------------------------------------- | 2023-09-25 15:41 | No
1185 | Yes | CM - SELF TEST FAILED (Device is EoL and replacement of part is not possible) | CM - SELF TEST FAILED (Device is EoL and replacement of part is not possible) | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : CM - SELF TEST FAILED+ CHILD KEY SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 0000000022283311 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. 發現問題時，該裝置是否正用於臨床用途？ 否 2. 是否有使用者或病患受傷？ 否 3. 如果裝置具備警報/警告功能，發現問題時，該裝置是否發揮應有的警報/警告功能？ 不適用 00000000是否為到貨不良？ 不適用 ----------------------------------------------------------------------------- | 2023-09-25 15:51 | No
1186 | Yes | alarms have no volume | alarms have no volume | Internal Remarks : Interaction Type: New Case Caller: Kira Walker IP & Issu: " 730000000080000000000000000 alarms have no volue" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created remote technical 000000000000000000000000001 roued call to quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-25 16:46 | No
1187 | Yes | It does not give the writing. | It does not give the writing. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008300000000DAVIDE VECCHI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: When doing the writing it says ful disk space. ugent. Injected patient. How often is issu occuring? : Continuuly Malfuctioning area: Image processing Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: When doing the writing it says ful disk space. ugent. Injected patient. uer Impact: yes Patient Impact: no SE 8.2.3000000005 Diagnostic performed by Engineer: contacted the department, even thouh they deleted all the patients the system shows ful disk uLLInternal Comments: uLL# and/or description of uLL When doing the writing it says ful disk space. ugent. Injected patient. Repair Action: Sw Resoluion: performed remote initialization of the image disk that was sucessfu restoring the fuctioning of the system that is back to clinical ue. Recontacted at 1:00000000 pm the system 000000000000000000000000001ks withou problems. I close the call. ----------------------------------------------------------------------------- Internal Remarks: Opened critical need becaue machine is down with injected patient. Contacted technician Davide Vecchi who took charge of the call ----------------------------------------------------------------------------- | 2023-09-25 16:51 | No
1188 | Yes | The foot pedal keeps losing contact | The foot pedal keeps losing contact | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009936 29 September 2000000003 DANIELSSON JESPER The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Doctor Device ue: in ue Eupected and ueupected Resuts: Wirelessfootswitch was installed on Friday, bu it doesn't seem to 000000000000000000000000001k as it shoud, keeps loosing connection and even if cutomer tries to restart it. uer Impact: Can't ue WFS Patient Impact: NA SE2.2.6 Resoluion: Replaced the wireless footswitch and base station. The footswitch is paired with the base station. Fuctionality tested thorouhly. Parts Consued: 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009516 JALAR JIM The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow-u 000000000000000000000000001k schedued Follow-u Requred: Wireless footswitch was installed on Friday, bu it doesn't seem to 000000000000000000000000001k as it shoud, keeps loosing connection and even if cutomer tries to restart it. It may be that either the WFS or the reciver in the table is bad and cauing this. I sugest replacement of both to make it fuly fuctional. Order parts for the cutomer and dispatch a FSE for replacement. Problem Description by Engineer: Cutomer Fuction/Role: Doctor Device ue: in ue Eupected and ueupected Resuts: Wirelessfootswitch was installed on Friday, bu it doesn't seem to 000000000000000000000000001k as it shoud, keeps loosing connection and even if cutomer tries to restart it. uer Impact: Can't ue WFS Patient Impact: NA SE2.2.6 Diagnostic performed by Engineer: Wireless footswitch was installed on Friday, bu it doesn't seem to 000000000000000000000000001k as it shoud, keeps loosing connection and even if cutomer tries to restart it. It may be that either the WFS or the reciver in the table is bad and cauing this. I sugest replacement of both to make it fuly fuctional. Order parts for the cutomer and dispatch a FSE for replacement. ----------------------------------------------------------------------------- | 2023-09-25 16:56 | No
1189 | Yes | does not detect air flow preventing its u. | does not deteuair flow preventing its u. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000083 29-Sep-2000000003 Víctor Rodríguz González The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Resoluion: Fan is checked and it is observed that the behavior is du to the alarm limits selected so the equpment restricts ventilation for safety reasons. The defaut ventilation valus are reset and proper operation is confirmed. The equpment is in correct condition of ue. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Respirator Eupected and ueupected Resuts: Operation uLL uer Impact: A different equpment is ued. Patient Impact: Impossible to ue the equpment. SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009829 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Respirator Eupected and ueupected Resuts: Operation uLL uer Impact: A different equpment is ued. Patient Impact: Impossible to ue the equpment. SE Diagnostic performed by Engineer: The client (Ana Roco Martin. eut:870000000062) indicates that the equpment does not detect air flow preventing its ue. There is no availability from the client to perform verifications, so the case is transferred onsite so that the equpment is checked by the field service. Resoluion: The case is transferred onsite. Follow u Requred: The client (Ana Roco Martin. eut:870000000062) indicates that the equpment does not detect air flow preventing its ue. There is no availability from the client to perform verifications. - Equpment located in THE INFANT MATERNAL HOSPITAL/2ND FLOOR/AREA D- PEDIATRIC OPERATING ROOMS - PEDIATRIC ICu - Assign fse - Check connections and equpment integrity. In case of needing to replace material, order: 989800000000617100000000and/or 00000000561512191 ----------------------------------------------------------------------------- | 2023-09-25 17:06 | No
1190 | Yes | Doesn't start up anymore | Doesn't start up anymore | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 Problem Description by Engineer: Device Doesn't start u Diagnostic performed by Engineer: Flash card holder board was not deep enouh in connector Resoluion: Flash card placed in place ----------------------------------------------------------------------------- | 2023-09-25 17:06 | No
1191 | Yes | Frozen screen | Frozen screen | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007800000000Koc, Migul The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: Frozen screen When was the first occurence? Within... : 18-Sep-2000000003 12:0000000000000000 How often is issu occuring? : uknown Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Frozen screen uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 8.2.17.00000000Diagnostic performed by Engineer: 28.00000000, FSE contacted the client, requsted to check it on site taking advantage of his visit, checking the uLLlog file, shows arc in tue and hd image and/or IPPC failue, field service engineer said that he will inform u tomorrow 29.00000000. 29.00000000, the client informed him that the equpment is in ue, he coud not do tue conditioning or recreate the image disk or evalute HDs and/or IPPC, we recommended the client to measue voltages, groud, impedance and adjut terminals of the electrical board (TE), also evalute Air Conditioning (AA). It is necessary for the field service to check te and AA, if it is OK, clean, silicone, adjut the pins of the high-energy terminals and measue continuty of the high-energy cables, replace cables if necessary, if it is ok, do conditioning and adaptation of the tue to allow the history of the same, otherwise evalute its replacement; Also ru HDs, IPPC and Host PC tests, check Post Resut and clean and reconnect cables and cards internally if necessary, replace PCs battery if necessary, ru consistency test of Database (BD) and repair BD if necessary, otherwise recreate image disk or scan and repair disks, if the failue persists, reinstall software or evalute replacement of HDs, IPPC and/or Host PC. Internal Comments: uLL# and/or description of uLL Frozen screen Repair Action: 28.00000000, FSE contacted the client, requsted to check it on site taking advantage of his visit, checking the uLLlog file, shows arc in tue and hd image and/or IPPC failue, FSE said that he will inform u tomorrow 29.00000000. 29.00000000, the client informed him that the equpment is in ue, he coud not do tue conditioning or recreate the image disk or evalute HDs and/or IPPC, we recommended the client to measue voltages, groud, impedance and adjut terminals of the electrical board (TE), also evalute Air Conditioning (AA). It is necessary for the field service to check te and AA, if it is OK, clean, silicone, adjut the pins of the high-energy terminals and measue continuty of the high-energy cables, replace cables if necessary, if it is ok, do conditioning and adaptation of the tue to allow the history of the same, otherwise evalute its replacement; Also ru HDs, IPPC and Host PC tests, check Post Resut and clean and reconnect cables and cards internally if necessary, replace PCs battery if necessary, ru consistency test of Database (BD) and repair BD if necessary, otherwise recreate image disk or scan and repair disks, if the failue persists, reinstall software or evalute replacement of HDs, IPPC and/or Host PC. Follow u Requred: uLL# and/or description of uLL Frozen screen Repair Action: 28.00000000, FSE contacted the client, requsted to check it on site taking advantage of his visit, checking the uLLlog file, shows arc in tue and hd image and/or IPPC failue, FSE said that he will inform u tomorrow 29.00000000. 29.00000000, the client informed him that the equpment is in ue, he coud not do tue conditioning or recreate the image disk or evalute HDs and/or IPPC, we recommended the client to measue voltages, groud, impedance and adjut terminals of the electrical board (TE), also evalute Air Conditioning (AA). It is necessary for the field service to check te and AA, if it is OK, clean, silicone, adjut the pins of the high-energy terminals and measue continuty of the high-energy cables, replace cables if necessary, if it is ok, do conditioning and adaptation of the tue to allow the history of the same, otherwise evalute its replacement; Also ru HDs, IPPC and Host PC tests, check Post Resut and clean and reconnect cables and cards internally if necessary, replace PCs battery if necessary, ru consistency test of Database (BD) and repair BD if necessary, otherwise recreate image disk or scan and repair disks, if the failue persists, reinstall software or evalute replacement of HDs, IPPC and/or Host PC. What are the neut Trouleshooting steps: It is necessary for the field service to check te and AA, if it is OK, clean, silicone, adjut the pins of the high-energy terminals and measue continuty of the high-energy cables, replace cables if necessary, if it is ok, do conditioning and adaptation of the tue to allow the history of the same, otherwise evalute its replacement; Also ru HDs, IPPC and Host PC tests, check Post Resut and clean and reconnect cables and cards internally if necessary, replace PCs battery if necessary, ru consistency test of Database (BD) and repair BD if necessary, otherwise recreate image disk or scan and repair disks, if the failue persists, reinstall software or evalute replacement of HDs, IPPC and/or Host PC. Technician Name: Koc, Migul ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000116100000000 Josimar, Velasquz Problem Description by Engineer: Problem description by engineer: Frozen screen When was the first occurence? Within... : 18-Sep-2000000003 12:0000000000000000 How often is issu occuring? : uknown Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Frozen screen uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 8.2.17.00000000Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2023-09-25 17:11 | No
1192 | Yes | 00he device is not going through the test. | 000The device is not going through the test. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000018 Buak Tezel The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: The inspection showed that the device did not pass the operation control test. Processor PCA and Cable part of the device are defective. In addition, the battery of the device has completed its service life and is defective. The battery shoud be suplied throuh the sales department. According to the failue detection, 000000005600000000311,0000000056000000009000000001 - YOZGAT CITY HOSPITAL SHOuD BE RECOMMENDED. (ali.u00onesans.com) Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Not ue. Eupected and ueupected Resuts: The device does not pass the test. uer Impact: No. Patient Impact: No. SE 2.00000000.21 Resoluion: The inspections did not pass the operation control test of the device. Processor PCA and Cable part of the device are defective. In addition, the battery of the device has completed its service life and is defective. The battery shoud be suplied throuh the sales department. According to the failue detection, 000000005600000000311,0000000056000000009000000001 - YOZGAT CITY HOSPITAL SHOuD BE RECOMMENDED. (ali.u00onesans.com) ----------------------------------------------------------------------------- | 2023-09-25 17:21 | No
1193 | Yes | Disk full, no space | Disk full, no space |  | 2023-09-25 17:26 | No
1194 | Yes | QAA0u FD0start failure | QAA0/u FD0start failure | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002228500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: QAA00000000/u uuuFD200000000 Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-09-25 17:36 | No
1195 | Yes | Device Doesn't start up anymore | Device Doesn't start up anymore | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000 Diagnostic performed by Engineer: Flashcard holder not deep enouh in connector Problem Description by Engineer: Device doesn't start u Resoluion: Flashcard holder placed in correct position ----------------------------------------------------------------------------- | 2023-09-25 17:41 | No
1196 | Yes | Fluoroscopy failure | Fluoroscopy failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: Fluroscopy failed last night Problem Reported by Cutomer: QAA00000000R/u uuuFD200000000Neut Step Action: Ding Shidu 1350000000038800000000 T2 Activities:Problem Description: Overview of cutomer problems: Fluroscopy failue Trouleshooting Action: Key faut codes or phenomena: A fluroscopy failue occured duing the night PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Go to the scene to check the event log and check whether the small focu of the tue is normal. Recommended spare parts uuand designation: " 989000000000000000085100000000, 00000000800000000651161" T2 Activities:Problem Description: Overview of cutomer problems: Fluroscopy failue Trouleshooting Action: Key faut codes or phenomena: A fluroscopy failue occured duing the night PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Go to the scene to check the event log and check whether the small focu of the tue is normal. Recommended spare parts uuand designation: 989,0000000000000000000000005,100000000,00000000,000000000000000000000000000000000000000000000000Safety Qustions: Case: 00000000222880000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 16.00000000.2000000003 00000000:00000000:57 uC (8600000000600000000) Service content (requred): Occasional fluroscopy fails, add high-pressue refrigeration oil and detector refrigerant. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None Internal Remarks: 16.00000000.2000000003 00000000:00000000:00000000 uC (8600000000600000000) "RDI Accuacy: Yes" $Q7:Confirmed reported problem.Intermittent fluroscopy failue. $Q8: On site Reviewed the uLLlog, Performed visul inspection, foud Refrigeration uit anomaly. $Q9: Add cooling oil and detector cooling flud. $Q100000000 Test and Inspection data refer to:PA.Device retuned to ful fuctionality. $Q11: Calibrated Tools refer to:PA ^Curent Tue installed date: 6/ 3/2000000002 ^Curent Tue uage: 251882.5luSAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-25 18:36 | No
1197 | Yes | 00T0Unit will not start | 00CDT0Unit will not start | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002829 27-Sep-2000000003 Sinkewicz Gary The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System will not boot u at all Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: eupect system to boot u, will not boot u at all uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion : Replaced Host PC in M-Cab. udated License Parts Consued : 1.00000000000000000000000000000000800000000200000000611,Nehalem Host PC Monoplane Rev_III NO_HDD Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000088 25-Sep-2000000003 Sinkewicz Gary The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace Host PC and HDD as Needed Problem Description by Engineer : Problem description by engineer : System will not boot u at all Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: eupect system to boot u, will not boot u at all uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Host PC will not Boot u. No power to HDD. No green LED's in Back. Ordered Host PC Nehalem& HDD Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000086 Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System will not boot u at all Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: eupect system to boot u, will not boot u at all uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Cutomer reports system will not boot u. Shu down, then watched boot cycle. No Bios seen at all. System down Internal Comments : uLL# and/or description of uLL: Will not boot u. likely HOST issu Repair Action : Verify issu. Likely HOST PC since no Bios seen or power to HOST PC. Repair as needed Possible parts: 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: Will not boot u. likely HOST issu Repair Action : Verify issu. Likely HOST PC since no Bios seen or power to HOST PC. Repair as needed Possible parts: 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Mcnamara Pau ----------------------------------------------------------------------------- Internal Remarks : Generated new license for FSE on Host replacement Internal Remarks : FSE needed part nuber for Host PC and HD. ----------------------------------------------------------------------------- | 2023-09-25 18:41 | No
1198 | Yes | Wireless footswitch dosen't work. Its green and WIFI symbol is visible but dosen't work. | Wireless footswitch dosen't work. Its green and WIFI symbol is visible but dosen't work. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000067 18-Oct-2000000003 Berrig Clau The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station Parts Consued: 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station Problem Description by Engineer: Cutomer Fuction/Role: MT Device ue: usuk Eupected and ueupected Resuts: Wireless foot switch is not 000000000000000000000000001king. Its green and WIFi symbol is visible bu doesn't 000000000000000000000000001k. uer Impact: WFS issus Patient Impact: NA SE2.2.u Resoluion: Base and foot switch replaced. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000086 JALAR JIM The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: Plan a visit rom FSE to replace the wireless footswitch and also the base station in table to get it to start 000000000000000000000000001king again. Parts added to follow u. Problem Description by Engineer: Cutomer Fuction/Role: MT Device ue: usuk Eupected and ueupected Resuts: Wireless foot switch is not 000000000000000000000000001king. Its green and WIFi symbol is visible bu doesn't 000000000000000000000000001k. uer Impact: WFS issus Patient Impact: NA SE2.2.u Follow u Requred: Plan a visit rom FSE to replace the wireless foot switch and also the base station in table to get it to start 000000000000000000000000001king again. Parts added to follow u. ----------------------------------------------------------------------------- | 2023-09-25 19:01 | No
1199 | Yes | PSA&PSE_QACASE for PCI imported product#00000the stent and guide wire images could not be seen under | PSA&PSE_QACASE for P0imported produu000000_the stent and guide wire images could not be seen under |  | 2023-09-25 19:04 | Yes
1200 | Yes | DFM0 failed the self-test | DFM0 failed the self-test | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Northwest Oilfield Emergency Center Cannot be ued after powering on, self-tests all pass Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 0000000022288583 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: jindi yang Last udated by: jindi yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-25 19:21 | No
1201 | Yes | Replace WFS TPR approver Patrick Helfer | Replace WFS TPR approver Patrick Helfer | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006000000005 00000000Oct-2000000003 Boncqut Sterling The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Parts Consued : 1.00000000000000000000000000000000800000000238212,Wireless Footswitch 000000002 Diagnostic performed by Engineer : Replaced WFS and tested it 000000000000000000000000001ked great. However after I rebooted the system wasn’t acting right. It was internally disconnecting every 300000000secs or so. I coudn’t log into service. I went to try and load a ghosts however no previou ghost was available. GE holds the contract and said they will call Avante in. The system was not 000000000000000000000000001king as it shoud when I left. ----------------------------------------------------------------------------- Internal Remarks : Sterling Boncqut - software loses connection to TSM/TSO's every 200000000seconds after wireless FSW replacement. FSE will restore ghost image on new HDD. ----------------------------------------------------------------------------- | 2023-09-25 19:31 | No
1202 | Yes | Post PC does not boot - System uLLinitializing connect to host but will not connect | Post PC does not boot - System uLLinitializing conneuto host but will not connect | Internal Remarks : Interaction Type: Cutomer Call Back Caller: marqu torres IP & Issu: ccb to check statu of qute Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created aa for qa to send qute asap as Eupedite service for 9/15 is being requsted Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: Cutomer Call back Caller: marqu torres 7000000003000000005200000000marqu.torres00rimedu.com IP & Issu: Post PC does not boot - System uLLinitializing connect to host bu will not connect Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: CREATED QuTE for afterhous cutomer is requsting service for 9/17/23 Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Planner Initials: CA Action(s) Taken: Called FSE Bryan Wallace again for scheduing Oucome: FSE Bryan Wallace stated he was going to speak to DSM Brent Chesley and then let me know how to proceed since qutes still hasn't sent eupedited qute to cutomer Neut Step Action(s): Waiting on CB from FSE Bryan Wallace for disposition of case. Last udated (Date & Time): 9/15 00:200000000m EST Internal Remarks : Interaction Type: Cutomer Call Back Caller: Marqu Torres IP & Issu: " 70000000053993 cutomer called to get statu udate on the case" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created AA for QA team to send qute Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Planner Initials: CA Action(s) Taken: Team messaged DSM Brent Chesley du to DSM Kevin Anderson being on PTO for FSE assist du to no available FSE’s Oucome: DSM Brent Chesley stated to still contact all FSE’s to see if they can move things aroud/Called and left msg for FSE Bryan Wallace to accept case Neut Step Action(s): Call FSE Bryan Wallace again for scheduing 002:25pm Last udated (Date & Time): 9/15 00:15pm EST Internal Remarks : **CuT REQuSTED ONSITE SuPORT Interaction Type: New Case Caller: marqu torres 7000000003000000005200000000marqu.torres00rimedu.com IP & Issu: Post PC does not boot - System uLLinitializing connect to host bu will not connect Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND QuTE Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : 00000000222860000000000000000 Brian Wallace FSE 000000000000000000000000001king on u 8.2 where system is not booting u. FSE says there is no hard drive activity on the host PC. FSE say he does not see BIOS either. Asked FSE if there was the LED at the bottom of the hard drive. FSE says yes. Advised FSE to stick in the ghost CD and see if it gives him the option to load the CD on the screen. FSE has a plan bu may call back. Possible bad host pc. ----------------------------------------------------------------------------- | 2023-09-25 19:36 | No
1203 | Yes | SPLIT-uDisable-See Reference case PR 00 | SPLIT-uDisable-See Reference case PR 000 |  | 2023-09-25 19:52 | No
1204 | Yes | SPLIT- Etco0isabled-See Reference case PR 00 | SPLIT- Etco0disabled-See Reference case PR 000 |  | 2023-09-25 19:54 | No
1205 | Yes | DFM0 machine is not switching on | DFM0 machine is not switching on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000066 Madhan, kuar The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : To replace 0000000056000000008951, DFM100000000 AC Power Modue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000008951, DFM100000000 AC Power Modue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000900000000Parts Consued : 1.0000000000000000000000000000000056000000008951,DFM100000000 AC Power Modue ----------------------------------------------------------------------------- | 2023-09-25 19:56 | No
1206 | Yes | Security_device frequently alarms Pacemaker patients cannot detect the uEvent alarm | Security_device frequently alarms Pacemaker patients cannot deteuthe uEvent alarm |  | 2023-09-25 20:24 | No
1207 | Yes | EC/00heck Vent: Primary Alarm Failed | EC/00-Check Vent: Primary Alarm Failed |  | 2023-09-25 20:26 | No
1208 | Yes | Screen blank on reboot, Fatal uLLC00 | Screen blank on reboot, Fatal uLLC00 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003798 26-Sep-2000000003 Holder Antwan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System will not boot u Loop on Bios Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replaced the Host PC Hard Drive. Reloaded System Backu and tested all Geometry and u-ray fuctions of the system. System is u and runing and fuly operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 25-Sep-2000000003 Holder Antwan The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Need Follow u. Problem Description by Engineer : Problem description by engineer : System will not boot u Loop on Bios Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Host PC keeps rebooting into a Blu Screen. Will order new Host PC and HD for tomorrow 100000000m delivery. I will be back onsite at that time to replace the parts. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001300000000Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: System will not boot Repair Action : Verify issu. Since Bios comes u, likley hard drive issu or PC itself. 000000000000000000000000001ld try new hard drive, restore ghost, then see if 000000000000000000000000001ks. If not, rpelace Host possibly Possible parts: 0000000080000000050000000000000000- Hard disk spare part 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD Follow u Requred : uLL# and/or description of uLL: System will not boot Repair Action : Verify issu. Since Bios comes u, likley hard drive issu or PC itself. 000000000000000000000000001ld try new hard drive, restore ghost, then see if 000000000000000000000000001ks. If not, rpelace Host possibly Possible parts: 0000000080000000050000000000000000- Hard disk spare part 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD Technician Name : Mcnamara Pau Problem Description by Engineer : Problem description by engineer : System will not boot u Loop on Bios Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Cutomer stated the system will not boot u. Shu down, and rebooted again, System starts with Bios screen, then reboots and starts over. Loops at bios and application does not start. ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022332679 FSE: Holder Antwan R8.2; Duing boot u keeps getting blu screen of death. Host PC keeps rebooting. Instruted FSE to place part order for 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD and 0000000080000000050000000000000000- Hard disk spare part. ----------------------------------------------------------------------------- | 2023-09-25 20:26 | No
1209 | Yes | EC 0 - battery failure | EC 00 - battery failure | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000100000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLL11200000000attery failed. Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLL11200000000attery failed. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/25/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Diagnostic code 11200000000CBIT) Battery failed - When the PM PCBA detects a battery uLL the uLLcontinus to ventilate and the battery charger tuns off. 1. Verify that the PM PCBA is equpped with PIC firmware version 1.300000000 2. Replace the internal battery. 3. Replace the PM PCBA. Biomed has advised the V600000000battery is a PartsSouce battery. Advised Biomed the approved battery for the V600000000mut be puchased from u. Biomed has advised they are contacting PartsSouce to puchase the battery from u. Part # 989800000000626900000000Assy Battery Li-Ion,11ah,V600000000List price is $771.00000000 u Medical Suplies 1-800000000-225-00000000300000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jutin 3166852221 u53700000000 IP & Issu: 1000000003511700000000receiving uLLmessage "11200000000Bit battery failue" and battery has been replaced bu still receiving uLLmessage Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: created remote 0000000000000000000000000010000000000000000100000000 dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-25 20:26 | No
1210 | Yes | CV*System down* System freezes at 0 | CV*System down* System freezes at 0 | 000000000000000000000000001O##: 00000000000000000000000000100000000120000000078 100000000Oct-2000000003 Islam, Abduaziz The system meets the specification for the service performed and is retuned to ue Problem description by engineer: Problem description by engineer: System freezes at 00000000 RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 25-Sep-2000000003 1000000006:00000000 RMT: Host PC POST failed - memory 3 25-Sep-2000000003 1000000003:22 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: no impact Patient Impact: no impact SE 8.2.17.00000000Parts Consued: 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD Diagnostic performed by Engineer: the part replaced sucessfuly and ghost image is created ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000160000000000000000-Oct-2000000003 Islam, Abduaziz The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: to replace IPPC Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Diagnostic performed by Engineer: the Host and HDD's replaced sucessfuly bu the IPPC was DEFOA so will order a new one Problem description by engineer: Problem description by engineer: System freezes at 00000000 RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 25-Sep-2000000003 1000000006:00000000 RMT: Host PC POST failed - memory 3 25-Sep-2000000003 1000000003:22 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: no impact Patient Impact: no impact SE 8.2.17.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009800000000Khaled Megud Fawzy The system does not meet the specification for the performed service and is not retuned to ue Problem description by engineer: Problem description by engineer: System freezes at 00000000 RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 25-Sep-2000000003 1000000006:00000000 RMT: Host PC POST failed - memory 3 25-Sep-2000000003 1000000003:22 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: no impact Patient Impact: no impact SE 8.2.17.00000000Diagnostic performed by Engineer: RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 25-Sep-2000000003 1000000006:00000000 RMT: Host PC POST failed - memory 3 25-Sep-2000000003 1000000003:22 Follow u Requred: uLL# and/or description of uLL Root Caue for RMT: Host PC POST failed - memory 3: 1. Host PC has a memory failed message in the Power On Self Test (POST). Since the Host PC has only 3 memory banks, one failing memory bank can affect the fuctionality of the system. Repair Action: Service actions for RMT: Host PC POST failed - memory 3: 1.1. Replace Host PC (TH) Technician Name: Khaled Megud Fawzy Resoluion: Onsite visit needed to check and replace Host PC and IIPPC, HDs Internal Comments: uLL# and/or description of uLL Root Caue for RMT: Host PC POST failed - memory 3: 1. Host PC has a memory failed message in the Power On Self Test (POST). Since the Host PC has only 3 memory banks, one failing memory bank can affect the fuctionality of the system. Repair Action: Service actions for RMT: Host PC POST failed - memory 3: 1.1. Replace Host PC (TH) ----------------------------------------------------------------------------- | 2023-09-25 20:41 | No
1211 | Yes | M0A - u00 The device cannot be ud normally during u | u The device cannot be ud normally during u |  | 2023-09-25 20:45 | No
1212 | Yes | M0A - u00 Battery Issue | u Battery Issue | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 Problem Description by Engineer : uuu- For overdu Service/Repair Resoluion : uon inspection one of the batteries foud to be fauty New battery fitted , NBP pup calibrated. EST attached uit tested and serviced to manuactuer's specifications. Parts Consued : 1.000000000000000000000000M3538A,Bateria ION LITIO Para DESF uuu 1.000000000000000000000000M3713A,HEARTSTART ADuT PLu PADS ----------------------------------------------------------------------------- | 2023-09-25 20:46 | No
1213 | Yes | Check Vent: Backup Alarm Failed" (Diagnostic code 00 | Check Vent: Backup Alarm Failed" (Diagnostic code 00) |  | 2023-09-25 20:54 | No
1214 | Yes | 00 CN00- Battery Unable to Charge - Battery Replaced | u- Battery Unable to Charge - Battery Replaced | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 27.00000000.2000000003 00000000:00000000:21 uC (860000000000000000100000000Department information: [Department: Dealer Contact Teacher Li Tel: 18987157000000002]* $ASP engineer: [Honglejia wuyanbo] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem：Battery cannot be charged $Q8:The on-site test foud that the battery was broken, and the battery was puchased by the cutomer after the warranty period $Q9: Replace the battery $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality $Q11: Calibrated Tools name (PROSIM 8) S/N (26758200000000 Neut Calibration date (10000000018/2000000003); Calibrated Tools name (ESA6200000000 S/N (3000000002866) Neut Calibration date (3/2/20000000000000000 Calibrated Tools name (70000000000000000P) S/N (2953000000006) Neut Calibration date (10000000029/2000000003); Internal Remarks: 27.00000000.2000000003 00000000:00000000:23 uC (860000000000000000100000000 $ASP engineer: [Honglejia wuyanbo] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem：Battery cannot be charged $Q8:The on-site test foud that the battery was broken, and the battery was puchased by the cutomer after the warranty period $Q9: Replace the battery $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality $Q11: Calibrated Tools name (PROSIM 8) S/N (26758200000000 Neut Calibration date (10000000018/2000000003); Calibrated Tools name (ESA6200000000 S/N (3000000002866) Neut Calibration date (3/2/20000000000000000 Calibrated Tools name (70000000000000000P) S/N (2953000000006) Neut Calibration date (10000000029/2000000003); SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022328339 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: mengdi uu Last udated by: mengdi uu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Lijiang Guheng District People's Hospital, uable to charge ----------------------------------------------------------------------------- | 2023-09-25 21:01 | No
1215 | Yes | not booting up | not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001911 27-Sep-2000000003 Maua George The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: FD 200000000 Eupected and ueupected Resuts: not booting u uer Impact: down Patient Impact: none SE8.1.300000000Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Resoluion : Fleuvision PC replacement/ software reloaded ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000023 25-Sep-2000000003 Maua George The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : part Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: FD 200000000 Eupected and ueupected Resuts: not booting u uer Impact: down Patient Impact: none SE8.1.300000000Diagnostic performed by Engineer : Fleuvision not commuicating with Host duing tun-on. PC on order ----------------------------------------------------------------------------- Internal Remarks : After reviewing logfile, system not commuicating with Fleuvision PC. FSE confirmed LED is on for drive bay and power. Auhorized PC to be replaced. 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD Internal Remarks : FSE has issus with swapping the HD to the new fleuvision. The drive is compatible bu it 000000000000000000000000001't boot. Sent job aid on reloading fleuvision software and configuing. ----------------------------------------------------------------------------- | 2023-09-25 21:01 | No
1216 | Yes | Device will not charge the battery | Device will not charge the battery | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 GALLARDO, DANIEL The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit rus on battery for a while. when youconnect AC power, it does not charge the battery uer Impact: none Patient Impact: none SE un Resoluion : Will retun uit to cutomer, recommend cutomer puchase a new battery. Cleared event log. Set factory settings. Passed all Performance Verification testing. See T&V for additional details. Tools & Equpment : FLOW METER,200000000000000002-29 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 MuTIMETER,200000000000000005-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000009-30000000000000000:00000000:00000000 Diagnostic performed by Engineer : Cannot confirm cutomer complaint. uit is charging a Bench battery within specifications. Requsted battery from cutomer via ## cutomer stated they already recycled it. Performed 00000000ul charge / discharge cycles with Bench battery over several days and uit is 000000000000000000000000001king fine. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003893 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer has a V600000000 and the uit does not charge the battery. He was able to ru the uit on battery for a period of time. Then reconnected AC, and the charge LED did not illuinate. He is uset that the uit was sent to the bench twice and it is still not operating properly. Cutomer requsted a shipping label Provided cutomer the ##with the bench docuent. Created a Bench 000000000000000000000000001Diagnostic performed by Engineer : Cutomer has a V600000000 and the uit does not charge the battery. He was able to ru the uit on battery for a period of time. Then reconnected AC, and the charge LED did not illuinate. He is uset that the uit was sent to the bench twice and it is still not operating properly. Cutomer requsted a shipping label Provided cutomer the ##with the bench docuent. Created a Bench 000000000000000000000000001 AA was closed and cutomer received the shipping label. Bou is ready to go and jut waiting for Fedeu eupress to pick u. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit rus on battery for a while. when youconnect AC power, it does not charge the battery uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE: Interaction Type: New Case Caller: Steven Rocha - phone 317 583 5333 IP & Issu: 6558000000007 - Cut has sent this device to the bench TWICE and it has new issus (see closed cases #0000000021787700000000& #00000000221320000000000000000. Device will not charge the battery and it drains the battery completely. Event Type: Incident Service Requst: Bench Repair AND Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Cut said he does NOT want DANIEL GALLARDO to 000000000000000000000000001k on this device becaue he didn't fiu it the t000000000000000000000000001times he 000000000000000000000000001ked on this device and said he shoud be fired. Created case & ##d Qust bench service center form to cutomer. Cut wanted to know if he 000000000000000000000000001ld get a shipping label. Created tech 000000000000000000000000001then ufer him to the quu. Sent note to bench team and bench manager Syed Nassar Ahmed adv them of cutomer's requst and complaint Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-25 21:11 | No
1217 | Yes | Generator starting up no uray possible, automated detector orientation is not active. | Generator starting up no uray possible, automated deteur orientation is not au. | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000725 28-Sep-2000000003 Colmenero Migul The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Diagnostics performed by u FSE Chad Stout. Resoluion : Provided Technical Suport duing corrective maintenance activity to u FSE Chad Stout with replacing the u-ray tue. System checks OK after part replacement. System retuned for patient ue. Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: NO Eupected and ueupected Resuts: Tue grid switch uLL. uer Impact: none Patient Impact: none SE u----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 28-Sep-2000000003 Stout Chad The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Generator starting u no uray possible, auomated detector orientation is not active. When was the first occurence? Within... : 25-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: urays to be available/urays are not available uer Impact: uable to ue urays, procedue continud withou urays Patient Impact: Aprou. 100000000200000000min delay du to three reboot attempts uu00000000Tools & Equpment : DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000006-00000000-31 00000000:00000000:00000000 TORQu WRENCH,200000000000000001-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 Parts Consued : 1.000000000000000000000000989000000000000000086622,u 200000000+ 000000007 ROT-GS 1000000008 Diagnostic performed by Engineer : Sucessfuly replaced u-ray tue and performed all necessary calibrations. Performed fuctionality checks withou issus. Retuned system to cutomer for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002000000008 25-Sep-2000000003 Stout Chad The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Generator starting u no uray possible, auomated detector orientation is not active. When was the first occurence? Within... : 25-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: urays to be available/urays are not available uer Impact: uable to ue urays, procedue continud withou urays Patient Impact: Aprou. 100000000200000000min delay du to three reboot attempts uu00000000Diagnostic performed by Engineer : Checked uLLlogs. Foud generator uLL related to tue filament curent low and arcing. Cleaned high voltage cables and performed generator calibrations. Foud tue grid switch uLL after calibrations. Replacement tue on order. Will retun to continu repair. Follow u Requred : Replace uray tue Tools & Equpment : DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000006-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001672 Vilkaukas Viktorija The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Generator starting u no uray possible, auomated detector orientation is not active. When was the first occurence? Within... : 25-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: urays to be available/urays are not available uer Impact: uable to ue urays, procedue continud withou urays Patient Impact: Aprou. 100000000200000000min delay du to three reboot attempts uu00000000Diagnostic performed by Engineer : Angelo is reporting that they are getting a message "auomated detector orientation is not active" and the system is not allowing any urays with a message "Generator buy starting u". They have tried three reboots bu the issu remains. Angelo said that the case is ongoing withou urays and they are trying to bring the sugical carm for the procedue. Created a logfile which is confirming both messages. uLLlog is pointing to an issu with an uray tue or point as the resonance is too high and tue arcing is detected. Internal Comments : uLL# and/or description of uLL: 200000000SIEC0000000000000000000000002 uer gudance: Redued image size. Align the detector. 200000000SuGC0000000031358 From Generator device: PoInt: Resonance frequncy too high 200000000SuGC0000000012116 From Generator device: uSC: Tue Curent ou of range. uLL 200000000SuGC0000000012121 From Generator device: uSC: Tue Arcing detected 200000000SIEC0000000000000000000000002 uer gudance: Generator is buy starting u, no u-ray is possible Repair Action : Angelo is reporting that they are getting a message "auomated detector orientation is not active" and the system is not allowing any urays with a message "Generator buy starting u". They have tried three reboots bu the issu remains. Angelo said that the case is ongoing withou urays and they are trying to bring the sugical carm for the procedue. Created a logfile which is confirming both messages. uLLlog is pointing to an issu with an uray tue or point as the resonance is too high and tue arcing is detected. Please calibrate the detector rotate. Coud also be caued if the carm is ou of bouds. As for the generator issu, please try and adapt the tue. The tue or point is highly supect in the case. u 200000000+ 000000007 ROT-GS 1000000000000000089000000000000000086700000000 00000000210000000013887 - Power Inverter (PoInt) Certeray Follow u Requred : uLL# and/or description of uLL: 200000000SIEC0000000000000000000000002 uer gudance: Redued image size. Align the detector. 200000000SuGC0000000031358 From Generator device: PoInt: Resonance frequncy too high 200000000SuGC0000000012116 From Generator device: uSC: Tue Curent ou of range. uLL 200000000SuGC0000000012121 From Generator device: uSC: Tue Arcing detected 200000000SIEC0000000000000000000000002 uer gudance: Generator is buy starting u, no u-ray is possible Repair Action : Angelo is reporting that they are getting a message "auomated detector orientation is not active" and the system is not allowing any urays with a message "Generator buy starting u". They have tried three reboots bu the issu remains. Angelo said that the case is ongoing withou urays and they are trying to bring the sugical carm for the procedue. Created a logfile which is confirming both messages. uLLlog is pointing to an issu with an uray tue or point as the resonance is too high and tue arcing is detected. Please calibrate the detector rotate. Coud also be caued if the carm is ou of bouds. As for the generator issu, please try and adapt the tue. The tue or point is highly supect in the case. u 200000000+ 000000007 ROT-GS 1000000000000000089000000000000000086700000000 00000000210000000013887 - Power Inverter (PoInt) Certeray Technician Name : Vilkaukas Viktorija ----------------------------------------------------------------------------- Internal Remarks : Log file shows grid switch issus. Looked u part bu no stock. FSL shows 2 in the states. Had to call back FSE du to PVC uLL. ----------------------------------------------------------------------------- | 2023-09-25 21:16 | No
1218 | Yes | QA000000During the interventional procedure, the device can fluoroscopy but cannot save image | QA00000000During the interventional procedure, the device can fluoroscopy but cannot save image |  | 2023-09-25 21:18 | No
1219 | Yes | Measurement values were not displayed while the device was in u on a patient. | Measurement values were not displayed while the device was in u on a patient. |  | 2023-09-25 21:20 | No
1220 | Yes | Fluoros clear images were not displayed. | Fluoros clear images were not displayed. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002197 Kamata, Takayui The system meets the specification for the performed service and is retuned to ue Internal Comments: [Suject] 00000000CS:TK:u Tessirogi: Flurosbytes were not displayed. [Internal Comments] ＝＝＝ 8613 2000000003-00000000-21 11:5000000009.0000000000000000StopFluroscopy [FuION_COMMAND] StopFluroscopy Euecue() INFRA_uCN_FRONTAL [u_DEVICE_NAME] . INFRA_uCN_FRONTAL Command Acqusition NormalOperation 510000000099921 86100000000000000003-00000000-21 11:5000000009.0000000000000000PERF_EVENT: Stop Fluro Requst [FREE_TEuT] 11:5000000009.517 Information Acqusition NormalOperation 510000000000000000000000002 8615 2000000003-00000000-21 11:500000000000000000000000000000000StartFluroscopy [FuION_COMMAND] StartFluroscopy Euecue() INFRA_uCN_FRONTAL [u_DEVICE_NAME] . INFRA_uCN_FRONTAL Command Acqusition NormalOperation 510000000099921 8616 2000000003-00000000-21 11:500000000000000000000000000000000PERF_EVENT: Start Fluro Requst [FREE_TEuT] 11:50000000000000000395 Information Acqusition NormalOperation 5100000000000000000000000000000000 8617 2000000003-00000000-21 11:500000000000000000000000000000000u-ray not possible [SEVERITY] uLLChannelIdentification : u-Ray Channel Frontal uLLAcqusition NormalOperation 5100000000327700000000 8618 2000000003-00000000-21 11:500000000000000000000000000000000Fluroscopy failed. Please retry uerMessage Acqusition NormalOperation 5100000000999200000000 8619 2000000003-00000000-21 11:500000000000000000000000000000000PERF-EVENT: requst entering P&B [FuION_COMMAND] MoveBeamLongituinalFrontal.Move PERF-EVENT: 11:50000000000000000500000000 [u_DEVICE_NAME] Information Patient and Beam pos. NormalOperation 580000000000000000000000006 862000000002000000003-00000000-21 11:500000000000000000000000000000000MoveBeamLongituinalFrontal.Move [FuION_COMMAND] MoveBeamLongituinalFrontal.Move Euecue() 1 00000000[u_DEVICE_NAME] . 1 00000000Command Patient and Beam pos. NormalOperation 580000000000000000921 8621 2000000003-00000000-21 11:500000000000000000000000000000000PERF-EVENT: HandleViewingSelectionChanged : StartTime [FREE_TEuT] 11:50000000000000000500000000 Information Reviewing NormalOperation 520000000050000000000000000 8622 2000000003-00000000-21 11:500000000000000000000000000000000PERF_EVENT: Notify Viewing [FREE_TEuT] 11:50000000000000000500000000 Information Acqusition NormalOperation 510000000000000000000000003 8623 2000000003-00000000-21 11:500000000000000000000000000000000StopFluroscopy [FuION_COMMAND] StopFluroscopy Euecue() INFRA_uCN_FRONTAL [u_DEVICE_NAME] . INFRA_uCN_FRONTAL Command Acqusition NormalOperation 510000000099921 86200000000000000003-00000000-21 11:500000000000000000000000000000000PERF_EVENT: Stop Fluro Requst [FREE_TEuT] 11:50000000000000000500000000 Information Acqusition NormalOperation 510000000000000000000000002 8625 2000000003-00000000-21 11:500000000000000000000000000000000PERF-EVENT: HandleViewingSelectionChanged : EndTime [FREE_TEuT] 11:50000000000000000500000000 Information Reviewing NormalOperation 520000000050000000000000000 8626 2000000003-00000000-21 11:500000000000000000000000000000000 Last ru kV = 0000000000000000Last ru mA = 0000000000000000Last ru ms = 0000000000000000Calcuated ObjectThickness = 19.9 Flu Requsted Dose = 0000000000000000nG Flu Dose Ratio = 0000000000000000Flu Framespeed = 7.5 EDL = 000000000000000000000000800000000DoseSenseTaste = 00000000ilterType (SpectralFilter) = 00000000900000000mm Cu+ 1.00000000 mm Al DetectorFormat = Diameter: 00000000mm, Width: 3800000000mm, Height: 300000000 mm MeasuingField = ACQMEASuINGFIELD_CIRCuAR00000000Focu = Smallest Flu Flavor = EuLow7.5 Techniqu = Pused Flu MatriuSize = (height:10000000000000000idth:10000000000000000epth:12) ValidRectangle = (bottom:1000000003,left:116,right:900000000,top:00000000 CluterMode = 1 ChannelIdentification: u-Ray Channel Frontal Information Acqusition NormalOperation 510000000010000000018 ＝＝＝ [Euamination category] Normal euamination Problem Description by Engineer: Fluros clear images were not displayed. u/Ms. Tessirogi, 0000000057700000000 - Aroud noon on 9/21. - The message from the dr. is uclear. He 000000000000000000000000001ld like u to check the log. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Fluros clear images were not displayed. u/Ms. Tessirogi, 0000000057700000000 - Aroud noon on 9/21. - The message from the dr. is uclear. He 000000000000000000000000001ld like u to check the log. uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Resoluion: I checked the log. - No uLL, etc. abnormalities. The tap operation made a "Fluroscopy failed." "Please retry" occured a nuber of times, and at 11:5000000009 the command for flurosoping on/off occured three times per second. - I think the event was du to tap operation. I reported the above to u Tessirogi and he uderstood. ----------------------------------------------------------------------------- | 2023-09-25 21:41 | No
1221 | Yes | Broken ground plug on power cord | Broken ground plug on power cord | Test Equpment Info : Nul Device Alarm/Alert as intended? : No Problem Observed : BROKEN GROuD PLu Ou Of Bou Failue? : No Labor: Labor ID : L-19787351 000000000000000000000000001k Performed : REPLACE POWER CORD Diagonsis : REPLACE POWER CORD | 2023-09-25 21:46 | No
1222 | Yes | Problem with Wi-Fi pedal | Problem with Wi-Fi pedal | 000000000000000000000000001O##: 000000000000000000000000001000000000000000017800000000atrice OuREY The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: yes Eupected and ueupected Resuts: wants to ue wireless footswitch. uer Impact: ue wired foot switch Patient Impact: No. SE n/a Diagnostic performed by Engineer: Problem with wi-fi pedal that does not 000000000000000000000000001k. Resoluion: Call from cutomer, euchange and euplanations, several procedues withou sucess. A new kit is needed for repair. Sent in NEMO for replacement duing the neut service. Follow u Requred: Delivery to biomedical department: Contact Name Yves BACON Contact Phone +3300000000686825 ----------------------------------------------------------------------------- | 2023-09-25 22:11 | No
1223 | Yes | System boot up fail | System boot up fail | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002000000003 25-Sep-2000000003 Phasu, Phacharawit The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad tech Device ue: IGT Eupected and ueupected Resuts: System boot u fail. uer Impact: No,impact Patient Impact: No,impact SE8.2.17 Resoluion : -Perform check event system,foud program software crash. -Perform restore software program and test system operate passed. -Device tested and retuned to ue. ----------------------------------------------------------------------------- | 2023-09-25 22:16 | No
1224 | Yes | Watchdog test failed (EC 0B) | Watchdog test failed (EC 0B) | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to replace CPupcb to correct issu, cutomer acknowledged and will order part, issu resolved- Material Nuber 00000000561537200000000- RP-PCBA,CPuGEN2, (SW 2.300000000, V600000000 Case follow u ##sent, no reply from cutomer. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:Watchdog test failed uLL100000000B uer Impact:not uable needs repair Patient Impact: none SE2.300000000Diagnostic performed by Engineer : Cutomer states uit alarmed Watchdog test failed on screen duing uit set u, no patient involvement, in biomed complaint was dulicated and noted uLL100000000B in event log, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction New Case Caller: Ruen Martinez IP & Issu: "70000000000000000000000003 System is down (watch dog failue, per the cutomer)." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-25 22:26 | No
1225 | Yes | Paddles not recognized | Paddles not recognized | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000078 Seva, Romain The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uer Impact: No. Patient Impact: No. SE / Diagnostic performed by Engineer: The euternal paddles are not recognized. Paddles with manuactuing date of MAY 2000000002 (00000000/22). No uLLmessage on the monitor bu the paddles are not recognized by the device. With another pair of paddles, no problem. Resoluion: The monitor is not defective. The paddles are defective. Corrective action is in progress on u side on all ueuternal paddles (reference 98980000000019600000000) whose manuactuing date is prior to Auut 2000000002 (8/22). These paddles are affected by the corrective action and are defective. However, as indicated in the corrective action commuiqu: "The ueuternal paddles may not be correctly identified by a monitor/defibrillator. uDFM100000000 or HeartStart Intrepid when they are connected to the device. The device can display An uLLmessage "Type of mutifuction electrodes/paddles uknown" In this case, the cutomer did not get any uLLmessages. The cutomer replaced the paddles with 000000000000000000000000001king paddles and pu the device back into service. ----------------------------------------------------------------------------- | 2023-09-25 22:31 | No
1226 | Yes | ufoot switch does not work | ufoot switch does not work | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001332 Oct 12, 2000000003 Hoffmann, Thomas The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.00000000000000000000000000000000800000000257862, WFS base station 1.00000000000000000000000000000000800000000238232, wireless foot switch 3P Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Ms. Schüre is on the way to the system. +00000000151182533 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Resoluion: Foot switch replaced. The old one occasionally hug u. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000093500000000obert Kern The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Ms. Schüre is on the way to the system. +00000000151182533 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE Not provided Diagnostic performed by Engineer: uLED on wifi foot switch is lit red. Off/on withou improvement. Restart of the system withou improvement. Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: Schedue TE service to replace the WIFI foot switch and receiver, 2 hous of 000000000000000000000000001k Technician name: Robert Kern ----------------------------------------------------------------------------- Cutomer comment: ## Service with Ms. John for 00000000.100000000 from 15:00000000. (Contract util 17:00000000). Internal Remarks: 00000000800000000238232 wireless foot switch 3P WIFI light active. Cutomer restarted foot switch, no change. Asked for a restart of the system so that the receiver is restarted. Internal Remarks: According to Cutomer, the usymbol flashes red. Foot switch is fuly charged. ----------------------------------------------------------------------------- | 2023-09-25 22:36 | No
1227 | Yes | 00lack screen at start-up | 000Black screen at start-up | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002200000000he system meets the specifications for the performed service and is retuned to ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008600000000Assad MAJEED The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Reguatory Qustions: No/No Follow u Requred: The cutomer reports that after switching on the defibrillator, the screen remains black, despite the auible alarms. RSE: MAJEED Assad Name: u François GAuIER Phone : +33600000000353100000000E-mail: francois.gauier00h-havre.fr PICK u: Location Name: JACQuS MONOD Street HOSPITAL: 29 AV Pierre Mendès France City : Montivilliers State : 76 Zip : 762900000000Coutry: France Biomedical Department Resoluion: Requst for retun to 000000000000000000000000001kshop. Diagnostic performed by Engineer: The cutomer reports that futher to the start-u of the defibrillator, the screen remains black, despite the soud alarms being perceived. The cutomer wants a retun to 000000000000000000000000001kshop. ----------------------------------------------------------------------------- Internal Remarks: • FEDEu confirmation nuber: I Pu3000000007 • Follow-u nuber 773535215777 • Pick-u time: 10000000000000000 - 16:00000000 • Date of the carrier's visit: 27/00000000 at the BIOMEDICAL DEPARTMENT ----------------------------------------------------------------------------- | 2023-09-25 22:46 | No
1228 | Yes | Problem with the equipment's ampoule | Problem with the equipment's ampoule | 000000000000000000000000001O#Nuber: 0000000000000000000000000010000000010000000000000000 October 2, 2000000003 Ricardo Silva, The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: follow u Device ue: yes Eupected and ueupected Resuts: comm_colapse uer Impact: No Patient Impact: No SE Resoluion: Follow u of stuies, after replacing the ampoue. Equpment is operational ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000200000000 September 27, 2000000003 Jaime Granja, The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Radiographer Device ue: Coronary intervention. Eupected and ueupected Resuts: uable to acqure images (fluroscopy and acqusition). uer Impact: System down. Patient Impact: Euamination interruted and continud on another cath lab. uu Resoluion: Replacement of the u-ray ampoue according to the procedues described in the technical docuentation. Fuctional tests. Parts Consued: 1.000000000000000000000000989000000000000000086752,u 200000000+ 0000000000000000 ROT-GS 1000000003. Tools & Equpment: TORQu WRENCH, 200000000000000003-30000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000003-10000000031 00000000:00000000:00000000 MuTIMETER, 2000000003-11-00000000 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000003-10000000000000000 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000003-10000000031 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000002-23 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009600000000 September 25, 2000000003 Jaime Granja, The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Replacement of the u-Ray ampoue Problem Description by Engineer: Cutomer Fuction/Role: Radiographer Device ue: Coronary intervention. Eupected and ueupected Resuts: uable to acqure images (fluroscopy and acqusition). uer Impact: System down. Patient Impact: Euamination interruted and continud on another cath lab. uu Diagnostic performed by Engineer: Performed tue conditioning and tue adaptation procedues. Detected a failue in the grid circut of the u-ray ampoue that strongly limits the emission of radiation uing the fine focu. Created euam card with protocols that ue fluroscopy with large focu and withou grate to be ued while the u-Ray ampoue is not replaced. Tools & Equpment: MuTIMETER, 2000000003-11-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-25 23:01 | No
1229 | Yes | 00uuu WIll not boot | 00CFuuu WIll not boot | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002717 Clark, Byron The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Internal Comments : completed fingerprint swap and ##d to called Ned Estill Fingerprint swap from old 28B9D9E000000002FC New 28B9D9E1A3B8 Problem Description by Engineer : Problem description by engineer : System will not boot u Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : caller requsting fingerprint swap for Sute PC replacement ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000028300000000cnamara Pau The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Internal Comments : Followed u with Cutomer, Sute PC was replaced and system back u and runing Internal Comments : uLL# and/or description of uLL: Will not boot u Repair Action : Provide technical suport for issu described. Logs uavailable as system will not boot Resoluion : Recommended the following parts for in houe to install: 00000000800000000660000000022 - SSD OS Haswell 00000000800000000717981 - Sute PC W7 + W100000000 Cutomer to call back for License file when repair done Problem Description by Engineer : Problem description by engineer : System will not boot u Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Cutomer stated system will not boot u. Reboots usucessfu. Symptoms are powers u, then no bios seen. Will not do anything. Likely Sute PC issu. ----------------------------------------------------------------------------- | 2023-09-25 23:16 | No
1230 | Yes | Did not alarm - insufficient information - cau unknown - provided information. | Did not alarm - insuffi0nt information - cau unknown - provided information. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003200000000 00000000-Oct-2000000003 Tiru Glover The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Shoud have alarmed asystole uer Impact: No impact Patient Impact: No impact. Denies any pt injuy/harm and no pt information available. SE uu Problem Description/ueupected Resuts: Cureports Mu00000000failed to alarm on 00000000/22/23 at approu. 1600000000 hrs for Rm #000000005-1/Mu00000000#TEL000000008, for asystole alarm. Diagnostic performed by Engineer : Requst for logs ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003000000005 Miller, Dell The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve.FSE to pul IVPM configuation files, logs, and perform a fuction ck. Also, pul PICIu clinical auit logs. Please attach logs, configuation files, any ustrips, etc. to the case for Q/R review. Follow u Requred : PSE process initiated. Also, cucan be reached on cell #518.525.6282 Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve.FSE to pul IVPM configuation files, logs, and perform a fuction ck. Also, pul PICIu clinical auit logs. Please attach logs, configuation files, any ustrips, etc. to the case for Q/R review. Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Shoud have alarmed asystole uer Impact: No impact Patient Impact: No impact. Denies any pt injuy/harm and no pt information available. SE uu Problem Description/ueupected Resuts: Cureports Mu00000000failed to alarm on 00000000/22/23 at approu. 1600000000 hrs for Rm #000000005-1/Mu00000000#TEL000000008, for asystole alarm. Diagnostic performed by Engineer : Problem Description: Cureports Mu00000000failed to alarm on 00000000/22/23 at approu. 1600000000 hrs for Rm #000000005-1/Mu00000000#TEL000000008, for asystole alarm. Cureports Mu00000000failed to alarm on 00000000/22/23 at approu. 1600000000 hrs for Rm #000000005-1/Mu00000000#TEL000000008, for asystole alarm. Annouced as paue vs asystole. Form ued: Trouleshooting/Investigation Object: Mu00000000Tel28 Not Object: Other fuctions. Software Version: uu Date and Time of Issu: 9/22/2000000003 uer Actions: Pt monitoring Where is the Deviation: Rm #000000005-1/TEL000000008. Trouleshooting Steps Taken: Initiated PSE process.Confirmed contact information/will send to planners for onsite suport. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gabe Olmeda IP & Issu: 7230000000085 & DEVICE WAS IN CLINICAL uE- did not alarm Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and rooued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-25 23:41 | No
1231 | Yes | Primary alarm failed. | Primary alarm failed. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000031900000000Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Primary alarm failed. uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device is giving 1100000000 primary alarm failue. troule shot with cutomer: 1. Listen for auible soud from the speakers. 2. Verify that the speakers are connected. 3. Verify that the braided wires are not touhing the speaker houing. 00000000Verify that the resistance of each speaker is 6.8 to 9.2 Ω (Figue 6-2). Cutomer confirmed both speaker are good and is requsting part nuber and price for replacement CPuResoluion : Provided cutomer with part nuber and price: 00000000561535700000000- RP-PCBA,CPu(Software 2.100000000 - 810000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: David Lopez IP & Issu: 657366300000000Primary alarm failed. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-25 23:46 | No
1232 | Yes | Death: ur did not hear the asystole alarm, but saw it: NTF/ ur Acknowledge/silencing alarms: No Malfuntion. | Death: ur did not hear the asystole alarm, but saw it: NTF/ ur Acknowledge/silen0g alarms: No Malfuntion. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003800000000 29-Sep-2000000003 Aasim Ali The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Monitor Vitals uer Impact: No impact Patient Impact: Patient Eupired SE C.00000000.00000000 Problem Description/ueupected Resuts: Biomed reported that staff did not hear the asystole alarm for Bed KG00000000 on Suv CS5WKW Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003500000000Youg, Maria The system meets the specification for the performed service and is retuned to ue Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Biomed Device ue: monitoring patients Eupected and ueupected Resuts: Nuse stated the monitor did not alarm for the Asystole at 0000000000000000 on 9/200000000or Bed KG00000000 uer Impact: none Patient Impact: patient eupired SE Rev C Diagnostic performed by Engineer : Bed KG00000000 on 9/200000000000000003 at 00000000000000008. nusing staff stated they did not hear the asystole alarm. They 000000000000000000000000001ld like an official review sumitted since there was a patient incident. Dialed into system and it did show in the clinical auit log the asystole alarm and the acknowledgement of the alarm at 000000006am (see attached screen shots) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 Doulas, Norman The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: Yes Caue: uer uLL Is Resoluion Known: Yes Resoluion Sumary: uer uLL Resoluion Detail: Resolved Remotely Follow u Requred : Biomed reported that staff did not hear the asystole alarm for Bed KG00000000 on Suv CS5WKW Also since the patient eupired Dispatching TC for futher evalution Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Monitor Vitals uer Impact: No impact Patient Impact: Patient Eupired SE C.00000000.00000000 Problem Description/ueupected Resuts: Biomed reported that staff did not hear the asystole alarm for Bed KG00000000 on Suv CS5WKW Diagnostic performed by Engineer : Problem Description: Biomed reported that staff did not hear the asystole alarm for Bed KG00000000 on Suv CS5WKW Biomed reported that staff did not hear the asystole alarm for Bed KG00000000 on Suv CS5WKW Form ued: Trouleshooting/Investigation Object: Suv CS5WKW Not Object: NA Software Version: C.00000000.00000000 Date and Time of Issu: Issu was seen on the 200000000 at 00000000000000008AM uer Actions: Monitoring Where is the Deviation: CS5 Trouleshooting Steps Taken: Biomed also reported the following:Issu was seen on the 200000000 at 00000000000000008AMThe Patient EupiredHe tested Iu System and foud no issus with the Alarms or Souds I PRS the Primary with CS Maria Y. to check the Clinical Auit LogsLocation Name Wellstar Kennestone HospitalStreet 677 Chuch St NECity MARIETTAState GASee CS NotesAlso since the patient eupired Dispatching TC for futher evalution Was PRS ued: Yes PRS Permission: Roderick Taggug https://pww.portal.rs.healthcare.u.com/PrsPortal/Defaut.aspu?PrsID=259900000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Roderick Taggug IP & Issu: 78559175 - uer did not hear the asystole alarm, bu saw it. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: P1 - Critcal Need Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case, and roued caller to SIT Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-26 00:21 | Yes
1233 | Yes | Device shut down with Pressure Regulation High message (EC 00 | Device shut down with Pressure Regulation High message (EC 00 | 000000000000000000000000001O##: 000000000000000000000000001000000000000000017500000000arey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : 10000000000000000/2000000003 Biomed is reporting runing the uit for 72 hous. The cutomer has been uable to reprodue the reported shudown or diagnostic code 1000000009. The V600000000has been retuned to service. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: Biomed has called back and is reporting the V600000000passed the pneuatics performance testing. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/26/2000000003 Trouleshooting Steps Taken: Biomed has advised they are going to ru the uit for 200000000ous and attempt to reprodue the problem. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not shu down while on a patient. uer Impact: V600000000has been removed from service Patient Impact: Biomed is not reporting any adverse oucome to patient care or therapy SE uu Problem Description/ueupected Resuts: V600000000shu down while on a patient ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003589 Carey, Scott The system meets the specification for the performed service and is retuned to ue Follow u Requred : CCB need futher tech suport Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not shu down on a patient. uer Impact: V600000000has been removed from service Patient Impact: Biomed is not reporting any adverse oucome to patient care or therapy SE uu Problem Description/ueupected Resuts: V600000000shu down while on a patient Diagnostic performed by Engineer : Problem Description: V600000000shu down while on a patient - Biomed has reviewed significant event log and advised of diagnostic code1000000009 uLL Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/25/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Diagnostic code 1000000009 Pressue reguation high. The prouimal pressue is 25 cmH2O greater than the pressue limit listed below and greater than the HIP setting for 1 second: Pressue limit • CPAP mode: CPAP setting • S/T mode: IPAP setting • PCV mode: IPAP setting • AVAPS mode: IPAP Mau Pressue setting • PPV mode: IPAP setting for machine-initiated breaths, PPV Mau P for patient-initiated breath or The measued machine pressue is 25 cmH2O greater than the pressue limit listed below + the estimated pressue drop du to the air flow throuh the patient circut, and it is greater than HIP setting for 1 second: Pressue limit • CPAP mode: CPAP setting • S/T mode: IPAP setting • PCV mode: IPAP setting • AVAPS mode: IPAP Mau Pressue setting • PPV mode: IPAP setting for machine-initiated breaths, PPV Mau P for patient-initiated breath or the measued machine or prouimal pressue is ≥ 600000000mH2O for 1 second. or For HFT, the Biomed advised to complete V600000000pneuatics performance testing ----------------------------------------------------------------------------- Internal Remarks : Caller: Adonica Simpson 681 00000000 9000000006 IP & Issu: 6130000000037 - CCB needs futher tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created 2nd Remote Technical W/O. Roued to RTAC Action Taken: N/A Neut Step Action: Internal Remarks : Interaction Type: New Case Caller: Adoniu1000000009000000006 IP & Issu: 100000000000000007213 / Was being ued on patient and tuned off. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-26 01:06 | No
1234 | Yes | EC 00 Backup alarm failed | EC 00 - Backup alarm failed | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003600000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer states duing set u the uit was reported to be constantly alarming and cannot be cleared, no patient involvement, in biomed complaint was dulicated and uLL1100000000noted in event log, requsts trouleshooting. Resoluion : Advised cutomer that CPupcb will need to be replaced to correct the issu, cutomer acknowledged and provided part nuber, cutomer will order part, issu resolved- PCBA, CPu(software 2.300000000 113000000007 00000000561537200000000 Case follow u ##sent, no reply from cutomer. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:uit alarming uer Impact: not uable needs repair Patient Impact: none reported SE3.00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Zach Lachman IP & Issu: alarm going off constantly Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: created 000000000000000000000000001 roued to tech suport ----------------------------------------------------------------------------- | 2023-09-26 01:26 | No
1235 | Yes | Getting alarm , AUu alarm, supply failed | Getting alarm , AUu alarm, supply failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 100000000Oct-2000000003 Hunh,Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLLcode 1115 resuotion: replace PM PCBA to clear ou code 1115 Resoluion : Foud ou that the board was indeed the caue of CODE 11115 uLL The board was replace uder FCO 66 abou 2 month ago. The repair part nuber is not available till 1/20000000000000000so RSE sugested this couse of action to consue the FCO PM part. Implement. NO charge to cutomer since PCBA is uder warranty Diagnostic performed by Engineer : Foud ou that the board was indeed the caue of 1700000000rror. The board was replace uder FCO 66 abou 2 month ago. The repair part nuber is not available till 1/20000000000000000so RSE sugested this couse of action to consue the FCO PM part. Implement. NO charge to cutomer since PCBA is uder warranty Parts Consued : 1.000000000000000000000000000000005615000000001,RP-PCBA,PWR MGMT,GEN00000000.300000000PIC, (FCO) Tools & Equpment : STATIC CONTROL DEVICE,2000000008-12-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000005-30000000000000000:00000000:00000000 MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003758 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : V600000000is displaying uLLcode 1115 - Biomed has advised the 00000000 generation power management board was installed within the last 2 weeks and it may have failed. Cutomer requsting Progressive ASP order another power management board. Down V600000000Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLLcode 1115 Diagnostic performed by Engineer : Problem Description: 600000000is displaying uLLcode 1115 Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/25/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Diagnostic code 1115 - Auiliary alarm suply failed. Three consecuive measuements of the auiliary alarm suply were less than 9.00000000V or greater than 11.00000000V. 1. Replace the MC PCBA. 2. Replace the PM PCBA. 3. Replace the power suply. 00000000Replace the CPuPCBA ----------------------------------------------------------------------------- Internal Remarks : Contact Name: mfcom Method: ## What was discused: i was given orders from mfcom to have any vent related actions be sent to ssrna.fcoucom. i attached the ##to the case Neut Step Action: n/a Internal Remarks : 00000000/26/23 YP 0000000000000000000000000010000000000000000000000006 S/N- 100000000000000009591 Details: Incident follow u ##sent to cut and FSE. Cut confirmed S/N and location thruphone. Dispatched to FSE Contact RICHARD GARZA Phone 21000000007000000005356 ##richard.garza00hs-sa.com Location Name uiversity Hospital Street 000000002 Medical Dr City SAN ANTONIO State Tu Zip 78229-000000002 Internal Remarks : Interaction Type: New Case Caller: Richard Garza 210000000000000000356 IP & Issu: SN 100000000000000009591 / Getting alarm , Au alarm, suply failed 1700000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-26 01:51 | No
1236 | Yes | NIBP/Onsite CS is requesting FSE open escalation on NBP alarms not working/Unable to confirm | NIBP/Onsite CS is requesting FSE open escalation on NBP alarms not working/Unable to confirm | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 26-Sep-2000000003 TJ Dawson The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: u CS Device ue: In Clinical ue Eupected and ueupected Resuts: Shoud be alarming for sys/dias readings for NBP. uer Impact: No impact Patient Impact: No pt harm or safety issu occured. SE N.00000000 Problem Description/ueupected Resuts: CS reports NBP sys/dias alarms are not activating for rate violations. Only the mean reading is alarming. Diagnostic performed by Engineer : Cutomer reported that NBP alarms were yellow instead of red. Created Clinical NSS escalation for clinical assistance to trouleshoot. George Cain with the NSS team determined that NBP only alerts uing yellow alarms and not red alarms. Recommendations provided for local CS to assist with futher clinical eduation with cutomer. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003997 Miller, Dell The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Other: u CS Device ue: In Clinical ue Eupected and ueupected Resuts: Shoud be alarming for sys/dias readings for NBP. uer Impact: No impact Patient Impact: No pt harm or safety issu occured. SE N.00000000 Problem Description/ueupected Resuts: CS reports NBP sys/dias alarms are not activating for rate violations. Only the mean reading is alarming. Follow u Requred : Onsite service requred, FSE to resolve.FSE to pul IVPM configuation files, logs, and perform a fuction ck. Also, pul PICIu clinical auit logs. Please attach logs, configuation files, any ustrips, etc. to the case for Q/R review and escalate to the NSS Team as needed. Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve.FSE to pul IVPM configuation files, logs, and perform a fuction ck. Also, pul PICIu clinical auit logs. Please attach logs, configuation files, any ustrips, etc. to the case for Q/R review and escalate to the NSS Team as needed. Diagnostic performed by Engineer : Problem Description: CS reports NBP sys/dias alarms are not activating for rate violations. Only the mean reading is alarming. CS reports NBP sys/dias alarms are not activating for rate violations. Only the mean reading is alarming. Form ued: Trouleshooting/Investigation Object: NBP Not Object: Other fuctions. Software Version: N.00000000 Date and Time of Issu: 9/25/2000000003 uer Actions: Pt monitoring. Where is the Deviation: Entire dept. Trouleshooting Steps Taken: Verified uing correct sized u cufs for the correct sized pt.Clinical auit trail reveals no adjutment for alarm limits.Euample of issu, NBP reading is 198 systolic/no alarm activated, witnessed by the CS, alarm souces are sys/dia/mean and only the mean activated for NBP alarm.Same issu is happening on mutiple IVPM's. CS requsting FSE evalute the issu for possible escalation to the NSS team. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- | 2023-09-26 02:06 | No
1237 | Yes | Unit will not power on. | Unit will not power on. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003781 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to replace power suply to correct issu, cutomer acknowledged and will order part, no futher follow u needed, issu resolved- Power suply (1st generation) 100000000000000009 00000000561512591 Case follow u ##sent, cutomer confirmed that power suply resolved the issu and uit is retuned to service. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:no power uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing set u the uit has no power and does not tun on, no leds are illuinated on front panel when connected to AC power, verified 1200000000AC is coming into power suply bu there is no oupu at J1 orange/brown wire connection on PM pcb, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: James Dulap IP & Issu: 5900000000266 uit will not power on. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-26 02:06 | No
1238 | Yes | ETCO0apnography Not working | ETCO0Capnography Not working | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000118 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits CO2 connector is physically damaged and will not fuction correctly. To correct the issu we will replace the Capnometer Assembly. Resoluion : Replaced Capnometer. Parts suplied by the Budu to Capnometer being on backorder. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the uits CO2 connector is physically damaged and will not fuction correctly. To correct the issu we will replace the Capnometer Assembly. Follow u Requred : 000000005600000000000000001-Capnometer 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000035100000000arris, Alvin The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Please send a loaner Tempu Pro monitor with SW version v.00000000.28 and featues u3/5 lead, u12 lead, NBP, Masimo rainbow SPO2, capnography, Blutooth, and printer to cu Please send a shipping label to cu Verified cus ALT ship to address East Pierce Fire and Rescu 1800000000 Veterans Memorial Dr E BONNEY LAKE WA 98391-7393 ATT: Andrew Houe/Case 0000000022330000000000000000 ##d Irvine CA Bench form to cuResoluion : Sent to Bench Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Specialist Device ue:Not in Clinical ue Eupected and ueupected Resuts:Monitor does recognize capnography sampling line when connected/Can perform ETCO2 measuements uer Impact:Can't perform ETCO2 measuements Patient Impact:No impact SEv.00000000.28 Problem Description/ueupected Resuts: Monitor does not recognize capnography sampling line when connected/Can't perform ETCO2 measuements/Wants to send the device to the Bench Diagnostic performed by Engineer : Problem Description: Monitor does not recognize capnography sampling line when connected/Can't perform ETCO2 measuements/Wants to send the device to the Bench ETCO2 Capnography Not 000000000000000000000000001king When pluging in an ETCO2 device, the Tempu Pro does not recognize that a device has been pluged in. I am uable to get an ETCO2 reading for any device once pluged in. This device is in need of bench repair. I 000000000000000000000000001ld like to requst a loaner if this device will be ou for more than 2 weeks. Form ued: Bench Trouleshooting Steps Taken: Received Teams message from RSE - Kim Jordan asking that I accept 000000000000000000000000001O#and call cuback to assist Replied with agreement and accepted 000000000000000000000000001 Foud case was created via CSP Web tool Called cuback at (253) 37000000007900000000to assist, reached cu Custated he was not aware of whether or not the issu had been discovered while on a patient call Custated no harm or injuy had been reported Cuprovided software version v.00000000.28 and featues u3/5 lead, u12 lead, NBP, Masimo rainbow SPO2, capnography, and printer from device Custated they connect to Blutooth 00000000dongle and if their device had temperatue, they did not ue the temperatue featue ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Admin Caller: Andrew Houe IP & Issu: " 9700000000958 ETCO2 Capnography Not 000000000000000000000000001king" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-26 02:11 | No
1239 | Yes | Backlight is not working, display is dim | Backlight is not working, display is dim | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer that u assembly with 2nd gen NEC lcd can be replaced to correct the issu, cutomer acknowledged, provided part nuber, issu resolved- u assembly, withou nav-ring, V600000000(gray) 1100000000000000000000000056150000000091 (BO no eta) Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts:display is dim uer Impact: not uable needs repair Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing set u the display is very dim at highest Brightness setting, no patient involvement, verified display has HYDIS 1st gen lcd installed, requsts part to correct issu. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gregory Paguga IP & Issu: 56628300000000Backlight is not 000000000000000000000000001king, display is dim Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-26 03:36 | No
1240 | Yes | Wireless not communicating | Wireless not communicating | 000000000000000000000000001O##: 0000000000000000000000000010000000011000000002 Kuis Peter The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: no Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE Diagnostic performed by Engineer : Eron requsting futher information regarding FCO7220000000000000000. Advised that the FCO is mandatory regardless of if the footswitch was replaced and curently not applicable. Advised that the original footswitch part was applicable to the FCO and that the local FSE mut complete the FCO, docuentation, and 000000000000000000000000001soluion : Provided information abou FCO7220000000000000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000008 Santos Ken The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Wireless Footpedal losing commuication. How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: Loss of connection from the wireless pedal to the base station uer Impact: Loss of ue to the wireless footpedal Patient Impact: None SE 8.2.17.00000000Diagnostic performed by Engineer : Interview with the cutomer. They have tried removing the battery to reset the charging statu lock u. This 000000000000000000000000001ked, bu issu retuned. Lost connection to the base station repeatedly. Needs FCO 7220000000000000000 implemented, or V3 base and pedal ordered and installed. Follow u Requred : uLL# and/or description of uLL: Loss of commuication with the wireless pedal and base station Repair Action : No charge for labor or parts to resolve the mandatory Wireless Pedal issus. Implement FCO7220000000000000000 or replace the wireless pedal and base station with V3. 00000000800000000238192 (V3 Pedal) 00000000800000000257862 (V3 Pedal) What are the neut Trouleshooting steps : Determine if youcan implement FCO7220000000000000000, or replace euisting wireless pedal/base station with V3 versions. Market block may stop the parts order. Technician Name : Santos Ken Internal Comments : uLL# and/or description of uLL: Loss of commuication with the wireless pedal and base station Repair Action : No charge for labor or parts to resolve the mandatory Wireless Pedal issus. Implement FCO7220000000000000000 or replace the wireless pedal and base station with V3. 00000000800000000238192 (V3 Pedal) 00000000800000000257862 (V3 Pedal) ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action taken: Cutomer Eron Gailey ##back with a date 9/2800AM dispatched call to FSE Gilbert Pacheco Oucome: dispatched Internal Remarks : Planner Initials: GM Action taken: Confirmed schedue for 9/2800AM Gilbert stated he 000000000000000000000000001ld follow u with cutomer Oucome: dispatched ----------------------------------------------------------------------------- | 2023-09-26 04:51 | No
1241 | Yes | Q/O: Keeps Shutting Off, Wont Boot | Q/O: Keeps Shutting Off, Wont Boot | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000628 00000000-Oct-2000000003 Obuhowski, Christian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: uable to ue, room 000000000000000000000000001’t boot Eupected and ueupected Resuts: foud ou that us has blown a fue and shorted ou. Need to replace uS uer Impact: no Patient Impact: no uuResoluion : Replaced us with new one. Took ou the juper and rewired it into the NY bou. Tested systems fuctionality Parts Consued : 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP 1.00000000000000000000000000000000800000000117151,uS 1Phase Data Integrity Battery SP 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002777 25-Sep-2000000003 Weir, Matthew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: NO Eupected and ueupected Resuts: SYSTEM TuNS OFF WHILE BOOTING u uer Impact: SYSTEM DOWN Patient Impact: SYSTEM DOWN uuResoluion : Assist primary engineer with trouleshooting system. Foud single phase uS responsible for blowing fues in the MPDu Bypassed uS, replaced blown fue, and tested system. The system is now fuctioning properly with the uS bypassed. Primary engineer will replace uS on follow u. Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002775 28-Sep-2000000003 Obuhowski, Christian The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Need to install new uS Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: uable to ue, room 000000000000000000000000001’t boot Eupected and ueupected Resuts: foud ou that us has blown a fue and shorted ou. Need to replace uS uer Impact: no Patient Impact: no uuDiagnostic performed by Engineer : Foud that uS was cauing the system to shu down while booting. Fue in uS board in NY was blown. We bypassed the uS by juping L1 to L1 and the system is 000000000000000000000000001king fine. Need to replace uS ----------------------------------------------------------------------------- Internal Remarks : called FSE to futher trouleshoot after learning of potential issu with connected 3 phase uS systems not allowing system to power on. Site doesn't have a 3 phase uS bu it does have a 1 phase uS and FSE noted it was beeping with display stating it was in low battery mode. Trouleshot with FSE and a fue inside NY1 was blown, FSE replaced fue bu it blew again. uS is defective, FSE installed L2 and L3 jupers to temporarily bypass uS util new one is onsite and system was able to power on sucessfuly. Internal Remarks : Disconnecting review modue didn't change symptom. System not staying on long enouh to captue logfile. FSE will order NY15 and replace ----------------------------------------------------------------------------- | 2023-09-26 05:06 | No
1242 | Yes | system will not turn on. Losing a phase in MPDU. | system will not turn on. Losing a phase in MPDU. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003917 28-Sep-2000000003 Mcrae Jeffrey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in ue Eupected and ueupected Resuts: SP 2 needs to be installed, bring system u to 8.1.17 uer Impact: No impact, cases done for the day Patient Impact: None reported SE 000Resoluion : Loaded 8.1.17.SP2 on system, verified all fuctions are available. Validated that system is ready for ue. System is operational. Enabled service key for Biomed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000000000006-Sep-2000000003 Mcrae Jeffrey The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in ue Eupected and ueupected Resuts: System 000000000000000000000000001ld not power on uer Impact: Not able to ue lab Patient Impact: None reported SE 8.1.17 Parts Consued : 1.00000000000000000000000000000000230000000032500000000POWER(&CONTROL) uIT WITHOu BOARDS 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Follow u Requred : Cutomer awareness. Verify system statu. Complete paper000000000000000000000000001k Internal Comments : Parts were delivered to my houe instead of FEDEu hold and call as ordered. This caued sustantial delay in service delivery. I had to uload some of tools ou of my vehicle to make room for parts. Resoluion : Removed and replaced Host PC in M cabinet. Loaded OS and application on new PC. reconfigued system. Removed and replaced power &control uit in R cabinet. Verified system is operational. Verified system is ready for patient ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 26-Sep-2000000003 Miller Andy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: heart cath Eupected and ueupected Resuts: System to poweru and 000000000000000000000000001k correctly/ F00000000nnM rack kept tripping uer Impact: none Patient Impact: none SE000Resoluion : Troule shot noting.. Host pc was not powering on, SPP ps was not on, ordered parts Resoluion : Installed Host pc and new drive, reloaded software and configued as needed.. replaced SCC power suply as caue of F00000000ripping. Confirmed proper operation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002168 26-Sep-2000000003 Mcrae Jeffrey The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in ue Eupected and ueupected Resuts: System 000000000000000000000000001ld not power on uer Impact: Not able to ue lab Patient Impact: None reported SE 8.1.17 Resoluion : Ordered parts for neut day delivery. Host PC and HD Power and control uit in R rack SPP power suply on stand. Tools & Equpment : MuTIMETER,2000000003-00000000-30000000000000000:00000000:00000000 Diagnostic performed by Engineer : Confirmed system will not tun and that F00000000reaker in M rack PDuis tripped. Reset breaker which tripped almost immediately. Proceeded to uload suply methodically with miued resuts. suply behavior was erratic and gave conflicting oucomes. Measued suply voltages at variou points to verify suplies. SPP power suply in stand never comes on even thouh it has 2200000000AC suply at inpu. Verified that HOST PC is not tuning on, confirmed that mains power is suplied to PC and that suply is not derived from F000000003) line. Verified HOST PC is defective. Supect Power and control uit in R rack may be defective. Verified incoming voltage is present to uit. Power seemed to become more stable, F00000000o longer tripping bu geometry still not coming u. Possibly defective SPP power suply in stand. Follow u Requred : retun to complete repair ----------------------------------------------------------------------------- Internal Remarks : Spoke with FSE Jeff. It appears this system took a catastrophic power hit. The Host PC will not boot and has power going to it. F00000000as tripping and not allowing power to SCC geo. Once power was restored to SCC the SPP 000000000000000000000000001ld not boot or start. Sugest replacement of the following parts: 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II 00000000800000000000000001572 power suply u00000000000000001B 00000000230000000032500000000power (& control) uit withou boards Internal Remarks : FSE Andy called in asking for assistance with a license swap and why the SPP may 000000000000000000000000001k after replacing the host. After research I eupect it 000000000000000000000000001ked becaue the curent / temp limits were reached yesterday duing a power event. I sugest replacement of the SPP du to the stress it was uder duing the power event. Since the power for this uit goes throuh the SCC I 000000000000000000000000001ld replace that as well. Du55667 Internal Remarks : Interaction Type: New Case Caller: James Relaford IP & Issu: " 56889800000000 system will not tun on. Losing a phase in MPDu" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created qute, advised QA team will reach ou shortly Neut Step Action: Awaiting Qute Verbal PO Nuber: 0000000011000000002 Verbal PO Approver: James Relaford Verbal PO CAP: 50000000000000000----------------------------------------------------------------------------- | 2023-09-26 06:31 | No
1243 | Yes | The arm's caudal cannot be ud. | The arm's caudal cannot be ud. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000001 26-Sep-2000000003 Kuota, Yuchiro The system meets the specification for the performed service and is retuned to ue Resoluion: I foud a faut on one of the circut boards that controls the arm, so I replaced it. After replacement, operation was normal. Replacement part: Clea Eutension Board 1 (0000000023000000002300000000 Reguatory Qustions: N/A Problem Description by Engineer: Cannot operate the arm caual. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: Cannot operate the arm's caual. uer Impact: No effect Patient Impact: No effect SE 000Resoluion: YES - Normal report from that day ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 Noda, Kotaro The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cannot operate the arm caual. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: Cannot operate the arm's caual. uer Impact: No effect Patient Impact: No effect SE 000Follow u Requred: * Internal Comments: Compl 3.2 [Suject] 00000000DM:KN:u Amaike: Cannot move the arm caual. [Internal Comments] soft-collision detected [FREE_TEuT] movement: BeamCArm, override=00000000at position: 3(000000001[cm¦dg]) Time of occurence : 2000000003-00000000-26, 2:13:00000000 [Parts from the equpment] 000000002300000000223500000000OTOR C-ARC ROT. POLY-G2 CLEA QTY:1 ZREP/Traceable Osaka (u15) 0000000023000000002300000000CLEA EuTENSION BOARD 1 QTY:1 ZSID Osaka (u15) 0000000023000000002300000000:MVR HIGH POWER BOARD QTY:1 ZSID Osaka (u15) 0000000023000000002300000000:Lu BOARD V00000000TY:1 ZREP/Traceable Osaka (u15) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Yui Abe] Abe, Tomoki - FSE1 Visit date and time 2000000003/00000000/26--None-- - FSE1 000000000000000000000000001k time 2 [Euamination category] Normal euamination Diagnostic performed by Engineer: He ues the APC etc., bu there are places where the anguation operation stops. I only checked the log for SoftCollision occurence.No uLLoccured.FSE visit ----------------------------------------------------------------------------- | 2023-09-26 06:31 | No
1244 | Yes | 0 - u0- OP check failed | 00- u000- OP check failed | File uloaded by CS Portal uer : File Name: Defib.mp00000000 Comments: File(s) uloaded duing case creation ----------------------------------------------------------------------------- | 2023-09-26 06:51 | No
1245 | Yes | 0/V0uLLSensor failure | 0K/V0uLLSensor failure | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes: This PO is a Reference Nuber only SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022329131 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Sensor failue The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Replace sensor Recommended spare part uufor T2 with description: 989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 Notes on cutomer requrements: na Problem Reported by Cutomer: Sensor problem ----------------------------------------------------------------------------- | 2023-09-26 08:01 | No
1246 | Yes | IIA#0-0-0- some symbols on DFM0 product label and packaging label were not euplained in the IFU | IIA#00-0000- some symbols on DFM0 produulabel and packaging label were not euplained in the IFU |  | 2023-09-26 08:22 | No
1247 | Yes | 0/u FD0unable euposure | 0R/u FD0unable euposure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Neut Step Action: PPD00000000210000000020000000091*1 Arranged to be sent to the venu within 72H FSE uuYonghu-86-133276613300000000AP Transfer: Case Transferred to SAP Internal Remarks: 28.00000000.2000000003 1000000003:31 uC (860000000013300000000 "RDI Accuacy: NA" $Q7:Confirmation reportedproblem, The device cannot be properly euposed. $Q8:The device cannot be euposed properly.The uLLlog indicates that the high-voltage commuication times ou. Futher check indicates that the high-voltage CuE H1 faut is blinking red and the CuE cannot be connected uing the service tool. The high voltage CuE is fauty(00000000210000000000000000319 CuE CVFD-5 ASSY). $Q9:Replace the high voltage CuE(00000000210000000020000000091 CuE CVFD-5 ASSY). $Q100000000Test and Inspection data refer to PA, Do a few patient, Observation of a certain period of time. The system is ok. Device retuned to ful fuctionality. $Q11:Calibrated tools NA. SAP Transfer: Case Transferred to SAP Euternal Remarks: 28.00000000.2000000003 1000000003:28 uC (860000000013300000000Faut phenomenon: The device cannot be euposed normally. Soluion/repair measues: The device cannot be euposed properly. The uLLlog indicated that the high-voltage commuication timed ou. Futher inspection revealed that the high-voltage CuE H1 faut light flashed red and the service tool coud not be ued to connect to the CuE.Determine the high voltage CuE failue.Replace the high voltage CuE. The equpment 000000000000000000000000001ks normally after service. T2 Activities:Repair Action: 00000000210000000020000000091 Cue CVFD 5 assy. Problem Reported by Cutomer: uable to eupose Safety Qustions: Case: 0000000022329200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: uable to eupose Detailed description of cutomer problem or faut phenomenon: Even perspective euposue 000000000000000000000000001’t 000000000000000000000000001k. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: It is invalid for the cutomer to restart the machine, Fluscopy not available reselect another application Image disk ful, deleted image is invalid, image playback is possible Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended to check the log Check the IPPC startu statu, it is recommended to replace the BIOS Check image disks smart data. If there is a problem with the hard disk, it is recommended to replace it. Reinstall IPPC software, bypass hard disk enclosue, recreate image disk If the log shows that the high-voltage generator connection failue Host(PC) is uable to commuicate via CAN with Velara Generator - Fronta, it is recommended to Disconnect the cable from NPu: 300000000 (frontal) in M-cabinet(Frontal) and connect it to NPu:300000000 ( lateral). Disconnect u 300000000 on the MPDucontrol board at the back of the main control cabinet and plu it into u 300000000 Recommended spare parts uuand designation: uu0000000080000000050000000093 IP PC Revised_II no HDD uu0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II Notes on cutomer requrements: Please contact the cutomer for on-site repair SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-26 08:26 | No
1248 | Yes | QA000000 - Device cannot turn on when rescuing the patient | QA00000000 - Device cannot turn on when rescuing the patient |  | 2023-09-26 08:42 | Yes
1249 | Yes | 0/u FD0fluoroscopy not possible | 0L/u FD0fluoroscopy not possible | BuQ&R Feedback: No SAP Cutomer Symptom Code: ACQP Cutomer Symptom Code: Imaging/Patient Recording/Acqusition Problem Reason: Corrective Maintenance Engineer Failue Code: Electrical/Electronic Failue Repair Activity Code: Reboot/Restart SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022329295 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: confirmed reported problem, fluroscopy not possible Trouleshooting Action: fluroscopy not possible, foud the IPPC POST failed Repair Action: instrut the uer to check IPPC indicator, Start the IPPC manully Service Content: If the ray cannot be emitted, check the image processing host and manully tun on the image host. Service resuts: After remote service, the equpment resues normal operation Test/Inspection: Device retuned to ful fuctionality. Test Equpment: N/A Problem Reported by Cutomer: Memory is ful Limited time manul dispatch confirmation ----------------------------------------------------------------------------- | 2023-09-26 08:56 | No
1250 | Yes | Fluoroscopy and Radiography not possible | Fluoroscopy and Radiography not possible | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000 Parts Consued : 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-09-26 09:36 | No
1251 | Yes | uis faulty (Problem solved by replacing the ulead set from cuomer stock) | uis faulty (Problem solved by repla0g the ulead set from cuomer stock) | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000096 Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Resoluion : Problem solved by replacing the ulead set from cutomer stock. Defective ucable lot no - 9-200000000Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : uis fauty and operation check failed ----------------------------------------------------------------------------- | 2023-09-26 10:11 | No
1252 | Yes | AED chirped for battery issue after pads replaced, unresolved by working battery | AED chirped for battery issue after pads replaced, unresolved by working battery | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007923 Bronte Westaway The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: CEO Device ue: defibrillation Eupected and ueupected Resuts: after pads replaced, device giving battery issu, battery 000000000000000000000000001ks in other device. uer Impact: can't ue Patient Impact: none SE none Diagnostic performed by Engineer : After pads replaced, AED giving battery requres replacement issu. Battery 000000000000000000000000001ks in other AEDs. Different good battery (not new) gives same resut in device. device is uder warranty, suplying euchange to check whether battery/device issu Follow u Requred : 1u 0000000056000000005291 Alan Mcvilly 000000007872000000005 Ravenhall Correctional Centre 97 Riding Boudary Rd Ravenhall Vic 3000000003 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003299 Parts Consued : 1.0000000000000000000000000000000056000000005291,FRu (C B) Defib, u English 3ed, Euch ----------------------------------------------------------------------------- | 2023-09-26 11:36 | No
1253 | Yes | 00-uissue | 0000-uissue | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 16335816-uissu Safety Qustions : Case: 0000000022329800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-26 11:36 | No
1254 | Yes | System does not go up | System does not go up | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005926 September 27, 2000000003 Baugarten, Rüdiger The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: RR Beate Kleber Phone 00000000000000001311720000000000000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1 Resoluion: The reported faut is confirmed. The following actions were performed for uLLcorrection: Trouleshooting in geometry, fiued interrution in the power suply. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Tools & Equpment: MuTIMETER, 2000000003-12-21 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005700000000Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: uLL# and/or description of uLL Fleu Vision dark/geometry not started Phase L1 in the M cabinet ou L1 switched on again - system still does not power u The orange LED on the power suply in the M cabinet is lit, all other LEDs do not light u Probably power suply defective Repair Action: Trouleshooting Replace power suply if necessary Problem Description by Engineer: Problem description by engineer: RR Beate Kleber Phone 00000000000000001311720000000000000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1 Diagnostic performed by Engineer: System hangs duing start-u at 00000000seconds start time Follow-u requred: uLL# and/or description of uLL Fleu Vision dark/geometry not started Phase L1 in the M cabinet ou L1 switched on again - system still does not power u The orange LED on the power suply in the M cabinet is lit, all other LEDs do not light u Probably power suply defective Repair Action: Trouleshooting Replace power suply if necessary Technician name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Cutomer comment: ## Service Ms. Kleber 27.00000000. from 9:00000000 hrs. Internal Remarks: Power suply NC in the M cabinet is off. Please clarify with FSE, on service withou parts makes sense. ----------------------------------------------------------------------------- | 2023-09-26 11:51 | No
1255 | Yes | 00utton issue | 000Button issue | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 0000000022329956 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : Buton issu SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-26 11:51 | No
1256 | Yes | uLS AED ¦ Hardware failure - pacing failure during PM | TempuLS AED ¦ Hardware failure - pa0g failure during PM | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005900000000 Problem Description by Engineer : Selhání hardwaruDiagnostic performed by Engineer : Nový přístroj Resoluion : Nový 000000005600000000000000001 RDT Tempu LS Defibrilator - AED, MAN ----------------------------------------------------------------------------- | 2023-09-26 12:01 | Yes
1257 | Yes | uLS AED ¦ hardware failure during PM | TempuLS AED ¦ hardware failure during PM | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005988 Problem Description by Engineer: Hardware failue. Diagnostic performed by Engineer: New device. ----------------------------------------------------------------------------- | 2023-09-26 12:21 | Yes
1258 | Yes | BZA0screen problem in the room: has already been replaced, so the problem persists (the screen switches off and we can | BZA0 screen problem in the room: has already been replaced, so the problem persists (the screen switches off and we can | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006919 September 28, 2000000003 MARTIN LITT The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Nuse headfeed Device ue: / Eupected and ueupected Resuts: The screen blinks sometimes. uer Impact: / Patient Impact: / SE 7.2.9 Resoluion: Replacement of the KVM transmitter (compuer side) Test OK Fiued problem Parts Consued: 1.00000000000000000000000000000000230000000016812, KVM TRANSMITTER (u Cab) ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008300000000 ANTOINE SANTANGELO The system meets the specifications for the performed service and is retuned to ue. Follow u Requred: Martin onsite 11/29 9-10000000000000000 Part: 00000000230000000016812 Delivery: Hospital Problem Description by Engineer: Cutomer Fuction/Role: Nuse headfeed Device ue: / Eupected and ueupected Resuts: The screen blinks sometimes. uer Impact: / Patient Impact: / SE 7.2.9 Diagnostic performed by Engineer: Continung trooleshooting by replacing MAV ----------------------------------------------------------------------------- | 2023-09-26 12:41 | No
1259 | Yes | 0/u FD0startup issue | 0R/u FD0startup issue | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 29.00000000.2000000003 1000000007:59 uC (86000000006252) Service content (requred): [Replace the host hard disk and reinstall the software, and re-do related calibrations.] Service resuts (requred): [The equpment can be ued normally after the repair is completed.] SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022330000000082 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Qingfeng Gu Last udated by: Qingfeng Gu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Neut Step Action: PPD FRuPFS3900000000(WD Veloc. SW Disk) 00000000800000000163852 1 This nuber has been confirmed by RSE and can be issud Pui-86-17712333138 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Boot failue Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: uable to enter window interface Remote diagnostics and test resuts: Supected system hard drive failue PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: It is recommended that engineers come to the site to replace the system hard disk and reinstall the software. Recommended spare parts uuand designation: "PPD: 00000000230000000028792 FRuPFS3900000000(WD Veloc. SW Disk)" SAP Transfer: Case Transferred to SAP Internal Remarks: 29.00000000.2000000003 1000000008:00000000 uC (86000000006252) RDI Accuacy: "yes" $Q7: (Confirm the reported problem & it can't boot u.) $Q8: (Check the hard disk, it was defective.) $Q9: (Replace the harddisk and reinstall the whole system. Then redo gererator and detector adjutments.) $Q100000000 (Test and Inspection data refer to PA & Device retuned to ful fuctionality) $Q11: (Calibration tools refer to PA) Problem Reported by Cutomer: uable to boot FSE Pucai 000000007712333138 call ----------------------------------------------------------------------------- | 2023-09-26 13:46 | No
1260 | Yes | 00 CN00- System battery uLLin Rumah Sakit Mayapada Bandung. | u- System battery uLLin Rumah Sakit Mayapada Bandung. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008827 Rahmat Kasyfu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Confirmed reported problem the battery was uLL Cutomer Fuction/Role: biomed engineer Device ue: no Eupected and ueupected Resuts: n/a uer Impact: no Patient Impact: no SE Rev D Resoluion : cross check with another battery, the device is retuned to fuly fuctionality sugest to replace the battery ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 ----------------------------------------------------------------------------- | 2023-09-26 13:56 | No
1261 | Yes | The installed udoes not pass the self-test | The installed udoes not pass the self-test | BuQ&R Feedback: No SAP Cutomer Symptom Code: TSTF Cutomer Symptom Code: Test Failue Diagnostics Reason: Corrective Maintenance Engineer Failue Code: Connection/Cable/Interconnect Repair Activity Code: Instruted Cutomer Problem Reported by Cutomer: 00000000N/uDFM100000000 Defibrillator/Monitor SAP Transfer: Case Transferred to SAP OneEMS internal remarks: When the agent installed the machine, it was foud that one DFM100000000 failed the self-test u and the other self-tests of the same batch were normal. If youwant to apply for a complete machine replacement, it is recommended that the agent go throuh the replacement process before sales, and close the order first. Safety Qustions: Case: 0000000022330000000099 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-26 14:01 | No
1262 | Yes | Equipment blocked at startup 0:0 u FD0 | Equipment blocked at startup 0:0 u FD0 | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006583 Dragosmin Bogdan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: dr Device ue: no Eupected and ueupected Resuts: FV not booting u uer Impact: low Patient Impact: no SE 8.2.17 Diagnosis performed by Engineer: It is reported that the equpment remains blocked when starting at "00000000:00000000". The Fleuvision compuer stuk in the boot loop is identified. Manul closing and starting is recommended. The equpment is fuctional and will be kept uder observation. ----------------------------------------------------------------------------- | 2023-09-26 14:06 | No
1263 | Yes | Alarms/Delayed Alarm: Tele0had a 0ec pau in alarm/Alarm telemetry faulty. NTF/ur Misinterpretation. Provide Info | Alarms/Delayed Alarm: Tele0 had a 0sec pau in alarm/Alarm telemetry faulty. NTF/ur Misinterpretation. Provide Info | 000000000000000000000000001O##: 000000000000000000000000001000000000000000063500000000Christiane Atwill The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: yes Eupected and ueupected Resuts: Asystolic not recognized in time uer Impact: no Patient Impact: no SE B.00000000.18 Resoluion: The logs of the system were saved and the configuation was checked. Alarm strips were made available. In the auit log, the following alarms are saved for the incident on 9/26: 26.00000000.2000000003 00000000:11:00000000 Tele200000000 Paue generated at 00000000:11:00000000. (HF = 00000000 26.00000000.2000000003 00000000:11:00000000 Tele200000000*HF 32 <35 generated at 00000000:11:00000000. (HF = 32) 26.00000000.2000000003 00000000:11:00000000 Tele200000000systlie generated at 00000000:11:00000000. (HF = 00000000 A report for the reported incident was created and forwarded to the factory for euamination. A separate report from the factory will then be created separately. ----------------------------------------------------------------------------- Cutomer comment: u Hofmann is on site. The problem is that the uers coincidentally discovered that the *alarm paue was recorded in the alarm log >00000000c in the alarm strip. 18.9.23 000000009:18 Tele18 00000000ec 26.9.23 3:11:00000000 Tele200000000Sek Alarm strips are present. There was no negative effect on the patient. It was discovered coincidentally and escalated to the management. Alarm strips requsted from MedSup. Cutomer comment: Alarm strips received (asysttolic 26.9. Tele 200000000Cutomer comment: Dear u Hofmann, Thank youfor the stripes. I will forward everything to ou factory for euamination. The alarm sequnce was*Paue (00000000:11:00000000), **HF low (00000000:11:00000000), uon asystlie (00000000:11:00000000 = 3 Sek after **HF low) How euactly is the alarm strip behavior uknown to u in Remote? The higher alarm overwrites the lower alarm, bu it is not clear to u whether the alarm strip label shoud be converted. No line or alarm is missing on you end since the asysttolic strip is also present. This is almost identical to the 1st paue strip, eucept for the 5 Sek at the beginning. For everything else abou timing, we have to wait for the report from the factory. If youhave futher qustions, please let me know. Kind regards, Christiane Atwill Cutomer comment: Andreas Rupel by ## Here is the form and 3 images from the control center. More data is ufortuately not available to me (SN: 2u810000000062) Internal Remarks: Contact data Reporting uer: 676675 Contact person: Med.2.3 Tel. No.: 2200000000Cutomer comment: Screenshots of the konfig and logs puled with u Hofmann on site via remote. Internal Remarks: System information is not saved Internal Remarks: CFMS created Cutomer comment: Nr - no answer. only annoucement Cutomer comment: It goes via Tustall to the alarm interface. ----------------------------------------------------------------------------- | 2023-09-26 14:51 | No
1264 | Yes | intermittent system not starting up - split | intermittent system not starting up - split |  | 2023-09-26 15:15 | No
1265 | Yes | 00 CN00 - Battery test failed of OP check | u- Battery test failed of OP check | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 28.00000000.2000000003 00000000:0000000056 uC (860000000000000000100000000 $ASP engineer: [Honglejia huhang] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem：Incompatible battery $Q8:Incompatible battery detected on site, machine cannot recognize battery $Q9: Disconnect the machine and reconnect the battery cable. After charging, perform an operation check. The operation check is passed, restart the device, and the device retuns to normal $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality $Q11: Calibrated Tools name (PROSIM 8) S/N (26758200000000 Neut Calibration date (10000000018/2000000003); Calibrated Tools name (ESA6200000000 S/N (3000000002866) Neut Calibration date (3/2/20000000000000000 Calibrated Tools name (70000000000000000P) S/N (2953000000006) Neut Calibration date (10000000029/2000000003); Neut Step Action: DR.pan 00000000898588700000000 call in Meitan Couty Traditional Chinese Medicine Hospital, Zuyi City, GuzhouProvince SAP Transfer: Case Transferred to SAP Euternal Remarks: 28.00000000.2000000003 00000000:0000000053 uC (860000000000000000100000000Department information: [Department: Emergency Department Contact Person: Teacher WuContact Nuber: 15285231772]* $ASP engineer: [Honglejia huhang] $Q1:N $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem：Incompatible battery $Q8:Incompatible battery detected on site, machine cannot recognize battery $Q9: Disconnect the machine and reconnect the battery cable. After charging, perform an operation check. The operation check is passed, restart the device, and the device retuns to normal $Q100000000Test and inspection data refer to PA checklist, Device retuned to ful fuctionality $Q11: Calibrated Tools name (PROSIM 8) S/N (26758200000000 Neut Calibration date (10000000018/2000000003); Calibrated Tools name (ESA6200000000 S/N (3000000002866) Neut Calibration date (3/2/20000000000000000 Calibrated Tools name (70000000000000000P) S/N (2953000000006) Neut Calibration date (10000000029/2000000003); Problem Reported by Cutomer: Battery failue reported SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: ASP Honglejia SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Operation check battery test failed Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: uplu and insert the battery and re-ru the operation check to trouleshoot the battery. Safety Qustions: Case: 0000000022331300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-09-26 15:41 | No
1266 | Yes | Order MS_u0SSEMBLY CBL u0P0peaker Assembly-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Order MS_u0ASSEMBLY CBL u0MP0Speaker Assembly-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000003 Bittlinsky, Simon The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: Nemo KVA Recipient/## see order/order BLIV / RSM: No Nuber of items/uuAT: 1u 0000000056000000008621 MS_u2 ASSY CBL u2/MP2 Speaker Assembly / no Other:- Diagnostic performed by Engineer: Nemo KVA Cutomer orders spare part for u2: 1u 0000000056000000008621 MS_u2 ASSY CBL u2/MP2 Speaker Assembly Did the louspeaker have no soud or coud a light, possibly distorted soud or a beeping soud be heard? • No soud, since contact on the louspeaker is defective Were any INOP uLLmessages displayed? If an INOP uLLwas displayed, was auio present duing this message? •No; uLLwas detected with device in partially disassembled state. Previouly there was no indication of a louspeaker problem. Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • on a host monitor Mu-500000000 When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? • The device was in the 000000000000000000000000001kshop becaue of another reported defect (NBP pup modue). Duing the repair 000000000000000000000000001k, a torn cable was noticed on the louspeaker uit (bust black caster with soldered contact). No louspeaker uLLwas reported before. Which software was installed on the device when the louspeaker malfuction occured? • L.00000000.15 Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: Parts Order uer Impact: no Patient Impact: no SE uknown ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000 Parts Consued: 1.0000000000000000000000000000000056000000008621.MS_u2 ASSEMBLY CBL u2/MP2 Speaker Assembly ----------------------------------------------------------------------------- Cutomer comment: Confirmed by ## Cutomer comment: How will it continu? - Estimate - Spare parts delivery Processing nuber 000000002233700000000 Kind regards, Eckart Helterhoff Cutomer comment: Cutomer was informed by ##that offer creation and transmission were assigned again. Cutomer comment: Sent an ##to the cutomer. Cutomer comment: Via ##sender<Eckart.Helterhoff00esudheitnord.de> Good day I answered the qustions below. S.a. Plant. Answers: Did the louspeaker have no soud or coud a light, possibly distorted soud or a beeping soud be heard? • No soud, since contact on the louspeaker is defective Were any INOP uLLmessages displayed? If an INOP uLLwas displayed, was auio present duing this message? •No; uLLwas detected with device in partially disassembled state. Previouly there was no indication of a louspeaker problem. Did the device 000000000000000000000000001k as an independent device or was it connected to a host monitor? • on a host monitor Mu-500000000 When was the problem with the louspeaker malfuction detected? Duing the setu, maintenance or duing application on the patient? • The device was in the 000000000000000000000000001kshop becaue of another reported defect (NBP pup modue). Duing the repair 000000000000000000000000001k, a torn cable was noticed on the louspeaker uit (bust black caster with soldered contact). No louspeaker uLLwas reported before. Which software was installed on the device when the louspeaker malfuction occured? • L.00000000.15 For futher qustions, youcan contact u again. Internal Remarks: Please create and send the requsted qute ----------------------------------------------------------------------------- | 2023-09-26 15:51 | No
1267 | Yes | 000- Paddle cables in poor condition | 0000- Paddle cables in poor condition | 000000000000000000000000001O##: 00000000000000000000000000100000000112600000000 200000000ov-2000000003 Hidalgo Flores, Lenin Otton The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Talk to suervisor Marc Trilla 9322750000000000000000search for 38322 Device ue: Not in clinical ue Eupected and ueupected Resuts: Blade cables in poor condition making it difficut to euchange blade cables patches uer Impact: None Patient Impact: None SE Not apply Resoluion: Bezel is received, installed to protect euternal connection. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007559 29-Sep-2000000003 Hidalgo Flores, Lenin Otton The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Part requred 98980000000019000000001 Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Talk to suervisor Marc Trilla 9322750000000000000000search for 38322 Device ue: Not in clinical ue Eupected and ueupected Resuts: Blade cables in poor condition making it difficut to euchange blade cables patches uer Impact: None Patient Impact: None SE Not apply Diagnostic performed by Engineer: Test cable protection requred Parts Consued: 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006700000000 Rau Cordero Calderon The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Follow u Requred: Talk to suervisor Marc Trilla 9322750000000000000000search for 38322 and he says that the paddle connector on the DFM100000000 is twisted and cannot properly connect the euternal paddles and the electrode cable. Assign FSE. Order part: 0000000056000000008971 - DFM100000000 ASSY Therapy Receptacle uH-E00000000-00000000-000000001 - Cardiac Hemodynamics Room hA Hybrid Problem Description by Engineer: Cutomer Fuction/Role: Talk to suervisor Marc Trilla 9322750000000000000000search for 38322 Device ue: Not in clinical ue Eupected and ueupected Resuts: Blade cables in poor condition making it difficut to euchange blade cables patches uer Impact: None Patient Impact: None SE Not apply Diagnostic performed by Engineer: Talk to suervisor Marc Trilla 9322750000000000000000searches for 38322 and says that the blade connector on the DFM100000000 is twisted and cannot connect the euternal blades and the electrode cable properly. Assign FSE. Order part: 0000000056000000008971 - DFM100000000 ASSY Therapy Receptacle uH-E00000000-00000000-000000001 - HSe speaks with suervisor Marc Trilla 9322750000000000000000searches for 38322 and says that the paddle connector on the DFM100000000 is twisted and cannot connect well either the euternal paddles and the electrode cable. Assign FSE. Order part: 0000000056000000008971 - DFM100000000 ASSY Therapy Receptacle uH-E00000000-00000000-000000001 - Cardiac Hemodynamics Room hA Hybrid Cardiac emodynamics Room hA Hybrid ----------------------------------------------------------------------------- | 2023-09-26 16:01 | No
1268 | Yes | Audio & Sounds/M00 Loudspeaker fault/Unknown | Audio & Sounds/M00: Loudspeaker fault/Unknown | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006889 Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Sending part(s) for replacement. Follow u Requred: Quntity(s): u1 per item Delivery: biomedical department Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer wants to order one or more parts. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007572 Parts Consued: 1.0000000000000000000000000000000056000000008621, MS_u2 ASSY CBL u2/MP2 Speaker Assembly ----------------------------------------------------------------------------- | 2023-09-26 16:01 | No
1269 | Yes | Alarm light ringing - device cannot be started | Alarm light ringing - device cannot be started |  | 2023-09-26 16:06 | No
1270 | Yes | System cannot acquire image | System cannot acquire image | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005238 ROSDI SHAM BIN ALIAS The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: radiographer Device ue: u u per FD 100000000100000000 Eupected and ueupected Resuts: Aqure no image uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : restart the system. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000111213 00000000-Oct-2000000003 MOHD FAIZSHADI ISHAK The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Radiographer. Device ue: Cutomer acqure image duing cases. Eupected and ueupected Resuts: Eupected image can be acqure bu cannot and uLL‘fluroscopy not possible, reselect application’ shown on screen. uer Impact: No impact. Patient Impact: impact. SE 8. Resoluion : Confirmed reported problem. System cannot acqure image. New IPPC, 2 image hardisk and Fan assembly for M-cabinet arrived and replaced accordingly. Recreated image storage database for lateral and frontal in term to make system fuly fuctional. System tested pass. No problem remain usolved. System retuned to ful fuctionality. Cutomer requst to standby for neut day. Parts Consued : 1.00000000000000000000000000000000800000000000000003352,SERVICE PART FAN uIT 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000520000000026-Sep-2000000003 ROSDI SHAM BIN ALIAS The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: u u’Per FD 100000000100000000 Eupected and ueupected Resuts: System cannot fluro/aqure uer Impact: none Patient Impact: none SE Follow u Requred : To replace ippc Diagnostic performed by Engineer : Checked and foud that lateral IPPC fauty.Need to replace ----------------------------------------------------------------------------- | 2023-09-26 16:16 | No
1271 | Yes | urhythm changed to a flat line. Alarm indicated Asystole. No Trouble Found-Evaluation Confirmed Device Working. | urhythm changed to a flat line. Alarm indicated Asystole. No Trouble Found-Evaluation Confirmed Device Working. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007937 Cumings, Craig The system meets the specification for the performed service and is retuned to ue Resoluion : The device and truk cable were inspected by Philps FSE onsite on 9/25. FSE monitored cases the week of 9/25. Ou FSE was uable to reprodue this issu. No futher issus were reported by ou cutomer. 1000000003000000002000000003 this case was closed Diagnostic performed by Engineer : FSE onsite the Monday after this was reported. FSE was uable to reprodue. FSE reported that u3, Hemo App and utruk cable are proper and properly connected. FSE was monitored for the week of 9/25/2000000003. No futher issu since the 9/21/2000000003 report. Case closed 1000000003000000002000000003. Problem Description by Engineer : Cutomer Fuction/Role: clinical manager Device ue: monitoring patients duing invasive procedue Eupected and ueupected Resuts: Cutomer reports the u rhythm line on the monitor screen changed to a flat line duing a case The arrhythmia detection alarm souded The INOPS alert message read asyst0000000000000000owever, ou cutomer reports the patient was talking. uer Impact: The EP2 Lab team reports they changed the uelectrodes (patches) and ulead wires and then changed the utruk cable ou twice (2 times) with 2 different known 000000000000000000000000001king truk cables before the u the issu was resolved - that the flat line was back to an u rhythm. Patient Impact: nothing was reported from ou cutomer SE Hemo App 1.3.1 ----------------------------------------------------------------------------- | 2023-09-26 17:06 | No
1272 | Yes | 00 CN000- Monitor is not switching on .. | u- Monitor is not switching on .. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000Gnanasekar, D The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Battery Got aged, Cutomer to procue the new battery. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : battery got aged, not powering with only battery. ----------------------------------------------------------------------------- | 2023-09-26 17:16 | No
1273 | Yes | Loudspeaker failure and audio failure cleared by audio test. | Loudspeaker failure and audio failure cleared by audio test. | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000789000000009-Sep-2000000003 Of Santos, Angel The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Constant monitor Eupected and ueupected Resuts: Louspeaker failue. uer Impact: Impossible to hear alarms. Patient Impact: The equpment is ued with limitations. SE Diagnostic performed by Engineer: Auio test is performed. The uLLis deleted. Verification that the speaker 000000000000000000000000001ks correctly. Verification of correct operation of the system. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007778 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Constant monitor Eupected and ueupected Resuts: Louspeaker failue. uer Impact: Impossible to hear alarms. Patient Impact: The equpment is ued with limitations. SE Diagnostic performed by Engineer: The client (MARIA TERESA MALAGA eut: 82500000000says that the equpment shows auio uLLand speaker failue. Material order is handled and the case is transferred onsite so that the equpment is checked by the field service. Resoluion: The case is transferred onsite. Follow u Requred: The client (MARIA TERESA MALAGA eut:82500000000says that the equpment shows auio uLLand speaker failue. - Equpment located in /PL. 00000000th FLOOR GENERAL RESIDENCE/uOLOGY FLOOR - assign Field 12:00000000 A.M. - Order the mentioned material ----------------------------------------------------------------------------- | 2023-09-26 18:01 | No
1274 | Yes | Problems with the anode, it does not produce u-rays | Problems with the anode, it does not produce u-rays |  | 2023-09-26 18:01 | No
1275 | Yes | 00T0ot booting up | 00CDT0not booting up | 000000000000000000000000001O##: 0000000000000000000000000010000000015000000001 00000000Oct-2000000003 Chase, Brian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: clinical Eupected and ueupected Resuts: will not boot u uer Impact: uable to ue Patient Impact: none SE Resoluion : Performed adaptation, yield, EDL and detector cals for frontal and lateral. GHS PACS added. Ghost backu performed. Retun to patient ue. Tools & Equpment : DOSE-KVP-LIGHT METER,2000000007-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000008-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001180000000000000000Oct-2000000003 Chase, Brian The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Final calibrations following software reload. Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: clinical Eupected and ueupected Resuts: will not boot u uer Impact: uable to ue Patient Impact: none SE Diagnostic performed by Engineer : Replaced ethernet card and commuication restored. Ful software configuation and reload of system completed. Additional calibrations and net000000000000000000000000001k verification to be performed neut BD. Parts Consued : 1.00000000000000000000000000000000230000000028771,FRuPFS388 (INTEL ETHERN.CARD) Tools & Equpment : MuTIMETER,200000000000000008-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000162 00000000-Oct-2000000003 Chase, Brian The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: clinical Eupected and ueupected Resuts: will not boot u uer Impact: uable to ue Patient Impact: none SE Diagnostic performed by Engineer : Replaced SBC and drives. System booting now however commuication problems preventing advancement in repairs. Follow u Requred : Continud 000000000000000000000000001k to resolve power u commuication problems. Parts Consued : 1.00000000000000000000000000000000800000000163853,PFS-00000000 CFG2 HDD 1.00000000000000000000000000000000230000000028782,Service part Single Board Compuer 1.00000000000000000000000000000000230000000028800000000,PFS391 PCBou backu disc Tools & Equpment : MuTIMETER,200000000000000008-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 28-Sep-2000000003 Stoy, John The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System will not boot u When was the first occurence? Within... : 26-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System will not boot u uer Impact: System is down Patient Impact: None/ no procedue SE 7.2.16.00000000Diagnostic performed by Engineer : Replaced MDucontrol board which temporarily seem to solve issu. Once booted, system became ustable (Application).SBC has intermittent VGA video ou. Plan is to replace SBC/hard drives/ reload system software. Case has been escalated to involve onsite NSS. Follow u 000000000000000000000000001O#created for secondary FSE. Tools & Equpment : MuTIMETER,2000000005-0000000030000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000920000000028-Sep-2000000003 Chase, Brian The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: operator Device ue: clinical Eupected and ueupected Resuts: dulicate 000000000000000000000000001seems to been created uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Dulicate 000000000000000000000000001O# Issu is being address on 0000000000000000000000000010000000000000000800000000. Closing dulicate. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008186 Whitaker, Steven The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System will not boot u When was the first occurence? Within... : 26-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: System will not boot u uer Impact: System is down Patient Impact: None/ no procedue SE 7.2.16.00000000Diagnostic performed by Engineer : On call case, called Simon back and he reports that the system will not boot and he is getting a no connection message on the Data monitor. The led blinks on the review modue bu no splash screen, rebooting did not resolve. Onsite was requsted, LOD was not available Internal Comments : uLL# and/or description of uLL: Not available Repair Action : Check power and fues, check SBC and HDD's, reseat , reload ghost if possible or replace defective parts foud Follow u Requred : uLL# and/or description of uLL: Not available Repair Action : Check power and fues, check SBC and HDD's, reseat , reload ghost if possible or replace defective parts foud Technician Name : Whitaker, Steven ----------------------------------------------------------------------------- Internal Remarks : Assisted with software loading after SBC and disk replacement. Internal Remarks : Interaction Type: New Case Caller: Simon Chayka IP & Issu: 555356 not booting u Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Case: 0000000022338500000000 FSE: Stoy, John System SW is locking u intermittently after loss of power. FSE states there no known good Ghost Back u becaue system was being maintained by third party. Instruted FSE to create backus and perform OS and App scratch load. ----------------------------------------------------------------------------- | 2023-09-26 18:01 | No
1276 | Yes | 00 CN00- FCO00 Failing Audio Test | u- FCO000 Failing Audio Test | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000800000000 00000000-Oct-2000000003 Andrew Hammond The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Field Service Engineer Device ue: not in ue Eupected and ueupected Resuts: The device is failing the auio test:- no soud is coming from the machine. Issud noted duing the implementation of FCO86100000000236A. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : u PCA Assembly replaced which fiued the faut, T&V completed, Device retuned to ue Parts Consued : 1.000000000000000000000000000000005600000000100000000,Intrepid u PCA ASSY Tools & Equpment : DEFIBRILLATOR ANALYZER,200000000000000002-21 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-11-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008337 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Engineer to attend site with the following parts for investigation and repair: 1u Intrepid Processor PCA ASSY* 0000000056000000001751 1u Intrepid u PCA ASSY* 000000005600000000100000000 1u Intrepid Speaker ASSY 0000000056000000001531 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Field Service Engineer Device ue: not in ue Eupected and ueupected Resuts: The device is failing the auio test:- no soud is coming from the machine. Issud noted duing the implementation of FCO86100000000236A. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Engineer to attend site with the following parts for investigation and repair: 1u Intrepid Processor PCA ASSY* 0000000056000000001751 1u Intrepid u PCA ASSY* 000000005600000000100000000 1u Intrepid Speaker ASSY 0000000056000000001531 ----------------------------------------------------------------------------- | 2023-09-26 18:06 | No
1277 | Yes | ur uLL RCP test was not verified (No fault found) | ur uLL RCP test was not verified (No fault found) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000093 FALASCA, IVO The system meets the specification for the performed service and is retuned to ue. Resoluion: Technician Alberto Peconio carried ou the fuctional test connecting the cable of the RCP sensor, with a positive oucome (Test passed). I close the call. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Reported sensor-Rcp gude test: not verified. The uer made a mistake when performing the fuctional test. uer Impact: no Patient Impact: no SE na ----------------------------------------------------------------------------- Internal Remarks: The technical report of the euecuion of the requsted intervention mut be necessarily transmitted to the following e-mails: rapportinidilavorosic00slbat.it; sic.aslbat00cspa.it; reporting the following requst details: 2000000003/00000000389 ----------------------------------------------------------------------------- | 2023-09-26 18:16 | No
1278 | Yes | 00 CN00- Power Supply uLLmessage | u- Power Supply uLLmessage | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Assad MAJEED The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: On-site intervention requred. Follow u Requred: Qute for on-site intervention for replacement of the Processor Board. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that the defibrillator displays a blu technical message at the top right of the screen: "power suply malfuction" and "faut detected on a non-critical device, maintenance requred". We carried ou a fuctional test which was sucessfu and the technical messages have disappeared. Regarding the analysis of the logs, it seems that the faut was located at the level of the "Clock Battery" of the Processor card. It is recommended to replace the Processor card. The cutomer wishes to carry ou an on-site intervention, becaue retuning to the 000000000000000000000000001kshop is not possible for the service. ----------------------------------------------------------------------------- Internal Remarks: Relauch to follow u on this file, I send a message via teams PCMS ----------------------------------------------------------------------------- | 2023-09-26 18:26 | No
1279 | Yes | Check Vent uLLMessage - Prouimal pressure sensor | Check Vent uLLMessage - Prouimal pressure sensor | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008719 00000000Oct-2000000003 Khouam, Fady The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Respiratory Manger. Device ue: ouside of ue Eupected and ueupected Resuts: Check Vent uLLMessage - Prouimal pressue sensor uer Impact: None. Patient Impact: None. SE 3.200000000Diagnostic performed by Engineer : The FSE puled the Event logs Checked it for uLL, nothing foud. Performed a preoperational check , no uLL foud. Prepared the device for patient ue. Device placed back into service and ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008627 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Check Vent uLLMessage - Prouimal pressue sensor uer Impact: system restricted Patient Impact: none reported SE uknown Follow u Requred : Cutomer states device was reported to have Check Vent uLLMessage - Prouimal pressue sensor. Cutomer is uable to verify in event logs. Resoluion : Creating on site service for FSE to repair device. Diagnostic performed by Engineer : Cutomer states device was reported to have Check Vent uLLMessage - Prouimal pressue sensor. Cutomer is uable to verify in event logs. Cutomer is requsting on site service to have FSE on site to repair device. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: David Gordon IP & Issu: 8600000000000000005/ Check Vent uLLMessage - Prouimal pressue sensor Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote W/O Roued to RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-26 18:36 | No
1280 | Yes | Alarms or Messages/ BP alarm rings only once & then it silences even if the value is out of range/Configuration | Alarms or Messages/ BP alarm rings only once & then it silences even if the value is out of range/Configuration | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001000000009 27-Sep-2000000003 Stoppielli, Lorenzo The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Performed alarm configuation as requsted by the department on the following monitors: inv. 69330000000069331-68700000000870000000068732-68719-69386-69365-68731-68717-68735-69300000000-69300000000-68738. The activity became necessary following the transfer of the monitors carried ou by hospital staff from the su intensive department to Intensive Care. Remember that in case of transfer, the monitors mut be configued according to the modes of ue therefore we ask to involve u in order to avoid safety problems related to the fuctioning of the alarms. The systems were left 000000000000000000000000001king and configued according to the requrements reported by the department. No faut of the u system ----------------------------------------------------------------------------- | 2023-09-26 19:01 | No
1281 | Yes | M0A - u00 Functional test failed | u Funuonal test failed | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009000000003 29-Sep-2000000003 Laezza Francesco The system meets the specification for the performed service and is retuned to ue. Resoluion: After analyzing the logs of the defibrillator it is reported that the malfuction and therefore the failue of the fuctional test was du to a wrong energy select by the operator duing the fuctional test. The correct fuctioning was restored. The internal tests and tests were carried ou with the positive oucome. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Rosario Lauita Device ue: Diagnosis Eupected and ueupected Resuts: problem with calibration uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na Clinical Ems Rosario Lauita reports calibration problems. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003187 MELONI, FABRIZIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Rosario Lauita Device ue: Diagnosis Eupected and ueupected Resuts: problem with calibration uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na Clinical Ems Rosario Lauita reports calibration problems. Diagnostic performed by Engineer: Clinical Engineering Rosario Lauita reports calibration problems, they requst the presence of a technician in order to recalibrate the def and then carry ou the fuctional tests, the system is uder contract, assign on site, equpped with sutable tools. Follow u Requred: Clinical Engineering Rosario Lauita reports calibration problems, they requst the presence of the technician in order to recalibrate the def and then carry ou the fuctional tests, the system is uder contract, assign on site, equpped with sutable tools. ----------------------------------------------------------------------------- | 2023-09-26 19:06 | No
1282 | Yes | Device will not power on | Device will not power on | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009100000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Parts/Suplies Information Provided Part ID and Price: Provided cutomer with the part # for the V600000000power suply Part # 00000000561512591 Rp - Power Suply, V600000000List price is $8900000000 Page 1600000000of version T service manul for referencing the repair. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Parts ID uer Impact: No impact Patient Impact: No impact SE uknown Problem Description/ueupected Resuts: Parts ID for V600000000power suply. Cutomer is reporting the uit will not power on. Diagnostic performed by Engineer : Problem Description: Parts ID Form ued: Parts Trouleshooting Steps Taken: Parts ID for V600000000power suply. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kerlin guman IP & Issu: 6000000002700000000/bad power suply Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-26 19:46 | No
1283 | Yes | Connector problem | Conneur problem | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009359 October 16, 2000000003 LAVENANT, Dorian The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The bolt was fastened becaue the port was moving. Operational and electrical tests were then performed. Resoluion: The defibrillator is 000000000000000000000000001king again and the tests were validated. Parts Consued: 1.0000000000000000000000000000000056000000008971, DFM100000000 ASSY Therapy Receptacle Tools and Equpment: PATIENT SIMuATOR, 200000000000000001-12 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000001-12 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000071700000000Assad MAJEED The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: On-site intervention requred. Follow u Requred: Qute. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that the therapy connector moved a lot. The cutomer wants on-site service to replace the therapy connector. ----------------------------------------------------------------------------- | 2023-09-26 19:51 | No
1284 | Yes | motorized movement unavailable | motorized movement unavailable | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008573 Vilkaukas Viktorija The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : motorized movement uavailable When was the first occurence? Within... : 26-Sep-2000000003 How often is issu occuring? : uknown Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.2 Diagnostic performed by Engineer : Jason is reporting that the system has a message "motorized movement uavailable". System is on PRS bu not reachable to retrieve any logfiles. Internal Comments : uLL# and/or description of uLL: "motorized movement uavailable" Repair Action : Jason is reporting that the system has a message "motorized movement uavailable". System is on PRS bu not reachable to retrieve any logfiles. Please check system logs to determine the root caue of the issu. Follow u Requred : uLL# and/or description of uLL: "motorized movement uavailable" Repair Action : Jason is reporting that the system has a message "motorized movement uavailable". System is on PRS bu not reachable to retrieve any logfiles. Please check system logs to determine the root caue of the issu. Technician Name : Vilkaukas Viktorija ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000900000000 26-Sep-2000000003 Rowles Michael The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : motorized movement uavailable When was the first occurence? Within... : 26-Sep-2000000003 How often is issu occuring? : uknown Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.2 Resoluion : System reported a speed controller uLL Asked the bio med to check to see if it was pressed on startu. He checked it was, he hug u the speed controller and rebooted and the issu was resolved. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Derek on behalf of Jason Boenig IP & Issu: motorized movement uavailable 58269936 Event Type: Incident Service Requst: onsite Did Cutomer Decline Remote Suport? yes Preferred Service Time: held for 9/26 Entitlement Options: entitled Reason for Entitlement Claim: N/A Action Taken: Created O000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-26 20:06 | No
1285 | Yes | M0A / u00/ Discharge button could not pop out normally (PSA / QA#000000. | M00A / u000/ Discharge button could not pop out normally (PSA / QA#00000000. |  | 2023-09-26 20:12 | No
1286 | Yes | uLS: DPM Hardware failure uLLmessage displayed during pacing training. | TempuLS: DPM Hardware failure uLLmessage displayed during pa0g training. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009386 Wilbu, Phillip The system meets the specification for the performed service and is retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Place Material 000000000000000000000000001for euchange uit Follow u Requred : Please Tu LIVINGSTON FIRE & RESCu Attention: Josh Thabalowski 00000000E Callender St Livingston MT 5900000000-2700000000 Internal Comments : T000000000000000000000000001failed attempts have been made to requst additional information from cutomer. Problem Description by Engineer : Cutomer Fuction/Role:Sales Device ue:Not in Clinical ue Eupected and ueupected Resuts: DPM Hardware failue uLLmessage displayed duing training. uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: DPM Hardware failue uLLmessage displayed duing training. Diagnostic performed by Engineer : Problem Description: DPM Hardware failue uLLmessage displayed duing training. Tempu LS: DPM Hardware failue uLLmessage displayed duing pacing training. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Sales Rep called to report Tempu LS: DPM Hardware failue uLLmessage displayed duing training. Cutomer refued to send log files. Cutomer says log files can be puled by u. The following information was requsted from cutomer: *Log and Rescu Files: Cutomer refued *Date of event? Information requsted. *Brief description of the event? DPM Hardware failue uLLmessage displayed duing pacing training. *Operator of the produt: Paramedic *Did the uit display any uLLmessages or warnings before the incident? DPM Hardware failue uLLmessage *Was the Patient or uer injued or caued any harm? NA *Patient Oucome? NA *Age of Patient? NA *Patient Gender? NA Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009700000000 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Josh 913-300000000-6316 firechief00ivingstonmontan.org IP & Issu: DPM failue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-26 20:16 | No
1287 | Yes | The equipment is working but it does not record. | The equipment is working but it does not record. | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022332931 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sthefany Pereira Last udated by: Sthefany Pereira 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: The machine is 000000000000000000000000001king bu does not record. Daniel Caetano de Faria Sector: Surinentos Tel: 61 38782831 Tel: 61 9997578100000000 daniel.faria00tcv.com.br CTCV - TREATMENT CENTER Address: Q SGAS QuDRA 613 CONJuTO C SN BRASILIA-DF 70000000000000000-7300000000 BR Repair Action: ESF Visit Schedued with client : yes Schedued date/time : 10000000011 from 1pm to 5pm Schedued date at the cutomer's requst? No SAP/1EMS priority changed? No Estimated duation : 00000000 Need machine stopped? yes Done smart schedue? Yes Name of ESF : Leonardo Reis FSE trained according to qulification 000000000000000000000000001ksheet? Yes SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.1000000002000000003 12:22:00000000 POLARIS (POLARIS) ========== Leonardo Reis - 0000000000000000573936330000000030000000000000000000000005 ========================== Cutomer visit. Requred parts were not available. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP OneEMS internal remarks: Problem description by engineer : Ou of space message What steps are taken duing the system diagnosis : Daniel informs that the equpment is not freeing u space to perform euams and presents an uLLmessage: Euposue not possible. Image disk ful; Checked in the Image disk 2 POST failue log; Reversed image disks 1 with 2 and continud failue on image disk 2; NOT OK Performed recreate and verification of the image store; OK IPPC and PC HOST check, STEP 1, 2 and 6 fail. Displays uLLon Image disk 2; NOT OK After restarting the equpment, it freed u more than 95,0000000000000000space for images; Performed image acqusition tests in Cardio and Vascuar and it 000000000000000000000000001ked normally. The client is advised to perform euams and at the end of each euam, save the entire euam, as there is a risk of the equpment failing du to lack of space; CONTACT: Daniel 61 9997578100000000 uLL# and/or description of uLL: IPu Euposue not possible. Image disk ful IPu Free disk space is too low. Repair Action "Internal" : Equpment malfuctioning on image disk 2; Requsted PART, euchange of IPPC; What are the neut Trouleshooting steps : Please schedue FSE + PART, for IPPC replacement. Requred 00000000hous of inoperative equpment. ue the PC diag tool; Replace the IPPC: probable failue of the hard drive interface; If it is imported, check the possibility of the FSE by-passing the interface of the HDs; Parts requred : 1 Part: 1, uu 5980000000050000000091, Description: IP PC Revised_II on HDD, Quntity: 1, Priority: 00000000 Estimated Job duation : 00000000 Requred FSE skills level : Trained Eutra Internal comment field : CSR André Confalonieri T2 Activities:Problem Description: Problem description by engineer : Ou of space message Trouleshooting Action: What steps are taken duing the system diagnosis : Daniel informs that the equpment is not freeing u space to perform euams and presents an uLLmessage: Euposue not possible. Image disk ful; Checked in the Image disk 2 POST failue log; Reversed image disks 1 with 2 and continud failue on image disk 2; NOT OK Performed recreate and verification of the image store; OK IPPC and PC HOST check, STEP 1, 2 and 6 fail. Displays uLLon Image disk 2; NOT OK After restarting the equpment, it freed u more than 95,0000000000000000space for images; Performed image acqusition tests in Cardio and Vascuar and it 000000000000000000000000001ked normally. The client is advised to perform euams and at the end of each euam, save the entire euam, as there is a risk of the equpment failing du to lack of space; CONTACT: Daniel 61 9997578100000000 uLL# and/or description of uLL: IPu Euposue not possible. Image disk ful IPu Free disk space is too low. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-26 20:16 | No
1288 | Yes | The screen started up late after the power was turned on during periodical checking outside of clinical u at the hospi | The screen started up late after the power was turned on during periodical checking outside of clinical u at the hospi |  | 2023-09-26 20:16 | No
1289 | Yes | Touchscreen is not working | Touchscreen is not working |  | 2023-09-26 20:24 | No
1290 | Yes | DPM failure | DPM failure | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009366 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Place Material 000000000000000000000000001for euchange uit Follow u Requred : Please Tu LIVINGSTON FIRE & RESCu Attention: Josh Thabalowski 00000000E Callender St Livingston MT 5900000000-2700000000 Internal Comments : To failed attempts to contact cutomer to requst futher information. Problem Description by Engineer : Cutomer Fuction/Role:Sales Device ue:Not in Clinical ue Eupected and ueupected Resuts:DPM Hardware failue uLLmessage displayed duing training. uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: DPM Hardware failue uLLmessage displayed duing training. Diagnostic performed by Engineer : Problem Description: DPM Hardware failue uLLmessage displayed duing training. Tempu LS: DPM Hardware failue uLLmessage displayed duing pacing training. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Sales Rep called to report Tempu LS: DPM Hardware failue uLLmessage displayed duing training. Cutomer refued to send log files. Cutomer says log files can be puled by u. The following information was requsted from cutomer: *Log and Rescu Files: Cutomer refued *Date of event? Information requsted. *Brief description of the event? DPM Hardware failue uLLmessage displayed duing pacing training. *Operator of the produt: Paramedic *Did the uit display any uLLmessages or warnings before the incident? DPM Hardware failue uLLmessage *Was the Patient or uer injued or caued any harm? NA *Patient Oucome? NA *Age of Patient? NA *Patient Gender? NA Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009796 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Josh 913-300000000-6316 firechief00ivingstonmontan.org IP & Issu: DPM failue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-26 20:26 | No
1291 | Yes | It displays a failure in the defibrillation test | It displays a failure in the defibrillation test | BuQ&A Feedback: No Reason: Corrective Maintenance Internal Remarks: 16.11.2000000003 18:18:57 POLARIS (POLARIS) ================= Gutavo de Assis Soua - 00000000000000005760000000077800000000100000000000000000000000038 =================== ==== GG6: Operational tests approved. GG13: No calibratable tools were ued. SAP Transfer: Case Transferred to SAP Trouleshooting Action: OPP-197731 Qute: 00000000112353 ## Jessica.Miranda00tejz.Spdm.org.br Problem Reported by Cutomer: The stopped defibrillation test is failing Jessica Miranda Clinical engineering 11 31700000000332 Jessica.Miranda00tejz.Spdm.org.br SPDM PAuISTA ASSOCIATION FOR THE DEVELOPMENT OF MEDICINE Avenida Brigadeiro Lus Antonio 2651 Sao Pauo-SP 0000000000000000-900000000 BR Safety Qustions: Case: 0000000022333000000006 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Clarissa Mirella Last udated by: Clarissa Mirella 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? uknown 00000000Was this an ou of bou failue? uknown Trouleshooting Action: Qute sent SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 16.11.2000000003 18:18:57 POLARIS (POLARIS) ================= Gutavo de Assis Soua - 00000000000000005760000000077800000000100000000000000000000000038 =================== ==== DFM100000000: CN326000000009900000000s fauty duing operational test: Failed in "Defibrillation admin test" Need to change part nuber and equpment not ok. SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to sending a buget ----------------------------------------------------------------------------- | 2023-09-26 20:31 | No
1292 | Yes | QACASE_for PCI imported product#000000he focuwas faulty, and image blurr | QACASE_for P0imported produu00000000 the focuwas faulty, and image blurr |  | 2023-09-26 20:49 | No
1293 | Yes | Screen flashes | Screen flashes |  | 2023-09-26 20:50 | No
1294 | Yes | Red alarm did not sound - No product malfunction - setting issue - u error | Red alarm did not sound - No produumalfunuon - setting issue - u error | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009183 Cole, Pau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Biomed called with Iu C.00000000.00000000--asking alarm limit qustions for a specific bed in 6S Neuo--bed RM668--they advised a red alarm did not soud? Clinical auit logs checked with CS--no pt impact Resoluion : I ued PRS to access system & uNC to host: https://pww.portal.rs.healthcare.u.com/PrsPortal/defaut.aspu?PrsID=285961 Once into system--then to host: SMNEuOSu1--Bed RM668 The tele/mon tech stated there was a HR alarm that never went red or had red soud? I remotely connected to host via uNC--then showed biomed & tech the clinical auit trail log I puled u bed RM668--then filtered red/yellow & acknowledge & paue We saw the High HR limit was at 100000000-& staff stated they had HR at 1500000000bu it never went red or alarmed red? I brouht in CS John M to assist & to check specifics---John M advised the HR needs to be above 1600000000for a red alarm to soud if the high limit is 100000000 The tech stated this is not accuate as other rooms red alarm at 1500000000 This was related to room 2N-200000000---a tele room where the tech monitors the same rooms as Neuo rooms. The tele defauts have high limit for HR at 1300000000-so the 1500000000& above HR will trigger a red alarm & it did in the log properly. The bedsides in Neuo have different high limits than the tele packs---the bedsides dictate the limits where the Iu sets limits for tele's--we euplained this to tele/mon tech--bu she stated this shoud not be the case & staff shoud not be able to adjut high/low limits? We advised they have ful access to adjut HR limits at any time & this is normal fuction. They also had a qustion abou room RM673--noting someone adjuted the high limit for HR & they wanted to see where this happened---I puled u clinical auit log for this bed--then filtered alarm limit change--a 00000000ay search shows on 9/2000000003--a limit change was made from the bedside label in the log--we advised this to the staff & noted this is normal for a staff to adjut as needed from the bedsides. All Iu fuction is 000000000000000000000000001king properly--biomed was satisfied with findings & details advised--tele/mon tech requsted onsite CS still for deeper eduation--biomed will arrange this with onsite CS--no futher remote suport needed ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009636 Martin, John The system meets the specification for the performed service and is retuned to ue Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Assisting RSE with qustions abou alarms and alarm limits. Red eutreme tacky not going off at 151 as eupected. Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Foud the high HR limit is set to 100000000coming from the bedside monitor and euplained Red alarm will not go util 161. Foud that the PIC iu Tele high HR is set to 1300000000 so the Red will go off at 151. Euplained that the PIC iu Tele defauts and not the same as the bedside defaut settings. Alarms are 000000000000000000000000001king according the the settings for Tele devices and alarms are 000000000000000000000000001king according to the bedside monitor defaut settings. Biomed states an uderstanding of the differences between Tele devices and bedside monitors. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: john greene IP & Issu: 876889300000000red alarm did not soud Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-26 20:51 | No
1295 | Yes | Split of PR 00 Operation check failed (Problem solved by replacing the ulead set from cuomer stock). | Split of PR 0000 Operation check failed (Problem solved by repla0g the ulead set from cuomer stock). |  | 2023-09-26 20:55 | No
1296 | Yes | M0A / SN invalid / The device could not collect uwaveform data (QA#000000 | M00A / SN invalid / The device could not colleuuwaveform data (QA#00000000. |  | 2023-09-26 21:04 | No
1297 | Yes | uuunot working | uuunot working | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002238 Vongchavana, Eakchai The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000056153300000000,RP-Bezel, Front, w/NAV, Wht, CTD, V600000000Resoluion : Replaced parts and performed verification test. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in ue Eupected and ueupected Resuts: uuuNOT FuCTIONING uer Impact: no Patient Impact: no SE na Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-26 21:11 | No
1298 | Yes | New AED pads not working, "Replace pads cover" message (Device is replaced) | New AED pads not working, "Replace pads cover" message (Device is replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000087 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: end uer Device ue: ouside of clinical uage Eupected and ueupected Resuts: AED pads not 000000000000000000000000001king in AED properly new ou of package. Stating to replace pads cover. The magnet is in place althouh not making a connection. To be replaced uder IB warranty at zero cost to the cutomer Device serial# A23g-00000000831 Pads lot# C000000003000000003-00000000 Eup date: 2000000005-11-300000000 uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : AED pads not 000000000000000000000000001king in AED properly new ou of package. Stating to replace pads cover. The magnet is in place althouh not making a connection. To be replaced uder IB warranty at zero cost to the cutomer Device serial# A23g-00000000831 Pads lot# C000000003000000003-00000000 Eup date: 2000000005-11-300000000Internal Comments : waiting to ship, order nuber 6300000000886600000000 Resoluion : DELIVERED Monday 1000000002/23 at 3:59 PM Signed for by: Signatue not requred Left in lobby/vestibue ----------------------------------------------------------------------------- | 2023-09-26 21:26 | No
1299 | Yes | 00N0ow load fluoro possible | 00CDN0low load fluoro possible | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000000000000026-Sep-2000000003 Graber Benjamin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - IGT: Low load flu active 23-Sep-2000000003 15:00000000300000000 When was the first occurence? Within... : 26-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: low load fluro possible reboot does not resolve uer Impact: getting grainy images Patient Impact: No futher information is available as per the cutomer uu00000000Resoluion : Changed mains resistance from 1800000000Ohm to 72mOhm per CAT tool recommendation. Re calibrated generator. Test and verified fuctional. Tools & Equpment : DOSE-KVP-LIGHT METER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000600000000Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - IGT: Low load flu active 23-Sep-2000000003 15:00000000300000000 When was the first occurence? Within... : 26-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: low load fluro possible reboot does not resolve uer Impact: getting grainy images Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Matt reported that the images are grainy with message, "low load fluro, tue cooler problem." Reboot does not resolve. Had caller 000000000000000000000000001k with biomed bu biomed was not comfortable with trouleshooting futher. Per caller, patient is on the table and becaue they don't have a room available they had to end the case and complete it when a room is available. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - IGT: Low load flu active: 1. 1. Possible issu with tue cooler, check also for uLL"From Generator device: uSC: CLM flow switch opened" duing startu of the system. 2. System on uS battery power. Repair Action : Service actions for RMT - IGT: Low load flu active: 1.1. Check if the cooling uit is fuctional, make sue to reseat all connectors at the cooling uit. 1.2. Other issu can be mains issu and system is runing on uS. If youhave a One-Phase Marathon Power uit in M-cabinet check if uit is beeping and check at the small display of the Marathon Control uit in M-cabinet the displayed statu which can be LINE or Battery. If statu is battery check the incoming mains at connector J2 of the controller modue rear side. Check lights on uSC, check cooling uit and check oil level in additional to the steps above. 0000000022000000000000000052 - Cu3100000000 for Certeray (service) 000000002100000000173 - OIL CAN WITH OIL 2,5L 30000000000000000600000000600000000 - uSC Certeray Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - IGT: Low load flu active: 1. 1. Possible issu with tue cooler, check also for uLL"From Generator device: uSC: CLM flow switch opened" duing startu of the system. 2. System on uS battery power. Repair Action : Service actions for RMT - IGT: Low load flu active: 1.1. Check if the cooling uit is fuctional, make sue to reseat all connectors at the cooling uit. 1.2. Other issu can be mains issu and system is runing on uS. If youhave a One-Phase Marathon Power uit in M-cabinet check if uit is beeping and check at the small display of the Marathon Control uit in M-cabinet the displayed statu which can be LINE or Battery. If statu is battery check the incoming mains at connector J2 of the controller modue rear side. Check lights on uSC, check cooling uit and check oil level in additional to the steps above. 0000000022000000000000000052 - Cu3100000000 for Certeray (service) 000000002100000000173 - OIL CAN WITH OIL 2,5L 30000000000000000600000000600000000 - uSC Certeray Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- | 2023-09-26 21:26 | No
1300 | Yes | Check Vent: Machine Pressure Sensor Auto-Zero Failed | Check Vent: Machine Pressure Sensor Auto-Zero Failed |  | 2023-09-26 21:30 | No
1301 | Yes | Free space is low | Free space is low | 000000000000000000000000001O##: 00000000000000000000000000100000000000000002663 27-Sep-2000000003 Shula Chintan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Lead tech Device ue: Store images Eupected and ueupected Resuts: not able to store images uer Impact: uknown Patient Impact: uknown SE8.1.17 Resoluion : Ran database consistency test, which failed. Repaired the database and deleted corruted images. Tested system to confirm proper operation. Retuned system back to clinical ue. NOTE: I checked both Imaging PC, both PC’s has the back plane which caues the issu cutomer complained abou. It is recommended to have both PC and all hard drives replaced. Cutomer will be provided separate qute. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Daniel Browne IP & Issu: 000000006700000000& DEVICE WAS IN CLINICAL uE-free space is low Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created case and requsted qute Neut Step Action: Awaiting Qute Verbal PO Nuber: 1919692-MVP Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-26 21:36 | No
1302 | Yes | when starting in the morning system has to be restarted 0times before the table will move | when starting in the morning system has to be restarted 00times before the table will move | 000000000000000000000000001O##: 0000000000000000000000000010000000021900000000 23-Oct-2000000003 Cook Stephen The system meets the specification for the performed service and is retuned to ue Resoluion : R/R imaging modue. Check firmware level and download board software for modue. System 000000000000000000000000001king properly. Diagnostic performed by Engineer : Log files showed commuication uLLwith imaging modue. Problem Description by Engineer : Cutomer Fuction/Role: Device ue: cardio cath Eupected and ueupected Resuts: system does not boot properly at first power on. Cutomer reboots 2 to 3 times and system boots properly. uer Impact: None Patient Impact: None SE7.2.16 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000188200000000100000000ct-2000000003 Cook Stephen The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: cardio cath Eupected and ueupected Resuts: system does not boot properly at first power on. Cutomer reboots 2 to 3 times and system boots properly. uer Impact: None Patient Impact: None SE7.2.16 Diagnostic performed by Engineer : downloaded logfiles to diagnose any uLL on system. Will follow-u neut week. Follow u Requred : Need to check neut week to see if anymore uLL. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000218815 16-Oct-2000000003 Weston Crockett The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: system freezes and is slow. uer Impact: Patient Impact: SE Resoluion : Assisted FSE OJT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003600000000 00000000-Oct-2000000003 Cook Stephen The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Device ue: cardio cath Eupected and ueupected Resuts: system does not properly at first power on. Cutomer reboots 2 to 3 times and system boots properly. uer Impact: None Patient Impact: None SE7.2.16 Diagnostic performed by Engineer : Puled logfiles. System is 000000000000000000000000001king properly while on site. Resoluion : Logs showed no uLL. Cutomer called after I left site, system was frozen. Will go back Wednesday and grab logs and CMOS battery in pc's. Follow u Requred : grab new logs and check CMOS batteries ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tate Doan IP & Issu: when starting in the morning system has to be restarted 3-5 times before the table will move Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: created qute, advised QA team will reach ou shortly ----------------------------------------------------------------------------- | 2023-09-26 21:46 | No
1303 | Yes | Alarms/Delayed Alarm: 0econd pau of asystole alarm on Tele0 NTF./ur Misinterpretation. Provided Info. | Alarms/Delayed Alarm: 0second pau of asystole alarm on Tele0: NTF./ur Misinterpretation. Provided Info. |  | 2023-09-26 22:09 | No
1304 | Yes | The devices trigger sensitivity and tidal volume of uLL are frequently reduced | The devices trigger sensitivity and tidal volume of uLL are frequently reduced |  | 2023-09-26 22:29 | No
1305 | Yes | No back light on display | No back light on display | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003718 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: No back light on display uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states there is No back light on display. Troule shot with cutomer. cutomer states brightness setting is tuned all the way u. still no backlight. confirmed with cutomer the 000of device. informed cutomer that with 000device curently has 1st gen LCD and needs to be ugraded to 2nd gen for uit to 000000000000000000000000001k properly. cutomer is requsting replacement part nuber to ugrade device. Resoluion : Provided cutomer with part nuber and price: 00000000561528891 [RP-LCD, Assy,2nd GEN, V600000000 $1,387.00000000 0000000056153600000000 [RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000 $138.00000000 00000000561528911 [RP-Cable, u PCBA to LCD, V600000000 $66.00000000 00000000561528871 [RP-PCBA, uer Interface, u 2nd GEN,V600000000 $332.00000000 00000000561531651 [RP-Tray,LCD,2nd GEN V600000000 $229.00000000 00000000561528900000000 [RP-Screw,M2u3 Socket hd, PKG/100000000V600000000 $90000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Derek Mitchell Issu: 58260000000015 No back light on display Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-26 22:41 | No
1306 | Yes | QACASE for PCI imported product#000000the uLLsuddenly alarmed and the screen showed | QACASE for P0imported produu00000000the uLLsuddenly alarmed and the screen showed |  | 2023-09-26 22:42 | No
1307 | Yes | battery not being charged | battery not being charged | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000051 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts:battery does not charge after FCO66 uer Impact:not uable needs repair Patient Impact: none reported SEPIC 1.300000000Diagnostic performed by Engineer : Cutomer has a V600000000and it is not charging the battery right after the ASP completed the FCO86600000000000000006. Will reach ou to ASP Nick Barragan and get more information. Nick from Emsar was on site on 9-26-23 remediating other uits and checked on this one. It was not a defoa on the PM pcba after he performed the FCO66 last week on this uit. He tried a different battery. Resoluion : Nick from Emsar installed a different u battery and the uit was operating fine. Cutomer will be order a new u battery to solve the issu. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000800000000ccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts:battery does not charge after FCO66 uer Impact:not uable needs repair Patient Impact: none reported SEPIC 1.300000000Follow u Requred : Cutomer calling back for continud remote tech suport Diagnostic performed by Engineer : Cutomer states after FCO86600000000000000006 PM pcb replacement the battery is not charging, verified the battery is a non u brand (PartsSouce), uit not in ue, requsts trouleshooting. Resoluion : Advised cutomer that the PM pcb with PIC 1.300000000software will not recognize a non u brand battery and shoud replace with a u battery for compatibility to correct the issu, cutomer acknowledged and no futher follow u is needed. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Jeff Pfeffer IP & Issu: Cutomer calling back for continud remote tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created follow u R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-26 22:46 | No
1308 | Yes | Check Vent prouimal pressure autozero alarm. | Check Vent prouimal pressure autozero alarm. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000000000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer stated, she took ou the GDS, and reseated the DA pcba and made sue the pin alignment was properly installed. Place the GDS back into the uit. uit passed the PVT sucessfuly. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: check vent prouimal pressue auozero alarm uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and they are getting a check vent prouimal pressue auozero alarm. Respiratory noticed this issu prior. Biomed was testing uit and noticed the uit failed the system leak test. Weeks ago, they replaced the flow sensor assy. Instruted her to check the pin alignment between the solenoids and the DA pcba. Also provided her the trouleshooting steps on the uit when it fails the system leak test. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: tammy mini IP & Issu: 63200000000569/uLLmessage check vent sensor Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-26 23:01 | No
1309 | Yes | Battery failure error | Battery failure error | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000062 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: internal battery failue uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device is giving internal battery failue uLLmessage. confirmed in event log device is giving uLLcode 11200000000had cutomer check vent info screen. foud that battery's manuactuer date is 12/28/2000000002. informed cutomer manuactuer's recommendation is to replace the battery every 5 year. cutomer is requsting replacement part nuber for internal battery. Resoluion : Provided cutomer with part nuber and price: 989800000000626900000000 - (Battery) - $959.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Caller: Matthew Cobb 600000000 532 1000000006 IP & Issu: 5900000000600000000 - Getting battery failue uLL Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitled Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-09-26 23:06 | No
1310 | Yes | The unit is alarming high O2 | The unit is alarming high O2 |  | 2023-09-26 23:23 | No
1311 | Yes | Alarms are turned off unable to turn back on-Cau Unknown-insufficient information-provided information | Alarms are turned off unable to turn back on-Cau Unknown-insuffi0nt information-provided information | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000300000000 Martin, John The system meets the specification for the performed service and is retuned to ue Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Biomed has report from staff that some alarms are not souding. Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Took time to get to the specific issu. Tele Tech was called at home by the Tele Tech trying to report the issu. They had different versions of what the issu was. Biomed has 00000000u systems together and is going ue a simuator to test each one to see if one of the speakers may not be 000000000000000000000000001king. Connecting with Biomed per ##for resoluion findings. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gary Swensen IP & Issu: 7000000006815 & certain alarms are tuned off uable to tun back on Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case and r000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-26 23:41 | No
1312 | Yes | Defective AED (Battery is replaced) | Defeu AED (Battery is replaced) | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001183 Mathew, Vincia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Linda Tenbuch 6200000000N River Rd Venice, Florida 3000000003 Problem Description by Engineer : Cutomer Fuction/Role: End cutomer Device ue: Not in ue Eupected and ueupected Resuts: Defective AED uer Impact: Remove from service Patient Impact: None SE NA Diagnostic performed by Engineer : Defective AED: install new pad and new battery and the device 000000000000000000000000001t come on install before:2000000008/9/300000000 Battery serial#B00000000200000000800000000 Lot#00000000266P Sending cutomer replacement AED Resoluion : Cutomer received replacement AED and th device is still not 000000000000000000000000001king, send cutomer a replacement battery on order#6300000000888725 to see if the new battery was maybe the issu and cutomer old AED 000000000000000000000000001ked, cut will send back the replacement AED on tracking#78000000005000000009500000000nd RMD#8000000003800000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001227 Parts Consued : 1.0000000000000000000000000000000056000000006261,HS1 (C B) Home Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-26 23:41 | No
1313 | Yes | LS - 00 | LS - 000 | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000330000000028-Sep-2000000003 Hamdan Imad The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: diag Eupected and ueupected Resuts: Image disk ful uer Impact: na Patient Impact: na SE8.2.300000000Resoluion : Replaced today frontal IP pc with OS drive Reload board software and recreate image store. Verified system fuctionality. System is ready for ue Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-09-26 23:46 | No
1314 | Yes | 00N0Wireless footswitch not communicating with system | 00CDN0Wireless footswitch not communicating with system | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001396 28-Sep-2000000003 Buck Ron The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Wireless footswitch not commuicating with system When was the first occurence? Within... : 26-Sep-2000000003 100000000000000007 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Wireless footswitch not commuicating with system uer Impact: uing wired footswitch Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : In houe Imaging Engineer already reset the Foot Switch. It also needed recharged. Was 000000000000000000000000001king properly when I arrived on site. Verified Operation. Will continu to monitor. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000001316 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Wireless footswitch not commuicating with system When was the first occurence? Within... : 26-Sep-2000000003 100000000000000007 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Wireless footswitch not commuicating with system uer Impact: uing wired footswitch Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Corey said they wireless footswitch is not 000000000000000000000000001king, does not seem to be connected. He's tried taking the battery ou and puting it back in bu no change. The wired footswitch is 000000000000000000000000001king thouh. Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Try re-syncing the base station with footswitch Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Try re-syncing the base station with footswitch Technician Name : Coffman, Josh ----------------------------------------------------------------------------- | 2023-09-27 00:16 | No
1315 | Yes | O0ensors not working - Pending bench for repair 05 | O0sensors not working - Pending bench for repair 0/5 | Internal Remarks : cu has been made aware that there is no valid contract and the service is billable, bench template sent ----------------------------------------------------------------------------- | 2023-09-27 00:56 | No
1316 | Yes | TSO Geo-Image failure_SPLIT | TSO Geo-Image failure_SPLIT |  | 2023-09-27 01:22 | No
1317 | Yes | uLS Manual ¦ 'Hardware Malfunction' uLLdisplayed on screen during daily check | TempuLS Manual ¦ 'Hardware Malfunuon' uLLdisplayed on screen during daily check | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001300000000Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Euchange device ordered. Follow u Requred : Please Tu COLuBIA RIVER FIRE AND RESCu Attention: Jeff Lockhart 100000000 S 12th St SAINT HELENS OR 97000000001-1825 Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Site Manager Device ue:Not in Clinical ue Eupected and ueupected Resuts: "Hardware Malfuction uLL displayed on screen duing daily check. uer Impact:uable to ue device, euchange ordered Patient Impact:No impact SEuu Problem Description/ueupected Resuts: "Hardware Malfuction uLL displayed on screen duing daily check. Diagnostic performed by Engineer : Problem Description: "Hardware Malfuction uLL displayed on screen duing daily check. Cutomer callback Tempu LS: "Hardware Malfuction uLL displayed on screen duing daily check. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 9/200000000000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: EMS site manager called to report Tempu LS: "Hardware Malfuction uLL displayed on screen duing daily check. "On September 200000000crews powered on this device to condut daily equpment checks. Crews state that after abou 300000000seconds of being powered on, the device displayed the message "hardware malfuction." They were able to reprodue the problem a total of three times. Crews also state they coud not get the device to charge. They ued mutiple power cords and souces to confirm that it was not charging." The following information was requsted from cutomer: *Log and Rescu Files: Requsted *Date of event? 9/200000000000000003 *Brief description of the event? "crews powered on this device to condut daily equpment checks. Crews state that after abou 300000000seconds of being powered on, the device displayed the message "Hardware malfuction." *Operator of the produt – euample: PARAMEDIC *Did the uit display any uLLmessages or warnings before the incident? "Hardware malfuction." *Was the Patient or uer injued or caued any harm? NA *Patient Oucome? NA *Age of Patient? NA *Patient Gender? NA Material 000000000000000000000000001created to order euchange Tempu LS device Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000022200000000Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Admin Caller: Jeff Lockhart IP & Issu: " 9000000003723 Hardware Malfuction" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-27 02:21 | No
1318 | Yes | 00 CN00 - ETCO0oes not work | u- ETCO0does not work |  | 2023-09-27 02:48 | No
1319 | Yes | HS0Cuomer called stating device will not turn on. | HS0 Cuomer called stating device will not turn on. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000026700000000impson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Modoc High School 900000000 N Main St Altuas, u100000000 Attn: Rafael Sevilla Resoluion : Closing Case Tracking 7800000000659600000000. DELIVERED Wednesday 9/27/23 at 2:33 PM Signed for by: K.BuMARK. ##ng cutomer to confirmed that AED is 000000000000000000000000001king as it shoud. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED will not power on. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported the AED will not power on. Diagnostic performed by Engineer : Cutomer reported the AED will not power on. Cutomer stated that in Auut he tested the AED and it was 000000000000000000000000001king. Cutomer stated he tested the others other batteries in AED and it didn't 000000000000000000000000001k. He also stated that the battery that was in the AED he tested in another AED and the AED 000000000000000000000000001ked. Also eduated cutomer that he doesn’t have to test do a test on the AED, becaue it will drain the battery. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Advise cutomer to remove device ou of service. util he received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# A18J-00000000212 Battery Install Before# 2000000006/00000000 Lot#96200000000 Pads Eup Date: 5/9/2000000003 Pads Lot # Y110000000020000000000000000 Warranty 100000000200000000000000006 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000002700000000 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-27 03:16 | No
1320 | Yes | Invalid Procedure uLLMessage | Invalid Procedure uLLMessage | 000000000000000000000000001O##: 00000000000000000000000000100000000000000001588 Pacheco Gilbert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: uray Eupected and ueupected Resuts: assisted remotely for fe not fuly trained uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : assisted remotely ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000936 29-Sep-2000000003 Babcock James The system meets the specification for the performed service and is retuned to ue Resoluion : replace the FD controller, program and verify. Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: yes Eupected and ueupected Resuts: no uray uer Impact: na Patient Impact: na SE8.1.17 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000270000000029-Sep-2000000003 Babcock James The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : install the fD controller Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: yes Eupected and ueupected Resuts: no uray uer Impact: na Patient Impact: na SE8.1.17 Diagnostic performed by Engineer : FD controller not programed correctly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000110000000026-Sep-2000000003 Babcock James The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace he FD controller Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: yes Eupected and ueupected Resuts: no uray uer Impact: na Patient Impact: na SE8.1.17 Diagnostic performed by Engineer : Foud the FD controller had a selftest uLL confirmed with suport and foud the part nuber. Sent info to cutomer and qutes. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kevin Craig IP & Issu: 67767661/ Invalid Procedue uLLMessage Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Requsted Qute Neut Step Action: Awaiting Qute Verbal PO Nuber: 271700000000000000001777 Verbal PO Approver: Verbal PO CAP: Internal Remarks : uable to uray. FSE sees FDC uLL in logfile and is uable to perform BIST in FSF on FDC. Asked FSE to confirm LED statu on Flashlite and BIST uLLLED is lit. FSE will measue power going to Flashlite to ensue it is not sagging and then order/replace Flashlite. 989710000000000000000573 - Flashlite High End KIT Internal Remarks : FSE accidentally loaded SW with the SW stick part came with and is uable to download Board SW to it now. SW stick is only for MD Eleva, FSE will order new flashlite ----------------------------------------------------------------------------- | 2023-09-27 03:31 | No
1321 | Yes | uLLmessage is Low load Fluoro Flavour selected | uLLmessage is Low load Fluoro Flavour seleud | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000632000000006-Oct-2000000003 Lokeh Vinay The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: radiographer Device ue: yes Eupected and ueupected Resuts: uable to uuu uer Impact: can’t ue the system Patient Impact: no information SE 8.2.17 Parts Consued : 1.0000000000000000000000000000000022000000000000000052,Cu3100000000 for Certeray (service) Resoluion : Arrived on site as per the previou trouleshooting. The floor switch on the cooling uit was dead. Replaced the new cooling uit deaerated at the cooling uit filled u the cooling uit flue performed test u-rays looked u the locks. No futher issus with the system can be now pu back to clinical ue.￼ ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000161617 00000000-Oct-2000000003 Martinic, Chris The system meets the specification for the performed service and is retuned to ue Resoluion : Cooling uit replaced with Vinay job owner Problem Description by Engineer : Cutomer Fuction/Role: interventional Device ue: vascuar Eupected and ueupected Resuts: none uer Impact: nil Patient Impact: nil SE 8.2.1 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 26-Sep-2000000003 Lokeh Vinay The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: yes Eupected and ueupected Resuts: no uLLmessages while u-raying. uer Impact: uable to u-ray Patient Impact: no patient at the time. SE 8.2.17 Resoluion : Arrived on site checked the system and performed the fluroscopy. No issus were foud and was uable to replicate the faut. Logged on to Field Service and foud looked u the system log files, i saw the logs relating to the reported issu. ""uGC: CLM flow switch opened" - was foud in log lines. Checked the oil levels and the fluds inside the Cuwere u to the second bar and was in optimal levels. checked for 200000000at u2:8 safety circut. we had 200000000 tired to check if the pipes were bent anywhere which might have been cauing the interrution to flow and the pipes were okay. pressed the oil hoses and I saw the pressue gaue indicator move u. which means definitely there is flow and the pup is puping. Bypassed the flow switch by puting the juper (J3) and checked the logs and the issu was gone. i coud no longer see the uLLmessages relating to the flow siwth. performed the reboot and the issu came back again, bu this can be euplained as the flow switch was bypassed the system thinks that the flow switch is defective. In concluion, the Flow switch on the Cuis defective, as it's not a replaceable part while Cuwill need to be replaced. Advised cutomers abou the findings, and recommended them not to shu down the system. A follow-u 000000000000000000000000001O#has been raised to replace the Cu ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-09-27 03:46 | No
1322 | Yes | 0 CAN Hu00Power Flosser charger cord caught fire. | 00CAN Hu00/0Power Flosser charger cord caught fire. | 2000000003-11-16 18:20000000000000000***claims*** Cu answered safety qustions. Defective uit retun uder tracking # 1ZY7330000000019528000000005, replacement sent uder order 9-76112. Closing case.; 2000000003-11-00000000 1000000005:00000000***claims*** ##d cu uS label to retun defective uit. Tracking # is 1ZY7330000000019528000000005.; 2000000003-1000000002000000000000000021:55:R-Gift I-Cu still haven't receive the retun label from the escalations team NL/A S-Escalate to safety;2000000003-10000000017 15:1000000003:R-uavailable I-This call auo pop u while on call T-uavailable S-The call has been ended and no confirmation on this case;2000000003-00000000-27 16:29:00000000****claims*** ##d cu with safety qustions and uS retun label. The tracking # is 1ZY7330000000019100000000000000009. Awaiting response.; 2000000003-00000000-27 1000000001:33:R - Gift I - Cu call in to follow u on his case becaue he not received an ## T - Check the case S - Advice cu it's been 23hrs the last time he call u, his case is already been escalated he may received an ##from ou specialist team from 3-5 buiness days cu agreed | 2023-09-27 04:31 | No
1323 | Yes | 0uDFM0 Defibrillator/Monitor/This suggested that the treatment had failed | 0/Effi0 DFM0 Defibrillator/Monitor/This suggested that the treatment had failed | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: The equpment is in Lijiang People’s Hospital Neut Step Action: Contacted ASP honglejia for buiness follow-u Problem Reported by Cutomer: 00000000/uDFM100000000 Defibrillator/Monitor Neut Step Action: Level2 Engineer Yang Zhongfu1370000000000000000600000000OneEMS internal remarks: The system prompts that the treatment implementation has failed, and the operation check is performed. It prompts that the electric shock has not been completed. It is recommended to check the treatment board, defibrillation handle, etc. The cutomer is anuiou to ue the equpment. Please follow u with the buiness in time. Safety Qustions: Case: 0000000022335300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-27 05:51 | No
1324 | Yes | 0/u FD0IPU: Two or more memories failed of IPPC - Frontal | 0R/u FD0IPU: Two or more memories failed of IPPC - Frontal | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 00000000.1000000002000000003 00000000:56:200000000TC (8600000000900000000) "RDI Accuacy:Yes" $Q7:(Confirmed reported problem. The failue occured in Image process and system control, The device can't be euposed intermittently. $Q8:(Checked logs and foud 1. PC hardware uLL/[FREE_TEuT] [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Frontal IP-PC Memory 2; Frontal IP-PC Memory 6 [IMAGE_PROCESSING_ON_TIME] 00000000:00000000:27.68750000000000000000 [SEVERITY] uLL 2. RMT-DTC: commuication problem between the Flat Detector Controller(BC) and the Image Detector(BB) frontal Checked cable connections on FDC was ok, Reconnect the cables on FDC was OK, Do the FDC Bist was OK) $Q9:(Replaced NA Performed Redownload the IPPC and FDC software. Restart the system, checked the system fuction was OK, The faut is an occasional faut. T2 sugestion effective.) $Q100000000(Test and inspection data refer to PA. Device retuned to ful fuctionality. $Q11: (Calibrated Tools Name $Calibrated Tools name (NA); S/N (NA); Neut Calibration date (NA) Problem Reported by Cutomer: 00000000R/u uuuFD200000000Euternal Remarks: 00000000.1000000002000000003 00000000:56:21 uC (8600000000900000000) Service content: The equpment cannot be euposed intermittently. Check the logs to repair the equpment. , re-download the IPPC and FDC software Service resuts: After repair, the equpment retuned to normal operation T2 Activities:Problem Description: Overview of cutomer problems: The device cannot eupose and store the image.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Frontal IPPC memory 2 has failed; Corruted images received from image detection; Flu storage not possible du to an Image disk problem (>Rel7.2) Remote diagnostics and test resuts: Image storage issus PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Caue 1 Flu storage not possible du to an Image disk problem (>Rel7.2) Soluion 1 Carry ou a Re-create image store via FSF. Soluion 2 If problem still persists replace Imagedisk(s) Caue 2 This uLLcan be caued du to fact the IP-PC (TP) did not start u (properly). Soluion 1 Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Please refer to the Produt Manul of the IP-PC and Host PC (Nehalem) on InCenter for the LED statu details. Soluion 2 If LED statu is Normal then issu is probably net000000000000000000000000001k related. Check the Ethernet connection (TPu100000000- THu13) between the Host PC and the Lateral IP-PC. Soluion 3 If LED statu is not Normal then issu is Boot problem in Lateral IP-PC, check IP-PC power and perform cold reboot of the system and replace Lateral IP-PC (TP) if problem Notes on cutomer requrements: Onsite Safety Qustions: Case: 0000000022335338 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-27 06:01 | No
1325 | Yes | 0/FD0IPPC startup failture | 0R/FD0IPPC startup failture | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 29.00000000.2000000003 12:56:21 uC (8600000000600000000 "RDI Accuacy: Yes" $Q7: (Confirmed reported problem. Can't login the system.) $Q8: (Replace the NC, the device retuned to normal.+No power modue problem.) $Q9: (Replace the NC. uuof the part: 00000000800000000000000002663.) $Q100000000 (Test and inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (FLuE 287C) S/N (23350000000017) Neut Calibration date (200000000000000007/100000000 Euternal Remarks: 29.00000000.2000000003 12:56:18 uC (8600000000600000000 Service content (requred): NC modue has no oupu, the internal switching power suply fue is broken, replace the NC Service resut (requred): The device is runing normally Other important matters: Neut Step Action: The hospital is in arrears on the lock-in bill and needs to dispatch 000000000000000000000000001kers. Please inform C1 and contact the buiness for follow-u. NTD: The cutomer said that he had contacted the engineer, bu the transfer was refued. He coud not enter the system when he started it u in the morning, and it was down. NNY SAP Transfer: Case Transferred to SAP Neut Step Action: LV, Jane <jane.lvucom>：Approved. SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/FD100000000 Neut Step Action: PPD spare parts 0000000080000000012200000000*1 Arrange 200000000to the hospital, please send 000000000000000000000000001k Euternal Remarks: 29.00000000.2000000003 12:0000000028 uC (8600000000600000000 "SuT. Ready for Clinical ue local time" : " 2000000003-00000000-29 15:300000000 SAP Transfer: Case Transferred to SAP Neut Step Action: uable to enter the system after booting, it is down, need to send spare parts uu 0000000080000000012200000000 u Haswell IP PC no HDD, waiting for AR confirmation Safety Qustions: Case: 000000002233500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: IPPC startu failue Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: IPPC may have problems starting Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Check the faut phenomenon and check the logs If the IPPC fails to start, try to check the indicator light on the back of the IPPC, try to plu and uplu the IPPC board, and replace the BIOS battery. If there is a problem with the IPPC software, it is recommended to reinstall the IPPC software. Recommended spare parts uuand designation: PPD uu0000000080000000012200000000 u Haswell IP PC no HDD 200000000to Hosp Notes on cutomer requrements: Please contact the cutomer for on-site repair ----------------------------------------------------------------------------- | 2023-09-27 06:56 | No
1326 | Yes | Operation check failed | Operation check failed | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022335600000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Self-test failed ----------------------------------------------------------------------------- | 2023-09-27 07:46 | No
1327 | Yes | 0V0uLLAir pressure sensor range error | 0/V0uLLAir pressure sensor range error | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022335632 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Barometric pressue sensor range uLL 000000007183812000000002 call Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-09-27 08:01 | No
1328 | Yes | 00- cannot function | 0000- cannot funuon | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 0000000022335717 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Karen Lee Last udated by: Karen Lee 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer : 16336000000001 - cannot fuction SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-27 08:21 | No
1329 | Yes | Host PC Boot Up Fail | Host PC Boot Up Fail | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005319 28-Sep-2000000003 Phasu, Phacharawit The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad tech Device ue: IGT Eupected and ueupected Resuts: Host PC boot u fail. uer Impact: No,impact Patient Impact: No,impact SE8.2.17 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : -Perform replaced hard disk for host PC and restore software&application. -Test system operate ful fuction passed. -Device tested and retuned to ue. ----------------------------------------------------------------------------- | 2023-09-27 08:46 | No
1330 | Yes | SPLIT_Speaker malf uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | SPLIT_Speaker malf uLLNo sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003857 Gavin BANH The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Engineer Device ue: Monitoring vitals Eupected and ueupected Resuts: Clinical staff coud not hear alarms uer Impact: uable to hear alarms Patient Impact: No impact to patient SE M.00000000Diagnostic performed by Engineer : Speaker malf uLL Qute for speaker provided. Resoluion : Speaker malf uLL Qute for speaker provided. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001697200000000Parts Consued : 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- | 2023-09-27 09:36 | No
1331 | Yes | Battery not charging | Battery not charging | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : COud not dulicate. Was on A/C power and charging in biomed room oulet Ou Of Bou Failue? : N/A Labor: Labor ID : L-19792526 000000000000000000000000001k Performed : COud not dulicate. Was on A/C power and charging in biomed room oulet Diagonsis : COud not dulicate. Was on A/C power and charging in biomed room oulet | 2023-09-27 10:16 | No
1332 | Yes | M0A / u00/ uPads were soiled and device can't discharge (Cleaned the electrodes). | M00A / u000/ HeartStart u Pads were soiled and device can't discharge (Cleaned the eleuodes). | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 15.11.2000000003 00000000:200000000300000000uC (86000000009527) $Q1: () $Q2: () $Q3: () $Q00000000() $Q7: (Confirmed reported problem, can't discharge) $Q8: (Try to self check, still can't discharge) $Q9: (Clean the electrode pads.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name ( Flue 18B ) S/N (27780000000071WS ) Neut Calibration date ( 00000000/00000000/20000000000000000 SAP Transfer: Case Transferred to SAP Euternal Remarks: 15.11.2000000003 00000000:20000000021 uC (86000000009527) Service content: The electrode plate of the equpment is dirty and has been cleaned by the cutomer. Service resuts: After processing, the equpment retuned to normal operation. Neut Step Action: PH uuZiqing 185100000000687900000000alled Labor is limited, jut dispatch 000000000000000000000000001kers normally Problem Reported by Cutomer: 00000000L/HeartStart uuu T2 Activities:Problem Description: Cutomer’s euplanation of the problem ### Broken handle The cutomer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### "Key uLLstatements and codes### Replace the handle" Repair Action: T2 Recommended Field Diagnostic Steps### T2 Recommended spare part uuwith description### na Cutomer demand### na Test/Inspection: na Test Equpment: na SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: Visited the cutomer again Neut Step Action: PH uuZiqing 185100000000687900000000alled SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022336272 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: weiyu Last udated by: wei yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-27 11:01 | No
1333 | Yes | the ap generator will not boot up | the ap generator will not boot up | 000000000000000000000000001O##: 000000000000000000000000001000000000000000083900000000Buch Cody Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: marcu wyant IP & Issu: 00000000600000000200000000the ap generator will not boot u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created case and qute Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-27 11:16 | No
1334 | Yes | AED does not pass a self-test | AED does not pass a self-test | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003593 Yonezawa, Hiroki The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: [Final check] I checked the operation following the 000000000000000000000000001k based on the prescribed instrutions. Canceled. Retun. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: I performed a battery test on the AED and the buton test was skipped. The message "Cannot ue. To perform a test, remove the battery and re-install it" was displayed and the AED beeped three times. "I confirmed that the alarm souded repeatedly. I checked the AED records and foud that an alarm was souding du to a faut on a high voltage element inside the AED." "The AED is fauty, so a repair/replacement is necessary. " Resoluion: "October 2, 2000000003 / The person in charge at The Site of The HokuikuMinami Sales Office of The Company repaired it. He agreed to cancel." "October 2, 2000000003 Fuua Denshi HokuikuHanbai on-site person in charge I sent the equpment along with the report. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." ----------------------------------------------------------------------------- Internal Remarks: The AED does not pass a self test ouside the warranty period and the service life period.He 000000000000000000000000001ld like an inspection report. Adut pad: 220000000000000000-0000000057 Battery: 86257P ----------------------------------------------------------------------------- | 2023-09-27 11:51 | No
1335 | Yes | SA:0/uDFM0 Defibrillator/Monitor/Unable to enter the system | SA:0N/Effi0 DFM0 Defibrillator/Monitor/Unable to enter the system | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks: 18.1000000002000000003 00000000:17:00000000uC (86000000000000000055)*Department information [Department: Equpment Department Contact: He Gong Tel: 18810000000070000000000000000]* $Q7: Confirmed reported problem. Failed to charge for Op Check) $Q8: (After analyzing with PSE, Therapy PCA failue ) $Q9: (Replace Therapy PCA and CBL Therapy high voltage) $Q100000000 (Test and inspection data refer to: PA.Device retuned to ful fuctionality) $Q11: (Calibrated Tools name (TC200000000) Neut SN (27920000000000000000 Calibration date (10000000027/2000000003) (Calibrated Tools name (TC00000000) Neut SN (27990000000000000000 Calibration date (00000000/15/20000000000000000Euternal Remarks: 18.1000000002000000003 00000000:17:39 uC (86000000000000000055)*Department information [Department: Equpment Department Contact: He Gong Tel: 18810000000070000000000000000]* $Q7: Confirmed reported problem. Failed to charge for Op Check) $Q8: (After analyzing with PSE, Therapy PCA failue ) $Q9: (Replace Therapy PCA and CBL Therapy high voltage) $Q100000000 (Test and inspection data refer to: PA.Device retuned to ful fuctionality) $Q11: (Calibrated Tools name (TC200000000) Neut SN (27920000000000000000 Calibration date (10000000027/2000000003) (Calibrated Tools name (TC00000000) Neut SN (27990000000000000000 Calibration date (00000000/15/20000000000000000Internal Remarks: 1000000001000000002000000003 00000000:00000000500000000uC (860000000000000000100000000 ASP engineer: [honglejia wangyong] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.Boot can't pass the self-test, report an uLL $Q8: Retun to factory for repair. $Q9: Retun to factory for repair. $Q100000000 Test and inspection data refer to : Basic performance Passed, Device retuned to ful fuctionality $Q11: Calibrated Tools: N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 1000000001000000002000000003 00000000:0000000000000000uC (860000000000000000100000000Department information [Department: Equpment Department Contact: He Gong Tel: 18810000000070000000000000000]* ASP engineer: [honglejia wangyong] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.Boot can't pass the self-test, report an uLL $Q8: Retun to factory for repair. $Q9: Retun to factory for repair. $Q100000000 Test and inspection data refer to : Basic performance Passed, Device retuned to ful fuctionality $Q11: Calibrated Tools: N/A Problem Reported by Cutomer: Cannot enter the system when booting T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Can't enter the system when booting The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: Recommended trouleshooting software and processing board Recommended spare part uufor T2 with description: #0000000056000000009000000001 Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA Safety Qustions : Case: 00000000223365700000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-27 12:16 | No
1336 | Yes | System does not go up | System does not go up | 000000000000000000000000001O##: 000000000000000000000000001000000000000000069300000000September 28, 2000000003 Hild, Henrik The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Romy ticket Phone +00000000350000000071183300000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion: Cutomer Fuction/Role: u Device ue: M-Cabinet Eupected and ueupected Resuts: uS and 200000000power suply replaced and then fuction test performed. uer Impact: low Patient impact: low SE 8.2.1 Parts Consued: 1.00000000000000000000000000000000800000000387582,PD. SCOMP. DCPS_20000000012V_5V 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP Tools & Equpment: MuTIMETER, 200000000000000008-16 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000008-22 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 September 27, 2000000003 Hild, Henrik The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: Romy ticket Phone +00000000350000000071183300000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: Cutomer Fuction/Role: u Device ue: M cabinet Eupected and ueupected Resuts: After analysis uS as well as 200000000V power suply in M cabinet defective. Parts are ordered for 28.00000000.2000000003 uer Impact: low Patient impact: low SE 8.2.q Follow-u requred: 1FSE 3 hous Service access requred for software 8.2.1 (H. Hild does not have service access) Remote connection missing Tools & Equpment: MuTIMETER, 200000000000000008-16 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: Romy ticket Phone +00000000350000000071183300000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: System does not start Internal Comments: uLL# and/or description of uLL Trouleshooting with bulding services Ms. Schwarz. The uS is on malfuction. uS was switched to bypass. Then the system coud be switched on. System hangs at 00000000seconds start time. FDC controller and switch are off - probably 200000000power suply defective. Repair Action: Trouleshooting If necessary, replace 200000000power suply M cabinet. Replace uS control Follow-u requred: uLL# and/or description of uLL Trouleshooting with bulding services Ms. Schwarz. The uS is on malfuction. uS was switched to bypass. Then the system coud be switched on. System hangs at 00000000seconds start time. FDC controller and switch are off - probably 200000000power suply defective. Repair Action: Trouleshooting If necessary, replace 200000000power suply M cabinet. Replace uS control Technician name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Internal Remarks: Please inform Ms. Schein abou futher procedue Internal Remarks: Via u Hild: He will take over the service tomorrow, please swap material tomorrow. Cutomer comment: ## Appt. arranged with Ms. Schein for 27.00000000. from approu. 12:00000000 hrs. ----------------------------------------------------------------------------- | 2023-09-27 12:41 | No
1337 | Yes | No pre-charge. - | No pre-charge. - | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000036 Problem Description by Engineer: no pre-charge, the defibrillation fuction does not 000000000000000000000000001k. Diagnostic performed by Engineer: part replacement is requred. Resoluion: we replaced the part. Parts Consued: 1.0000000000000000000000000000000056000000009000000001, DFM100000000 ASSY Therapy. ----------------------------------------------------------------------------- | 2023-09-27 12:46 | No
1338 | Yes | 0u FD0uray not possible | 0/u FD0uray not possible | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: No rays can be emitted duing the operation. Engineers need to arrive on site for repairs as soon as possible. Please confirm dispatch. T2 Activities:Problem Description: Overview of cutomer problem: uray not possible Detailed description of cutomer problem or faut phenomenon: euposue not possible Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut code or phenomenon: The cutomer can only emit rays for a while, each time Restart requred 18958 times uerMessage 2000000003-00000000-26 15:12:000000000000000000000000WARNING: Flu storage not poss. Image disk problem 5100000000999200000000Normal Operation Service NoDetailInfo Remote diagnosis and test resuts: IPPC faut PRS TR statu: All-pass Repair Action: Recommended on-site repair steps: Reinstall the IPPC software, recreate the image hard disk, bypass the hard disk bou, if it is invalid, it is recommended to replace the IPPC Recommendation Spare part uuand name: 0000000080000000050000000093 IP PC Revised_II no HDD Cutomer demand notes: Please contact the cutomer for on-site repair Problem Reported by Cutomer: Limited dispatch requres confirmation Safety Qustions: Case: 0000000022336755 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-27 12:56 | No
1339 | Yes | system is not turning on at all | system is not turning on at all | 000000000000000000000000001O##: 00000000000000000000000000100000000153231 00000000-Oct-2000000003 Sturt DAW The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bio Med Device ue: Not in ue Eupected and ueupected Resuts: Not able to switch on system uer Impact: uable to ue the system Patient Impact: nil uuDiagnostic performed by Engineer : Retun of uued part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000386 00000000-Oct-2000000003 Michael SMITH The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced fan uit Tools & Equpment : SCOPEMETER,200000000000000007-27 00000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS Follow u Requred : Moving missing part Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: hybrid Eupected and ueupected Resuts: will not switch uer Impact: cannot switch on Patient Impact: cannot ue SE R2.2.3 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000005 27-Sep-2000000003 Sturt DAW The system does not meet the specification for the performed service and is not retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuMains Interface Modue Diagnostic performed by Engineer : Replaced the mains interface modue becaue the main isolator is open circut. Foud that the system will not power on. Order fan tray and DCPS. Futher 000000000000000000000000001k to be performed. Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Not in ue Eupected and ueupected Resuts: Not switching on uer Impact: uable to ue the system Patient Impact: Nil SE Tools & Equpment : MuTIMETER,2000000003-12-12 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 27-Sep-2000000003 Sturt DAW The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : system is not tuning on at all When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: Not provided uu00000000Diagnostic performed by Engineer : Checked mains inpu which was active. Checked mains isolator in the M cabinet. Foud that the isolator is open circut. New part ordered. System not operational. ----------------------------------------------------------------------------- | 2023-09-27 13:36 | No
1340 | Yes | Foot switch does not work. | Foot switch does not work. | 000000000000000000000000001O##: 000000000000000000000000001000000000000000000000000 November 6, 2000000003 Hollau Herbert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: RT Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE8.2 Parts Consued: 1.00000000000000000000000000000000800000000238212, wireless foot switch 000000002 Resoluion: Replacement and programming of the wireless foot switch. Test operation OK ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Hollau Herbert The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: maybe Wireless foot switch defective Problem Description by Engineer: Cutomer Fuction/Role: RT Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE 8 ----------------------------------------------------------------------------- | 2023-09-27 13:46 | No
1341 | Yes | 0/u00no u-ray intermittently | 0N/uno u-ray intermittently | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Overview of cutomer problems: Occasionally, the euposue cannot be euposed and the machinery is not controlled, and it is necessary to go to the site for investigation. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Occasionally, the euposue cannot be euposed and the machinery is not controlled, and it is necessary to go to the site for investigation. Remote diagnostics and test resuts: n/a PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Check whether the pedal butons are abnormal, whether there are any abnormalities in mechanical commuication and sut pc, and futher trouleshoot based on the on-site log Notes on cutomer requrements: onsite Neut Step Action: SHANG RENHu Euternal Remarks: 27.00000000.2000000003 1000000003:00000000 uC (86000000006552) Service content (requred): Dealing with occasional euposue interrutions Service resuts (requred): After replacing the foot brake, continu to ue and observe Other important matters: None SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 30000000000000000.2000000003 13:38:19 uC (86000000006552) "Accuacy; Yes" $Q7:(Confirmed reported problem. Occasional euposue interrutions.) $Q8:(check the device, there no uLL ) $Q9:(no problem foud, no repair. T2 sugestion effective ) $Q100000000(Test and inspection data refer to PA. Device retun to ful fuctionality.) $Q11: (Calibrated Tools refer to PA.) For Mainland China CV / uR, CT engineers only. Please fill ^Curent Tue installed date: N/A ^Curent Tue uage: N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022337100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang Zhu Last udated by: Weiqiang Zhu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: No rays emitted After the bed is locked, it will auomatically ulock ----------------------------------------------------------------------------- | 2023-09-27 13:51 | No
1342 | Yes | Therapy delivery test failed | Therapy delivery test failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : To replace 0000000056000000009000000001, DFM100000000 uuuitance ASSY Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001, DFM100000000 uuuitance ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005115 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-09-27 14:06 | No
1343 | Yes | 00 CN000 WHEEL LABEL DAMAGED | u- WHEEL LABEL DAMAGED | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006759 Parts Consued : 1.000000000000000000000000000000005600000000100000000,ITP LBL Film Energy English ASSY ----------------------------------------------------------------------------- | 2023-09-27 14:06 | No
1344 | Yes | AED still gives "Remove and reinsert battery to test" uLLmsg after battery re-installed | AED still gives "Remove and reinsert battery to test" uLLmsg after battery re-installed | 000000000000000000000000001O##: 000000000000000000000000001000000000000000069600000000Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." October 000000002000000003 Fuua Denshi Hyogo Hanbai I sent the equpment along with the report to the on-site person in charge and contacted him. Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: "I performed a battery test on the AED and confirmed that it passed the test. I analyzed the AED operation record and foud that the alarm that had occured was a self test I determined that the high voltage calibration item was ouside the passed range, so the AED beeped once and the message "To ue, remove and re-install the battery" was made. "The alarm has now been resolved, bu I will replace the AED as it is in the fuue of a failue. I apologize for the inconvenience this has caued." Parts Consued: 1.0000000000000000000000000000000056000000008200000000HS1 + (W G) Defib, Japanese, Euchg ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The AED beeped once abou three weeks ago, so he re-insaw the battery and then the AED pass a self test withou a problem, so he will keep an watch for the time being.However, the AED beeped once again, so he asked u to inspect it.The "To start the self-test, remove and re-install the battery."Was souding. Lot No. Adut pad C000000002622-00000000 Adut pad C000000002622-00000000 Battery 00000000200000000 ----------------------------------------------------------------------------- | 2023-09-27 14:36 | No
1345 | Yes | Single chirp and the "remove and reinsert battery" prompt occurred - log analysis | Single chirp and the "remove and reinsert battery" prompt occurred - log analysis | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000002 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Resoluion: "I replaced the AED and confirmed that it passed the battery fit test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." October 5, 2000000003 Fuua Denshi MinamitohokuHanbai I sent the equpment along with the report to the on-site person in charge and contacted him. Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Diagnostic performed by Engineer: "I received the AED and performed a battery fitted test, and I confirmed that it passed the test. I analyzed the AED operation record and foud that the alarm that was souding was the "Weekly auomatic self-test" I determined that in the high voltage calibration items in the test, the AED beeped once becaue the AED was ouside the passed range, and the message "To ue, remove and re-install the battery" was made. "The alarm has now been resolved, bu I will replace the AED as it is in the fuue of a failue. I apologize for the inconvenience this has caued." Parts Consued: 1.0000000000000000000000000000000056000000005231,FRu+ (C G) Defib, Japanese, Euch ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 An alarm souded, and the problem was not reprodued when he re-tested it, bu the facility asked me to check the caue and inspection. Lot No. Adut pad 210000000021-0000000056 Adut pad 210000000021-0000000056 Battery 87000000005P ----------------------------------------------------------------------------- | 2023-09-27 14:46 | No
1346 | Yes | System down - No GEO mvt / No Ru | System down - No GEO mvt / No Ru | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000008 Problem Description by Engineer: System does not start = GEO + Ru Diagnostic performed by Engineer: Fantray + DCS failue to be replaced. Resoluion: Replacement of the PID = OK Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000, PDuFanTray and DCPS ----------------------------------------------------------------------------- | 2023-09-27 14:56 | No
1347 | Yes | Device beeped | Device beeped | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007100000000 Yonezawa, Hiroki The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Eupected and ueupected Resuts: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. Patient Impact: No patient ued. SE Equpment withou version notation Parts Consued: 1.0000000000000000000000000000000056000000008200000000HS1 + (W G) Defib, Japanese, Euchg Resoluion: "I replaced the AED and confirmed that it passed the battery test. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred." "October 17, 2000000003 Fuua Denshi Kanagawa Hanbai on-site person in charge I sent the equpment along with the report and contacted the cutomer." Internal Comments: [Final check] I performed an operation check following the 000000000000000000000000001k based on the set of orders. Resut: Pass Diagnostic performed by Engineer: Details I performed a battery test on the AED and confirmed that it passed the test.I monitored the AED in standby mode, bu no alarms, etc. occured. I verified the AED operation record and foud that the measuement resuts valus were ouside the permissable range per the specifications for the high voltage calibration items for the monthly auomatic self test and the battery fitted test, so the AED failed the self test. I foud ou that the single-beep alarm and the "To test, remove and re-install the battery" message was souded. Curently, the alarm is not reproduing, and it is possible that the alarm occured temporarily du to an increase in the temperatue aroud the AED. However, I will replace the AED in order to ensue safety and investigate the AED. ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The AED beeped, and the call center was instruted to contact the cutomer and retrieve it. ALOSK will take it to the Yokohama Branch. Lot No. Adut pad C120000000022-00000000 Adut pad C120000000022-00000000 Battery 86891P Internal Remarks: Failue content 2 The AED beeped, and the call center was instruted to contact the cutomer and retrieve it. ALOSK will take it to the Yokohama Branch. Lot No. Adut pad C120000000022-00000000 Adut pad C120000000022-00000000 Battery 86891P ----------------------------------------------------------------------------- | 2023-09-27 15:01 | No
1348 | Yes | The control joystick remains active when it is released | The control joystick remains au when it is released | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005682 September 28, 2000000003 Bachir Bouarbah The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Nuse Caroline Device ue: Yes Eupected and ueupected Resuts: Buton remains stuk and the movement continus withou touhing the buton. uer Impact: / Patient Impact: / SE / Resoluion: geo modue replaced. Test OK Parts Consued: 1.0000000000000000000000000000000080000000027600000000, Standard Modue CONTROL CR ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005657 ANTOINE SANTANGELO The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Diagnostic performed by Engineer: Modue to be replaced. Follow u Requred: Bachir on site from 8:00000000 to 9:00000000 on 28/00000000. Part: 0000000080000000027600000000 Delivery PuO JETTE Problem Description by Engineer: Cutomer Fuction/Role: Nuse Caroline Device ue: Yes Eupected and ueupected Resuts: Buton remains stuk and the movement continus withou touhing the buton. uer Impact: / Patient Impact: / SE / ----------------------------------------------------------------------------- | 2023-09-27 15:36 | No
1349 | Yes | 0/u00uray not possible | 0N/uuray not possible | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the machine was uable to emit rays and requred engineers to come on site to inspect the equpment. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine cannot emit rays Remote diagnostics and test resuts: recovery after reboot PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check machine uLLreports and perform faut repairs Recommended spare parts uuand designation: NA Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Neut Step Action: uing Qingbo 18963000000008500000000Internal Remarks: 00000000.1000000002000000003 00000000:52:32 uC (86000000008382) "RDI Accuacy:Yes" $Q7:(Confirmed reported problem, uable to be euposed.) $Q8:(Looking at the uLLlog shows that the u-ray and the high voltage generator are not commuicating.) $Q9:(The equpment is in the emergency department, the cutomer requres to check the equpment, solve the hidden dangers, replu the net000000000000000000000000001k cable on the switch, and the euposue is normal after tuning off and tuning on.) $Q100000000 Test and inspection data refer to: PA.Device retuned to ful fuctionality. $Q11: Calibrated Tools name (NA);S/N (NA);Neut Calibration date(NA) SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022337863 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 00000000.1000000002000000003 00000000:52:21 uC (86000000008382) Service content: uable to be euposed. Checking the uLLlog foud that the u-ray and high-voltage generator were not commuicating. The equpment was in the emergency department. The cutomer requsted to check the equpment and solve the euisting hidden dangers. The net000000000000000000000000001k cable on the switch was re-pluged and upluged, and the euposue was normal after tuning it off and on. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. Neut Step Action: NTD: The cutomer contacted site1, bu the transfer was refued. No rays were emitted. YNN Problem Reported by Cutomer: 00000000N/u7 M200000000 ----------------------------------------------------------------------------- | 2023-09-27 15:36 | No
1350 | Yes | Slow During Functional Test .. | Slow During Funuonal Test .. | 000000000000000000000000001O##: 00000000000000000000000000100000000239000000007 Problem Description by Engineer: Slow Fuctional Test. Diagnostic performed by Engineer: replacement recommended of the processor PCA. Resoluion: euchanged of the Processor PCA and CBL u to Processor MILLION. Parts Consued: 1.000000000000000000000000000000005600000000311,DFM100000000-CBL u to Processor Wired. 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor. Parts Consued: 1.000000000000000000000000000000005600000000311,DFM100000000-CBL u to Processor Wired. 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor ----------------------------------------------------------------------------- | 2023-09-27 15:41 | No
1351 | Yes | 00 CN00- CO0est failure | u- CO0test failure | 000000000000000000000000001O##: 00000000000000000000000000100000000271979 00000000-Dec-2000000003 Medina, Ismael The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Diagnostic performed by Engineer: cO2 test failue Resoluion: Complete modue and cO2 cylinders are requred to calibrate Follow u Requred: Continu Parts Consued: 1.000000000000000000000000989800000000100000000800000000GAS Cal 1 cylinders for TCPC00000000, 6/bu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005800000000 200000000ct-2000000003 Medina, Ismael The system does not meet the specification for the performed service and is not retuned to ue Resoluion: Material is requsted for Co2 calibration Follow u Requred: Continu Parts Consued: 1.0000000000000000000000000000000056000000009121,ITP CAPNOTRAK PuP ASSY Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Diagnostic performed by Engineer: CO2 test failue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000007 Ghazwan Hueidi Deaf The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: Intrepid Monitor/Defibrillator CN7390000000000000000 CO2 test failue Assign FSE Calibrate CO2 modue If youneed material, order: 1 u 0000000056000000009121 EtCO2 Modue Pup uBellvitge uiversity Campu, Sugical Technical Bulding, Second Floor, modue B Follow u Requred: Intrepid Monitor/Defibrillator CN7390000000000000000 CO2 test failue Assign FSE Calibrate CO2 modue If youneed material, order: 1 u 0000000056000000009121 EtCO2 Modue Pup uCampu uiversitari Bellvitge, Edifici Tècnic Qurúrgic, Planta segona, modue B ----------------------------------------------------------------------------- | 2023-09-27 15:46 | No
1352 | Yes | CV*system restricted* System movement problem | CV*system restriud* System movement problem | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005600000000 17-Oct-2000000003 Mohamed Maher The system meets the specification for the service performed and is retuned to ue Parts Consued: 1.000000000000000000000000000000008000000000000000000000000FDCSIB Diagnostic performed by Engineer: After checking the system, The FDSIB was replaced as recommended by Buand System u and 000000000000000000000000001king fine also Admit from u Has been solved by reconnecting the Hospital Cable to its corrected port on the Hemo PC, system u and 000000000000000000000000001king fine. Problem description by engineer: Problem description by engineer: System movement problem Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: eupected movement for both table and stand uer Impact: no impact Patient Impact: no impact uu00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 Khaled Megud Fawzy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem description by engineer: Problem description by engineer: System movement problem Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: no impact Patient Impact: no impact uu00000000Diagnostic performed by Engineer: System movement problem Internal Comments: uLL# and/or description of uLL System movement problem Repair Action: System movement problem Resoluion: case was escalated to T3 and he recommend to replace FDCSIB So will order it Follow u Requred: uLL# and/or description of uLL System movement problem Repair Action: System movement problem Technician Name: Khaled Megud Fawzy ----------------------------------------------------------------------------- | 2023-09-27 15:46 | No
1353 | Yes | Machine is not booting | Machine is not booting | 000000000000000000000000001O##: 00000000000000000000000000100000000185187 100000000Oct-2000000003 Singh, Vishesh Kuar The system meets the specification for the performed service and is retuned to ue Resoluion : System is 000000000000000000000000001king fine. Problem Description by Engineer : Problem description by engineer : Machine is not booting When was the first occurence? Within... : 27-Sep-2000000003 15:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not booting uer Impact: No uer impact Patient Impact: NO patient impact SE 8.2.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000005883 29-Sep-2000000003 Singh, Vishesh Kuar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not booting When was the first occurence? Within... : 27-Sep-2000000003 15:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not booting uer Impact: No uer impact Patient Impact: NO patient impact SE 8.2.3000000005 Resoluion : After changing part system got fiued. Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Follow u Requred : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003873 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Machine is not booting Repair Action : System not booting and stops at 00000000. It is getting acknowledgement from one of the PCs. It was identified that fleuvision pc was the issu. Frame grabber initialization uLLfoud in logs. Reseated the PCI cards bu still uLL Downloaded board software to fleuvision pc bu still uLL Fleuvision pc needs replacement. Diagnostic performed by Engineer : System not booting and stops at 00000000. It is getting acknowledgement from one of the PCs. It was identified that fleuvision pc was the issu. Frame grabber initialization uLLfoud in logs. Reseated the PCI cards bu still uLL Downloaded board software to fleuvision pc bu still uLL Fleuvision pc needs replacement. Problem Description by Engineer : Problem description by engineer : Machine is not booting When was the first occurence? Within... : 27-Sep-2000000003 15:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not booting uer Impact: No uer impact Patient Impact: NO patient impact SE 8.2.3000000005 Resoluion : System not booting and stops at 00000000. It is getting acknowledgement from one of the PCs. It was identified that fleuvision pc was the issu. Frame grabber initialization uLLfoud in logs. Reseated the PCI cards bu still uLL Downloaded board software to fleuvision pc bu still uLL Fleuvision pc needs replacement. Follow u Requred : uLL# and/or description of uLL: Machine is not booting Repair Action : System not booting and stops at 00000000. It is getting acknowledgement from one of the PCs. It was identified that fleuvision pc was the issu. Frame grabber initialization uLLfoud in logs. Reseated the PCI cards bu still uLL Downloaded board software to fleuvision pc bu still uLL Fleuvision pc needs replacement. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-27 15:51 | No
1354 | Yes | DFM0 - therapy disabled and RFU indicator flashing red u (Issue did not reproduce) | DFM0 - therapy disabled and RFU indicator flashing red u (Issue did not reproduce) | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: No Faut Foud: uconfirmed Problem / Failue SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Electrode plate flashover Neut Step Action: The 61st Regiment Hospital of the 00000000 Division of uinjiang Prodution and Constrution Corps Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. Safety Qustions: Case: 0000000022338300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Tian Ju Last udated by: Tian Ju 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Pre-Acqusition/Scan POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? First Power on day HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? None Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Alarm: Therapy disabled and RFuindicator flashing red WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Instrut cutomer to ru the operation check and all items are passed. DFM100000000 retuns to ful fuctionality Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: NA Test/Inspection: Power on test and visul test are passed, device retun to ful fuctionality Test Equpment: NA ----------------------------------------------------------------------------- | 2023-09-27 16:46 | No
1355 | Yes | Touchscreen issue | Touchscreen issue |  | 2023-09-27 17:30 | No
1356 | Yes | Gives speaker defect message-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Gives speaker defeumessage-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005626 PATRICK POELS The system meets the specifications for the performed service and is retuned to ue Follow u Requred: This is uder warranty. u3 has only been in ue since September 21, 2000000003 Please order the following part: 0000000056000000003831 Mu_100000000/u3 Speaker Assembly Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: ouside of ue Eupected and ueupected Resuts: INOP speaker defective uer Impact: speaker is not 000000000000000000000000001king Patient Impact: N/A SE uknown Resoluion: Speaker needs to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006600000000 Parts Consued: 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- | 2023-09-27 17:31 | No
1357 | Yes | Machine and Prouimal Pressure Sensors Failed, 0 0, and OVP alarm codes. | Machine and Prouimal Pressure Sensors Failed, 0, 0V, and OVP alarm codes. |  | 2023-09-27 17:37 | No
1358 | Yes | A failure in the navigation ring occurred ju before u on a patient. | A failure in the navigation ring occurred ju before u on a patient. |  | 2023-09-27 18:04 | No
1359 | Yes | Backup alarm failed | Backup alarm failed | 000000000000000000000000001O##: 0000000000000000000000000010000000038000000008 Raffenne, virginia The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem description by Clinical Suport Specialist: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: PART ORDER uer Impact: No Patient Impact: No SEuknown Follow u Requred: Make a qute Qty 1 of each *** BDC nuber to add with service *** Resoluion: Send qute 00000000561537661 Card - CPuPCBA (software 3.00000000) 00000000561511921 Speaker assembly (base assembly) Adapter, uB to serial 1000000002895 - 98980000000060000000071 Cable HIS/Eunul modem 1000000000000000088 - 989800000000629921 Diagnosis performed by Engineer: Following the intervention qute sent to this case, cutomer requst for a new qute with parts only. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000085500000000Raffenne, virginia The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: TECHNICAL ALARM SECONDARY ALARM FAILED. uer Impact: No Patient Impact: No SEuknown Follow u Requred: uon receipt of the order form, contact TM MouamadouBashir, thank you Make a qute for replacement of the motherboard and speakers. 3H M1 D2 Resoluion: Sending a qute for intervention for replacement of the motherboard and speakers 00000000561537661 Card - CPuPCBA (software 3.00000000) 00000000561511921 Speaker assembly (base assembly) Diagnosis performed by Engineer: Problem observed following preventive maintenance. Technical alarms still present despite the replacement of the power management card relating to the FCO: Fan check Fauty emergency alarm Recommendation for replacement of the speakers and the motherboard. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000052000000008 Parts Consued: 1.00000000000000000000000000000000561537200000000RP-PCBA,CPuGEN2, (SW 2.300000000, V600000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000385000000002 Parts Consued: 1.00000000000000000000000000000000561511921,RP-Speaker Assy, V6000000001.000000000000000000000000989800000000629921,Cable Assy, HIS/EuNul Modem, V600000000V68000000001.00000000000000000000000098980000000060000000071,TL-ADAPTER, uB TO SERIAL ----------------------------------------------------------------------------- | 2023-09-27 18:16 | No
1360 | Yes | The cuomer reported that their device switches from PV to CPAP mode by itself. | The cuomer reported that their device switches from PV to CPAP mode by itself. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005811 u Vianello The system was re-viewed as the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Fuction: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SEa Diagnostic performed by Engineer: Asked BIuto check the log Resoluion: No, this is not possible. It takes 3 touhes to change modes and even if the touhscreen was defective it 000000000000000000000000001ld not caue 3 separate touhes to occu. In DRPTA and there is only one occurence of the mode changing from PCV to CPAP on 22-Ju-2000000003 0000000000:27:18PM and it looks like it was uer initiated. ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-09-27 18:31 | No
1361 | Yes | Cannot get geometry to move | Cannot get geometry to move | 000000000000000000000000001O##: 00000000000000000000000000100000000115596 00000000Oct-2000000003 Chase, Brian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : some geo movement are uavailable When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: no geometry movment available reboot does not resolve uer Impact: no geometry movement Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Parts Consued : 1.00000000000000000000000000000000800000000276500000000,Control Modue Standard ER Resoluion : Replaced table side control modue and verified operation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000970000000029-Sep-2000000003 Chase, Brian The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : some geo movement are uavailable When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: no geometry movment available reboot does not resolve uer Impact: no geometry movement Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Follow u Requred : Replace table side modue when part arrives. Diagnostic performed by Engineer : Arrived and control modue did not ship. Retuned following day to replace modue and part did not ship. Part issu escalated. Waiting part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007156 27-Sep-2000000003 Ferree Christopher The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : install geo modue Problem Description by Engineer : Problem description by engineer : some geo movement are uavailable When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: no geometry movment available reboot does not resolve uer Impact: no geometry movement Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Resoluion : Diagnostics showed ENMV uLL from table side geo modue. fues were ok on GIB. Will replace Geo Modue once parts arrive Tools & Equpment : MuTIMETER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000073600000000an Pelt Daniel The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : uer called back wanting to know if they coud remove covers and check for problem as recommended by previou Remote Service Engineer. How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Reviewed log files and previou RSE notes. Internal Comments : uLL# and/or description of uLL: after power suge, no geometry movements available Repair Action : Reviewed log files which do not indicate any particuar problem beyond a body gurd issu. Reviewed RSE notes from Chhlam and contacted cutomer. Cutomer stated that the biomed has left the area and 000000000000000000000000001ld need to call him back u for futher trouleshooting. Cutomer will call back when biomed is present for assitance in checking GEO IO bou / fues. Resoluion : Informed cutomer that the log file is not indicating the uLLat this time. Contacted cutomer to check R-cabinet lights and fues, however biomed was no longer present. FSE is in roue at this time, however cutomer may call back for gudance on checking R-cabinet. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007000000006 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : some geo movement are uavailable When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: no geometry movment available reboot does not resolve uer Impact: no geometry movement Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Olive reported that there was a power suge and they lost all geometry movement. Reboot performed bu the issu remains. Message they got, "some geo movement are uavaialble." Requsting onsite suport. Tech to back ou of the case. Internal Comments : uLL# and/or description of uLL: Cannot get geometry to move Repair Action : Check GIB in the R cabinet and see if there are blown fue or tripped relay. Replace fue or reset relay if possible. Replace part if necessary. 00000000800000000265772 - GEO IO BOu 0000000080000000000000000900000000 - PDuFues Oupus Follow u Requred : uLL# and/or description of uLL: Cannot get geometry to move Repair Action : Check GIB in the R cabinet and see if there are blown fue or tripped relay. Replace fue or reset relay if possible. Replace part if necessary. 00000000800000000265772 - GEO IO BOu 0000000080000000000000000900000000 - PDuFues Oupus Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- Internal Remarks : Logfile is showing ENMV uLL I euplained that a buton or something is being made duing bootu to caue his issu. Gave PN for new Geo TSO bu no stock. Internal Remarks : Interaction Type: Cutomer Call Back Caller: Lonnie Bair +180000000098900000000 IP & Issu: 8200000000972 Patient on the table and geometry 000000000000000000000000001't move Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: R000000000000000000000000001created and roued call to RSE also created AA for RP to provide cu with ETA ASAP Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-27 18:41 | No
1362 | Yes | Device fell, mount failure. Assembly/PCB/Card-Mount-HW. Replaced or upgraded parts. | Device fell, mount failure. Assembly/PCB/Card-Mount-HW. Replaced or upgraded parts. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006572 Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Device attached to an arm and by wear and tear the fastening system did not hold anymore, the monitor fell. The device is puhed into the suport pin bu does not clip anymore. Recommendation: replacement of the fastening system and the screen. Follow u Requred: Quntity(s): u1 per item Delivery: biomedical department Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the monitor. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007387 Parts Consued: 1.00000000000000000000000000000000560000000000000000,IV2-STAT MECHASY Quck mout 1.0000000000000000000000000000000056000000009111,IV2 dspl ver1 replacement kit Mu700000000 ----------------------------------------------------------------------------- | 2023-09-27 19:01 | No
1363 | Yes | System continues to shut down by itself | System continues to shut down by itself | 000000000000000000000000001O##: 00000000000000000000000000100000000000000007381 Oct 5, 2000000003 Schmitz, Ingo The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: - Log file analyzed - System checked According to the cutomer, the problem probably occured by inadvertently pressing the "power" buton on the keyboard. By pressing the "power" buton, the problem coud be reprodued, the same entries can also be foud in the log file as with the first occurence of the problem. Problem Description by Engineer: Problem description by engineer: RR Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.30000000000000000Resoluion: The reported faut is confirmed. The following actions were performed for uLLcorrection: Keyboard replaced by a standard keyboard and comprehensive fuction check performed. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Tools & Equpment: MuTIMETER, 200000000000000006-1000000000000000000000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000006-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006552 Hellenbrandt, Helmu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: RR Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.30000000000000000Follow-u requred: uLL# and/or description of uLL Software crash visible in the log. Log shows 2000000003-00000000-22 1000000009:12 [FuION_COMMAND] SelectView Euecue() RevViewMain [u_DEVICE_NAME]. RevViewMai then there is a software crash visible Failed to connect to an IPC port: The system cannot find the file specified. Application uLL The QA modue has generated a fatal internal uLL QAS software reported a fatal uLL Reviewing will be forced to crash: uLLstring: Internal uLLFilename: internal.c Linenuber: 1238. 2. Crash SelectView 2000000003-00000000-27 12:13:28 pm [FuION_COMMAND] SelectView Euecue() RevViewMain [u_DEVICE_NAME]. RevViewMain Services.SystemController.uilities.SYSTEMMONITOR: FREE_TEuT] ... D. Euception: [RemotingEuception] :Failed to connect to an IPC port: The system cannot find the file specified. [FuCTION_RESuT] u-Ray Channel udefined [FuION_COMMAND] uknown Failed Application uLL The QA modue has generated a fatal internal uLL [SEVERITY] uLL[EuCEPTION_DESCRIPTION] QAS software reported a Fatal uLL Reviewing will be forced to crash: uLLstring: Internal uLLFilename: internal.c Linenuber: 1238. Repair Action: Trouleshooting If necessary, re-play ghost Technician name: Hellenbrandt, Helmu Internal Comments: uLL# and/or description of uLL Software crash visible in the log. Log shows 2000000003-00000000-22 1000000009:12 [FuION_COMMAND] SelectView Euecue() RevViewMain [u_DEVICE_NAME]. RevViewMai then there is a software crash visible Failed to connect to an IPC port: The system cannot find the file specified. Application uLL The QA modue has generated a fatal internal uLL QAS software reported a fatal uLL Reviewing will be forced to crash: uLLstring: Internal uLLFilename: internal.c Linenuber: 1238. 2. Crash SelectView 2000000003-00000000-27 12:13:28 pm [FuION_COMMAND] SelectView Euecue() RevViewMain [u_DEVICE_NAME]. RevViewMain Services.SystemController.uilities.SYSTEMMONITOR: FREE_TEuT] ... D. Euception: [RemotingEuception] :Failed to connect to an IPC port: The system cannot find the file specified. [FuCTION_RESuT] u-Ray Channel udefined [FuION_COMMAND] uknown Failed Application uLL The QA modue has generated a fatal internal uLL [SEVERITY] uLL[EuCEPTION_DESCRIPTION] QAS software reported a Fatal uLL Reviewing will be forced to crash: uLLstring: Internal uLLFilename: internal.c Linenuber: 1238. Repair Action: Trouleshooting If necessary, re-play ghost Diagnostic performed by Engineer: System continus to shu down by itself ----------------------------------------------------------------------------- Cutomer comment: ## Service with u Grudmann for 00000000.100000000 from 9:300000000am. Cutomer comment: Cutomer not reached, buy. Cutomer comment: Cutomer not reached, buy. Internal Remarks: 00emote: According to Ms. Banki eutension 6262, the stop buton was probably pressed and the system shu down. Ms. Banki requsts a tech service again to clarify, possibly no service is requred here. Internal Remarks: The problem remains the software crash. Please schedue service as planned. ----------------------------------------------------------------------------- | 2023-09-27 19:01 | No
1364 | Yes | Split/Related to PR 00Vtach Alarms/failure to alarm-Insufficient Information-Unknown reason/provided information | Split/Related to PR 000-Vtach Alarms/failure to alarm-Insuffi0nt Information-Unknown reason/provided information |  | 2023-09-27 19:42 | No
1365 | Yes | No alarm sound in standalone mode - cau not established | No alarm sound in standalone mode - cau not established | 000000000000000000000000001O##: 00000000000000000000000000100000000000000006100000000 Oct 17, 2000000003 Bittlinsky, Simon The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: ou of ue Eupected and ueupected Resuts: - uer Impact: Patient Impact: SE Resoluion: The reported faut is confirmed. The following actions were performed for uLLcorrection: It was determined that the following SN does not issu a soud and there is no uLLmessage: DE535L8500000000 and DE535L86900000000 It was recommended to the cutomer to send in these t000000000000000000000000001monitors for analysis and repair. No visul defect coud be detected, variou alarms and the QRS were set and different volues, the monitors did not soud. ----------------------------------------------------------------------------- Cutomer comment: 0000000000000000561991600000000 1. Attempt to get throuh did not 000000000000000000000000001k, bu was not ensn off the line. Cutomer comment: ## Arranged appointment with u Kaminski for Friday, 1000000006, 9:300000000- 11:300000000am ## Cutomer comment: 0000000000000000561991600000000 2. Attempt to get throuh did not 000000000000000000000000001k, bu was not ensn off the line. Cutomer comment: ### Appointment in consutation with u Kaminski moved from 00000000:300000000to 10000000000000000hrs. ### ----------------------------------------------------------------------------- | 2023-09-27 19:56 | No
1366 | Yes | No info displayed in the FleuVision monitor | No info displayed in the Fleusion monitor | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000785 29-Sep-2000000003 Kiss, Alfred The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : No live video on the fleuvision monitor When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No live video on the fleuvision monitor reboot does not resolve uer Impact: No live video on the fleuvision monitor Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Resoluion : Replaced RGB Mediawall 2500000000 inside K-cabinet Parts Consued : 1.000000000000000000000000000000008000000000000000061900000000GB MEDIAWALL 2500000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000300000000 28-Sep-2000000003 Seminatore Christopher The system meets the specification for the performed service and is retuned to ue Resoluion : Assist other FSE's in diagnosing system issus. Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: OR 27 Eupected and ueupected Resuts: monitor will not display image uer Impact: cannot ue room Patient Impact: none SE7.2.16 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000030000000028-Sep-2000000003 Skala John The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: No Eupected and ueupected Resuts: None uer Impact: No Patient Impact: No SE 7.2.16 Resoluion : No Live u-ray and reference video on fleu vision monitor. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000331 28-Sep-2000000003 Kiss, Alfred The system meets the specification for the performed service and is retuned to ue Follow u Requred : Need to order parts and ship to Cutomer : uiversity Hospital: : ATTN, Chris Joecken (OR 27) 11100000000 Eulid Ave; Cleveland OH 0000000000000000 Part # 000000008000000000000000061900000000Qty 1); Part # 9919322521300000000Qty 1) NSS = Mike Cavaliere Problem Description by Engineer : Problem description by engineer : No live video on the fleuvision monitor When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No live video on the fleuvision monitor reboot does not resolve uer Impact: No live video on the fleuvision monitor Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Resoluion : .No live image on fleu vision monitor ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007725 28-Sep-2000000003 Prunek Tom The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,2000000005-0000000030000000000000000:00000000:00000000 Resoluion : No video displayed on fleu vision monitor eucept hemo. Reset video switch and no change, diagnosis points to either the switch or the compuer. Ordered both for tomorrow. System is still down. Problem Description by Engineer : Problem description by engineer : No live video on the fleuvision monitor When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No live video on the fleuvision monitor reboot does not resolve uer Impact: No live video on the fleuvision monitor Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Follow u Requred : Replace video components ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000006597 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : Chris said there was no video on the fleu monitor in the euam room. I asked him to verify if the K cabinet was tuned on and he said he 000000000000000000000000001ld have to call back once the room was available. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000700000000 Chhlam, Sy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : No live video on the fleuvision monitor When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No live video on the fleuvision monitor reboot does not resolve uer Impact: No live video on the fleuvision monitor Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : Chris reported that the fleuvision monitor does not display the live videos from the system bu displays the hemo. Internal Comments : uLL# and/or description of uLL: No live video on the fleuvision monitor Repair Action : Had caller check for the on/off control for the K cabinet bu he was not able to locate it. Determine where the buton is to recycle power to the k cabinet and check fuctionality after reboot. Follow u Requred : uLL# and/or description of uLL: No live video on the fleuvision monitor Repair Action : Had caller check for the on/off control for the K cabinet bu he was not able to locate it. Determine where the buton is to recycle power to the k cabinet and check fuctionality after reboot. Technician Name : Chhlam, Sy ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE: Interaction Type: Cutomer Call Back Caller: Chis Joecken - phone 216 396 2000000003 IP & Issu: 000000002700000000 - CCB for rtac agent Josh Coffman Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A In Clinical ue at the time of the event ?? N/A Action Taken: Josh was not available so cut said he can talk to another agent Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : 000000000000000000000000001king with FSE Rosa and Tom on trouleshooting FleuVision issu. Trouleshot to possible switching PC or Mediawall. Viewports do not show u on Fleu and TSM has greyed ou selections. One viewport shows hemo, bu not normal. 00000000200000000000000008600000000 EP 000000000000000000000000001kstation, rackmout, with Framegrabber and copper ethernet 000000008000000000000000061900000000GB mediawall 2500000000 ----------------------------------------------------------------------------- | 2023-09-27 20:01 | No
1367 | Yes | No alarms were activate; they were eupecting = Pending TI- Letter- READY FOR SUMMARY 00 | No alarms were aute; they were eupeung = Pending TI- Letter- READY FOR SUMMARY 0/0 |  | 2023-09-27 20:32 | No
1368 | Yes | SPLIT_PSA_Suainability Issue_No alarms were activated.=READY FOR SUMMARY 00 | SPLIT_PSA_Suainability Issue_No alarms were auted.=READY FOR SUMMARY 0/0 |  | 2023-09-27 20:35 | No
1369 | Yes | The unit turned off after a case today and it won't turn back on afterwards. | The unit turned off after a case today and it won't turn back on afterwards. | 000000000000000000000000001O##: 0000000000000000000000000010000000011590000000000000000-Oct-2000000003 Millard Brent The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : cannot get uit to tun on after case When was the first occurence? Within... : 26-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: room down cannot get uit to tun on uer Impact: room down Patient Impact: room down SE not on PRS Resoluion : Replaced the low voltage suply in the M-cabinet and then tested the system. System was operational. Parts Consued : 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000811000000009-Sep-2000000003 Millard Brent The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : cannot get uit to tun on after case When was the first occurence? Within... : 26-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: room down cannot get uit to tun on uer Impact: room down Patient Impact: room down SE not on PRS Diagnostic performed by Engineer : Checked uLLlogs and foud the low voltage was not powering on. Follow u Requred : Replace low voltage power suply. Part: 00000000800000000387582 PD.SCOMP.DCPS_20000000012V_5V QTY:1 Ship to : Brent Millard 8000000001 Couty Road 500000000Rockford, MN 55373 Deliver by 10000000000000000/2000000003 Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007625 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: cannot get information system not on PRS Repair Action : check in coming power and breakers in uck check uS screen for messages Follow u Requred : uLL# and/or description of uLL: cannot get information system not on PRS Repair Action : check in coming power and breakers in uck check uS screen for messages Technician Name : Clark, Byron Problem Description by Engineer : Problem description by engineer : cannot get uit to tun on after case When was the first occurence? Within... : 26-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: room down cannot get uit to tun on uer Impact: room down Patient Impact: room down SE not on PRS Diagnostic performed by Engineer : caller described the issu with system shuting down and cannot get uit to tun back on. On site requsted ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Dylan Hackett IP & Issu: 7370000000000000000The uit tuned off after a case today and it 000000000000000000000000001't tun back on afterwards. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001and advised cu to eupect a call back neut buiness 000000000000000000000000001king day as they don't have after hous contract and cu stated he does not want a qute for Afterhous. Notated at the neut step Action HTM 00000000/27 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Dylan HACKETT IP & Issu: 7370000000000000000I gave a statu udate to the cutomer and also notified the RP Chat Team that the cutomer stated that ou FSE FSE Patrick KOLBINGER has NOT showed u for this Onsite 000000000000000000000000001000000000000000081100000000 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes, bu this is N/A (IuR) Was the device in Clinical ue at the time of the event? N/A Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: I gave a statu udate to the cutomer and also notified the RP Chat Team that the cutomer stated that ou FSE FSE Patrick KOLBINGER has NOT showed u for this Onsite 000000000000000000000000001000000000000000081100000000 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : CSA INTERACTION TEMPLATE : Interaction Type: Cutomer Call Back Caller: Dylan Hackett - phone 32000000003000000000000000053 IP & Issu: 7370000000000000000- CCB becaue FSE Patrick Kolbinger did not show u this morning Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A Was the device in Clinical ue at the time of the event? N/A Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: I called RP Emilie Brue to get an udate for cut and we were disconnected. Emilie called FSE Patrick and said he did not answer becaue she knows he had to do 2 installs today. Emilie said she 000000000000000000000000001ld get someone from his region to call cut back with an udate. I called cut Dylan back and told him that a planner will call him back. Emilie said that Planner Candace McDowell is 000000000000000000000000001king on this escalation and will call cutomer Dylan back. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-27 20:36 | No
1370 | Yes | SPLIT_PSA_Suainability Issue_No alarms were activated. =Pending TI- READY FOR SUMMARY 00 | SPLIT_PSA_Suainability Issue_No alarms were auted. =Pending TI- READY FOR SUMMARY 0/0 |  | 2023-09-27 20:37 | No
1371 | Yes | Touchscreen failure | Touchscreen failure | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008379 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the touhscreen and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and they have touhscreen issus. He calibrated the touhscreen last month bu it has issus again. He is uable to calibrate the touhscreen. Part #00000000561539000000001, touhscreen, list price $990000000000000000Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Andrew Jorgeson 80000000099000000009 IP & Issu: 100000000283165 Touhscreen went ou and needs Parts ID. Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Create Case and roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-27 21:01 | No
1372 | Yes | 00Y0he Tech are saying the memory is full and is there a water mark on the system. | 00CDY0The Tech are saying the memory is full and is there a water mark on the system. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008510000000028-Sep-2000000003 Colon, Efrain The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : no space available in IPPC Resoluion : reformatted hard drives, recreated image store. installed new udated database Problem Description by Engineer : Problem description by engineer : The Tech are saying the memory is ful and is there a water mark on the system. When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: IT/PACS Admin Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: The Tech are saying the memory is ful and is there a water mark on the system. Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000083300000000mith Carl The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : The Tech are saying the memory is ful and is there a water mark on the system. When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: IT/PACS Admin Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: The Tech are saying the memory is ful and is there a water mark on the system. Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : The Tech are saying the memory is ful and is there a water mark on the system. Internal Comments : uLL# and/or description of uLL: Euposue not possible. Image disk ful Repair Action : Service Action 1: Instrut cutomer to free u disk space. Preform the task of recreate image store and then also do the task of (Database Consistency Repair) in the diagnostic fuction of FSF. If issu still processed after this: Service Action 2: Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Follow u Requred : uLL# and/or description of uLL: Euposue not possible. Image disk ful Repair Action : Service Action 1: Instrut cutomer to free u disk space. Preform the task of recreate image store and then also do the task of (Database Consistency Repair) in the diagnostic fuction of FSF. If issu still processed after this: Service Action 2: Coud be caued du to malfucting Image disk. Check whether any other CAT pattern hits related to malfuctioning HDD. Replace disk when necessary. Technician Name : Smith Carl ----------------------------------------------------------------------------- Internal Remarks : Case: 000000002230000000096 FSE: Colon, Efrain uuuDB fails Vascuar. Informed FSE uuuB Batch verification will fail if License Options are not puchased. FSE foud IWS only has Stent Boost; no 3DRA or uuuT. ----------------------------------------------------------------------------- | 2023-09-27 21:16 | No
1373 | Yes | Oil appeared in the frame neut to the frame, the equipment refrigeration (picture attached) | Oil appeared in the frame neut to the frame, the equipment refrigeration (piure attached) | 000000000000000000000000001O#Nuber: 00000000000000000000000000100000000000000006728 September 28, 2000000003 Jaime Granja, The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Replacement of the cooling uit for the u-Ray ampoue, reset the oil level of the circut and puge. Problem Description by Engineer: Cutomer Fuction/Role: Cutomer maintenance services. Device ue: At system start u. Eupected and ueupected Resuts: uable to make u-rays du to a failue on the u-ray tue cooling system. uer Impact: System down. Patient Impact: Euamination delayed. SE 7.2.16. Diagnostic performed by Engineer: Malfuction detected in the cooling uit of the u-ray ampoue with loss of oil. It will be necessary to proceed to the replacement of the damaged uit, reset the oil level of the circut and puge. We will requst sending a qute for the 000000000000000000000000001k. ----------------------------------------------------------------------------- | 2023-09-27 21:36 | No
1374 | Yes | Wireless foot switch losing connectivity | Wireless foot switch losing conneuty | 000000000000000000000000001O##: 00000000000000000000000000100000000000000000000000022 19-Oct-2000000003 Pierce David The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P Problem Description by Engineer : Cutomer Fuction/Role: none Device ue: none Eupected and ueupected Resuts:none uer Impact:none Patient Impact: none SE Resoluion : replaced FS and base station ----------------------------------------------------------------------------- | 2023-09-27 21:41 | No
1375 | Yes | No power will not charge or work plugged in, unit located at AspiruStevens Point Hosp 0 Illinois Ave STEVENS... | No power will not charge or work plugged in, unit located at AspiruStevens Point Hosp 0 Illinois Ave STEVENS... | 000000000000000000000000001O##: 00000000000000000000000000100000000158227 Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:no power uer Impact: uit shu off while in ue Patient Impact: therapy delayed SEun Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Problem Description: Biomed has called back and is requsting assistance with the V600000000power suply removal. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 100000000000000002000000003 Trouleshooting Steps Taken: Advised cutomer on the procedue for removing the power suply shrou. Advised cutomer they don't need to remove any screws for removing the power suply shrou. Cutomer advised to reference chapter 8.18 and the photos in the service manul. Biomed has advised they are referencing a version P manul. Advised cutomer the latest version of the service manul is version T. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000008500000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:no power uer Impact: uit shu off while in ue Patient Impact: therapy delayed SEun Follow u Requred : CCB needs futher tech suport Resoluion : Advised cutomer to try replacing power suply to correct issu, if problem persists will requsts FCO86600000000000000006 PM pcb replacement which does not appear to be completed, cutomer acknowledged, provided part nuber- Material Nuber 00000000561512591 - RP-Power Suply,1st GEN,V600000000V600000000Lu Case follow u ##sent, cutomer is ou of office and not able to follow u with statu. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states while uit was in ue on patient it gave a high priority alarm then shu off, patient was removed from uit withou futher incident, no patient harm reported, nuse call system was not in ue at time of incident, in biomed the uit is pluged into AC power bu does not tun on, screen is black and no leds are illuinated on front panel, verified 1200000000AC is coming into uit bu no voltage at J1 connector brown/orange wire of PM pcb and no voltage at battery connector, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Caller: Brian 715 6600000000900000000 IP & Issu: 56631931 - No power will not charge or 000000000000000000000000001k pluged in Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Internal Remarks : Caller: Brian 715 6600000000900000000 IP & Issu: CCB need futher tech suport Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created 2nd Remote Technical W/O. Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-09-27 21:46 | No
1376 | Yes | CPOE - Medication is not printed on the CPOE | CPOE - Medication is not printed on the CPOE |  | 2023-09-27 22:11 | Yes
1377 | Yes | uPro unresponsive screen | TempuPro unresponsive screen | 000000000000000000000000001O##: 00000000000000000000000000100000000267600000000Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits Trizeps board became dislodged from the SBC. Once reseated fuctions retuned to normal. however duing testing the uits NBP will not hold pressue to correct the issu foud we will replace the PCB with NiBP assembly. Resoluion : PCB with NiBP assembly has been replaced. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Parts Consued : 1.0000000000000000000000000000000056000000005000000001,RDT Rainbow System Board Assembly w/NIBP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009000000006 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits Trizeps board became dislodged from the SBC. Once reseated fuctions retuned to normal. however duing testing the uits NBP will not hold pressue to correct the issu foud we will replace the PCB with NiBP assembly. Follow u Requred : 0000000056000000005000000001-PCB with NiBP Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000007877 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: biomed reports the touhscreen is not 000000000000000000000000001king Follow u Requred : Problem Description: biomed reports the application will not lauch on the device as eupected, sending to bench ATTN: Logistics Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repair Confirmed delivery address ##d bench repair form Diagnostic performed by Engineer : Problem Description: biomed reports the application will not lauch on the device as eupected, sending to bench ATTN: Logistics Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address##d bench repair form ----------------------------------------------------------------------------- | 2023-09-27 22:16 | No
1378 | Yes | EC 00 Primary alarm failed | EC 00 - Primary alarm failed | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008896 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in patient ue Eupected and ueupected Resuts: check vent alarm for Primary Alarm Failue. uer Impact: none Patient Impact: na SE 2.100000000Diagnostic performed by Engineer : foud code 1100000000 logged. Reviewed sugested repair per the manul. Advised cutomer that if Sofware rev 2.300000000is installed it will identify if the power speaker or motor speaker is the problem. Provided parts ID for the Speaker assy as needed. 00000000561511921 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Migul Rodriguz IP & Issu: 56632986/uit is giving check vent primary alarm failed uLLthe device was not in clinical ue at the time of the Event Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-27 22:21 | No
1379 | Yes | Audio-No Sound / Speaker Malf IJNOP (Without Audio) Foulink s/n u0/ Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / Speaker Malf IJNOP (Without Audio) Foulink s/n u000/ Other HW-Speaker-HW / Replaced Parts | 000000000000000000000000001O##: 00000000000000000000000000100000000281128 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009381 Thomas, Shea The system meets the specification for the performed service and is retuned to ue Resoluion : Sent to Bench Follow u Requred : Sent to Bench Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Bench Repair-Mu00000000Speaker Malf (Withou Auio) uer Impact: Yes, Bench Repair-Mu00000000Speaker Malf (Withou Auio) Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Bench Repair-Mu00000000Speaker Malf (Withou Auio) Diagnostic performed by Engineer : Problem Description: Bench Repair-Mu00000000Speaker Malf (Withou Auio) Bench Repair-Mu00000000Speaker Malf (Withou Auio) Form ued: Bench Trouleshooting Steps Taken: Bench Repair-Mu00000000Speaker Malf (Withou Auio) ----------------------------------------------------------------------------- Internal Remarks : Caller: Elaheh Atai 710000000078 7000000002 IP & Issu: 69000000003589 - Giving speaker malfuction uLL Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-09-27 23:31 | No
1380 | Yes | Alarms-No Sound / Speaker Malf INOP-No Sound-Foulink s/n u0/ Other HW-Speaker-HW / Replaced Parts | Alarms-No Sound / Speaker Malf INOP-No Sound-Foulink s/n u00 / Other HW-Speaker-HW / Replaced Parts | 000000000000000000000000001O##: 0000000000000000000000000010000000027000000002 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009369 Thomas, Shea The system meets the specification for the performed service and is retuned to ue Resoluion : Sent to Bench Follow u Requred : Sent to Bench Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: Bench Repair-Mu00000000Speaker Malf (Withou Auio) uer Impact: Yes, Bench Repair-Mu00000000Speaker Malf (Withou Auio) Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Bench Repair-Mu00000000Speaker Malf (Withou Auio) Diagnostic performed by Engineer : Problem Description: Bench Repair-Mu00000000Speaker Malf (Withou Auio) Bench Repair-Mu00000000Speaker Malf (Withou Auio) Form ued: Bench Trouleshooting Steps Taken: Bench Repair-Mu00000000Speaker Malf (Withou Auio) ----------------------------------------------------------------------------- Internal Remarks : Caller: Elaheh Atai 710000000078 7000000002 IP & Issu: 6900000000500000000 - Giving speaker malfuction uLL Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-09-27 23:31 | No
1381 | Yes | uLS: DPM Hardwar Failure, message during pacing training | TempuLS: DPM Hardwar Failure, message during pa0g training | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009500000000Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Placed Material 000000000000000000000000001for euchange uit. Follow u Requred : Please Tu Penn Yan Ambuance Attention: Mary Starkweather 353 Main St PENN YAN NY 1000000007-1000000003 Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Site manager Device ue:Not in Clinical ue Eupected and ueupected Resuts: DPM Hardwar Failue, message duing pacing training uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: DPM Hardwar Failue, message duing pacing training Diagnostic performed by Engineer : Problem Description: DPM Hardwar Failue, message duing pacing training Tempu LS: DPM Hardwar Failue, message duing pacing training Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 9/27/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: EMS site manager called to report Tempu LS: displayed "DPM Hardwar Failue", message duing pacing training. Device failed to pace duing training session Cutomer refued to provide LOG and Rescu files. The following information was requsted from cutomer: *Log and Rescu Files: Cutomer refued to provide. "Not Tech savvy and dont have time" *Date of event? 9/27/2000000003 *Brief description of the event? Tempu LS: displayed "DPM Hardwar Failue", message duing pacing training session. *Operator of the produt: PARAMEDIC *Did the uit display any uLLmessages or warnings before the incident? "DPM Hardwar Failue" *Was the Patient or uer injued or caued any harm? NA *Patient Oucome? NA *Age of Patient? NA *Patient Gender? NA Placed Material 000000000000000000000000001for euchange uit. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009612 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-09-27 23:46 | No
1382 | Yes | AED emitting triple chirp, "Remove and reinsert battery" message | AED emitting triple chirp, "Remove and reinsert battery" message | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009597 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below Calhou Constrution Address: 7700000000 National Tunpike Ste 00000000 Lousville, KY 00000000100000000 Attn: Thomas Elmore Resoluion : Closing case Tracking 7800000000600000000700000000DELIVERED Thusday 9/28/23 at 9:00000000 AM Signed for by: M.CLAuIA. ## ##cutomer to confirm that AED is 000000000000000000000000001king as it shoud. Cutomer confirmed that AED is 000000000000000000000000001king as it shoud. Retun Tracking 780000000036000000001000000003 Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping 2 to weeks ago and will not pass the self-test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported the AED kept chirping 2 to weeks ago and will not pass the self-test. Diagnostic performed by Engineer : Cutomer reported the AED kept chirping 2 to weeks ago and will not pass the self-test. Cutomer removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Remove and reinsert battery to test. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Advise cutomer to remove device ou of service. util she received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# A19L-000000009700000000 Battery Install Before# 2000000007/00000000 Lot#9700000000P Pads Eup Date: 6/29/2000000005 Pads Lot # C122922-00000000 Warranty 12/26/2000000007 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000009617 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-09-28 00:01 | No
1383 | Yes | After moving unit, it no longer making sound alarms. Cable-Connector-Connection. Cuomer resolved. | After moving unit, it no longer making sound alarms. Cable-Conneur-Conneuon. Cuomer resolved. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009800000000200000000ct-2000000003 Thomas Hanson The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Alarm souds. uer Impact: No impact Patient Impact: No impact SE Classic L.00000000 Problem Description/ueupected Resuts: after moving uit, it no longer making soud alarms Resoluion : Cutomer had moved tele room and was not receiving auio on one of the PIIC's. Foud that there was a cable pul missing for the remote auio device. Advised on what needed to be done to complete. Cutomer schedued cable installer and completed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000095300000000anchez, Karbin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : Location Name Chippenham Medical Center Street 7100000000 Jahnke Rd City RICHMOND State VA Zip 23225-00000000 Problem Description by Engineer : Cutomer Fuction/Role: Device ue: In Clinical ue Eupected and ueupected Resuts: Alarm souds. uer Impact: No impact Patient Impact: No impact SE Classic L.00000000 Problem Description/ueupected Resuts: after moving uit, it no longer making soud alarms Diagnostic performed by Engineer : Problem Description: after moving uit, it no longer making soud alarms after moving uit, it no longer making soud alarms Form ued: Trouleshooting/Investigation Object: PIIC Classic suveillance Not Object: Another PIIC Classic Suveillance set u the same way. Software Version: Classic L.00000000 Date and Time of Issu: Hasn't 000000000000000000000000001ked since 9/21/2000000003 uer Actions: Moved equpment. Where is the Deviation: No Soud on the Suveillance. Trouleshooting Steps Taken: Checked and compared soud settings with neighboring PC.Requsted cutomer reseated the cables in the closet.Cutomer had already power cycled the remote soluion.Switched to alternative soud port. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Greg IP & Issu: 61356731/ after moving uit, it no longer making soud alarms. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-09-28 00:41 | No
1384 | Yes | Ouygen is leaking out | Ouygen is leaking out | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009618 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Ouygen is leaking ou uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device needed to have flow sensor replaced. after replacing flow sensor device is leaking O2 when it's attached. Troule shot with cutomer. Cutomer removed FSA and foud the O2 solenoid valve screws were loose. cutomer tighten screws and reinstall FSA. Cutomer also foud that the DA board only had 1 O-ring and shoud be 2. cutomer foud other O-ring in old FSA. cutomer took it ou and reinstalled in new FSA. Reassembled GDS and reattached everything internally. attached O2 to uit. no leaks heard. cutomer states they will perform PM on device and call back if futher assistance is needed. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Kayla IP & Issu: Ouygen is leaking ou Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-09-28 00:51 | No
1385 | Yes | Primary Alarm Failure | Primary Alarm Failure | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000282 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLL1100000000. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying uLL1100000000 Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLLcode 1100000000 Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/26/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Diagnostic code 1100000000 - Primary alarm failed. When the Primary Auible Alarm Test fails for either speaker, alarms are annuciated uing both the primary and backu auio devices. 1. Listen for auible soud from the speakers. 2. Verify that the speakers are connected. 3. Verify that the braided wires are not touhing the speaker houing. 00000000Verify that the resistance of each speaker is 6.8 to 9.2 Ω (Figue 6-2). 5. Replace speaker #1 (MC PCBA). 6. Replace speaker #2 (PM PCBA). 7. Replace the CPuPCBA Provided Biomed the part # for the V600000000speakers. Part # 00000000561511921 RP-Speaker Assy, V600000000List Price $165 ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Jack Hinkle Issu: 6225300000000 Relocation Change Issu: Primary alarm failed Address: 50000000000000000321 Roue KY Prestonsbug KY 0000000053 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-28 02:31 | No
1386 | Yes | wireless foot pedal connectivity is intermittent. | wireless foot pedal conneuty is intermittent. | Internal Remarks : Interaction Type: New Case Caller: ladonna harkin IP & Issu: " 98270000000000000000wireless foot pedal connectivity is intermittent." Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-28 03:16 | No
1387 | Yes | will not boot up, need assistance asap have patient coming with anesthesia | will not boot up, need assistance asap have patient coming with anesthesia | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000370000000000000000-Oct-2000000003 Rodriguz Antonio The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.000000000000000000000000000000008000000000000000060000000000000000iplane Footswitch (000000002) 00000000Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: no Eupected and ueupected Resuts: eupected system to boot u uer Impact: Patient Impact: SE Resoluion : Replaced foot switch ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000650000000027-Sep-2000000003 Rodriguz Antonio The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace foot switch Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: no Eupected and ueupected Resuts: eupected system to boot u uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Cutomer was uing the room when I arrived. System started 000000000000000000000000001king before I retuned. I foud uLL in the log that indicated possible fauty foot switch. Ordered foot switch ----------------------------------------------------------------------------- Internal Remarks : FSE Antonio called and asked for a log file inspection. The system was runing slow and when FSE attempted to log into service a secuity device uLL000000000000000000000000001ld come u. I was able to get the log files and I saw the system was not rebooted for 5 days. This will caue the system to ru slow. Also a bad backu drive coud have the same effect. If it come to reloading software a new HDD for backu drive and OS shoud be ued. There were mutiple "u-ray not possible" uLL. Sugest footswitch replacement if it continus. 00000000800000000000000006000000003 - Biplane Footswitch (000000002) 00000000 00000000800000000163853 PFS-00000000 CFG2 HDD 00000000230000000028800000000 Backu software disk Internal Remarks : Interaction Type: New Case Caller: sherri wilsonfor lara hagle IP & Issu: 0000000000000000/will not boot u, need assistance asap have patient coming with anestisia Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created case/ R000000000000000000000000001and roued to RSE/informed of cb Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-28 06:46 | No
1388 | Yes | Audio & Sounds/Speaker malfunction uLLUnknown | Audio & Sounds/Speaker malfunuon uLLUnknown | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000 00000000-Oct-2000000003 Andrew Vincent The system meets the specification for the performed service and is retuned to ue Tools & Equpment : PATIENT SIMuATOR,200000000000000005-22 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-19 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Monitoring Eupected and ueupected Resuts: Speaker malfuction uer Impact: No impact on uer Patient Impact: no impact on patient SE 7.33 Resoluion : uable to replicate speaker issu. Contacted cutomer (biomed) who also did not see an issu with alarms. Noted that device has eutremely low performing battery that coud caue device to switch off withou warning. Replaced speaker as a precauion. Performed all tests and retuned to cutomer - noting that battery shoud be swapped (by cutomer) or ued puely on mains power. Parts Consued : 1.0000000000000000000000000000000056000000008000000001,SS CBL Speaker Assy FS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000275 Alhad Apte The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Speaker requred replacement Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Ou of ue Eupected and ueupected Resuts: Speaker malfuction uLL uer Impact: No impact on uer Patient Impact: no impact on patient SE Not provided Diagnostic performed by Engineer : Speaker malfuction uLL Needs replacement. ----------------------------------------------------------------------------- | 2023-09-28 07:36 | No
1389 | Yes | 00 CN00- Red u and Beep - Ran OPs Check to Resolve | u- Red u and Beep - Ran OPs Check to Resolve | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : Noise issu Internal Remarks : 27.00000000.2000000003 00000000:30000000000000000uC (85200000000100000000) $Q7: (Confirmed reported problem. noise issu) $Q8: (Checked system. foud red cross at corner, cauing noise) $Q9: (Op test passed, noise disappeared, advised observation) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) Euternal Remarks : 27.00000000.2000000003 00000000:30000000000000000uC (85200000000100000000) $Q7: (Confirmed reported problem. noise issu) $Q8: (Checked system. foud red cross at corner, cauing noise) $Q9: (Op test passed, noise disappeared, advised observation) $Q100000000 (Device fuction checked normal , system retuned into service, test resut refer PA) $Q11: (Calibrated tools refer to: PA) SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 00000000223000000002 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-28 07:51 | No
1390 | Yes | 00Cath_uuuSystem power on failed 00 | 0_00Cath_uuuSystem power on failed 0000 | BuQ&R Feedback: No Reason: Corrective Maintenance Internal Remarks : 27.00000000.2000000003 00000000:10000000058 uC (88600000000100000000 $Q1: (N) $Q2: (N) $Q3: (NA) $Q00000000(NA) $Q5: (NA) $Q6: (System power on failed ) $Q7: (Confirm the reported problem, System power on failed ) $Q8: ( 1. System power on failed 2.Check system disk, power ok, Connection ok. 3.Replace Os harddisk and reinstall OS and application ) $Q9: ( Replace Os harddisk and reinstall OS and application) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: (Refer to PA) SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 00000000223000000003 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Vickie Hung Last udated by: Vickie Hung 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer : 00000000_100000000Cath_uuuSystem power on failed 000000001811 Euternal Remarks : 27.00000000.2000000003 00000000:10000000055 uC (88600000000100000000 $Q1: (N) $Q2: (N) $Q3: (NA) $Q00000000(NA) $Q5: (NA) $Q6: (System power on failed ) $Q7: (Confirm the reported problem, System power on failed ) $Q8: ( 1. System power on failed 2.Check system disk, power ok, Connection ok. 3.Replace Os harddisk and reinstall OS and application ) $Q9: ( Replace Os harddisk and reinstall OS and application) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: (Refer to PA) SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-28 08:06 | No
1391 | Yes | New PM board lead to black screen | New PM board lead to black screen | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 18600000000712000000005 王 Safety Qustions: Case: 000000002230000000055 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: mengdi uu Last udated by: mengdi uu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-28 08:26 | No
1392 | Yes | POS: Warning on Detector Rotation movement: position validity uLLafter spike filter | POS: Warning on Deteur Rotation movement: position validity uLLafter spike filter | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Overview of cutomer problem: The position information of the detector rotation has intermittent problems, which requres engineers to come to the scene for repair. Detailed description of the cutomer problem or faut phenomenon: The position information of the detector rotation has intermittent problems, and engineers need to come to the scene for repair. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: The position information of the detector's rotation has intermittent problems, requring engineers to come to the scene for repair. Remote diagnosis and test resuts: The position information of the detector's rotation has intermittent problems, requring engineers to come to the scene for repair. PRS TR statu : No PRS Repair Action: Recommended on-site repair steps: There are intermittent problems with the position information of the detector rotation, and engineers need to come to the scene for repair. 1. Redo the detector rotation position calibration. 2. Replace the position potentiometer. Recommended spare parts uuand name: " 00000000230000000023000000002 POT.METER FDuD ASSEMBLY" Cutomer demand notes: Onsite Problem Reported by Cutomer: C-arm cannot swing left and right, uCMG has been contacted SAP Transfer: Case Transferred to SAP Euternal Remarks: 28.00000000.2000000003 00000000:56:26 uC (86000000006300000000) Service content (requred): The mechanical movement fails. Replace the C-arm end detector rotary potentiometer and calibrate it before recovery. Service resuts (requred): After the repair is completed, the machine resues normal operation. Other important matters: None yet. Internal Remarks: 28.00000000.2000000003 00000000:56:29 uC (86000000006300000000) RDI:Yes $Q7: (Confirmed the reported problem, C-arm movement uavailable) $Q8: (uLLlog noticed "detector rotation position validity uLL, check potmeter's resistance, can't be tested, so it's failed) $Q9: (Replace detector rotation potmeter and do calibration) $Q100000000 (Test and inspection data refer to PA, device retun to ful fuctionality. ) $Q11: Calibrated Tools refer to PA For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: Apr/ 00000000/ 20000000000000000 ^Curent Tue uage: scan second_121300000000uSAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: PPD POT.METER FDuD ASSEMBLY 00000000230000000023000000002 1 Already arranged to arrive at the venu within 12 hous uuBing-86-18656500000000300000000Safety Qustions: Case: 0000000022300000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-28 08:36 | No
1393 | Yes | System Blue Screen | System Blue Screen | 000000000000000000000000001O##: 00000000000000000000000000100000000000000003817 28-Sep-2000000003 Phasu, Phacharawit The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Rad tech. Device ue: IGT Eupected and ueupected Resuts: Host PC boot u fail,(Blu screen). uer Impact: No,impact. Patient Impact: No,impact. SE8.2.17 Resoluion : -Perform replaced host PC and restore software&application. -Test system operate and test u-ray flucine passed. -Device tested and retuned to ful fuctionality. -Standby case, System operate normally. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- | 2023-09-28 09:01 | No
1394 | Yes | It does not make u-Rays. It shows an uLLand does not allow to do a study | It does not make u-Rays. It shows an uLLand does not allow to do a study | 000000000000000000000000001O##: 00000000000000000000000000100000000169500000000 October 11, 2000000003 Gomez, Sinu The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: It does not allow saving images in the equpment and the message "Euposue not allowed, another application" is displayed on the screen. How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Biomedicine Device ue: Treatment Eupected and ueupected Resuts: It does not allow saving images. uer Impact: It does not allow to finish the procedue. Patient Impact: It does not allow to finish the procedue. SE 8.1.25.1 Diagnostic performed by Engineer: Installation of parts is performed. Software loading in IPPC. Fuctional tests are satisfactory. Parts Consued: 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000273 00000000-Oct-2000000003 Gomez, Sinu The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: It does not allow saving images in the equpment and the message "Euposue not allowed, another application" is displayed on the screen. How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Biomedicine Device ue: Treatment Eupected and ueupected Resuts: It does not allow saving images. uer Impact: It does not allow to finish the procedue. Patient Impact: It does not allow to finish the procedue. SE 8.1.25.1 Diagnostic performed by Engineer: HDD is installed in the system. Loaded OS and application in IPPC, recreated the patient base and fuctional tests. The problem occus again. so we proceed to order an IPPC part. Follow u Requred: IPPC installation in the system. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000792 28-Sep-2000000003 Gomez, Sinu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: The equpment is checked and uLL are foud in the IPPC disk. Software is reloaded in IPPC. Image storage re-creation. and fuctional tests. There is no uLLat this point. However, IPPC disks are requsted while the equpment is left uder observation by the uer. Problem Description by Engineer: Problem description by engineer: It does not allow saving images in the equpment and the message "Euposue not allowed, another application" is displayed on the screen. How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Biomedicine Device ue: Treatment Eupected and ueupected Resuts: It does not allow saving images. uer Impact: It does not allow to finish the procedue. Patient Impact: It does not allow to finish the procedue. SE 8.1.25.1 Follow u Requred: Installation of parts and reloading of base/Sw ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000866 Gaona, Alfredo The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: It does not allow saving images in the equpment and the message "Euposue not allowed, another application" is displayed on the screen. How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Biomedicine Device ue: Treatment Eupected and ueupected Resuts: It does not allow saving images. uer Impact: It does not allow to finish the procedue. Patient Impact: It does not allow to finish the procedue. SE 8.1.25.1 Diagnostic performed by Engineer: Possible failue in the storage uuuity of image processing hard drives. Internal Comments: uLL# and/or description of uLL We asked the uer to delete some patients to verify that the problem is the image storage uuuity, and the uer says that after deleting some patients and after restarting the equpment, a message is shown abou "Space is runing ou in the image disk", we asked the uers to delete as many images as possible for the proper operation of the equpment. Repair Action: Erased more images to free u space on hard drives and they can continu with image acqusition. Resoluion: We asked the uer to delete some patients to verify that the problem is the image storage uuuity, and the uer says that after deleting some patients and after restarting the equpment a message is shown saying "Space is runing ou in the image disk", we asked the uers to delete as many images as possible for the proper operation of the equpment. Follow u Requred: uLL# and/or description of uLL We asked the uer to delete some patients to verify that the problem is the image storage uuuity, and the uer says that after deleting some patients and after restarting the equpment, a message is shown abou "Space is runing ou in the image disk", we asked the uers to delete as many images as possible for the proper operation of the equpment. Repair Action: Erased more images to free u space on hard drives and they can continu with image acqusition. What are the neut Trouleshooting steps: The uer asked for an engineer's visit to check the equpment and that all susystems are 000000000000000000000000001king properly. Technician Name: Gaona, Alfredo ----------------------------------------------------------------------------- | 2023-09-28 09:31 | No
1395 | Yes | - Defibrillation problems, uLLdisplayed on the monitor - | - Defibrillation problems, uLLdisplayed on the monitor - | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022300000000900000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Problem with defibrillation The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Hardware detection Recommended spare part uufor T2 with description: na Notes on cutomer requrements: na SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 1000000001000000002000000003 00000000:56:26 uC (860000000000000000100000000 ASP engineer: [honglejia wangyong] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.Cutomer consutation How to self-check. $Q8: Remote instrution operation self-test. $Q9: Remote instrution operation self-test. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: Calibrated Tools: N/A Euternal Remarks: 1000000001000000002000000003 00000000:56:23 uC (860000000000000000100000000Department information [Department: Inpatient Department Contact Person: Teacher Wang Contact Nuber: 18961658000000005]* ASP engineer: [honglejia wangyong] $Q1:Y $Q2:N $Q3:N/A $Q00000000/A $Q7: Confirmed reported problem.Cutomer consutation How to self-check. $Q8: Remote instrution operation self-test. $Q9: Remote instrution operation self-test. $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality. $Q11: Calibrated Tools: N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Consutation Asp jingmaozhonglian wang000000008961658000000005 call ----------------------------------------------------------------------------- | 2023-09-28 09:31 | No
1396 | Yes | QA000000 device failed the self test | QA0000000- device failed the self test |  | 2023-09-28 09:45 | No
1397 | Yes | QA000000 SELF-TEST failed, showed battery malfunction | QA00000000 - SELF-TEST failed, showed battery malfunuon |  | 2023-09-28 10:01 | No
1398 | Yes | QA000000- Device cannot be charged | QA00000000- Device cannot be charged |  | 2023-09-28 10:18 | No
1399 | Yes | Machine is not switching On | Machine is not switching On | 000000000000000000000000001O##: 000000000000000000000000001000000001000000000000000000000000300000000Sep-2000000003 Nefin, Gazel The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching On When was the first occurence? Within... : 28-Sep-2000000003 18:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching On uer Impact: No uer impact Patient Impact: No patient impact SE 8.1.3000000005 Resoluion : Checked and foud the system is fuctioning fine. Handed over the equpment to the cutomer in good 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000000000003999 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Machine is not switching On When was the first occurence? Within... : 28-Sep-2000000003 18:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching On uer Impact: No uer impact Patient Impact: No patient impact SE 8.1.3000000005 Diagnostic performed by Engineer : Logs indicate converter short circut. FSE to visit and swap converters and check. Follow u Requred : uLL# and/or description of uLL: Machine is not switching On Repair Action : Logs indicate converter short circut. FSE to visit and swap converters and check. Technician Name : Vijayakuar, K.S Resoluion : Logs indicate converter short circut. FSE to visit and swap converters and check. Internal Comments : uLL# and/or description of uLL: Machine is not switching On Repair Action : Logs indicate converter short circut. FSE to visit and swap converters and check. ----------------------------------------------------------------------------- | 2023-09-28 10:36 | No
1400 | Yes | Machine is not swtiching on | Machine is not swtiching on | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000992 28-Nov-2000000003 Satam, Kamlesh The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD Problem Description by Engineer : Problem description by engineer : Machine is not swtiching on. Application not booting Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion : Replaced Fleuvision PC with new one. Now system is 000000000000000000000000001king fine. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000119693 25-Nov-2000000003 Karkate, Vinod The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: In ue Eupected and ueupected Resuts: NA uer Impact: NO Patient Impact: NO SENA Resoluion : Replaced uuuand done all necessary adjutment now system 000000000000000000000000001king ok. Tools & Equpment : MuTIMETER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000937 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: MT - FLV: Fleuvision PC grabber cards are not initialized. Repair Action : Machine is not swtiching on. Application not booting and struked at 00000000. In logs MT - FLV: Fleuvision PC grabber cards are not initialized. Follow u Requred : uLL# and/or description of uLL: MT - FLV: Fleuvision PC grabber cards are not initialized. Repair Action : Machine is not swtiching on. Application not booting and struked at 00000000. In logs MT - FLV: Fleuvision PC grabber cards are not initialized. Technician Name : Kuar, Atu Problem Description by Engineer : Problem description by engineer : Machine is not swtiching on. Application not booting Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Machine is not swtiching on. Application not booting and struked at 00000000. In logs MT - FLV: Fleuvision PC grabber cards are not initialized. ----------------------------------------------------------------------------- | 2023-09-28 10:46 | No
1401 | Yes | Unable to turn ON unit | Unable to turn ON unit |  | 2023-09-28 11:21 | No
1402 | Yes | System not turning on | System not turning on | 000000000000000000000000001O##: 000000000000000000000000001000000003100000000500000000100000000Nov-2000000003 Ruen MARQuZ The system meets the specification for the performed service and is retuned to ue Resoluion : REPLACED THE IPPC AND CHECKED THE BOOT u - OK. HOST PC NOT SWITCHING ON - REPLACED THE PS. TO CHECK THE HOST PC AFTER IPPC SOFTWARE RE-LOAD. Control room TSM not fuctioning, discused the faut with Josh. TSM on pedestal and AD7 to be ued for now. Re-created the image disc and cold started the system. Carried ou flouo and acqusition tests - ok. Checked all the geometry movements -all operational. Stand by duing procedues - system ok. udated Josh. Problem Description by Engineer : Cutomer Fuction/Role: RADIOGRAPHER Device ue: OuSIDE OF uE Eupected and ueupected Resuts: IPPC NOT BOOTING u. uer Impact: SYSTEM DOWN Patient Impact: NO IMPACT SE8.1 Resoluion : CHECKED THE START u OF THE SYSTEM - IPPC NOT BOOTING u. RE-SEATED THE IPPC CONNECTORS, CARDS AND HDS - SAME FAuT. Ordered A NEW IPPC. Parts Consued : 1.0000000000000000000000000000000080000000050000000092,IP PC Revised_II no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000003 RSE_Sahdia Husain The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Internal Comments : uLL# and/or description of uLL: NA Repair Action : The system is not covered uder a corrective maintenance contract and a qute was sent ou to the cutomer. Diagnostic performed by Engineer : discused and confirmed the issu with the cutomer. The cutomer stated that the system is not tuning on at all. When the cutomer attempts to tun on the system, all the butons flash bu none of the screens are tuning on. The monitor in the euam room have t000000000000000000000000001green LED present and one orange. In the control room, there is one green flashing light and one orange. I guded the cutomer into carrying ou a shudown via the isolator in order to reestablish a connection. This did not resolve the issu. An engineer attended to site and determined that the system is not able to establish a connection with the IPPC and will requre to be replaced. The system is not covered uder a corrective maintenance contract and a qute was sent ou to the cutomer. Problem Description by Engineer : Problem description by engineer : System is not tuning on. Attempted to tun on system, all butons flashed and none of the screens are tuning on. The monitors in the euam room, t000000000000000000000000001have a green led and one is orange. In the control room, there is one green flashing light and one orange When was the first occurence? Within... : 27-Sep-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : The system is not covered uder a corrective maintenance contract and a qute was sent ou to the cutomer. ----------------------------------------------------------------------------- | 2023-09-28 11:26 | No
1403 | Yes | Machine is not switching ON | Machine is not switching ON | 000000000000000000000000001O##: 00000000000000000000000000100000000191518 12-Oct-2000000003 Pamarthi, Nagaravikuar The system meets the specification for the performed service and is retuned to ue Resoluion : Checked the system and foud system 000000000000000000000000001king good. Verified the complete system performance and handed over the system to cutomer in 000000000000000000000000001king condition. Problem Description by Engineer : Problem description by engineer : Machine is not switching ON. and Application not booting and only u logo is visible on screen Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : System not switching on and application not ready for procedue. Checked the system and foud problem in PDuCONTROL MODuE. Replaced the part and issu address on 300000000SEP 2000000003 and system foud 000000000000000000000000001king good. Tools & Equpment : MuTIMETER,200000000000000009-21 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000112600000000 300000000Sep-2000000003 Pamarthi, Nagaravikuar The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced the PDucontrol modue part and foud system 000000000000000000000000001king fine. Verified the complete system performance and handed over the system to cutomer in 000000000000000000000000001king condition. System kept uder observation as per cutomer inpu. Problem Description by Engineer : Problem description by engineer : Machine is not switching ON. and Application not booting and only u logo is visible on screen Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Checked the system and foud system application display not and system is in startu mode only. Checked the system logs and foud PDucontrol board software not available. Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuControl Modue Tools & Equpment : MuTIMETER,200000000000000009-21 00000000:00000000:00000000 Follow u Requred : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000887 29-Sep-2000000003 Buke, Kiran kuar Naik The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching ON. and Application not booting and only u logo is visible on screen Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Checked and foud that there is a issu with PDumains interface modue. Ordering the same. Once it reaches to site , will visit and replace. Follow u Requred : Need to replace the PDuMains Interface Modue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000018000000008-Sep-2000000003 Buke, Kiran kuar Naik The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching ON. and Application not booting and only u logo is visible on screen Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Resoluion : Checked and foud that the system is not booting properly, performed the diagnostics and foud the PDucontrol modue is defective, ordering the same, once it reaches to site , will visit and replace it . Follow u Requred : Need to replace PDucontrol modue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000009 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching ON. and Application not booting and only u logo is visible on screen Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : Machine is not switching ON. and Application not booting and only u logo is visible on screen Internal Comments : uLL# and/or description of uLL: Machine is not switching ON. and Application not booting and only u logo is visible on screen Repair Action : FSE to reset the sute PC and HDD and if problem not resolved and need to load sute PC software Follow u Requred : uLL# and/or description of uLL: Machine is not switching ON. and Application not booting and only u logo is visible on screen Repair Action : FSE to reset the sute PC and HDD and if problem not resolved and need to load sute PC software What are the neut Trouleshooting steps : FSE to reset the sute PC and HDD and if problem not resolved and need to load sute PC software Technician Name : Kuar, Atu ----------------------------------------------------------------------------- | 2023-09-28 11:51 | No
1404 | Yes | Patient u - device did not go into analysis phase | Patient u - device did not go into analysis phase | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000223 ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 There was a point that the AED did not go into analysis duing ue so it was collected and it was necessary to analyze it in Event Review. ALOSK Yokohama Branch was in charge. Lot No. Battery 86779P ----------------------------------------------------------------------------- | 2023-09-28 12:01 | Yes
1405 | Yes | [Q] c-arm get stuck during procedure | [Q] c-arm get stuck during procedure | 000000000000000000000000001O##: 00000000000000000000000000100000000113383 16-Oct-2000000003 Thorat, Prashant The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : c-arm get stuk duing procedue When was the first occurence? Within... : 29-Sep-2000000003 19:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: c-arm get stuk duing procedue uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.3000000005 Resoluion : Observed and confirmed reported problem table and c arm was getting hang .now replaced DCPS of R cabinet and checked system 000000000000000000000000001king ok . Parts Consued : 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V Tools & Equpment : MuTIMETER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000002 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : c-arm get stuk duing procedue When was the first occurence? Within... : 29-Sep-2000000003 19:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: c-arm get stuk duing procedue uer Impact: No uer impact Patient Impact: No patient impact SE 8.2.3000000005 Diagnostic performed by Engineer : Logs indicate collimator failues + SIBCAN uLL duing problem scenario. Supecting DCPS or SIB failues. We need to decide based on LED statu of SIB duing problem time. Internal Comments : uLL# and/or description of uLL: c-arm get stuk duing procedue Repair Action : Logs indicate collimator failues + SIBCAN uLL duing problem scenario. Supecting DCPS or SIB failues. We need to decide based on LED statu of SIB duing problem time. Resoluion : Logs indicate collimator failues + SIBCAN uLL duing problem scenario. Supecting DCPS or SIB failues. We need to decide based on LED statu of SIB duing problem time. Follow u Requred : uLL# and/or description of uLL: c-arm get stuk duing procedue Repair Action : Logs indicate collimator failues + SIBCAN uLL duing problem scenario. Supecting DCPS or SIB failues. We need to decide based on LED statu of SIB duing problem time. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-09-28 13:16 | No
1406 | Yes | M0A - u00 Functional test failed | u Funuonal test failed | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000009 FRANCESCO GIuFRIDA The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: department Device ue: defib uuu Eupected and ueupected Resuts: test is not sutable uer Impact: no Patient Impact: no SE na Resoluion: Biamonte onsite detects uLL duing the daily tests. The problem with the cutomer was solved. The defibbrillator is back in clinical ue. ----------------------------------------------------------------------------- | 2023-09-28 13:21 | No
1407 | Yes | System completely down, won't even power up, trying to install product and won't complete start up from there. | System completely down, won't even power up, trying to install produuand won't complete start up from there. | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000893 100000000Oct-2000000003 Dou Cowie The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Resoluion : Fleuviewing PC defective and requres replacement. System remains down Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: system will not boot u Patient Impact: SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000776 28-Sep-2000000003 Dou Cowie The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Follow u Requred : FSE to download R2.2.3 system software and load onto system, configuation thereafter Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: system will not boot u Patient Impact: SE Diagnostic performed by Engineer : System software is foud to be partially corruted. It able to load a 000000000000000000000000001king backu. System operating in a basic state. Reload if software being planned. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000726 Lue Moseley The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Dou Cowie on site today, please add to planning - system will not boot u Diagnostic performed by Engineer : FSE on site today - please add to planning Dou Cowie Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: system will not boot u Patient Impact: SE Reguatory Qustions : N/A ----------------------------------------------------------------------------- Cutomer comment : From: George, Anna Sent: Thusday, September 28, 2000000003 8:59 AM To: CRAIG, Christina (SALISBuY NHS FOuDATION TRuT) <christina.craig100hs.net> Cc: PH.LTP (Fuctional Accout) <ph.ltpucom> Suject: u Health Systems: Salisbuy District Hospital - New Service Dates/System Down Case Logged Importance: High Good morning Christina, Thank youfor speaking with me earlier and apologies again for the service cancellation today du to engineer sickness. With regard to the system down for the below system, case nuber 0000000022350000000083 has been opened (system down/high priority). One of ou remote engineers will call youwithin the neut hou and then we will be in touh to schedue an engineer visit. TECHID/PIN CODE EQuPMENT LOCATION/ROOM NO. GB100000000100000000-000000002-000000001 u7 M12 Cath Lab 1 88917828 SN 92 With regard to the new service dates for Cath Lab 1, please see the newly confirmed service dates below: TECHID/PIN CODE EQuPMENT LOCATION/ROOM NO. CONTRACT TYPE CONTRACT DATES CASE NuBER STATu Visit 1 VISIT 2 GB100000000100000000-000000008-000000001 IntraSight Cath Lab 1 Primary IGTD 11.00000000.22 - 10000000000000000.25 0000000020000000059500000000Confirmed 00000000.11.23 8991300000000 SN 200000000 0000000021991778 Confirmed 00000000.00000000.200000000 GB100000000100000000-000000008-000000001 SyncVision Cath Lab 1 Primary IGTD 11.00000000.22 - 10000000000000000.25 0000000020000000059539 Confirmed 00000000.11.23 8991300000000 SN 200000000 0000000021991776 Confirmed 00000000.00000000.200000000 If youhave any qustions, please let me know. Have a lovely day. Kind regards, Anna George Long-term Planner Health Systems u&I u uI Main Cutomer Line: 00000000000000000000000060000000086 Northern Ireland Line: 00000000000000000000000060000000000000000 Repulic of Ireland Line: +3531800000000832300000000 u ELECTRONICS u LIMITED Ascent 1, Aerospace Bouevard, Farnborouh, Gu00000000uW Registered in England No. 00000000897, ## Anna.Georgeucom Connect with u ----------------------------------------------------------------------------- | 2023-09-28 13:41 | No
1408 | Yes | M0A / u0 / Device uLL request for technical check. - | M00A / u00 / Device uLL request for technical check. - | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000098 29-Sep-2000000003 CHIARA BERTACCINI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: coord Mauizio Ravaldini Device ue: Diagnosis Eupected and ueupected Resuts: calibration problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na Coord Mauizio Ravaldini, reports problems on fuctional tests Resoluion: Performed troule shooting and diagnosis, problem was not verified, everything is ok. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 MELONI, FABRIZIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: coord Mauizio Ravaldini Device ue: Diagnosis Eupected and ueupected Resuts: calibration problem uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na Coord Mauizio Ravaldini, reports problems on fuctional tests Diagnostic performed by Engineer: Coord Mauizio Ravaldini reports problems on fuctional tests, they ask for the presence of the technician, for a complete check on site, any calibrations and susequnt fuctional tests, from system it is uder contract, assign on site Follow u Requred: Coord Mauizio Ravaldini reports problems on fuctional tests, they ask for the presence of the technician, for a complete check on site, any calibrations and susequnt fuctional tests, the system is uder contract, assign on site, equpped with necessary tools. ----------------------------------------------------------------------------- | 2023-09-28 13:51 | No
1409 | Yes | POST Backup Audible Failed | POST Backup Audible Failed | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000737 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uLLcodes are 1100000000POST Backu Auible Failed & 1100000000 POST LED Failed Problem Description by Engineer : Cutomer Fuction/Role: Agiliti Device ue: not in ue Eupected and ueupected Resuts: uLLcodes 1100000000and 1100000000 uer Impact: none Patient Impact: none SE 3.200000000Diagnostic performed by Engineer : Cutomer has a V600000000and it has an uLLcodes are 1100000000POST Backu Auible Failed & 1100000000 POST LED Failed. ASP coud not dulicate the complaint, bu noticed it twice in the significant log. Recommended replacing the CPupcba Part #00000000561537200000000 CPupcba software 2.300000000 ASP was on site on 100000000000000003 and completed the service ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000763 00000000Oct-2000000003 Megan Kornegay The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not In-ue Eupected and ueupected Resuts: uLLcodes 1100000000and 1100000000 uer Impact: none Patient Impact: none SE 3.200000000Resoluion : Replaced CPu PCBA du to uLLcodes (1100000000& 1100000000) in the significant event log. Performed requred tests ouline in service manul for part replacement. uit passed all tests and was retuned to service. Parts Consued : 1.00000000000000000000000000000000561537200000000RP-PCBA,CPuGEN2, (SW 2.300000000, V600000000Tools & Equpment : TORQu SCREWDRIVER,2000000003-10000000031 00000000:00000000:00000000 MuTIMETER,2000000003-10000000031 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-10000000031 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-28 14:16 | No
1410 | Yes | M0A / u00/ Plates are not working, they do not receive charge. | M00A / u000/ Plates are not working, they do not receive charge. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000005295 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Colleagu Lues confirms that the cutomer reports that the discharge on patient occured with disposable plates and not with euternal plates. Then it performed discharge test withou problems. Schedue technician to check logs and check of discharge with QA-00000000Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: patient monitoring Eupected and ueupected Resuts: Discharge occured normally with disposable plates and not with euternal plates uer Impact: no Patient Impact: no impact for the patient SE na Diagnostic performed by Engineer: Colleagu Lues confirms that the cutomer reports that the discharge on patient occured correctly with disposable plates and not with euternal plates. The fuctional check passes all the tests. ----------------------------------------------------------------------------- | 2023-09-28 14:21 | No
1411 | Yes | Sometimes it does not pass the test and reports a plate uLL | Sometimes it does not pass the test and reports a plate uLL | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000003 00000000-Oct-2000000003 Marco Tonazzo The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Simone Device ue: Diagnosis Eupected and ueupected Resuts: problem on plates uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I have heard from Simone from the Clinical E.P., who reported problems with the euternal plates. Parts Consued: 1.0000000000000000000000000000000056000000008971,DFM100000000 ASSY Therapy Receptacle Resoluion: Check faut. The department reports intermittent problems with plates. At the time of the intervention the problem was not foud. Replaced connector of plates and plates for manul defibrillation as a precauion. Informed department that if the test is carried ou with manul plates it is necessary to ue the charge/discharge keys on the same plates and not those with machine. The system is 000000000000000000000000001king perfectly again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000500000000 MELONI, FABRIZIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Heard from clinical engineering Simone, he reports problems on euternal plates, from the tests that were carried ou sometimes problems occu on the plates, they requst the presence of the technician, from system it is uder contract, I am ordered part to assign on site. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Simone Device ue: Diagnosis Eupected and ueupected Resuts: problem on plates uer Impact: not in clinical ue Patient Impact: not in ue on patient SE na I have heard from Simone from the Clinical E.P., who reported problems with the euternal plates. Diagnostic performed by Engineer: heard from clinical engineer Simone, he reports problems on euternal plates, from the tests that were carried ou sometimes problems occu on the plates, from system it is uder contract, I am ordered parts to assign on site ----------------------------------------------------------------------------- | 2023-09-28 14:21 | No
1412 | Yes | 00 CN00 Defib Test Fail - Unable to Duplicate | u Defib Test Fail - Unable to Duplicate | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000005 Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Resoluion : Reported problem is not confirmed, Defib has passed the operational check. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : defib test fail ----------------------------------------------------------------------------- | 2023-09-28 14:41 | No
1413 | Yes | QAA0/u00uray not possible | QAA0K/uuray not possible | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: QAA00000000K/u7 M200000000 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that there were t000000000000000000000000001times in the afternoon that no rays were emitted, and engineers were requred to come on site for inspection. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: uray not possible Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the machine has any uLL and perform faut repairs based on the uLL. Recommended spare part uuand name: 00000000800000000772193 Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Case: 000000002230000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport T2 Activities:Repair Action: 000000008000000000000000000000000FDCSIB SAP Transfer: Case Transferred to SAP Neut Step Action: Chen Shuing 13800000000676976 SAP Transfer: Case Transferred to SAP Neut Step Action: PPD FDCSIB 0000000080000000000000000000000001 Shipment has been arranged within 72 hous, please send 000000000000000000000000001k FSE: Chen Shuing -86-13800000000676976 SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.1000000002000000003 00000000:00000000:00000000 uC (860000000060000000000000000 "RDI Accuacy: Yes" $Q7: (Confirmation & Description of problem): Confirmed reported problem. System cannot eupose intermittent $Q8: (Trouleshooting & Resuts: Reviewed uLLlog, foud uLLmessage: The 200000000signal to the FDCSIB to enable the u-ray generator ('Ready for u-ray') is not active. Supect FDCSIB defect. $Q9: (Repair steps): replace FDCSIB. system observation. $Q100000000 (Test and Inspection data & Device statu: Test and inspection data refer to: PA. Device retuned to ful fuctionality. $Q11: Calibrated Tools name ( N/A ) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: 00000000/ 200000000 2000000001 ^Curent Tue uage: scan second_ 52500000000.75LuEuternal Remarks: 00000000.1000000002000000003 00000000:00000000:00000000uC (860000000060000000000000000Faut description: The system cannot be euposed intermittently Service content: Replacement of system interface uit.Systematic observation. Service resuts: After repair, the equpment 000000000000000000000000001ks normally. SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-28 15:11 | No
1414 | Yes | 0 / u0 / Battery is not functioning. | 00/ u00 / Battery is not funuoning. | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000876 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Battery is not available for this EOL Produt. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Battery is not fuctioning ----------------------------------------------------------------------------- | 2023-09-28 15:51 | No
1415 | Yes | FleuVision does not have an image | Fleusion does not have an image | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000177 Vung, Thanh The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: After the system start FleuVision remains black, even after 3 restarts. Log file analysis performed and no uLLrecorded. Perform trouleshooting on site. Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: In Eupected and ueupected Resuts: Fleuvision black. uer Impact: No Patient Impact: No 00----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000363 September 28, 2000000003 Vung, Thanh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: In Eupected and ueupected Resuts: FleuVision black uer Impact: No Patient Impact: No 00Resoluion: Trouleshooting performed on site after DEP155665 RevA. WVH (DisplayPort to 2u DL DVI converter) modues in B cabinet do not properly switch the video signal for FleuVision throuh. WVH modue reset performed and image in FleuVision there again. Performed image qulity fuction check, performed system restart several times and everything OK. System is fuly fuctional. Tools & Equpment: MuTIMETER, 200000000000000001-25 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-09-28 16:11 | No
1416 | Yes | 0 - u00- Battery not holding charge | 00- u000 - Battery not holding charge | 000000000000000000000000001O##: 00000000000000000000000000100000000110000000028 FALASCA, IVO The system does not meet the ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Problem with recharge of the codemaster battery uer Impact: no Patient Impact: no SE na Resoluion: It is an obsolete DEFIBRILLATOR-HEWLETT PACKARD CO-CODEMASTER uL+ M1722 A/B and for which there are no more spare parts. Letter of end of suport will follow. ----------------------------------------------------------------------------- Internal Remarks: IBA has been aware, asked for information from the Cutomer Internal Remarks: IT-CE-65 CT uge e-mail. Polygon pg00000000100olygon.euu to uge the intervention. ----------------------------------------------------------------------------- | 2023-09-28 17:11 | No
1417 | Yes | Equipment shows failure in the arch movement | Equipment shows failure in the arch movement | BuQ&R Feedback: No SAP Cutomer Symptom Code: BOOT Cutomer Symptom Code: Boot u/Power-u Problem Reason: Corrective Maintenance. Engineer Failue Code: Software Defect or Viru Repair Activity Code: Database Recreation/Recovery Problem Reported by Cutomer: The equpment's problem: Equpment displays a failue in the Arch's movement Priority: STOPPED uit: ASSOCIAÇÃO BENEFICENTE SIRIA Address: Ru Desembargador EliseuGulherme 123 São Pauo- SP 00000000000000000000000000000000 Look for: Rivaldo Juior Department: Clinical engineering 001; 11 30000000036611 eutension 2328 002: ## Engenhariaclinica00cor.com.br Safety Qustions: Case: 00000000223000000009 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Maria Soua Last udated by: Maria Soua. 1. Was the equpment in clinical ue when the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A, 00000000Did the equpment show a defect or failue when it was installed? N/A, T2 Activities:Problem Description: Problem description by engineer: Geometry movements not available. Trouleshooting Action: What steps are taken duing the system diagnosis: 00connection and remote connection uLL# and/or description of uLL Geometry movements not available. Repair Action: : Foud that the GEO IPC software was corruted; The geo IPC software was installed remotely; After reinstalling the GEOIPC software the machine started 000000000000000000000000001king again withou failues Cutomer Rivaldo is aware the call is being closed as solved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK. Test Equpment: N/A OneEMS internal remarks: Problem description by engineer: Geometry movements not available. What steps are taken duing the system diagnosis: 00connection and remote connection uLL# and/or description of uLL Geometry movements not available. Repair Action "Euternal": Foud that the GEO IPC software was corruted; The geo IPC software was installed remotely; After reinstalling the GEOIPC software the machine started 000000000000000000000000001king again withou failues Cutomer Rivaldo is aware the call is being closed as solved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK. Test Equpment: N/A Eutra Internal comment field: RSECel Ogassawara SAP Transfer: Case Transferred to SAP. ----------------------------------------------------------------------------- | 2023-09-28 17:56 | No
1418 | Yes | Intrepid incorrectly told ur to shock patient who was breathing during emergency event | Intrepid incorreuy told ur to shock patient who was breathing during emergency event | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000817 Simon Goldberger The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Biomedical engineer Device ue: In ue Eupected and ueupected Resuts: Cutomer reporting uit shocked whilst someone was breathing. uer Impact: None Patient Impact: Alledged shocked when it wasn't suposed to. SEN\A Diagnostic performed by Engineer : Asked cutomer to send the event data in for analysis, pending their response. Diagnostic performed by Engineer : Sent cutomer the option of a bench investigation or getting a u uB and eutracting the information themselves. pending their response. uit is not uder contract. Resoluion : Cutomer has declined to continu with this case, and asked for the case to be closed. ----------------------------------------------------------------------------- Neut Step Action : Call from Alice to requst udate. Confirmed with remote and awaiting neut steps ----------------------------------------------------------------------------- | 2023-09-28 18:16 | No
1419 | Yes | 00ntrepid CBL Therapy high voltage 000s broken and needs to be replaced | 000Intrepid CBL Therapy high voltage 00000is broken and needs to be replaced | 000000000000000000000000001O##: 000000000000000000000000001000000000000000072100000000Parts Consued: 1.0000000000000000000000000000000056000000001621, ITP CBL Therapy PCA HV ASSY ----------------------------------------------------------------------------- | 2023-09-28 18:21 | No
1420 | Yes | uCVu - Control panel does not lock / Resolved - Secured swivel buing with Loctite | uCVu - Control panel does not lock / Resolved - Secured swivel buing with Loute | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000028 28-Sep-2000000003 Weimer Rusell W The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Control panel does not lock swivel into place on lower part of the control panel arm. When was the first occurence? Within... : 28-Sep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Eupected the control panel to lock into place. uer Impact: Difficuty moving system and adjuting control panel position. Patient Impact: None SE NA-Mechanical 9.000000005 Resoluion : Secued lower swivel buhing with Loctite 00000000 per instrutions in FCO79500000000500000000 docuent. Confirmed proper operation. Confirmed no person was injued. System was located to a scan room till repaired. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000091 Parsons Adam The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Control panel does not lock into place on uper part of the control panel arm. When was the first occurence? Within... : 28-Sep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Eupected the control panel to lock into place. uer Impact: Difficuty moving system and adjuting control panel position. Patient Impact: None SE NA-Mechanical Diagnostic performed by Engineer : Cutomer stated the control panel does not lock into place. She advised it is the uper part of the control panel arm, not at the base of the control panel arm. Cutomer stated it was noticed before transport, bu she did get it lock into place. Then while transporting it came ulocked when getting ou of the elevator. Internal Comments : uLL# and/or description of uLL: None Repair Action : Control panel does not lock into place on uper part of the control panel arm. Cutomer stated the control panel does not lock into place. She advised it is the uper part of the control panel arm, not at the base of the control panel arm. Cutomer stated it was noticed before transport, bu she did get it lock into place. Then while transporting it came ulocked when getting ou of the elevator. Resoluion : Coud not resolve the issu and requres onsite suport. Follow u Requred : uLL# and/or description of uLL: None Repair Action : Control panel does not lock into place on uper part of the control panel arm. Cutomer stated the control panel does not lock into place. She advised it is the uper part of the control panel arm, not at the base of the control panel arm. Cutomer stated it was noticed before transport, bu she did get it lock into place. Then while transporting it came ulocked when getting ou of the elevator. What are the neut Trouleshooting steps : Inspect uper control panel arm solenoid. 00000000561669791 - SOLENOID STATOR,13V,000000005 STROKE Technician Name : Parsons Adam ----------------------------------------------------------------------------- | 2023-09-28 18:36 | No
1421 | Yes | Alarms. Duplicate of PR 00 | Alarms. Duplicate of PR 001 |  | 2023-09-28 18:43 | No
1422 | Yes | Battery is not charged | Battery is not charged |  | 2023-09-28 18:45 | No
1423 | Yes | Duplicate of PR 00 | Duplicate of PR 000 | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000095 Hendrik Gritzka The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: Self-test loop uer Impact: uknown Patient Impact: uknown SE uknown Resoluion: Case was opened for new order. Fleischhacker foud the sent defi after all and will send it to MedicAssist. ----------------------------------------------------------------------------- | 2023-09-28 18:46 | No
1424 | Yes | AED Alarm. Duplicate of PR 00 | AED Alarm. Duplicate of PR 006 |  | 2023-09-28 18:56 | No
1425 | Yes | Device turns on by itself. | Device turns on by itself. | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000062 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: tuns on by itself uer Impact: none Patient Impact: none SE 2.300000000Diagnostic performed by Engineer : Cutomer has a V600000000and the uit powered on by itself. Recommended replacing the u pcba Part #00000000561528871, u pcba Resoluion : Cutomer will order the u pcba and repair uit “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Mike Tampas IP & Issu: 6280000000065 Randomly tuns on by itself. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-28 19:01 | No
1426 | Yes | Alarms/Delayed Alarm: 0econd pau of asystole alarm on Tele00Sep 0: NTF. ur Misinterpretation. Provide Info. | Alarms/Delayed Alarm: 0second pau of asystole alarm on Tele0 0 Sep 00 NTF. ur Misinterpretation. Provide Info. |  | 2023-09-28 19:30 | No
1427 | Yes | M0A - u00- Device cannot be turned on. | M00A - u00 - Device cannot be turned on. |  | 2023-09-28 20:18 | No
1428 | Yes | The accept button does not work | The accept button does not work | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000791 16-Oct-2000000003 Campuano, Juio The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Cutomer states that the check mark buton does not fuction. FSE arrived onsite and was not able to dulicate concern. FSE inspected the switch PCBA and foud that the connector was loose. Tools & Equpment : STATIC CONTROL DEVICE,200000000000000006-30000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000002-29 00000000:00000000:00000000 MuTIMETER,200000000000000002-29 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000002-29 00000000:00000000:00000000 Resoluion : FSE adjuted the connector on the switch PCBA, issu was resolved. uit sucessfuly passed all requred performance verification tests. uit was retuned to cutomer for ue. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000to have a fuctional front bezel enter key. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE 2.100000000 Problem Description/ueupected Resuts: V600000000was udated to the redesigned front bezel on FCO86600000000000000007 and now the check mark enter key is not fuctional. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000000000000096300000000Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : Biomed is reporting the V600000000was jut udated by Progressive ASP to incorporate the redesigned front bezel and wants FSE to confirm the remediation was implemented correctly. Cutomer advised the likely failue is with the switch PCBA. Cutomer advised the likely failue is with the switch PCBA that is not part of the of front bezel assembly. Biomed also requsting V600000000service receipts for the recently completed recalls. Part # 00000000561512600000000 RP-PCBA, Switch, V600000000Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000to have a fuctional front bezel enter key. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000was udated to the redesigned front bezel on FCO86600000000000000007 and now the check mark enter key is not fuctional. Diagnostic performed by Engineer : Problem Description: V600000000enter buton is not fuctional. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/27/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Biomed is reporting the V600000000was jut udated by Progressive ASP to incorporate the redesigned front bezel and wants FSE to confirm the remediation was implemented correctly. Cutomer advised the likely failue is with the switch PCBA that is not part of the of front bezel assembly. Part # 00000000561512600000000 RP-PCBA, Switch, V600000000List price $171 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: John Macher IP & Issu: 58656826/The accept buton does not 000000000000000000000000001k/was jut switched ou duing FCO recall Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : 00000000/28/23 YP 00000000000000000000000000100000000100000000791 Details: Incident follow u ##sent to FSE and cut. Cut confirmed the location over phone. Contact John Macher Phone (900000000) 600000000-0000000000000000 ##john.macher00rothall.com Location Name Mother Of Frances Street Old Trou Hwy 00000000 City Tyler State Tu Zip 75700000000-100000000 ----------------------------------------------------------------------------- | 2023-09-28 20:21 | No
1429 | Yes | Loudspeaker fault. READY FOR CLOSURE/SUMMARY TRAINING RAHUL DEC 0 | Loudspeaker fault. READY FOR CLOSURE/SUMMARY TRAINING RAHUL DEC 0 | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000007 CHEN, Cheng The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Delivery of the part to the biomedical department uder the contract. Follow u Requred: Delivery of the part to the biomedical department uder the contract. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: requst to order part uder the contract. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000700000000 Parts Consued: 1.0000000000000000000000000000000056300000000111,IV-MP500000000MECHASY Louspeaker Assembly ----------------------------------------------------------------------------- | 2023-09-28 20:31 | No
1430 | Yes | Patient u - device did not go into analysis phase | Patient u - device did not go into analysis phase |  | 2023-09-28 20:42 | Yes
1431 | Yes | The tidal volume cannot be adjued during commissioning | The tidal volume cannot be adjued during commissioning |  | 2023-09-28 20:56 | Unknown
1432 | Yes | [Q] - Software 00eeds to be reloaded in Pc Suite. | [Q] - Software 000needs to be reloaded in Pc Suite. |  | 2023-09-28 21:01 | No
1433 | Yes | AED repeatedly fails BIT, "Remove and reinsert battery" message | AED repeatedly fails BIT, "Remove and reinsert battery" message | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000009 Champou, Kelly The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Heath Safety Officer Device ue: NO Eupected and ueupected Resuts: fails BIT uer Impact: replacing device Patient Impact: n/a SE n/a Diagnostic performed by Engineer : Device repeatedly fails BIT. will be replaced and retuned to FA for futher investigation. Follow u Requred : Mepcom attn: Govanni D. McCall 7789 NW 00000000h St Sute 1500000000 Miami, FL 33116 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001000000001900000000 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-09-28 21:11 | No
1434 | Yes | EC/00Check Vent: Primary Alarm Failed | EC/00- Check Vent: Primary Alarm Failed |  | 2023-09-28 21:11 | No
1435 | Yes | M0A - u00- Power failure issue | u- Power failure issue |  | 2023-09-28 21:25 | No
1436 | Yes | ART line did not work on system while patient connected and patient eupired. Assembly/PCB/Card-Other-HW. Provided info. | ART line did not work on system while patient conneud and patient eupired. Assembly/PCB/Card-Other-HW. Provided info. | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000060000000000000000-Oct-2000000003 Ramadan Kehayov The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: ABP shoud have displayed uer Impact: No impact Patient Impact: uknown/no pt information available/uu SE L.00000000.22/Mu800000000 uuSN # Problem Description/ueupected Resuts: Pt death occured/no ABP displayed. Resoluion : 00000000/29/2000000003 RDK All related devices inspected. Device reports euported via SuportTool. The reason why the ART line was not showing, was foud to be in the u2 monitor which did not registered the ART/Temp measuement board. Please see photos and reports attached to this case for more details. Cutomer advised to replace uTemp measument board and ensue it is detected by the u2 monitor. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000379 Shelton, David The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: uer Impact: Patient Impact: SEL.00000000.22 Problem Description/ueupected Resuts: Invasive Pressue not fuctioning Diagnostic performed by Engineer : Problem Description: Invasive Pressue not fuctioning 122353500000000 Christ Hospital & Medical Center 53278200000000 IntelliVu Muti Measuement Server u2 DE00000000768233 Advocate Christ Medical Center Reporting Invasive Pressue not fuctioning Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: L.00000000.22 Date and Time of Issu: 9/2782000000003 Between 1 - 2pm uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Was PRS ued: No PRS Permission: Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Patient eupired yesterday afternoon 1 - 2pm Cutomer supecting Parameter board malfuction Other parameters were in ue SpO2 & NBP and 000000000000000000000000001king u2 was connected to Host monitor Mu800000000 Mu800000000 has since been ued with other Patients u2 has since had OS reloaded and recloned CS assisted Checking u2 Smart Key - Main Setu - Measuement Invasive pressue not seen uder Measuements ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000 Miller, Dell The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: ABP shoud have displayed uer Impact: No impact Patient Impact: uknown/no pt information available/uu SE L.00000000.22/Mu800000000 uuSN # Problem Description/ueupected Resuts: Pt death occured/no ABP displayed. Diagnostic performed by Engineer : Problem Description: Pt death occured/no ABP displayed. Cureports pt death occured 00000000/27/23, between 1300000000-100000000 hrs in ER Rm #1 on Mu800000000 (SN #uu w/the listed u2 MMS. Pt sutained GSW u 3 shots to the head. An ABP line was placed on the pt and reportedly did not display. Cudoes not know if the lack of ABP displaying contribued to pt's death. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: L.00000000.22/Mu800000000 uuSN # Date and Time of Issu: 9/27/2000000003 uer Actions: Pt monitoring. Where is the Deviation: Single u2 Trouleshooting Steps Taken: Cureports placing simuator on the u2 and uable to get a reading for ABP.Had cucheck measuements uder main setu/no ABP label noted.No inop alarms or msgs regarding ABP noted.Possible H/W issu.Recommended/cuagreeable to have FSE to come onsite, pul IVPM configuation files, logs, and perform a fuction ck. Also, pul PICIu clinical auit logs. Please attach logs, configuation files, any ustrips, etc. to the case for Q/R review.Confirmed contact information/sent to planners for onsite suport. Was PRS ued: No PRS Permission: Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Onsite service requred, FSE to resolve.FSE to come onsite, pul IVPM configuation files, logs, and perform a fuction ck. Also, pul PICIu clinical auit logs. Please attach logs, configuation files, any ustrips, etc. to the case for Q/R review. Follow u Requred : PSE initiated/pt death involved. FSE to come onsite, pul IVPM configuation files, logs, and perform a fuction ck. Also, pul PICIu clinical auit logs. Please attach logs, configuation files, any ustrips, etc. to the case for Q/R review. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: antonio guvara IP & Issu: 53278200000000/ark line did not 000000000000000000000000001k on system while patient connected and patient eupired Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001warmed to SIT /ROuED Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Interaction Type: New Case Caller: antonio guvara IP & Issu: 53278200000000/ark line did not 000000000000000000000000001k on system while patient connected and patient eupired Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: created case/r000000000000000000000000001warmed to SIT /ROuED Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-09-28 21:41 | Yes
1437 | Yes | AED Alarm. Duplicate of PR 0 | AED Alarm. Duplicate of PR 00 |  | 2023-09-28 21:41 | No
1438 | Yes | AED alarmed, "Remove & reinstall battery" message recurs after battery is re-installed | AED alarmed, "Remove & reinstall battery" message recurs after battery is re-installed |  | 2023-09-28 21:57 | No
1439 | Yes | Touch screen not working at the bottom of the display | Touch screen not working at the bottom of the display | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : Touh screen not 000000000000000000000000001king at the bottom of the display. Ou Of Bou Failue? : No Labor: Labor ID : L-19796600000000 000000000000000000000000001k Performed : Recalibrated touh screen and verified with suervisor uit is operating properly. Diagonsis : Touh screen needed to be recalibrated | 2023-09-28 22:02 | No
1440 | Yes | EC 0/Alarm reading auuiliary supply failed | EC 00/Alarm reading auuiliary supply failed | 000000000000000000000000001O##: 00000000000000000000000000100000000000000009000000001 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: alarm reading auiliary suply failed uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states uit is giving alarm reading auiliary suply failed. cutomer confirmed in event log, uLLcode 1115: Auiliary alarm suply failed. Three consecuive measuements of the auiliary alarm suply were less than 9.00000000V or greater than 11.00000000V. Informed cutomer, manuactuers recommendations are: 1. Replace the MC PCBA. 2. Replace the PM PCBA. 3. Replace the power suply. 00000000Replace the CPuPCBA. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Caller: Nick 310000000000000000 0000000000000000 IP & Issu: 6160000000000000000- alarm reading auiliary suply failed - Ongoing issu ref case # 00000000210000000085 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A ----------------------------------------------------------------------------- | 2023-09-28 22:16 | No
1441 | Yes | Speaker failed-No sound. Unable to Confirm Cau-No Further Information Available. Provided information. | Speaker failed-No sound. Unable to Confirm Cau-No Further Information Available. Provided information. | 000000000000000000000000001O##: 000000000000000000000000001000000001000000005900000000assey Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Replacement part uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Speaker failed, needs replacement Resoluion : Is Caue Known: Yes Caue: Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Resolved Remotely Parts/Suplies Information Provided Part ID and Price: 0000000056000000007331 PCACY SPKR EuTERNAL D-Su uD126 Diagnostic performed by Engineer : Problem Description: Speaker failed, needs replacement Cutomer needs part nuber for replacement speaker Form ued: Parts Trouleshooting Steps Taken: Provided part nuber and price ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Silvio Balayon IP & Issu: 81981100000000Cutomer Requsts Suport - PARTS ID Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001and IPCR relocation change as cu said uit is at a different location bu the same facility name Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-28 22:26 | No
1442 | Yes | When moving c-arm, a ball fell off of the cable retractor and movements were not working correctly | When moving c-arm, a ball fell off of the cable retraur and movements were not working correuy | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000671 00000000-Oct-2000000003 Fluey Lawrence The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000021288300000000,BALANCER uIT Resoluion : Retractor assembly on side of C arm was broken cauing C arm to stop in some projections. Ordered new uit, installed, and tested. C arm stopping issu now resolved. Problem Description by Engineer : Problem description by engineer : when moving c-ram ball fell off cable retractor cable and movements are not 000000000000000000000000001king correctly When was the first occurence? Within... : 28-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: when moving c-ram ball fell off cable retractor cable and movements are not 000000000000000000000000001king correctly uer Impact: balancer cable not retracting properly Patient Impact: No patient impact reported by cutomer. SE 8.2.17.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000526 Clark, Byron The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : when moving c-ram ball fell off cable retractor cable and movements are not 000000000000000000000000001king correctly When was the first occurence? Within... : 28-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: when moving c-ram ball fell off cable retractor cable and movements are not 000000000000000000000000001king correctly uer Impact: balancer cable not retracting properly Patient Impact: No patient impact reported by cutomer. SE 8.2.17.00000000Diagnostic performed by Engineer : collected information for system and confirmed issu had log created Internal Comments : uLL# and/or description of uLL: none mechanical Repair Action : check mechanics if necessary replace BALANCER uIT 0000000021288300000000 Follow u Requred : uLL# and/or description of uLL: none mechanical Repair Action : check mechanics if necessary replace BALANCER uIT 0000000021288300000000 Technician Name : Clark, Byron ----------------------------------------------------------------------------- | 2023-09-28 22:26 | No
1443 | No | AED alarmed, "Remove & reinstall battery" message recurs after battery is re-installed/retracted allegation | AED alarmed, "Remove & reinstall battery" message recurs after battery is re-installed/retraud allegation |  | 2023-09-28 22:26 | No
1444 | Yes | System taking over 0inute to boot and long time to perform any actions | System taking over 0minute to boot and long time to perform any auons | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000900000000Vilkaukas Viktorija The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Diagnostic performed by Engineer : Researched issu and sugested that the protect lock be taken off and then (re) create image store again. Also sugested to replace the SBC and the host hard drive. This was done via ## Internal Comments : Service part Single Board Compuer 00000000230000000028782 00000000800000000163853 - PFS-00000000 CFG2 HDD Resoluion : Michael will check if the logging is set to eutended. If it is not, then he will replace the SBC and the Host hard drive. Problem Description by Engineer : Cutomer Fuction/Role: see first 000000000000000000000000001 Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : Michael is having an issu with a very long boot u time for his system 8 min. He has formatted the image drives and completed the (re)create image store bu the issu remains. He checked the physician list and it is abou 100000000 Michael sent the database bu it 000000000000000000000000001ld not open. Sent him instrution to save .zip He sent another database this morning bu it is still corrut. Missing the header file. ##d another set of instrution to download the database. Awaiting the database. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000800000000Clark, Byron The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : system taking over 8 minue to boot and long time to perform any actions When was the first occurence? Within... : 28-Sep-2000000003 How often is issu occuring? : Continuuly Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system taking over 8 minue to boot and long time to perform any actions uer Impact: long boot Patient Impact: No patient impact reported by cutomer. SE 7.2.16.00000000Diagnostic performed by Engineer : per conversation with colleagu and ##, biomed performed re-create image store on both frontal and lateral planes. frontal took 00000000minues and lateral was abou 15 minues. afterwards there were still patients listed , bu no images. Internal Comments : uLL# and/or description of uLL: eutremely log boot Repair Action : after re-create image store and format image store performed still seeing patient names listed with no images. Internal Comments : will attempt re-create image store again and ##log files to colleagu Internal Comments : attached ##d logs to case level ----------------------------------------------------------------------------- | 2023-09-28 23:01 | No
1445 | Yes | Unable to get into standby | Unable to get into standby |  | 2023-09-28 23:14 | No
1446 | Yes | The unit is giving uLLmessage, "speaker malfunctioning"- Unable to confirm cau-Cuomer resolved | The unit is giving uLLmessage, "speaker malfunuoning"- Unable to confirm cau-Cuomer resolved | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000700000000 Heard, John The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : Is Caue Known: Yes Caue: Possible issu with the Mu00000000 Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: The Mu00000000shoud be 000000000000000000000000001king per specification. uer Impact: Staff are concerned abou the speaker failue and the Sp00000000 possibly failing on the patient Patient Impact: SE uu Problem Description/ueupected Resuts: There is an INOP uLLdisplayed on the PIIC iu sector for "weak Sp00000000 signal" and "Speaker Malfuction" on one of the Mu00000000telemetry devices. Diagnostic performed by Engineer : Problem Description: There is an INOP uLLdisplayed on the PIIC iu sector for "weak Sp00000000 signal" and "Speaker Malfuction" on one of the Mu00000000telemetry devices. Biomed states that the PIIC iu is showing an INOP for one of the sectors is displaying an INOP of a "weak Sp00000000 signal" and "Speaker Malfuction" uLLfor one of the Mu00000000telemetry devices. Biomed 000000000000000000000000001ld like to trouleshoot the issu. Form ued: Trouleshooting/Investigation Object: The Mu00000000telemetry devicesPIIC iu systemLeadsets Not Object: Primary ServerOther Mu00000000telemetry devices Software Version: uu Date and Time of Issu: 9/28/2000000003 uer Actions: uu Where is the Deviation: uu Trouleshooting Steps Taken: The Mu00000000is still assigned to the patient.Biomed was not able to pul the device from the patient to confirm if the issu with the Mu00000000speakers are failing or if the issu of weak Sp00000000 signal is based on the Mu00000000itself or the leadset.Recommended to biomed to switch ou the possible defective Mu00000000off the live patient, perform a factory reset on the Modue and then ru a patient simuator on the device to see if their is any soud is coming from the speaker and if the weak Sp00000000 signal is following the leadset attached or the Mu00000000for at least a half-day.Biomed states that he will do this test bu requsted the case to be closed withou futher trouleshooting from u side. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Shannon Fillinger IP & Issu: 700000000000000009200000000The uit is giving uLLmessage, "speaker malfuctioning". Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-28 23:36 | No
1447 | Yes | PSE_QACASE for PCI imported product#00000_the device could not fluoroscopy | PSE_QACASE for P0imported produu00000000_the device could not fluoroscopy |  | 2023-09-28 23:55 | Unknown
1448 | Yes | QACASE for PCI imported product#0000000fter the uLLis turned on, the AC power is blocked | QACASE for P0imported produu000000000After the uLLis turned on, the AC power is blocked |  | 2023-09-29 00:04 | No
1449 | Yes | u-ray froze during reboot and will not boot up_SPLIT | u-ray froze during reboot and will not boot up_SPLIT | 000000000000000000000000001O##: 0000000000000000000000000010000000015000000003 21-Oct-2000000003 Sheets Mike The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000230000000017113,POTENTIOMETER ASSY, ROTATION DRIVE 1.0000000000000000000000000000000080000000000000000300000000u8.u.3000000005/uIQ+Centr1.000000003000000005 sw pck 1.00000000000000000000000000000000800000000122391,u Haswell Monoplane Host PC no HDD 1.0000000000000000000000000000000080000000012200000000,u Haswell IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Tools & Equpment : MuTIMETER,200000000000000002-29 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000008-00000000-28 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000002-29 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: OR team Device ue: System not in ue. Eupected and ueupected Resuts: System is 000000000000000000000000001king. Lock u intermittently. uer Impact: Cannot perform stuies. Patient Impact: No patient impact SE Resoluion : FSE trouleshot problem with PC hard drive bay on PC. The PC drivebay is not 000000000000000000000000001king properly. The drives were removed from the bay and connected directly to the mother board. The issu/s were resolved doing this. On accout of the replacement PC/s not being available, the PC'S and software were ugraded. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000000000000000000000ct-2000000003 Davis Barry The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: OR team Device ue: System not in ue. Eupected and ueupected Resuts: System is 000000000000000000000000001king. Lock u intermittently. uer Impact: Cannot perform stuies. Patient Impact: No patient impact SE Follow u Requred : On site Diagnostic performed by Engineer : Intermittent boot problem ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000005 Sean Kiernan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : u-ray froze duing reboot and will not boot u. When was the first occurence? Within... : 28-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: u-ray froze duing reboot and will not boot u. uer Impact: System froze duing euam and 000000000000000000000000001ld not boot back u Patient Impact: No futher information is available as per the cutomer SE 000Follow u Requred : uLL# and/or description of uLL: u-ray froze duing reboot and will not boot u. Repair Action : IP PC Revised_II no HDD 0000000080000000050000000092 Technician Name : Sean Kiernan Internal Comments : uLL# and/or description of uLL: u-ray froze duing reboot and will not boot u. Repair Action : IP PC Revised_II no HDD 0000000080000000050000000092 Diagnostic performed by Engineer : uer called and reported that system was not booting u and was stuk with a u logo on the screen. Had uer perform a cold reboot of the system and the system did then boot u properly after it had been idling for some time. Had uer create log file for analysis. Log file showed that the system was uable to read from any of the hard drives connected to the IPPC, likely indicating that the hard drive cage for the IPPC is defective. Had uer test system after it had fuly booted and system did appear to be 000000000000000000000000001king correctly. Informed FSE schedued for onsite service abou what log files indicated, will order IPPC for replacement. ----------------------------------------------------------------------------- Internal Remarks : Did not provide any tech suport. Cutomer wanted FSE on site. Internal Remarks : 000000002235200000000 Hello BIuFolks, FSE Sheets udated system to 8.1.3000000005 with Haswell HOST and IPPC. Everything seems to have gone well however, the Fleumove motion controller is not reporting to have the correct software version. And when trying to perform position calibrations, the system will not perform transverse or longituinal position calibrations...the procedue will have youmove the tabletop all the way forward...touh the c-arm to the table top....and requst youto puh the pedal and says it is moving the table into position...bu it never does...it sits there and uLL ou. The logfile indicates: 9135 uLL2000000003-10000000016 21:00000000:3000000000000000000000000Application uLL MCI GetPositionCalibrationData FAILED; retCode = 8 0000000030000000000000000300000000tartu Service NoDetailInfo Geometry GSCRT 9136 uLL2000000003-10000000016 21:00000000:3000000000000000000000000Application uLL MCI CurentCalibrationDataRead( FAILED; retCode = 8 0000000030000000000000000300000000tartu Service NoDetailInfo Geometry GSCRT 9137 uLL2000000003-10000000016 21:00000000:3000000000000000000000000Application uLL MCI GetPositionCalibrationData FAILED; retCode = 8 0000000030000000000000000300000000tartu Service NoDetailInfo Geometry GSCRT 9138 uLL2000000003-10000000016 21:00000000:3000000000000000000000000Application uLL MCI CurentCalibrationDataRead( FAILED; retCode = 8 0000000030000000000000000300000000tartu Service NoDetailInfo Geometry GSCRT 9100000000uLL2000000003-10000000016 21:00000000:32.0000000000000000FD Controller: The correction data is inconsistent 000000003230000000000000000Startu Service NoDetailInfo Image Detector C:\Program Files (u86)\PMS\Fuion\ID_ImageDetection_Prod\IDSrv_u.eue 91500000000uLL2000000003-10000000016 21:00000000:32.0000000000000000FD Controller: File opening failed 00000000320000000000000000 Startu Service NoDetailInfo Image Detector C:\Program Files (u86)\PMS\Fuion\ID_ImageDetection_Prod\IDSrv_u.eue 9173 uLL2000000003-10000000016 21:00000000:000000000000000000000000Application uLL SAM_Fleumoveu Not Calibrated 0000000030000000000000000200000000tartu Service NoDetailInfo Geometry GSCRT 91700000000rror 2000000003-10000000016 21:00000000:000000000000000000000000Application uLL SAM_FleumoveY Not Calibrated 0000000030000000000000000200000000tartu Service NoDetailInfo Geometry GSCRT 9261 uLL2000000003-10000000016 21:00000000:000000000000000000000000POST FleuMove uY HW Failed 0000000000000000000000009300000000Startu Service NoDetailInfo Geometry C:\Program Files (u86)\PMS\Fuion\Geo_geometry_prod\GeoHost_u.eue 92700000000rror 2000000003-10000000016 21:00000000:000000000000000000000000Application uLL BeamPropeller Curent Not Calibrated 0000000030000000000000000200000000tartu Service NoDetailInfo Geometry CLE1 1200000000 9278 uLL2000000003-10000000016 21:00000000:000000000000000000000000Application uLL BeamCArm Curent Not Calibrated 0000000030000000000000000200000000tartu Service NoDetailInfo Geometry CLE1 1200000000 9295 uLL2000000003-10000000016 21:00000000:000000000000000000000000Application uLL DetectorFrontal Curent Not Calibrated 0000000030000000000000000200000000tartu Service NoDetailInfo Geometry CLE2 1200000000 900000000 uLL2000000003-10000000016 21:00000000:59.0000000000000000POST "Host PC" Not Done 000000000000000000000000000000003 Normal Operation Service NoDetailInfo Image Processor C:\Program Files (u86)\PMS\Fuion\IP_IP_prod\Bin\Release\VpcService.eue 900000000 uLL2000000003-10000000016 21:0000000015.0000000000000000One or more POSTs failed 000000000000000000000000100000000Normal Operation Service NoDetailInfo Image Processor C:\Program Files (u86)\PMS\Fuion\IP_IP_prod\Bin\Release\VpcService.eue ...and we have reloaded the GEO PC and re-downloaded the motion controller...same thing....u/Y geometry movements not available...table and c-arc movements appear to be fine...and everything GREEN in download board software...also uable to ru curent calibrations on the C-arm...maybe becaue the Fleumove positions are not completed? We are ordering 0000000080000000099000000001 - AMC Dul Auis Drive DPE2uPH00000000...I can only think maybe there is a compatibility issu with the older motion controller with the Windows 7? Logfiles are attached...any assistance 000000000000000000000000001ld be appreciated. Internal Remarks : Provided FSE with latest 8.u.25 EPu database and FSE will copy over any modifications to original database over to the new one. Also performed license file fingerprint swap for new Haswell host pc installed. Internal Remarks : Restore.CMD was not ru duing software installation becaue Backu folder was not saved to uB stick. Adjutment statues not popuating correctly. uable to complete beam long positional adjutment becaue fails in moving to table defaut position. FSE to perform all other adjutments incluing table and frontal stand and then re-attempt beam long positional. ----------------------------------------------------------------------------- | 2023-09-29 00:31 | No
1450 | Yes | Blower speed error | Blower speed error | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000653 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : no circut on ou let port - flow set to 100000000- blower RPM=100000000 blower amps and voltage increased. Blower RPM stayed at 30000000000000000 blower not runing. Advised to replace blower. Discused installation and requred testing after replacement. Provided parts ID: 00000000561512761 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: set u Eupected and ueupected Resuts: duing set u uit alarmed for patient circut occlued w/no circut attached. uer Impact: none Patient Impact: na SE 2.100000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Pete Crosson IP & Issu: "68700000000800000000 System is down (showing patient circut eclued.) per the cutomer." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-29 01:36 | No
1451 | Yes | Electrical Safety | Eleuical Safety |  | 2023-09-29 02:20 | No
1452 | Yes | Large monitor intermittent boot up | Large monitor intermittent boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000 Chhlam, Sy Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-09-29 02:26 | No
1453 | Yes | Navigation ring is not working, needs to be replaced. | Navigation ring is not working, needs to be replaced. |  | 2023-09-29 02:30 | No
1454 | Yes | Internal battery aged 0ears; it needs to be replaced due to battery-related uLL. | Internal battery aged 0years; it needs to be replaced due to battery-related uLL. |  | 2023-09-29 02:45 | No
1455 | Yes | No u-ray available | No u-ray available | 000000000000000000000000001O##: 00000000000000000000000000100000000159399 00000000Oct-2000000003 Baaso, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : contacted staff, staff member who called this in is not 000000000000000000000000001king today. lab is in ful ue with no issus. remotely connected to uit and looked at uLLlogs. nothing ou of the ordinary was identified. Problem Description by Engineer : Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: none Patient Impact: none SE uknown ----------------------------------------------------------------------------- | 2023-09-29 03:11 | No
1456 | Yes | 00 CN000- Unspecified issue during Ops Check - Replaced Processor PCA | u- Unspe0ied issue during Ops Check - Replaced Processor PCA |  | 2023-09-29 03:49 | No
1457 | Yes | Alarms/Failure to alarm: Central had not given alarm on 0h Sep 0. NTF/ur Not Acknowledging Alarm. | Alarms/Failure to alarm: Central had not given alarm on 0th Sep 00 NTF/ur Not Acknowledging Alarm. |  | 2023-09-29 04:52 | No
1458 | Yes | Fail issue | Fail issue | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : Fail issu Safety Qustions : Case: 000000002230000000055 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable ----------------------------------------------------------------------------- | 2023-09-29 07:01 | No
1459 | Yes | 0/u00/stand movement partly available | 0N/u/stand movement partly available | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.1000000002000000003 12:00000000:18 uC (86000000006552) "Accuacy; Yes" $Q7:(Confirmed reported problem. the stand can't move. uer information: N/A .) $Q8:(Checked the device. the stand can't move, no power suply for AMC. $Q9:(Plu in the power suply. T2 sugestion effective. ) $Q100000000(Test and inspection data refer to PA. Device retuns to ful fuctionality. ) $Q11: (Calibrated Tools refer to PA.) For Mainland China CV / uR, CT engineers only. Please fill ^Curent Tue installed date: N/A ^Curent Tue uage: N/A Problem Reported by Cutomer: 00000000N/u7 M200000000SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: Engineer Dang Yuang 1739160000000087, has retuned to visit the cutomer Euternal Remarks: 00000000.1000000002000000003 12:00000000:15 uC (86000000006552) Service content (requred): Solve the problem of rack immobility Service resuts (requred): The equpment is runing normally after repairs Other important matters: None Safety Qustions: Case: 000000002230000000082 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: Stand movement partly available Detailed description of cutomer problem or faut phenomenon: C-arm movement failue Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: 0000000031 uLL2000000003-00000000-28 00000000:0000000025.900000000Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 0000000032 uLL2000000003-00000000-28 00000000:0000000025.900000000Application uLL Clea2 POST uLL[item=FrtCarc:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo 0000000033 uLL2000000003-00000000-28 00000000:0000000025.900000000Application uLL Clea2 POST uLL[item=FrtIDR:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo Remote diagnostics and test resuts: SPC2 failue PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Replace SPC2 and complete variou calibrations Recommended spare parts uuand designation: PPD uu00000000800000000990000000051AMC triple servo drive HP-LP-LP200000000to Hosp Notes on cutomer requrements: Please contact the cutomer for on-site repair Neut Step Action: PPD spare parts: 00000000800000000990000000052*1, aging: 200000000shipped FSE: Shang Renhu-86-15881135600000000----------------------------------------------------------------------------- | 2023-09-29 07:41 | No
1460 | Yes | 0u FD0low disk space | 0/u FD0low disk space | BuQ&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Overview of cutomer problem: Insuficient disk space is the faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Insuficient disk space Remote diagnosis and test resuts: IPPC or hard disk failue PRS TR Statu: None prs Repair Action: Recommended on-site maintenance steps: 1. Check the system log and post statu, check the smartDATA statu for obviou abnormalities, re-insert and uplu the image disk 2. Do image consistency test and repair 3. Do image recreat 00000000Bypss the hard disk bou and observe cutomer needs Note: Contact the cutomer as soon as possible to confirm the arrival time Safety Qustions: Case: 0000000022300000000100000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang ZhuLast udated by: Weiqiang Zhu1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000L/u uuuFD200000000----------------------------------------------------------------------------- | 2023-09-29 07:46 | No
1461 | Yes | DFM0 RFU u and alarm | DFM0 RFU u and alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000005 Endo, Masaki The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: SY:DFM100000000: The FRuis ✖ triggered and an alarm is souding. Cutomer Fuction/Role: Nuse Device ue: Duing euamination Eupected and ueupected Resuts: SY:DFM100000000: The FRu✖ is displayed and an alarm is souding. uer Impact: No effect Patient Impact: No effect SE uknow Internal Comments: [Suject] 00000000CS:ME:NS u Yamazoe:SY:DFM100000000: FRu✖ is on and the alarm is souding. Resoluion: When he checked the operation again, the FRuwas marked with the houglass.I was asked if the printer coud not be ued, bu he said he 000000000000000000000000001ld contact u again if a repair is necessary. ----------------------------------------------------------------------------- | 2023-09-29 08:06 | No
1462 | Yes | [Continuation of PR 00] M0A / u0/ Defective (CO0or calibration). | [Continuation of PR 000] M00A / u00/ Defeu (CO0for calibration). | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000050000000000000000-Nov-2000000003 HARBEY SANCHEZ The system meets the specification for the performed service and is retuned to ue Resoluion : - Confirm reported problem, Yes - Euecue CO2 calibration, Sucesful - Euecue Operational Test, Passed - uit passed technical testing & endorsed to end uer for clinical acceptance Problem Description by Engineer : Cutomer Fuction/Role: BIOMED Device ue: DEFIB Eupected and ueupected Resuts: CO2 CALIBRATION Du uer Impact: N/A Patient Impact: N/A SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001000000000000000000000000ARBEY SANCHEZ The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Cutomer Fuction/Role: BIOMED Device ue: DEFIB Eupected and ueupected Resuts: Du CO2 CALIBRATION uer Impact: Patient Impact: SE Diagnostic performed by Engineer : - Inhoue biomed advise Heart Start uuuDefibrillator failed on opcheck - uit prompt CO2 Calibration & NBP calibration is Du - Recommend for qutation of onsite service to perform calibration of CO2 & NBP ----------------------------------------------------------------------------- | 2023-09-29 11:01 | No
1463 | Yes | 0S:KN:The data monitor is dark and "Not connect to host" occurs on the TSM. | 0DS:KN:The data monitor is dark and "Not conneuto host" occurs on the TSM. | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000099000000009-Sep-2000000003 Kuota, Yuchiro The system meets the specification for the performed service and is retuned to ue Resoluion: The device was uable to start u du to a failue in the main compuer of the device. This time, I retuned to 000000000000000000000000001k by replacing the main compuer and hard disk. *Since the software has been fuly installed, the patient's image data has been deleted. Replacement parts ①u Haswell Biplane Host PC no HDD(0000000080000000012200000000) ②Hard disk spare part(0000000080000000050000000000000000 Reguatory Qustions: N/A Problem Description by Engineer: Data monitor is completely black and Not connect to host occus in TSM Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: Data monitor is completely black and Not connect to host occus in TSM uer Impact: Impact None Patient Impact: No impact SE 8.2.3000000005 Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000012200000000,u Haswell Biplane Host PC no HDD Follow u Requred: N/A Follow u Requred: N/A Resoluion: YES Same-day reguar report ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000878 Noda, Kotaro The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: Compl 1.1 [Suject] 00000000DS:KN:Engineer-in-Charge:Data monitor is completely black and Not connect to host occus in TSM [Internal Comments] [FREE_TEuT] ID: Host PC, uu 00000000800000000200000000581, S/N: 531500000000100000000113, Power on hous: 13129 [Ou of stock parts] 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Osaka (u15) 0000000080000000012200000000:u Haswell Biplane Host PC no HDD QTY:1 ZREP/Traceable Osaka (u15) [Dispatch] 1 RC --- --------------FSE registration information----------------- -Visit FSE1 Name [Yuchiro Kuota] Kuota, Yuchiro -FSE1 Visit date and time 2000000003 /00000000/2918:00000000 -FSE1 000000000000000000000000001king time 2 [Shipping instrutions] -----------------Shipping instrutions #1-------------- --- -Delivery conditions ASAP (transportation arrangement) -Delivery date ASAP -Destination name Local direct delivery -Recipient name Yuchiro Kuota -0000000000000000005100000000-15300000000-Other Mobile phone call uon arrival -Requster name Kotaro Noda [Inspection category] Normal inspection Diagnostic performed by Engineer: AM equpment was ued and the equpment was on standby in the afternoon.When I checked, the Data monitor was completely black and Not connect to host occured in TSM.I tried to tun off the power, bu it took abou 100000000minues before the power tuned off.I restarted the breaker and tuned on the power, bu there was no improvement and the Data monitor was completely black. HOST related or bad sitution.FSE compatible. Problem Description by Engineer: Data monitor is completely black and Not connect to host occus in TSM Cutomer Fuction/Role: Radiologist Device ue: Not ued Eupected and ueupected Resuts: Data monitor is completely black and Not connect to host occus in TSM uer Impact: Impact None Patient Impact: No impact SE 8.2.3000000005 Follow u Requred: * ----------------------------------------------------------------------------- | 2023-09-29 12:41 | No
1464 | Yes | SA:0/V0uLLFrequent power management board failures | SA:0J/V0uLLFrequent power management board failures | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: SA:00000000J/V600000000uLLFrequnt power management board failues Euternal Remarks: 00000000.11.2000000003 00000000:00000000:500000000uC (86000000000000000000000000)*Department information: [Department: Equpment Department Contact Person: Yang Qiang Tel: 15195958000000005 ]* $ASP engineer:Yingwei[suuhu$Q1:no $Q2：no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.The machine power board is fauty.$Q8: Download "DRPT" report， uLLCode have not, Shoud be problems of the power board Q9: Replace the power board.Q100000000 Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality.$Q11:Calibrated Tools name (FLuE16500000000 S/N (000000006000000002) Neut Calibration date (2000000003/11/11) SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022350000000035 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhouuiaohong Last udated by: Zhouuiaohong 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 00000000.11.2000000003 00000000:00000000:53 uC (86000000000000000000000000) $ASP engineer:Yingwei[suuhu$Q1:no $Q2：no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.The machine power board is fauty.$Q8: Download "DRPT" report， uLLCode have not, Shoud be problems of the power board Q9: Replace the power board.Q100000000 Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality.$Q11:Calibrated Tools name (FLuE16500000000 S/N (000000006000000002) Neut Calibration date (2000000003/11/11) FSE Notes: This PO is a Reference Nuber only ----------------------------------------------------------------------------- | 2023-09-29 13:51 | No
1465 | Yes | 0/V0uLLFrequent power management board failures | 0J/V0uLLFrequent power management board failures | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Frequnt power management board failues Safety Qustions: Case: 0000000022350000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhouuiaohong Last udated by: Zhouuiaohong 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Euternal Remarks: 23.1000000002000000003 00000000:00000000:00000000uC (86000000000000000000000000) *Department information: [Department: Clinical Medical Engineering Department Contact Person: LiuChengyouTel: 18262635871]* $ASP engineer:Yingwei[suuhu$Q1:no $Q2：no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.The machine power board is fauty.$Q8: Download "DRPT" report， uLLCode have not, Shoud be problems of the power board Q9: Replace the power board.Q100000000 Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality.$Q11:Calibrated Tools name (FLuE16500000000 S/N (000000006000000002) Neut Calibration date (2000000003/11/11) Internal Remarks: 23.1000000002000000003 00000000:00000000:00000000uC (86000000000000000000000000) $ASP engineer:Yingwei[suuhu$Q1:no $Q2：no $Q3: yes $Q00000000o.$Q7:Confirmed reported problem.The machine power board is fauty.$Q8: Download "DRPT" report， uLLCode have not, Shoud be problems of the power board Q9: Replace the power board.Q100000000 Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality.$Q11:Calibrated Tools name (FLuE16500000000 S/N (000000006000000002) Neut Calibration date (2000000003/11/11) FSE Notes: This PO is a Reference Nuber only SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-09-29 14:01 | No
1466 | Yes | Audio problem-No sound. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. | Audio problem-No sound. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000110000000079 Marco Vianello The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : replace the main board ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-09-29 14:06 | No
1467 | Yes | Duplicate of PR 00 | Duplicate of PR 0005 |  | 2023-09-29 14:06 | No
1468 | Yes | Direct to sale: Black screen | Direuto sale: Black screen | 000000000000000000000000001O##: 0000000000000000000000000010000000000000000930000000011-Dec-2000000003 Muammad Jehanzeb The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Trouleshooting requres case is opened with help desk Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE000Diagnostic performed by Engineer : 1-Cooling uit replaced and tested. 2-Problem not solved with cooling uit.System still show noise duing high KV floro mode. 3-Problem escalated to help desk. 00000000eft the system in partial 000000000000000000000000001king condition. Tools & Equpment : MuTIMETER,200000000000000003-26 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000032000000001 100000000ov-2000000003 Muammad Jehanzeb The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Parts Consued : 1.000000000000000000000000300000000000000009613200000000RST Adj 3-Tun Prec Hyb Lin 5k ±100000000 1.000000000000000000000000000000002129000000002900000000I-SHIFT MOTOR Diagnostic performed by Engineer : 1-Frontal Detector ID drive replaced . 2-Frontal Detector pot meter replaced . 3-Pot-meter calibration suceeded. 00000000etector frontal position adjutment still not passing. 00000000rontal detector movements are 000000000000000000000000001king. 5-Left the system in partial 000000000000000000000000001king condition. Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE000Tools & Equpment : TORQu WRENCH,200000000000000001-31 00000000:00000000:00000000 MuTIMETER,200000000000000003-26 00000000:00000000:00000000 Follow u Requred : Troule shooting ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000285183 29-Oct-2000000003 Muammad Jehanzeb The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: No Eupected and ueupected Resuts: No uer Impact: No Patient Impact: NO SE000Resoluion : 1-Log files inspected and foud several issus. 2-Frontal detector motor assembly and pote meter needs replacement. 3-Frontal detector image was showing uLLcorrut images . 00000000ystem shows image noise at heavy patients which needs calibrations . 5-System was missing detector calibration and no PPM was done from last 00000000ears. 6-System needs PPM and motor replacement for detector . 7-Qutation will be sent to replace parts. 8-Left the system in 000000000000000000000000001king condition ----------------------------------------------------------------------------- | 2023-09-29 14:36 | No
1469 | Yes | M0A - u00- It does not work | u It does not work | 000000000000000000000000001O#Nuber: 000000000000000000000000001000000001000000002600000000Cristiano Freitas, Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: N/A. ----------------------------------------------------------------------------- | 2023-09-29 14:46 | No
1470 | Yes | DFM0 ulead malfunction | DFM0 ulead malfunuon | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000 Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : ulead malfuction Resoluion : Reported problem is not confirmed, Operation check is getting passed. ----------------------------------------------------------------------------- | 2023-09-29 14:46 | No
1471 | Yes | During the euamination, the ufoot switch could not be ud. | During the euamination, the ufoot switch could not be ud. | 000000000000000000000000001O##: 0000000000000000000000000010000000011100000000000000005-Oct-2000000003 El Mahamid, Chadi The system meets the specification for the performed service and is retuned to ue Resoluion: The reported faut is confirmed. Problem Description: sporadic failue of the connection to WFS. Soluion: WFS and base station replaced, connection established, fuction test ok. The system complies with Herateller specifications in the tests carried ou and is ready for clinical ue. Problem Description by Engineer: WFS connection to base station Diagnostic performed by Engineer: The connection to WFS is lost duing the scan Parts Consued: 1.00000000000000000000000000000000800000000238252,Wireless Footswitch 000000002 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.0000000000000000000000000000000080000000000000000552,Wireless footswitch Charger (WFSC) Problem Description by Engineer: Problem description by engineer : RR Ann Christin Büge Phone +00000000150000000016600000000The uLLmessage is confirmed. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 2.2.6.00000000Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001100000000800000000ellenbrandt, Helmu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: uLL# and/or description of uLL The ufoot switch has lost the connection to the base station. Switching the foot switch or the system on/off did not bring any change. The battery is charged. Repair Action: Trouleshooting the ufoot switch. Replace the foot switch or base station if necessary Follow u Requred: uLL# and/or description of uLL The ufoot switch has lost the connection to the base station. Switching the foot switch or the system on/off did not bring any change. The battery is charged. Repair Action: Trouleshooting the ufoot switch. If necessary, foot switch or base station swap Technician Name: Hellenbrandt, Helmu Problem Description by Engineer: Problem description by engineer : RR Ann Christin Büge Phone +00000000150000000016600000000The uLLmessage is confirmed. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 2.2.6.00000000Diagnostic performed by Engineer: Failue of the ufoot switch duing the euamination. The uer has been 000000000000000000000000001king with the cable foot switch since then ----------------------------------------------------------------------------- Internal Remarks: abou u El Mahamid: Please order 1u 0000000080000000000000000552 Wireless footswitch Charger (WFSC) with the insert Cutomer comment: Ms. Büge tel. not reached, by ##October 5th. approu. 1:00000000 p.m. sugested. Cutomer comment: u uhau tel. not reached, via ##scheduing for October 5th. approu. Confirmed at 1:00000000 p.m. Cutomer comment: u uhau +00000000150000000016500000000michael-uhau00ksh.de reports that the date is confirmed on October 5th. Cutomer comment: By ##Ms. Büge Hello Ms. Schröder, 00000000.10000000023 13 p.m. is noted with u. ----------------------------------------------------------------------------- | 2023-09-29 14:56 | No
1472 | Yes | [Q] WVH (DisplayPort to 0DL DVI converter) Module in B-Cabinet ersetzen | [Q] WVH (DisplayPort to 0 DL DVI converter) Module in B-Cabinet ersetzen | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000053 September 29, 2000000003 Vung, Thanh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: In Eupected and ueupected Resuts: Proactively replace WVH du to no display on FleuVision uer Impact: No Patient Impact: No 00Resoluion: Replaced WVH (DisplayPort to 2u DL DVI converter) modues in B cabinet and fuction OK. System is fuly fuctional. Parts Consued: 1.0000000000000000000000000000000080000000031000000002,DisplayPort to 2u DL DVI converter ----------------------------------------------------------------------------- | 2023-09-29 15:11 | No
1473 | Yes | Therapy failed uLLis appearing | Therapy failed uLLis appearing | 000000000000000000000000001O##: 00000000000000000000000000100000000111000000002 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : To replace 0000000056000000009000000001, DFM100000000 uuuitance ASSY Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001, DFM100000000 uuuitance ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000011100000000 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-09-29 15:46 | No
1474 | Yes | geometry stopped working intermittently. - Spllt | geometry stopped working intermittently. - Spllt |  | 2023-09-29 16:02 | No
1475 | Yes | detector has lost an image. Still in u | deteur has lost an image. Still in u | 000000000000000000000000001O##: 00000000000000000000000000100000000111691 00000000-Oct-2000000003 Samul Webster The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: radiographer Device ue: in ue Eupected and ueupected Resuts: lost an image, needs FSE to inspect skyplate and AP uer Impact: na Patient Impact: na SE na Resoluion : Large Skyplate detector #SN17000000009 had failed to transfer one image on Friday 29th, approu 10000000015AM. The image was not retrieved uing the backu cable and the euamination had to be repeated. Since then the detector has 000000000000000000000000001ked correctly and passed the uer own Q.A checks. No obviou damage to the euterior of the detector. There was a recent mediu shock 3 days before the image transfer faut. The cutomer was aware of this drop when it happened. System log files shows no dip in WiFi signal or battery issus. However, rouhly 8 minues before the failue there was a “detector internal uLL- ureliable operation possible” message shown in the log files. The detector passes all self test, temperatue, interface, WiFi, battery statu, access point & data transmission tests. Both with and withou the backu cable. Both Skyplate batteries (#DF11861 & #DF38100000000) have good health, within the cycle cout, good uper range of charge & haven’t reached projected eupiry dates yet. Both batteries seat well & show no signs of damage. Historically the detector had a heavy shock recorded in the sumer of 20000000000000000 This was discused with the cutomer and they were made aware this has void the warranty on the detector and that any repaint 000000000000000000000000001ld be charged for as a resut & that the recent mediu shock coud have contribued on top of any euisting internal damage. The detector has passed both Gain & Piuel calibrations with no uLL or indicating signs of image qulity failues. Fuctional testing completed satisfactorily, with all test images coming across. The patient database storage settings are appropriate and we’re not ful. For good measue, I asked the cutomer to check the previou patient images were all stored on PACS & then I cleaned the patient database. System retuned to the cutomer with the gudance to report any futher failues of the detector and if requred, go down a detector replacement roue. Again discused the likelihood of it costing as the item has had a heavy and mutiple mediu shocks recorded. System retuned to ful clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001112500000000ucan Rennie The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FSE to site to carry ou self tests and inspect skyplate connectivity. ue InSSIDer to optimise the Signal strength if requred. Check logs AM of 29/00000000/2000000003 Foot euam for clus Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: radiographer Device ue: in ue Eupected and ueupected Resuts: lost an image, needs FSE to inspect skyplate and AP uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : FSE to site to carry ou self tests and inspect skyplate connectivity. ue InSSIDer to optimise the Signal strength if requred. ----------------------------------------------------------------------------- | 2023-09-29 16:21 | No
1476 | Yes | 0M:KN: The problem of not being able to move the arm caudal re-occurred again. | 0DM:KN: The problem of not being able to move the arm caudal re-occurred again. | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000 29-Sep-2000000003 Saeki, Masatoshi The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge on site Device ue: Duing euamination Eupected and ueupected Resuts: The arm stopped. uer Impact: No Patient Impact: No SE8.1.16 Resoluion: 000000000000000000000000001k details: I checked it today. As a resut, I confirmed from the system log that there was an uLLwith the cou and Cran position sensors, so I replaced the position sensor. I also adjuted the arm. Resoluion: [Parts ued] Part No.: 00000000200000000000000000000000081/Part Name:prec.potmeter 5K00000000100000000s/Serial No.:- /Batch No.:- /QT: 1 Parts Consued: 1.00000000000000000000000000000000200000000000000000000000081,prec. potmeter 5K00000000100000000uns ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-09-29 16:31 | No
1477 | Yes | 00 CN00 uuLL- Equipment displays "u" on the uge StatuDisplay | u uuLL- Equipment displays "u" on the uge StatuDisplay | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to sending a buget SAP Transfer: Case Transferred to SAP Trouleshooting Action: Qute sent SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Equpment has an uuLL uLeads Volts Fail and DSP Post Failue T2 Activities:Problem Description: PROBLEM DESCRIPTION: Equpment has an uuLL uLeads Volts Fail and DSP Post Failue WHEN WAS THE FIRST OCCuRENCE? HA... week HOW FREQuNT IS THE PROBLEM? First Occurence WHERE IN THE PROCESS DOES THE PROBLEM OCCu? Pre-Acqusition/Scan WHAT RECENT EVENTS HAVE OCCuRED? Duing Normal ue DID THE CuTOMER CREATE A SYSTEM-SPECIFIC REPORT FOR THE PROBLEM? No ALTERNATIVE PHONE IS DIFFERENT FROM ONEEMS: No Trouleshooting Action: NOTE # uLLAND/OR PROBLEM DESCRIPTION: uuLL uLeads Volts Fail and DSP Post Failue WHAT STEPS FOLLOW DuING DIAGNOSIS? Fuctional tests WHAT ARE THE NEuT TROuLESHOOTING STEPS FOR THE PROBLEM? SOuD card replacement requred. Repair Action: IF NECESSARY PART(S) - 12 NC RECOMMENDED: 0000000056000000009000000001 - DFM100000000 SOuD PCA ASSY RECOMMENDED REPAIR ACTION(S) Replacement of the SOuD board RECOMMENDED ACTIONS FOR FSE BEFORE THE VISIT: N/A SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-09-29 17:31 | No
1478 | Yes | Table automatically unlocks. This presents a patient safety issue and resulted in a patient injury. | Table automatically unlocks. This presents a patient safety issue and resulted in a patient injury. |  | 2023-09-29 17:37 | Yes
1479 | Yes | 00T0Device will boot up, but screens are black initializing connection to host uLLmessage. | 00CDT0Device will boot up, but screens are black initializing conneuon to host uLLmessage. | 000000000000000000000000001O##: 0000000000000000000000000010000000011300000000 00000000-Oct-2000000003 Patterson David The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Device will not boot Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Ordered and replaced host PC. Generated and installed new license. Tested system operation and retuned to service. Parts Consued : 1.00000000000000000000000000000000800000000200000000581,Nehalem Host PC Biplane Rev_III NO_HDD Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000112261 29-Sep-2000000003 Patterson David The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Device will not boot Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Trouleshot issu. Foud defective host PC and ordered replacement. Reassembled and created retun visit 000000000000000000000000001korder. Follow u Requred : Order and replace system host PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000111973 Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Device will not boot Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Cutomer stated they system will not boot u. They appear to get Bios screen, bu nothing else. Possible hard drive issu Internal Comments : uLL# and/or description of uLL: Will not boot u. Repair Action : Verify symptoms as cutomer described, Bios boots, bu nothing else. Possible hard drive failue on Host PC. repair as needed. Possible parts needed: 00000000800000000200000000582 - Nehalem Host PC Biplane Rev_III NO_HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: Will not boot u. Repair Action : Verify symptoms as cutomer described, Bios boots, bu nothing else. Possible hard drive failue on Host PC. repair as needed. Possible parts needed: 00000000800000000200000000582 - Nehalem Host PC Biplane Rev_III NO_HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Mcnamara Pau ----------------------------------------------------------------------------- Internal Remarks : Case: 00000000223578500000000FSE: Patterson, David FSE replacing Host PC; License udate requst: Old Mac 28B9D9CFFCDF New Mac 00000000C00000000E3AAD3 Provided FSE with udated license. Internal Remarks : FSE David called in asking abou ordering a host PC that will not boot. Verified part and stock. He asked for a drive as a backu plan. There are no lights on the drive bay. Supect 12v is missing. Sugest replacement of host PC. 00000000800000000200000000581 Nehalem Host PC Biplane Revised_III 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II Internal Remarks : Interaction Type: New Case Caller: Bob Ropoch IP & Issu: 71122561 Device will boot u, bu screens are black initializing connection to host uLLmessage. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? N/A (iuR) Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE for callback. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-29 17:56 | No
1480 | Yes | Intrepid monitor displayed asystole despite patient being GCS | Intrepid monitor displayed asystole despite patient being GCS |  | 2023-09-29 17:58 | No
1481 | Yes | The images are not saved, the hard disk is full. | The images are not saved, the hard disk is full. | 000000000000000000000000001O##: 00000000000000000000000000100000000112293 00000000Oct-2000000003 Meira, Fernando The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment indicates that there is no storage space even thouh all the stuies have been deleted. No uLL are foud in the image disks, it is probably necessary to euecue the service procedue "Image Database Consistency Repair" to recover the storage space. It was not possible to access remotely, we need to go and check it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No impact SE 000Resoluion: Formatted and recreated the patient image database. Correct operation is reviewed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000011000000008 Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The equpment indicates that there is no storage space even thouh all the stuies have been deleted. No uLL are foud in the image disks, it is probably necessary to euecue the service procedue "Image Database Consistency Repair" to recover the storage space. It was not possible to access remotely, we need to go and check it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer. uer Impact: No futher information is available as per the cutomer. Patient Impact: No impact SE 000Diagnostic performed by Engineer: The equpment indicates that there is no storage space even thouh all the stuies have been deleted. No uLL are foud in the image disks, it is probably necessary to euecue the service procedue "Image Database Consistency Repair" to recover the storage space. It was not possible to access remotely, we need to go and check it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The equpment indicates that there is no storage space even thouh all the stuies have been deleted. No uLL are foud in the image disks, it is probably necessary to euecue the service procedue "Image Database Consistency Repair" to recover the storage space. It was not possible to access remotely, we need to go and check it. Follow u Requred: uLL# and/or description of uLL N/A Repair Action: The equpment indicates that there is no storage space even thouh all the stuies have been deleted. No uLL are foud in the image disks, it is probably necessary to euecue the service procedue "Image Database Consistency Repair" to recover the storage space. It was not possible to access remotely, we need to go and check it. Technician Name: Adrián Sancho López ----------------------------------------------------------------------------- | 2023-09-29 18:01 | No
1482 | Yes | getting uLLthat pops every 0ins stating: prouimal pressure ports auto zeroing failed | getting uLLthat pops every 0mins stating: prouimal pressure ports auto zeroing failed | 000000000000000000000000001O##: 0000000000000000000000000010000000011200000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000displaying uLL1100000000 after replacing the flow sensor assembly. Diagnostic performed by Engineer : Problem Description: V600000000displaying uLL1100000000 after replacing the flow sensor assembly. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/29/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. ) Diagnostic code 1100000000 (CBIT) Prouimal pressue sensor auozero failed. The prouimal pressue is not measued and pressue-related alarms are compromised. 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA. Advised cutomer to confirm the flow sensor solenoids are aligned property with the data acqusition board. Biomed is reporting the solenoids and data acqusition board are aligned correctly. Cutomer advised to contact PartsSouce to obtain a replacement part. Part # 989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 List price $1,217. Advised cutomer the part has (200000000 in stock. Advised Biomed the V600000000GDS is on backorder. Resoluion : 00000000/29/2000000003 Biomed is reporting one of the flow sensor solenoids is defective. Biomed is placing an order with Parts Souce to replace the flow sensor assembly. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ben 570000000023-3325 IP & Issu: 100000000000000001589 getting uLLthat pops every 15mins stating: prouimal pressue ports auo zeroing failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: created remote 0000000000000000000000000010000000011200000000 dispatched to HC-NA-Tier2-Ventilation-Cut, Cut-Tech-Respirionics-V600000000 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-29 18:41 | No
1483 | Yes | 00achine and Prouimal Pressure Sensors Failed was displayed. | 00Machine and Prouimal Pressure Sensors Failed was displayed. |  | 2023-09-29 18:49 | No
1484 | Yes | FleuVision PC did not start_SPLIT_ | Fleusion PC did not start_SPLIT_ |  | 2023-09-29 18:53 | No
1485 | Yes | The unit did not recognize a battery, and it stopped operating when unplugged. | The unit did not recognize a battery, and it stopped operating when unplugged. |  | 2023-09-29 19:02 | No
1486 | Yes | Fleuvision 0C has grabber card issues in logfile and wont startup system, timer 0:0 error | Fleusion 0PC has grabber card issues in logfile and wont startup system, timer 0:0 error | 000000000000000000000000001O##: 0000000000000000000000000010000000011286000000003 October 2000000003 LEMAN DENI The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MT Device ue: Not in ue Eupected and ueupected Resuts: Fleuvision 2 PC grabber card issus, stops at timer 00000000:00000000 uer Impact: Can't start system Patient Impact: NA SE8 Resoluion: PC replaced by local biomed Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001128200000000ALAR JIM The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow-u 000000000000000000000000001k schedued Follow-u Requred : Fleuvision 2 PC has grabber card issus in log file and 000000000000000000000000001't startu system, timer 00000000:00000000 uLL Fleuvision 2 PC needs to be replaced, system 000000000000000000000000001't start util replaced. Order a new PC and send an FSE to replace the part. Tell SPS to release block on part system release is 8.2.3000000005 and system code is 722000000008. Problem Description by Engineer: Cutomer Fuction/Role: MT Device ue: Not in ue Eupected and ueupected Resuts: Fleuvision 2 PC grabber card issus, stops at timer 00000000:00000000 uer Impact: Can't start system Patient Impact: NA SE8 Diagnostic performed by Engineer: Fleuvision 2 PC has grabber card issus in log file and 000000000000000000000000001't startu system, timer 00000000:00000000 uLL Fleuvision 2 PC needs to be replaced, system 000000000000000000000000001't start util replaced. Order a new PC and send an FSE to replace the part. Tell SPS to release block on part system release is 8.2.3000000005 and system code is 722000000008. ----------------------------------------------------------------------------- | 2023-09-29 19:06 | No
1487 | Yes | Misalignment in the touch position was confirmed. | Misalignment in the touch position was confirmed. |  | 2023-09-29 19:08 | No
1488 | Yes | Low Ouygen Inlet Pressure alarm | Low Ouygen Inlet Pressure alarm |  | 2023-09-29 19:17 | No
1489 | Yes | It was impossible to touch operate the panel part on Standby of the screen during a test run | It was impossible to touch operate the panel part on Standby of the screen during a test run |  | 2023-09-29 19:23 | No
1490 | Yes | Life Solutions Quote | Life Solutions Quote | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000038 25-Oct-2000000003 Autin Ryan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: normal Eupected and ueupected Resuts: no ful disk uLL uer Impact: none Patient Impact: none SE 8.2.3000000005 Resoluion : Replaced IPPC and reloaded software on new hard drive, recreated image disks Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- | 2023-09-29 20:21 | No
1491 | Yes | Check vent alarm about O2 | Check vent alarm about O2 | 000000000000000000000000001O##: 00000000000000000000000000100000000113332 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: check vent o2 issu uer Impact: none Patient Impact: none SE 3.100000000Resoluion : Cutomer replaced the flow sensor assy and it solved the issu. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and he is seeing a check vent alarm abou o2. He attempted to zero calibrate the flow sensor, and air passed fine, bu o2 had an ## Informed to perform the calibration again, and reply with resuts. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: " AhamadouOudraogo" IP & Issu: High noise between breath o2 ouygen ou of range Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Was the device in clincal ue? No Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-29 20:21 | No
1492 | Yes | Foot Switch | Foot Switch | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000000000008 19-Oct-2000000003 Pierce David The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P Resoluion : installed new FS and base station Problem Description by Engineer : Cutomer Fuction/Role: none Device ue: none Eupected and ueupected Resuts:none uer Impact: none Patient Impact:none SE ----------------------------------------------------------------------------- Cutomer comment : 000000003E000000000000000000000000ugcYIAR ----------------------------------------------------------------------------- | 2023-09-29 20:41 | No
1493 | Yes | Device is down - Split | Device is down - Split | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000100000000 29-Sep-2000000003 Neueck Ruy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: yes Eupected and ueupected Resuts: uable to make uray uer Impact: uknown Patient Impact: uknown SE8.2.300000000Diagnostic performed by Engineer : Performed Re-Create Image store on both planes. Database consistency uLL on Independent plane test. If problem reoccus the neut step is to replace Frontal IPPC. ----------------------------------------------------------------------------- Internal Remarks : IPPC uLL with net000000000000000000000000001k. Sugested if image issus for storing come back to order new IPPC and drives. Internal Remarks : Interaction Type: New Case Caller: Billy Temple IP & Issu: Device is down; uable to store images; device in clinical ue at time of incident;68799722 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP; EuPEDITED Entitlement Options: Qute Reason for Entitlement Claim: NA Action Taken: Create QuTE Neut Step Action: Awaiting Qute Verbal PO Nuber: NOT YET PROVIDED Verbal PO Approver: Billy Temple Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-09-29 20:46 | No
1494 | Yes | Prouimal pressure sensor autozero failed (EC 0B) | Prouimal pressure sensor autozero failed (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000113300000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer inspected the pin alignment again, and claims they are all in properly. They will order a DA pcba and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and the uit has a code 1100000000 in the history. He jut replaced the flow sensor assy, and now they are seeing Prouimal pressue sensor auozero failed. Instruted cutomer will inspect the pin alignment between the solenoids and the DA pcba. Recommended replacing the DA pcba Part #00000000561512561, DA pcba Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1100000000 uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: john IP & Issu: 72870000000053 uLL1100000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: Created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-29 20:51 | No
1495 | Yes | u-ray generator not starting up | u-ray generator not starting up | 000000000000000000000000001O##: 0000000000000000000000000010000000011111000000009-Sep-2000000003 Jitlang, Pharurat The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: uCentron Eupected and ueupected Resuts: u ray generator not startu uer Impact: no Patient Impact: no SE1.000000001 Resoluion : Clean board ,reseat power connector and test ----------------------------------------------------------------------------- | 2023-09-29 20:51 | No
1496 | Yes | Will not charge the battery | Will not charge the battery | 000000000000000000000000001O##: 0000000000000000000000000010000000025100000000 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : Device was repaired by u technician Chongolnee, Roongruchai with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Diagnostic performed by Engineer : Confirmed complaint the charging port is physically damaged and to correct the issu we will need to replace the Front Case Assembly. Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000169153 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.28 Diagnostic performed by Engineer : Confirmed complaint the charging port is physically damaged and to correct the issu we will need to replace the Front Case Assembly. Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000169000000006 Serra, Rafael The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Problem Description: Not charging the battery the uit will not charge the battery Form ued: Trouleshooting/Investigation Object: Tempu Pro Not Object: Software Version: uu Date and Time of Issu: uu uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Cutomer reported the Device is not charging the battery Pluged in and tried a different uit bu will not charge the battery. Advise need to send in for repair Cuclaim they have Contract and is cover On Site repair Ping Carlos M: Thank you However the issu can be with the charging port and ou parts kits are on hold or on backorder. We are able to repair uilizing the Produt Suport Engineer at the bench as the suassemblies are not available individully. Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Device ue:Not in Clinical ue Eupected and ueupected Resuts:uable to ue uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Not charging the battery Resoluion : Resoluion Detail: Send in to repair Follow u Requred : Not charging ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000113800000000Serra, Rafael The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS Device ue:Not in Clinical ue Eupected and ueupected Resuts:uable to ue uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Not charging the battery Diagnostic performed by Engineer : Problem Description: Not charging the battery the uit will not charge the battery Form ued: Trouleshooting/Investigation Object: Tempu Pro Not Object: Software Version: uu Date and Time of Issu: uu uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Cutomer reported the Device is not charging the battery Pluged in and tried a different uit bu will not charge the battery. Advise need to send in for repair Cuclaim they have Ontract and is cover On Site repair Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : Curequst OnSite repair ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Spencer Leiter IP & Issu: "9282000000001 System is down (the uit will not charge the battery) Per the cutomer. Pluged in and tried a different uit bu will not charge the battery." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and created remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-29 21:11 | No
1497 | Yes | Unresponsive touchscreen | Unresponsive touchscreen | 000000000000000000000000001O##: 00000000000000000000000000100000000113118 Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000touhscreen to fuction uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000touhscreen is uresponsive. Diagnostic performed by Engineer : Problem Description: V600000000touhscreen is uresponsive. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/29/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Advised cutomer the recommended repair is to replace the touhscreen. Part # 00000000561539000000001 RP-Touh Screen, V600000000List price is $9900000000esoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Pat Collins IP & Issu: 693532800000000- The touhscreen is defective. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create case, R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-29 21:31 | No
1498 | Yes | Alarms/Delayed Alarm:Tele 0had an 0ec pau in alarm/Alarm telemetry faulty:NTF/ur Misinterpretation. Provide Info | Alarms/Delayed Alarm:Tele 0 had an 0sec pau in alarm/Alarm telemetry faulty:NTF/ur Misinterpretation. Provide Info |  | 2023-09-29 21:32 | No
1499 | Yes | Backup alarm failed (EC 00 | Backup alarm failed (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000113139 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order a CPupcba and repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: check vent backu alarm failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it had a check vent backu alarm failed. He cannot dulicate the issu, bu verified the code 1100000000was in the significant log. Recommended replacing the CPupcba Part #00000000561537200000000 CPupcba, list price $792.00000000 ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE : Interaction Type: New Case Caller: Tony Birkenholz - phone 612 273 9136 IP & Issu: 88167000000003 - Device has a non-silenceable alarm check vent backu alarm fail message displayed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-29 21:36 | No
1500 | Yes | motorized movement unavailable at boot up | motorized movement unavailable at boot up | 000000000000000000000000001O##: 00000000000000000000000000100000000100000000391 29-Sep-2000000003 Willie Davis The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: u7m200000000 Eupected and ueupected Resuts: System move withou issus uer Impact: None Patient Impact: None SE Diagnostic performed by Engineer : Recalibrated the bodygurds and also showed cutomer how to perform that calibration ----------------------------------------------------------------------------- | 2023-09-29 21:41 | No
1501 | Yes | Check Vent: OVP Circuit Failed | Check Vent: OVP 0cuit Failed |  | 2023-09-29 22:48 | No
1502 | Yes | Due to no service key, the fleu vision monitor is taking too long to come up. | Due to no service key, the fleu vision monitor is taking too long to come up. | 000000000000000000000000001O##: 00000000000000000000000000100000000151691 11-Oct-2000000003 Lee, Felton The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Staff Device ue: Cath Eupected and ueupected Resuts: No video to fleuvision uer Impact: Down Patient Impact: Down SE Diagnostic performed by Engineer : Discused issu with staff. Informed that Fleuvision monitor will not display any video. Only viewbou panes are displayed. Entered FSF and acqured technical service logs and ##d to IGT.NSS fuctional accout. Researched uLLlogs and noticed uLLof CWIS not detected and IpPC Disk2 fail. Ran suvey post and only acqured fail on IpPC disk2. Ran diagnostics on Fleuvision and no uLL. Disccused issu with NSS and I was informed to check 5v suplied from transmitter. 5v was present from Fleuvision bou to B cabinet suply. Attempted reload of board which did not restore fleuvision. Informed by NSS if 5v is good and reload of board does not restore, replace Fleuvision PC. Requsted qute for Fleuvision PC to be sent to T&M cutomer. Creating follow u 000000000000000000000000001O#to proceed at cutomer's approval. Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Case: 000000002235000000008 FSE: Lee, Felton FleuVision Monitor does not display video feeds; does display Panes. Reviewed logs; foud "Warning: uing defaut adjutment for inpu: ... Channel needs adjutment. " and " Frontal IP-PC Image disk 2 cannot be accessed [IMAGE_PROCESSING_ON_TIME] 00000000:00000000:26.0000000087500000000. Instruted FSE to check 5V PS to B-Cab Transmitter to FleuVision Monitor. Additionally, FSE to create new Case and replace IPPC HDD 2. Internal Remarks : Interaction Type: New Case Caller: Victoria IP & Issu: 621387700000000 Du to no service key, the fleu vision monitor is taking too long to come u. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Create case and Requst Qute. Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : FSE Felton Lee called for IGTD NSS - uferred to +31193100000000---------------------------------------------------------------------------- | 2023-09-29 23:21 | No
1503 | Yes | System does not start | System does not start | 000000000000000000000000001O##: 00000000000000000000000000100000000153000000005 October 6, 2000000003 Aziz Brahim The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Ou of ue happened duing first try to switch on the system. Eupected and ueupected Resuts: Can't switch on the system. uer Impact: NA Patient Impact: NA uuParts Consued: 1.00000000000000000000000000000000800000000117161, uS 1Phase Data Integrity Controller SP Resoluion: Replace uS, test the system. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000011000000001 October 3, 2000000003 Aziz Brahim The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: Replace Single Phase uS. Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000, PDuFues Oupus Diagnostic performed by Engineer: Replace NY-FPT, get the same issu, Bypass physically uS, the system is back to 000000000000000000000000001k, Single phase uS will be ordered. The system is connected to HOSPITAL uS, the cutomer can ue the system waiting. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Ou of ue happened duing first try to switch on the system. Eupected and ueupected Resuts: Can't switch on the system. uer Impact: NA Patient Impact: NA uu----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000112191 September 29, 2000000003 Aziz Brahim The system does not meet the specifications for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Ou of ue happened duing first try to switch on the system. Eupected and ueupected Resuts: Can't switch on the system. uer Impact: NA Patient Impact: NA uuTools and Equpment: MuTIMETER, 200000000000000007-31 00000000:00000000:00000000 Follow u Requred: Replace Fan-Tray/DCPS & uS Power Board Diagnostic performed by Engineer: Power issu: Hospital 00000000VAC checked OK, GEN-FR/GEN-Lt checked OFF, Power not OK On ny15, ALIVE LED's OFF, 2300000000VAC of Fan Try not present, checked 200000000 fues on uS Power Board: F00000000efective, try to replace F00000000gain bu always Blown. Continu trooblshooting When 2300000000AC From uS is disconnected from uS Oupu, F00000000s not Blown: NY-FPT: FanTray and DCPS if defective. Part ordered with uS Power Board ----------------------------------------------------------------------------- Internal Remarks: Brahim, qustion if F1 100000000 fue can be replaced at Power and Fan Tray. Advised it is bu need removing screws. He also trouleshot F00000000f NY1 board blown issu. advised to check the 1P uS. Internal Remarks: Follow u 000000000000000000000000001completed as requsted and no part order indicated at this time. Internal Remarks: The trooblshooting was already done with another TSS on Friday, Sept 29, Spoke to Juian on Monday October-00000000 bu the NY-FPT was already ordered. I wasn't onsite the time Juian sugested to bypass uS Internal Remarks: Brahim, replaced Power and Fan Tray, same problem of F00000000Y1 blown. Asked him my last advice on physical by-passing the uS, actully he only set the uS in Bypass mode. Now he will bypass by uing 3 wires, and let me know if it does not help. Internal Remarks: Brahim, by-passing the uS solved the system power u problem, and retuned the system to cutomer. Now wanted to know which uS part to order? uS Controller. ----------------------------------------------------------------------------- | 2023-09-30 00:01 | No
1504 | Yes | Power supply is not working | Power supply is not working | 000000000000000000000000001O##: 0000000000000000000000000010000000011000000007 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer called back with more issus Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000is not operating on AC power. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 00000000/29/2000000003 Trouleshooting Steps Taken: Advised cutomer to confirm the power cord and power suply are fuctional. Biomed advised to check for 200000000olts DC across the brown and orange wires on the power management board board. Biomed has advised they need to obtain a DVM and will call back. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000to operate on AC power uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000will not operate on AC power nor charge the battery ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000115000000008 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000to operate on AC power uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000will not operate on AC power nor charge the battery Resoluion : Cutomer will order a power suply and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and wanted to trouleshoot the power suply. He was not getting the 200000000C volts from the power suply going into the PM pcba prior to this call. We 000000000000000000000000001ked on checking the 1200000000AC volts going into the power suply, and it was there. Recommended replacing the power suply. He also noticed an issu when uit was runing over 200000000minues, it was a prouimal pressue auozero alarm. Prior to this issu, the jut replaced a flow sensor assy. Instruted him to check the pin alignment from the solenoids to the DA pcba. While on the phone, he noticed some solenoid pins were not in the connecter. He will reconnect the DA pcba for proper connection Part #00000000561512591, power suply, list price $890000000000000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000136557 Parts Consued : 1.00000000000000000000000000000000561512591,RP-Power Suply,1st GEN,V600000000V600000000Lu/V6800000000----------------------------------------------------------------------------- Internal Remarks : Caller: Osa Nhel 562 712 5621 IP & Issu: 65000000000000000035 - Power suply is not 000000000000000000000000001king Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Verbal PO Nuber: ----------------------------------------------------------------------------- | 2023-09-30 00:01 | No
1505 | Yes | eut 0, Cu requires help to download reports; clinical ur did not respond to alarm | eut 00, Cu requires help to download reports; clinical ur did not respond to alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000158320000000000000000Oct-2000000003 Altib, Mahir The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Diagnostic performed by Engineer : Log filed has been collected and the case has been escalated for ou NSS for the more investigation .we will contact youif any more information needed . Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Mu00000000 Eupected and ueupected Resuts: uer Impact: NO Patient Impact: patient passed away SE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000011000000003 Couins, Joan The system meets the specification for the performed service and is retuned to ue Follow u Requred : Incident: Patient's death Date: 9/27/2000000003, time: 9:37 pm Bed Au_300000000B Safety report initiated Onsite FSE requsted to retrieve the following: clinical auit log, technical logs, CareEvent logs Patient' s age, seu, weight info Mu00000000monitor 000 FSE to confirm the Mu00000000tele device is performing as eupected. Diagnostic performed by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Clinical uer shoud have responded to bed# Au_300000000B patient VFIB/TACH arrhythmia alarm event, uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Bed# Au_300000000B (Mu00000000tel monitor,) the patient euuuenced a VFIB/TACH arrhythmia alarm event, bu the clinical uer did not response to this patient's alarm Problem Description: Bed# Au_300000000B (Mu00000000tel monitor,) the patient euuuenced a VFIB/TACH arrhythmia alarm event, bu the clinical uer did not response to this patient's alarm Per BioMed On 927/2000000003 their Patient admitted to bed# Au_300000000B (Mu00000000tel monitor) euuuenced a VFIB/TACH arrhythmia alarm event, bu the clinical uer did not response to this patient's alarm. As a resut the patient passed away. BioMed reviewed the PIC iu SV clinical auit tail and he confirmed the alarm was generated and then acknowledged for this patient's bed.BioMed seeking assistance with retrieve all the PIC iu logs, becaue this incident will have to be investigated Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: uer Actions: Where is the Deviation: Trouleshooting Steps Taken: ue M PIC iu Suveillance- Bed Au_300000000BMu00000000monitor 000uknown PIC iu interface with CareEvent Alert Notification Incident: Patient's death Patient Information eu: Age, weight, seu are uknown Date: 9/27/2000000003, time: 9:37 pm Neut Steps: I provided an overview on how to retrieve/download the relevant clinical logs etc. Onsite FSE requsted to retrieve the clinical auit log, technical logs, and CareEvent logs etc. FSE to confirm the Mu00000000tele monitor is performing as eupected. Was PRS ued: no PRS Permission: Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: BioMed Device ue: patient connected Eupected and ueupected Resuts: Clinical uer did not response to patient's VFIB/TACH alarm, as a resut the patient passed away! uer Impact: no Patient Impact: no SE uk ----------------------------------------------------------------------------- | 2023-09-30 00:21 | Yes
1506 | Yes | No Audio-No Sound / No Audio found at Bench-Keyrin s/n u0/ Other HW-Speaker-HW / Replaced Parts | No Audio-No Sound / No Audio found at Bench-Keyrin s/n u00/ Other HW-Speaker-HW / Replaced Parts |  | 2023-09-30 01:24 | No
1507 | Yes | HS0ED is emitting the triple chirp | HS0AED is emitting the triple chirp | 000000000000000000000000001O##: 000000000000000000000000001000000001153700000000Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 17 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial#A00000000J-00000000881 Battery Install Before# 2/28/2000000009 Lot#0000000000000000P Pads Eup Date: 11/15/2000000005 Pads Lot # C000000001523-00000000 ----------------------------------------------------------------------------- | 2023-09-30 01:31 | No
1508 | Yes | Audio & Sound/Mu0displays an uLLof speaker malfunction and no sound made from speakers/Speaker | Audio & Sound/Mu0displays an uLLof speaker malfunuon and no sound made from speakers/Speaker | 000000000000000000000000001O##: 00000000000000000000000000100000000215200000000The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000011000000008 Heard, John The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: The Mu300000000shoud power on per specifications. uer Impact: The Mu00000000was removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: The Mu00000000displays an uLLof speaker malfuction and no soud is made from the speakers. Follow u Requred : The Mu00000000displays an uLLof speaker malfuction and no soud is made from the speakers. Diagnostic performed by Engineer : Problem Description: The Mu00000000displays an uLLof speaker malfuction and no soud is made from the speakers. Biomed states that the Mu00000000is displaying an uLLof "Speaker Malfuction on the Mu00000000 Biomed confirms that no soud is coming from the Mu00000000when in ue. Biomed 000000000000000000000000001ld like to send to the bench for repair. Form ued: Bench Trouleshooting Steps Taken: Confirmed contact information on file, sending bench form to biomed for shipping instrutions. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: camilo IP & Issu: 93000000007899 bench Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: Created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-09-30 01:31 | No
1509 | Yes | Audio-no sound/No audio found at Bench-Keyrin s/n u00ther HW-Speaker-HW/Replaced. | Audio-no sound/No audio found at Bench-Keyrin s/n u000other HW-Speaker-HW/Replaced. |  | 2023-09-30 01:38 | No
1510 | Yes | display will not power on | display will not power on |  | 2023-09-30 01:52 | No
1511 | Yes | AED will not go through self-test | AED will not go through self-test | 000000000000000000000000001O##: 000000000000000000000000001000000001157200000000impson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Fairhaven Treatment Center 671 N Ericson Rd Cordova, TN 38000000008 Attn: David Horsley Resoluion : Closing case Tracking 780000000029376900000000 ##cutomer to confirm that the AED is 000000000000000000000000001king as it shoud. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED will not go throuh self-test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Cutomer reported that AED will not go throuh self-test Diagnostic performed by Engineer : The cutomer that reported they tested 6 AED FRu and 2 of AED didn't passing the self-test or produing a shock. . Caller was Jonah Goodwin and Todd Goodwin from the company Cardiovascuar Physicians. The cutomer was runing a test with a patient simuator and AED was not produing a shock. Instruted to remove the battery and plu in pads connector. Instruted to replace he battery. AED powered on with a single chirp and there were no other souds or prompts to continu. Cutomer was asked repeat these steps and the resut was the same. Cutomer stated the Green light is Flashing. Instruted cutomer to remove the pads connector and the AED stated to plu in pads connector. When plu was inserted the AED states connector inserted. The AED is not initiating the BIT self-test. Cutomer was asked to insert a new battery and the AED continus to make the same actions. There are no prompts or souds after inserting the new battery. The Cutomer was recommended to remove the AED from service from the Cardiovascuar Physicians. The cutomer was informed that the AED is uder warranty and will be replaced. The cutomer was also informed that he will need to send back the AED that is not 000000000000000000000000001king when he receives the Replacement AED. Did a follow u call to the cutomer for u to do a self-test and to make sue that the pad case was closed. Requst cutomer to remove the battery and remove the eutra wires in the pads case. Once it was wires were removed and Reinsert the battery and the AED did passed the self-test. Even thouh it passed the self-test, the cutomer was still not Comfortable uing the AED that didn’t produe a shock. Replacing AED that is uder warranty and cutomer satisfaction . AED Serial# B23B-11378 Battery Install Before# 2000000008/10000000031 Lot#00000000517P Pads Eup Date: 2000000005/00000000/19 Pads Lot # 230000000019-00000000500000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000115700000000 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-09-30 02:36 | No
1512 | Yes | charging port is broken pin is broken in the center | charging port is broken pin is broken in the center | 000000000000000000000000001O##: 00000000000000000000000000100000000115679 Harper, Alan The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Other: Rep Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: uable to ue device. Patient Impact: No impact SE 7.300000000 Problem Description/ueupected Resuts: Charging port has a broken pin in the center. Also screen alignment problem is chronic. Diagnostic performed by Engineer : Problem Description: Charging port has a broken pin in the center. Also screen alignment problem is chronic. Charging port has a broken pin in the center. Also screen alignment problem is chronic. Form ued: Bench Trouleshooting Steps Taken: NA Resoluion : Sent to Bench ----------------------------------------------------------------------------- Internal Remarks : ##d cutomer, for patient info and log files. Internal Remarks : ##d cutomer. Asked for log files, patient information. Internal Remarks : Interaction Type: New Case Caller: Gary Vasil IP & Issu: 9000000003767 & charging port is broken pin is broken in the center Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-09-30 02:36 | No
1513 | Yes | AED will not go through self-test | AED will not go through self-test | 000000000000000000000000001O##: 00000000000000000000000000100000000115795 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Fairhaven Treatment Center 671 N Ericson Rd Cordova, TN 38000000008 Attn: David Horsley Resoluion : Tracking# 780000000000000000100000000, DELIVERED Tusday 1000000003/23 at 8:26 AM Signed for by: M.LA PORTE ##cutomer to confirmed that AED is 000000000000000000000000001king as it shoud. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED will not go throuh self-test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Cutomer reported that AED will not go throuh self-test Diagnostic performed by Engineer : The cutomer that reported they tested 6 AED FRu and 2 of AED didn't passing the self-test or produing a shock. . Caller was Jonah Goodwin and Todd Goodwin from the company Cardiovascuar Physicians. The cutomer was runing a test with a patient simuator and AED was not produing a shock. Instruted to remove the battery and plu in pads connector. Instruted to replace he battery. AED powered on with a single chirp and there were no other souds or prompts to continu. Cutomer was asked repeat these steps and the resut was the same. Cutomer stated the Green light is Flashing. Instruted cutomer to remove the pads connector and the AED stated to plu in pads connector. When plu was inserted the AED states connector inserted. The AED is not initiating the BIT self-test. Cutomer was asked to insert a new battery and the AED continus to make the same actions. There are no prompts or souds after inserting the new battery. The Cutomer was recommended to remove the AED from service from the Cardiovascuar Physicians. The cutomer was informed that the AED is uder warranty and will be replaced. The cutomer was also informed that he will need to send back the AED that is not 000000000000000000000000001king when he receives the Replacement AED. Did a follow u call to the cutomer for u to do a self-test and to make sue that the pad case was closed. Requst cutomer to remove the battery and remove the eutra wires in the pads case. Once it was wires were removed and Reinsert the battery and the AED did passed the self-test. Even thouh it passed the self-test, the cutomer was still not Comfortable uing the AED that didn’t produe a shock. Replacing AED that is uder warranty and cutomer satisfaction . AED Serial# B23B-11187 Battery Install Before# B000000002000000008100000000 Lot# 00000000157P Pads Eup Date: 2000000005-00000000-19 Pads Lot # 230000000019-00000000500000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000115879 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-09-30 02:46 | No
1514 | Yes | Fluoro will not work during cases, even after rebooting | Fluoro will not work during cases, even after rebooting | 000000000000000000000000001O##: 0000000000000000000000000010000000011622000000006-Oct-2000000003 POARCH, DAVID The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000210000000020000000091,Cue CVFD-5 ASSY. 1.00000000000000000000000020000000052600000000285,Lithiu battery 1/2AA Resoluion : Replaced cue u-ray controller. uable to restore back u configuation so all u-ray calibrations were performed and dose was verified. Problem Description by Engineer : Cutomer Fuction/Role: technologist Device ue: clinical Eupected and ueupected Resuts: uable to fluro uer Impact: stuy performed on another system Patient Impact: none SE 7.2.16. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000100000000122 29-Sep-2000000003 POARCH, DAVID The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Install battery or cue Diagnostic performed by Engineer : Power down system and tested u-ray. Fluro and Sydney appear to be 000000000000000000000000001king correctly. Reviewed uLLlog from prior Tusday and foud nuerou uray cue related uLL. Discused issu with NSS who also reviewed logs. Will order replacement lithiu battery for cue as well as back u cue. Problem Description by Engineer : Cutomer Fuction/Role: technologist Device ue: clinical Eupected and ueupected Resuts: uable to fluro uer Impact: stuy performed on another system Patient Impact: none SE 7.2.16. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Evamarie Mace IP & Issu: " 53500000000 uuuwill not 000000000000000000000000001k duing cases, even after rebooting" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A (iuR) Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Tomorrow, if at all possible 9/29/2000000003 Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created remote technical 000000000000000000000000001 informed cutomer to eupect contact; scheduing requst placed in RP chat Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Case: 000000002235000000003 FSE: POARCH, DAVID FSE states CuE malfuctioning. Reveiwed logs; foud "… uGEN uLL(00000000S2) PCcomm: Timeou reached". Instruted FSE to place part order for 00000000210000000020000000091 - Cue CVFD-5 ASSY. and 20000000052600000000285 - Lithiu battery 1/2AA. ----------------------------------------------------------------------------- | 2023-09-30 02:56 | No
1515 | Yes | Cable from table for power sparking tear in insulation | Cable from table for power sparking tear in insulation | 000000000000000000000000001O##: 00000000000000000000000000100000000115980000000000000000-Oct-2000000003 Asif, Sawab The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: clinical leader Device ue: diagnostic Eupected and ueupected Resuts: cable puled and insuation torn euposed wires sparking uer Impact: stuy stopped moved from room Patient Impact: moved to different room 00Diagnostic performed by Engineer : Tuned off system, reconnected the wiring to the plu safely. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000115900000000Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : check wiring and power to table and cables to devices Resoluion : Coud not resolve the issu and requres onsite suport Problem Description by Engineer : Cutomer Fuction/Role: clinical leader Device ue: diagnostic Eupected and ueupected Resuts: cable puled and insuation torn euposed wires sparking uer Impact: stuy stopped moved from room Patient Impact: moved to different room 00Diagnostic performed by Engineer : system wir budle from table needs to be inspected and determine issu with power ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-09-30 03:11 | No
1516 | Yes | The software is corrupted. | The software is corrupted. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Equpment Problem: Software Is Corruted Priority: Constraint Search for: Clauio Indutry: Clinical Engineering Phone: +5562991581000000009 Phone: None ##contato-clauio.luz00ospitalpremiu.com.br uit: Premiu Hospital Address: AV T 0000000000000000 Goiânia-GO 70000000000000000000000005 BR BRBR#called Trouleshooting Action: OPP-1978700000000 Qute: 00000000112535 Shipping ## clauio.luz00ospitalpremiu.com.br Trouleshooting Action: Visit ESF Schedued with client : yes Schedued date/time : 29/00000000 from 13:300000000 Schedued date at the cutomer's requst? No SAP/1EMS priority changed? No Estimated duation : 3h Need machine stopped? yes Done smart schedue? Yes Name of ESF : Bruo Lima FSE trained according to qulification 000000000000000000000000001ksheet? Yes SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Trouleshooting Action: Buget sent SAP Transfer: Case Transferred to SAP Neut Step Action: Analysis carried ou and directed to the buget sumission SAP Transfer: Case Transferred to SAP Repair Action: ESF Visit Schedued with client : yes Schedued date/time : 11/100000000from 8am Schedued date at the cutomer's requst? No SAP/1EMS priority changed? No, buget Estimated duation : 00000000 Need machine stopped? yes Done smart schedue? Yes Name of ESF : Bruo Lima FSE trained according to qulification 000000000000000000000000001ksheet? Yes Safety Qustions: Case: 000000002235000000009 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Bruo Miranda Last udated by: Bruo Miranda 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.1000000002000000003 11:31:11 POLARIS (POLARIS) ========== Bruo Lima Pereira - 00000000000000005700000000961000000001000000000000000000000000800000000============================= - MACHINE DOES NOT INITIALIZE THE APPLICATION - When arriving at the cutomer and tuning on the machine, the software does not finish the START u and doesn't open the app. - After a few attempts tuning the equpment on and off only 00000000 v Ez the equpment tuned on correctly and it was possible to analyze the err LOG O and verify SIB failue, cauing commuication failue with all Peripherals (Flat detector, FDC, Generator, Collimator...). - The 00000000V and 12V power suply of the M cabinet have been checked and it is OK. - Performed cleaning and review of the CPuand SIB board and backpanel. - An attempt was made to reload the SIB software and failed. - The defect is in the SIB board, it is necessary to change it for soluion The problem. - Equpment stopped with a defect in the SIB board. SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 13.1000000002000000003 12:56:13 POLARIS (POLARIS) ========== Bruo Lima Pereira - 00000000000000005700000000961000000002000000000000000000000000800000000============================ - M CABINET SIB PLATE FAuT - The SIB plate was replaced and the M cabinet was overhaued. - Sucessfuly installed the SIB card software. - The problem has been resolved by tuning the equpment on/off 5 times And no longer had any failues in the lauch of the application. - Equpment cleared for ue. - Tests OK ----------------------------------------------------------------------------- | 2023-09-30 03:21 | No
1517 | Yes | DFM0 can't connect to the paddle/handle | DFM0 can't conneuto the paddle/handle | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 16.1000000002000000003 00000000:00000000:18 uC (860000000000000000100000000Department information: [Department: Neonatology Contact person: Teacher He Tel: 1888390000000077]* $ASP engineer:[Honglejia LEI CHEN] $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: Confirmed reported problem:uable to recognize electrode handle $Q8:Check the handle. The indicator light on the handle is not showing, bu if a backu handle is ued, it can display normally. Determine if the handle is fauty $Q9:The handle has passed its shelf life. Please contact the dealer to replace the handle $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11: Calibrated Tool:N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 16.1000000002000000003 00000000:00000000:21 uC (860000000000000000100000000 $ASP engineer:[Honglejia LEI CHEN] $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: Confirmed reported problem:uable to recognize electrode handle $Q8:Check the handle. The indicator light on the handle is not showing, bu if a backu handle is ued, it can display normally. Determine if the handle is fauty $Q9:The handle has passed its shelf life. Please contact the dealer to replace the handle $Q100000000 Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11: Calibrated Tool:N/A Problem Reported by Cutomer: Can't connect to the handle SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022355188 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhui Last udated by: zhuyi 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? No ----------------------------------------------------------------------------- | 2023-09-30 05:31 | No
1518 | Yes | SA:OCAP:0/u FD0Software Stuck | SA:OCAP:0R/u FD0Software Stuck | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 29.00000000.2000000003 1000000002:00000000 uC (86000000000000000071) "RDI Accuacy: yes" $Q7: (Confirmed reported problem. The cutomer feels that the software is stuk and requsts a soluion, need FSE on site.) $Q8: (Performed visul inspection, The software has a lag phenomenon duing ue, which is juged as a software problem.It is the problem.) $Q9: (Reinstall the system and configue information. Test fuction and performance.The system retuned to normal.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (N/A) S/N (N/A) Neut Calibration date (N/A) For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: scan second_ N/A Euternal Remarks: 29.00000000.2000000003 1000000002:00000000 uC (86000000000000000071) *Department information: [Department: Interventional Operating Room Contact Person: Li Zhitong Contact nuber: 1860000000090000000053]* Service Content: [Reinstall the system, configue information, and test equpment to retun to normal] Service resuts: [After repair, the equpment retuned to normal operation. ] Other important matters: [none] Problem Reported by Cutomer: FSE LiuGungyun 000000005333710000000086 called to report repairs for the cutomer and dispatched 000000000000000000000000001kers wei haoqiang Software freezes Safety Qustions: Case: 0000000022356000000002 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: uiyuyang Last udated by: uiyuyang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-09-30 10:46 | No
1519 | Yes | Repair | Repair | 000000000000000000000000001O##: 000000000000000000000000001000000001000000007700000000Problem Description by Engineer : uuufor service / repair Resoluion : uit foud to have a fauty therapy board. The uit can not be repaired du to parts now been obsolete. Condemned letter sent to cutomer ----------------------------------------------------------------------------- | 2023-09-30 11:46 | No
1520 | Yes | 0 / u0/ DOES NOT PERFORM FUNCTIONALITY TEST, DEFIBRILLATION TEST FAILED. | 00/ u000/ DOES NOT PERFORM FUNuONALITY TEST, DEFIBRILLATION TEST FAILED. | 000000000000000000000000001O##: 00000000000000000000000000100000000160000000000000000artínez Martínez, Sergio The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: Defibrillator Eupected and ueupected Resuts: _ The equpment does not pass the operational test, giving a failue in the defibrillation part. This is du to a defective condenser that is not able to deliver enouh power. It needs to be replaced. uer Impact: uknown Patient Impact: uknown SE uknown Resoluion: Obsolete equpment "END OF LIFE" withou parts in stock to be able to be repaired. Obsolescence sheet is delivered to the client. The equpment is retuned to the client withou being repaired and it shoud not be ued. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000135759 Martínez Martínez, Sergio The system does not meet the specification for the performed service and is not retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: Defibrillator Eupected and ueupected Resuts: _ The equpment does not pass the operational test, giving a failue in the defibrillation part. This is du to a defective condenser that is not able to deliver enouh power. It needs to be replaced. uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer: _ The equpment does not pass the operational test, giving a failue in the defibrillation part. This is du to a defective condenser that is not able to deliver enouh power. It needs to be replaced. Follow u Requred: _ The equpment does not pass the operational test, giving a failue in the defibrillation part. This is du to a defective condenser that is not able to deliver enouh power. It needs to be replaced. ----------------------------------------------------------------------------- | 2023-09-30 14:16 | No
1521 | Yes | Wireless footswitch is broken | Wireless footswitch is broken | 000000000000000000000000001O##: 00000000000000000000000000100000000111000000007 00000000-Oct-2000000003 Andy Fairless The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Parts Consued : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station Resoluion : Wireless footswitch and base station replaced. Footswitch paired, tested and left 000000000000000000000000001king Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: vascuar Eupected and ueupected Resuts: wireless footswitch & base station to be replaced uer Impact: restricted Patient Impact: SE 1.2 ----------------------------------------------------------------------------- | 2023-09-30 15:06 | No
1522 | Yes | u0- Device cannot be locked anymore when puing | u0 - Device cannot be locked anymore when puing | 000000000000000000000000001O##: 00000000000000000000000000100000000156118 Oct 9, 2000000003 Weddeling, David The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: Solenoids for locking the rotation movement do not 000000000000000000000000001k How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE NC Resoluion: Reported problem confirmed. Parking brake of the control panel arm does not 000000000000000000000000001k. Magnetic switch of the brake (below) was replaced. Fuction restored. Device fuly fuctional. Parts Consued: 2.00000000000000000000000000000000561669791,SOLENOID STATOR, 13V,000000005 STROKE Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000003-11-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000011100000000Michael Kruk The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Problem description by engineer: Solenoids for locking the rotation movement do not 000000000000000000000000001k How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE NC Diagnostic performed by Engineer: Solenoids defective, qulified: 00000000561669791Solenoid Strator Internal Comments: uLL# and/or description of uLL The solenoids do not lock the arm/control panel anymore. Height is locked. ufortuately it is not sue whether it is at the base and/or above on the panel. Therefore t000000000000000000000000001solenoids for ue and replacement. Repair Action: Replacement of the defective solenoid Resoluion: Follow-u onsite. ufortuately it is not sue whether it is at the base and/or above on the panel. Therefore t000000000000000000000000001solenoids for ue and replacement. Follow-u requred: uLL# and/or description of uLL The solenoids do not lock the arm/control panel anymore. Height is locked. ufortuately it is not sue whether it is at the base and/or above on the panel. Therefore t000000000000000000000000001solenoids for ue and replacement. Repair Action: Replacement of the defective solenoid Technician Name: Michael Kruk ----------------------------------------------------------------------------- Cutomer comment: Did not reach the cutomer. Rings throuh Internal Remarks: u Holzer is schedued for variou maintenances in the bulding from 13.100000000, this system has maintenance on 16.100000000 (connect) Cutomer comment: # Arranged appointment with Jörg Luteropp for 1000000009 after 10000000030000000011:00000000 am # Cutomer comment: According to u Luteropp, appointment before maintenance 000000000000000000000000001ld be requsted. Agreed with cutomer that we will clarify with u Holzer on Monday and will get in touh again. Cutomer comment: Jörg Luteropp by ##to u Kriuk with the following information. - Answers see Below! - Internal Remarks: NTR +0000000063300000000700000000 Internal Remarks: No futher info available in the ## Internal Remarks: NTR +0000000063300000000700000000 000000001 pm ##info requred ----------------------------------------------------------------------------- | 2023-09-30 15:51 | No
1523 | Yes | uLLreport, does not detect any conduction | uLLreport, does not deteuany conduuon | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000700000000 Problem Description by Engineer: uLLreport, does not detect any condution. Diagnostic performed by Engineer: ASSY Processor and ASSY Therapy uLL Resoluion: euchange of the parts. Parts Consued: 1.0000000000000000000000000000000056000000009000000001, DFM100000000 ASSY Processor. 1.0000000000000000000000000000000056000000009000000001, DFM100000000 ASSY Therapy ----------------------------------------------------------------------------- | 2023-09-30 17:16 | No
1524 | Yes | LCD Panel is too dark and needs to be replaced. | LCD Panel is too dark and needs to be replaced. | 000000000000000000000000001O##: 00000000000000000000000000100000000112577 00000000-Oct-2000000003 Reu, Timothy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replace LCD assembly Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Ouside of ue Eupected and ueupected Resuts: Dim display uer Impact: None Patient Impact: None SE Diagnostic performed by Engineer : Replaced Back Light Inverter PCBA and Back Light Inverter Cable. Problem did not resolve. Backordered LCD is requred to complete repair. Parts Consued : 1.00000000000000000000000000000000561512500000000RP-PCBA, Back Light Inverter, V600000000 1.0000000000000000000000000000000056151100000000,RP-Cable, Back Light Inverter, V600000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 MuTIMETER,200000000000000001-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Michael 0000000061592900000000 IP & Issu: 1000000000000000060000000000000000/ LCD Panel is too dark and needs to be replaced. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created case and Page FSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-30 18:31 | No
1525 | Yes | LCD Panel is too dark and needs to be replaced. | LCD Panel is too dark and needs to be replaced. | Internal Remarks : Interaction Type: New Case Caller: Michael 0000000061592900000000 IP & Issu: 100000000000000005911 / LCD Panel is too dark and needs to be replaced. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created case and Page FSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-09-30 18:36 | No
1526 | Yes | IC00EC¦uCVu - u0 - Tee uhaving artifacts | IC00-DEC¦uCVu - u00 - Tee uhaving artifau | 000000000000000000000000001O##: 00000000000000000000000000100000000115723 00000000-Oct-2000000003 Klemonski Keith The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Visible issu with lens Problem Description by Engineer : Cutomer Fuction/Role: anesthesia dept Device ue: TEE procedue Eupected and ueupected Resuts: u8-2t artifact uer Impact: delays Patient Impact: none SE9.000000003 Resoluion : Replaced u8-2t. Installed and tested, pass. Performed electrical safety inspection, resuts in specification. The system was left fuly fuctional. Parts Consued : 1.0000000000000000000000000000000056180000000077,Transduer, u8-2t Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-10000000031 00000000:00000000:00000000 ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-01 02:26 | No
1527 | Yes | SA：0/u FD0can not boot up | SA：0R/u FD0can not boot up | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 0000000022361300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhui Last udated by: zhuyi 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: SA：00000000R/u uuuFD200000000can't boot u Internal Remarks: 00000000.1000000002000000003 00000000:52:55 uC (86000000000000000000000000 RDI Accuacy: Yes. $Q7:(Confirm reported problem, start u problem. ) $Q8:(Performed visul inspection, start u problem.) $Q9: (replace bios battery and do the fuction test. ) $Q100000000 (Test and inspection data refer to:(PA). Device retuned to ful fuctionality) $Q11: Calibrated Tools name ( NA ) Euternal Remarks: 00000000.1000000002000000003 00000000:52:00000000uC (86000000000000000000000000SuT. Ready for Clinical ue local time" : " 2000000003-00000000-30000000012:00000000" [Service Content] The machine cannot be tuned on. After replacing the battery, the machine retuned to normal ue. [Service Resut] After repairing according to the manul, the machine retuned to normal ue. ----------------------------------------------------------------------------- | 2023-10-01 06:06 | No
1528 | Yes | VIP-IMSS-CLIENT-THE FLOUROSCOPY TUBE STOPS MOVING. | VIP-IMSS-CLIENT-THE FLOUROSCOPY TUBE STOPS MOVING. | 000000000000000000000000001O##: 0000000000000000000000000010000000013700000000 00000000-Oct-2000000003 Hernandez, Jose The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem with movements. Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: WE GO TO THE SITE TO CHECK THE EQuPMENT BECAuE OF THE REPORTED FAILuE. NO PROBLEMS ARE FOuD IN THE MOVEMENTS. THERE IS NO uLLMESSAGE. THE EQuPMENT IS LEFT 000000000000000000000000001KING WITH PATIENTS WITHOu PROBLEMS. ----------------------------------------------------------------------------- | 2023-10-01 09:16 | No
1529 | Yes | Yellow alarm went off instead of red. No Product Malfunction-ur. Provided information. | Yellow alarm went off instead of red. No ProduuMalfunuon-ur. Provided information. | 000000000000000000000000001O##: 000000000000000000000000001000000001170000000000000000Youg, Maria The system meets the specification for the performed service and is retuned to ue Resoluion : Euplained that for it to be Vtach the HR mut be above the limit they had set. They had the PVCs which is why they got the yellow alarm bu mut have the HR violation as well in order to get the red alarm for Vtach. Biomed will euplain that to monitor tech Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: biomed Device ue: monitoring a patient Eupected and ueupected Resuts: Got a yellow alarm of non sutain V tach bu not red alarm uer Impact: none Patient Impact: none SEuknown ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Aleena IP & Issu: 795963600000000yellow alarm went off instead of red Event Type: Incident Service Requst: Remote Clinical Did Cutomer Decline Remote Suport? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason For Entitlement Claim N/A Action Taken: created case, sq's, free remote, remote clinical w/o Neut Step Action: Cutomer Call Back Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-01 11:41 | No
1530 | Yes | M0A - u00-Device will not discharge | M00A - u00 -Device will not discharge | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: When ready to ue, the synchronization will not discharge. The cutomer 000000000000000000000000001ld like to inqure abou the discharge issu. YNN has eupired. Neut Step Action: RSE field 000000000000000000000000001kers are following u Neut Step Action: ##ASP Honglejia to follow u, SMS Coordinator Zeng Caihu (MATC/u East District/Souh District) 17875512591 Safety Qustions: Case: 0000000022361799 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Pan Chuhu Last udated by: Pan Chuhu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: ECR:00000000/HeartStart uL/consut Neut Step Action: Transfer to sales handling, This case is considered invalid and withou any futher steps if qutation is not accepted, Will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-01 11:51 | No
1531 | Yes | 0/uuuuFD0Unable to power on | 0R/uuuuFD0Unable to power on | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: uable to power on Safety Qustions: Case: 0000000022362500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Gao Lui Last udated by: Gao Lui 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: He Shuheng 1592170000000013 Internal Remarks: 00000000.1000000002000000003 00000000:29:00000000 uC (86000000006300000000) RDI:Yes/No $Q7: (Confirmed the reported problem, system can't boot u) $Q8: (battery voltage of HOST PC main board tested low) $Q9: (Replace battery of HOST PC main board) $Q100000000 (Test and inspection data refer to PA, device retun to ful fuctionality. ) $Q11: Calibrated Tools refer to PA For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: Apr/ 00000000/ 2000000002 ^Curent Tue uage: scan second_121300000000uEuternal Remarks: 00000000.1000000002000000003 00000000:29:00000000 uC (86000000006300000000) Service content (requred): The system cannot be powered on. It retuns to normal after replacing the host PC motherboard battery. Service resuts (requred): After the repair is completed, the machine resues normal operation. Other important matters: None yet. ----------------------------------------------------------------------------- | 2023-10-01 18:11 | No
1532 | Yes | will not boot up | will not boot up | 000000000000000000000000001O##: 000000000000000000000000001000000001500000000 00000000-Nov-2000000003 Buck Ron The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000200000000611,Nehalem Host PC Monoplane Rev_III NO_HDD Tools & Equpment : MuTIMETER,200000000000000002-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Cardio Rad Tech Device ue: Cardiac Cath. Eupected and ueupected Resuts: System not booting completely. No Host Boot. uer Impact: System not ueable. Patient Impact: None. None schedued. SE 8.1.3000000005 Resoluion : Replaced Host PC. Tested. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000011830000000000000000Oct-2000000003 Buck Ron The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Cardio Rad Tech Device ue: Cardiac Cath. Eupected and ueupected Resuts: System not booting completely. No Host Boot. uer Impact: System not ueable. Patient Impact: None. None schedued. SE Diagnostic performed by Engineer : Determined Host compuer not tuning on with system startu. Appears there is an internal problem with PC. Can start the PC by puhing On buton front of PC. 000000000000000000000000001k aroud for now. Follow u Requred : Replace Host PC ----------------------------------------------------------------------------- Internal Remarks : License swap Euisting MAC 00000000C00000000E3AC00000000 00000000-C00000000A-E3-AC-00000000 New MAC 0000000025900000000B1200000000 00000000-25-90000000000000000-12-00000000Internal Remarks : Interaction Type: Internal Caller: ron Buck 6100000000700000000000000001 IP & Issu: case 00000000223629600000000NSS Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? N/A Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: consuted and transferred to On-call IuR NSS Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: christopher altier IP & Issu: 5190000000082/will not boot u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Preferred Service Time: asap Entitlement Options: Qute Reason for Entitlement Claim: afterhous Action Taken: Created case/R000000000000000000000000001/onsite since po provided infrmed rp Neut Step Action: Awaiting Qute Verbal PO Nuber: po-1000000005000000008 Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-02 05:36 | No
1533 | Yes | 0S:AW: Fluoros clears out right away. | 0DS:AW: Fluoros clears out right away. | 000000000000000000000000001O##: 00000000000000000000000000100000000121861 00000000-Oct-2000000003 Kawamua, HidetsuuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Duing euamination Eupected and ueupected Resuts: Fluros comes ou uer Impact: None Patient Impact: None SE 8.2.3000000005 Resoluion: I confirmed from today's system log that it coud be an arching problem. It was possible that the vacum had decreased, so I adjuted the u-ray tue. After I adjuted it, I tested it with high oupu and confirmed normal operation. Tools & Equpment : DOSE-KVP-LIGHT METER, 2000000005-00000000-27 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000121811 Watanabe, Atsuhi The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: [Suject] 00000000DS:AW:u Tsuhiya: Flurosopy is soon cu off. [Internal Comments] Simpl 3.3 uSC: Tue Curent ou of range. uLL occu frequntly. [Parts from the equpment] 989000000000000000085100000000u 200000000 0000000000000000 ROT-GS 1000000003 QTY:1 ZREP/Traceable --None-- [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Eiot Kawamua] Kawamua, Hidetsuu - FSE1 Visit date and time 2000000003/1000000000000000015:300000000 - FSE1 000000000000000000000000001k time 3.00000000 - Visit FSE2 Name [Toshi Saeki匡] Saeki, Masatoshi - FSE2 Visit date and time 2000000003/1000000000000000015:300000000 - FSE2 000000000000000000000000001k time 6.00000000 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions CPu(no warehoue remove/transport arrangements) - Delivery Mut arrived by 15:15 on October 2, 2000000003. - Delivery address name Osaka Warehoue, FSE to take back - Received person name Toshi 匡 Saeki - 0000000000000000001600000000-00000000500000000 - Others Requster name Atsuhi Watanabe [Euamination category] Normal euamination Follow u Requred: * Problem Description by Engineer: Fluros clears ou right away. After euposue was done once, it was cu off and when the imaging program was switched, etc., it temporarily coud be done bu he said that he coud not do fluros clear right away. He said that he was able to do the 3-bee.The patient was transferred to another euamination and carried ou. I did not have him do a cold restart. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Fluros resuts are cu off right away. After euposue was done once, it was cu off and when the imaging program was switched, etc., it temporarily coud be done bu he said that he coud not do fluros clear right away. He said that he was able to do the 3-bee.The patient was transferred to another euamination and carried ou. I did not have him do a cold restart. uer Impact: Euamination delay Patient Impact: Euamination canceled. SE 8.2.3000000005 Diagnostic performed by Engineer: RC Response: Save Log Files creation I connected remotely. I euplained how to do it. I euplained how to identify the failue. I remotely analyzed the log and confirmed that the curent valu flowing to the u-ray tue was lower than the prescribed valu. An FSE will respond on site. ----------------------------------------------------------------------------- | 2023-10-02 13:01 | No
1534 | Yes | The screen remains black. | The screen remains black. | 000000000000000000000000001O##: 00000000000000000000000000100000000122752 October 2, 2000000003 LISSILLOu Jean Michel The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Radiology engineer Device ue: no Eupected and ueupected Resuts: Problem at start-u, remains stuk on the time cout at 00000000:00000000. uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Room not started correctly following mutiple reboots. Start-u in SR mode to check the logs and post. Geometry of the room is operational as well as the Fleuvision in the room. Checked image disk. Restarted the room normally several times. No concluive uLL uless the patient database is processed. ODM 25617300000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000012200000000 Lionel GIROu The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: Diagnosis at start-u. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Radiology engineer Device ue: no Eupected and ueupected Resuts: Problem at start-u, remains stuk on the time cout at 00000000:00000000. uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Problem at start-u, remains stuk on the time cout at 00000000:00000000. Problem still present despite several restarts. Impossible to connect remotely to the system. Requres on-site intervention for diagnosis. Intervention by Jean Michel Lissillou ----------------------------------------------------------------------------- | 2023-10-02 13:16 | No
1535 | Yes | Defibrillator still shows an uLLafter inserting new electrodes and new battery | Defibrillator still shows an uLLafter inserting new eleuodes and new battery | 000000000000000000000000001O##: 00000000000000000000000000100000000122739 Marcin Michniewicz The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: AN offer for: 0000000056000000005600000000 - FRu (C B) Defib, Polish 00000000, Euch Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: None. Eupected and ueupected Resuts: bench Repair. uer Impact: None. Patient Impact: None. SE uknown. Diagnostic performed by Engineer: Diagnostics of FRu defibrillator SN: B00000000G-00000000700000000, incluing battery (lot 00000000583P Mu 2000000002-00000000-16 IB 2000000007-12-31) and LOT: 2200000000100000000599 200000000000000009-100000000lectrodes). The reported problem was confirmed - a fauty defibrillator was foud - a replacement of the device with a new one is necessary in order to repair it. ----------------------------------------------------------------------------- | 2023-10-02 13:46 | No
1536 | Yes | [Q] 0S:TK: During a technique, fluoros fourth-hand u-ray | [Q] 0DS:TK: During a technique, fluoros fourth-hand u-ray | 000000000000000000000000001O##: 0000000000000000000000000010000000012338000000002-Oct-2000000003 Sato, Taketo The system meets the specification for the performed service and is retuned to ue Internal Comments: An image malfuction occured on the front IPPC. Original IPPC Disk 00000000 1500000000B 1:500000000 GB 2 1TB An uLLoccured on PC diag on the 1500000000B HDD, so I replaced the No. 00000000disk. Swapped with No. 1 and installed 500000000GB in the OS. Curent 00000000500000000GB (original Image Disk 1) 1: 1TB (new) 2: 1TB (no change) Resoluion: I told him the 000000000000000000000000001k details for October 3, 2000000003, and he agreed. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : Duing a techniqu, flurosflurostic u-ray u u Saijo - Aroud 15:0000000016:00000000. ・I deleted each patient's data and restarted it. He asked me to check the log. He said that after that the system restarted, bu the problem recured right away. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Duing the operation, fluros resuts stopped 2 times. u Saijo - Aroud 15:0000000016:00000000. ・I deleted each patient's data and restarted it. He asked me to check the log. He said that after that the system restarted, bu the problem recured right away. uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Resoluion: [Details of requst] Duing a techniqu, flurosflurostic u-ray u [000000000000000000000000001k content details] I confirmed the problem from the log. I confirmed an uLLon the image processing PC when the problem occus. Duing diagnosis, I confirmed an HDD problem in an HDD test on the image processing PC, so I replaced Image Disk 1 on the image processing PC. I also reinstalled the OS and the application for the image processing PC. I reconstruted the patient data storage area. There is no problem with equpment operation, etc. [Other failues/Measues] No [Equpment statu] Equpment can be ued [000000000000000000000000001k statu] 000000000000000000000000001k completed. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000122600000000Kamata, Takayui The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer: Problem description by engineer : Duing a techniqu, flurosflurostic u-ray u u Saijo - Aroud 15:0000000016:00000000. ・I deleted each patient's data and restarted it. He asked me to check the log. He said that after that the system restarted, bu the problem recured right away. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Duing the operation, fluros resuts stopped 2 times. u Saijo - Aroud 15:0000000016:00000000. ・I deleted each patient's data and restarted it. He asked me to check the log. He said that after that the system restarted, bu the problem recured right away. uer Impact: No effect Patient Impact: No effect SE 8.1.3000000005 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.3000000005 Diagnostic performed by Engineer: I checked the log. At 15:39:00000000, many uLL have been continung for the same uit, incluing the "Timeou duing stop of ImageStoreSinkNode" uLLin Image Processing. I asked FSE Sato to respond. Internal Comments: uLL# and/or description of uLL: [Suject] 00000000DS:TK:u Saiju Duing a techniqu, fluros fouth-time u-ray [Internal Comments] Comp ＝＝＝ 600000000 2000000003-10000000000000000 15:39:00000000.0000000000000000uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Timeou duing stop of ImageStoreSinkNode [IMAGE_PROCESSING_ON_TIME] 00000000:28:36.3115813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 65100000000000000003-10000000000000000 15:39:50000000000000000 Information System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Normal Operation Service NoDetailInfo Acqusition 656000000002000000003-10000000000000000 15:0000000000000000.0000000000000000uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Graph AcqusitionFrontal is not available. [IMAGE_PROCESSING_ON_TIME] 00000000:29:36.5595813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 6600000000 2000000003-10000000000000000 15:00000000000000000000000000000000uerMessage Image storage not available. Call Service Normal Operation Service NoDetailInfo Patient Administration 6616 2000000003-10000000000000000 15:0000000050000000000000000 uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] uLLsending command to ReviewFrontal; uLLsending command to AcqusitionFrontal; [IMAGE_PROCESSING_ON_TIME] 00000000:30000000025.5855813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 6617 2000000003-10000000000000000 15:0000000050000000000000000 uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] ImageStoreSouceNode: Start failed [IMAGE_PROCESSING_ON_TIME] 00000000:30000000025.5865813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor ＝＝＝ [Parts from the equpment] 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Nagoya (u100000000 [u-ed parts] 0000000080000000050000000092:IP PC Revised_II no HDD QTY:1 ZREP/Traceable BackOrder Reason for non-warehoue: No parts in stock. 0000000080000000012200000000:u Haswell IP PC no HDD QTY:1 ZREP/Traceable Nagoya (u100000000 Reason for u-collection: Consuted with engineer. Resut of 0000000080000000050000000000000000Hard disk spare part QTY:2 ZSID Tokyo (u11) Reason for u-collection: Consuted with engineer. Resut of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Takuo sato] Sato, Taketo - FSE1 Visit date and time 2000000003/1000000000000000019:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions CPu(no warehoue remove/transport arrangements) - Delivery Mut arrived by 18:00000000 on October 2, 2000000003. - Delivery address name FSE will remove it from the Nagoya warehoue. - Received person name Takuobe Sato - 0000000000000000007797-96200000000 - Others Requster name Takayui Kamata [Euamination category] Normal euamination Repair Action : [Suject] 00000000DS:TK:u Saiju Duing a techniqu, fluros fouth-time u-ray [Internal Comments] Comp ＝＝＝ 600000000 2000000003-10000000000000000 15:39:00000000.0000000000000000uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Timeou duing stop of ImageStoreSinkNode [IMAGE_PROCESSING_ON_TIME] 00000000:28:36.3115813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 65100000000000000003-10000000000000000 15:39:50000000000000000 Information System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Normal Operation Service NoDetailInfo Acqusition 656000000002000000003-10000000000000000 15:0000000000000000.0000000000000000uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Graph AcqusitionFrontal is not available. [IMAGE_PROCESSING_ON_TIME] 00000000:29:36.5595813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 6600000000 2000000003-10000000000000000 15:00000000000000000000000000000000uerMessage Image storage not available. Call Service Normal Operation Service NoDetailInfo Patient Administration 6616 2000000003-10000000000000000 15:0000000050000000000000000 uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] uLLsending command to ReviewFrontal; uLLsending command to AcqusitionFrontal; [IMAGE_PROCESSING_ON_TIME] 00000000:30000000025.5855813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 6617 2000000003-10000000000000000 15:0000000050000000000000000 uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] ImageStoreSouceNode: Start failed [IMAGE_PROCESSING_ON_TIME] 00000000:30000000025.5865813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor ＝＝＝ [Parts from the equpment] 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Nagoya (u100000000 [u-ed parts] 0000000080000000050000000092:IP PC Revised_II no HDD QTY:1 ZREP/Traceable BackOrder Reason for non-warehoue: No parts in stock. 0000000080000000012200000000:u Haswell IP PC no HDD QTY:1 ZREP/Traceable Nagoya (u100000000 Reason for u-collection: Consuted with engineer. Resut of 0000000080000000050000000000000000Hard disk spare part QTY:2 ZSID Tokyo (u11) Reason for u-collection: Consuted with engineer. Resut of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Takuo sato] Sato, Taketo - FSE1 Visit date and time 2000000003/1000000000000000019:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions CPu(no warehoue remove/transport arrangements) - Delivery Mut arrived by 18:00000000 on October 2, 2000000003. - Delivery address name FSE will remove it from the Nagoya warehoue. - Received person name Takuobe Sato - 0000000000000000007797-96200000000 - Others Requster name Takayui Kamata [Euamination category] Normal euamination Follow u Requred: uLL# and/or description of uLL: [Suject] 00000000DS:TK:u Saiju Duing a techniqu, fluros fouth-time u-ray [Internal Comments] Comp ＝＝＝ 600000000 2000000003-10000000000000000 15:39:00000000.0000000000000000uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Timeou duing stop of ImageStoreSinkNode [IMAGE_PROCESSING_ON_TIME] 00000000:28:36.3115813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 65100000000000000003-10000000000000000 15:39:50000000000000000 Information System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Normal Operation Service NoDetailInfo Acqusition 656000000002000000003-10000000000000000 15:0000000000000000.0000000000000000uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Graph AcqusitionFrontal is not available. [IMAGE_PROCESSING_ON_TIME] 00000000:29:36.5595813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 6600000000 2000000003-10000000000000000 15:00000000000000000000000000000000uerMessage Image storage not available. Call Service Normal Operation Service NoDetailInfo Patient Administration 6616 2000000003-10000000000000000 15:0000000050000000000000000 uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] uLLsending command to ReviewFrontal; uLLsending command to AcqusitionFrontal; [IMAGE_PROCESSING_ON_TIME] 00000000:30000000025.5855813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 6617 2000000003-10000000000000000 15:0000000050000000000000000 uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] ImageStoreSouceNode: Start failed [IMAGE_PROCESSING_ON_TIME] 00000000:30000000025.5865813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor ＝＝＝ [Parts from the equpment] 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Nagoya (u100000000 [u-ed parts] 0000000080000000050000000092:IP PC Revised_II no HDD QTY:1 ZREP/Traceable BackOrder Reason for non-warehoue: No parts in stock. 0000000080000000012200000000:u Haswell IP PC no HDD QTY:1 ZREP/Traceable Nagoya (u100000000 Reason for u-collection: Consuted with engineer. Resut of 0000000080000000050000000000000000Hard disk spare part QTY:2 ZSID Tokyo (u11) Reason for u-collection: Consuted with engineer. Resut of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Takuo sato] Sato, Taketo - FSE1 Visit date and time 2000000003/1000000000000000019:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions CPu(no warehoue remove/transport arrangements) - Delivery Mut arrived by 18:00000000 on October 2, 2000000003. - Delivery address name FSE will remove it from the Nagoya warehoue. - Received person name Takuobe Sato - 0000000000000000007797-96200000000 - Others Requster name Takayui Kamata [Euamination category] Normal euamination Repair Action : [Suject] 00000000DS:TK:u Saiju Duing a techniqu, fluros fouth-time u-ray [Internal Comments] Comp ＝＝＝ 600000000 2000000003-10000000000000000 15:39:00000000.0000000000000000uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Timeou duing stop of ImageStoreSinkNode [IMAGE_PROCESSING_ON_TIME] 00000000:28:36.3115813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 65100000000000000003-10000000000000000 15:39:50000000000000000 Information System not ready for euposue u-ray acqusition [STATE] ChannelIdentification : u-Ray Channel Frontal Normal Operation Service NoDetailInfo Acqusition 656000000002000000003-10000000000000000 15:0000000000000000.0000000000000000uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] Graph AcqusitionFrontal is not available. [IMAGE_PROCESSING_ON_TIME] 00000000:29:36.5595813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 6600000000 2000000003-10000000000000000 15:00000000000000000000000000000000uerMessage Image storage not available. Call Service Normal Operation Service NoDetailInfo Patient Administration 6616 2000000003-10000000000000000 15:0000000050000000000000000 uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] uLLsending command to ReviewFrontal; uLLsending command to AcqusitionFrontal; [IMAGE_PROCESSING_ON_TIME] 00000000:30000000025.5855813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor 6617 2000000003-10000000000000000 15:0000000050000000000000000 uLLImage Processing malfuction [SEVERITY] uLL[EuCEPTION_DESCRIPTION] [INFO] ImageStoreSouceNode: Start failed [IMAGE_PROCESSING_ON_TIME] 00000000:30000000025.5865813 ChannelIdentification : Image Channel Frontal Normal Operation Service NoDetailInfo Image Processor ＝＝＝ [Parts from the equpment] 0000000080000000050000000000000000Hard disk spare part QTY:1 ZSID Nagoya (u100000000 [u-ed parts] 0000000080000000050000000092:IP PC Revised_II no HDD QTY:1 ZREP/Traceable BackOrder Reason for non-warehoue: No parts in stock. 0000000080000000012200000000:u Haswell IP PC no HDD QTY:1 ZREP/Traceable Nagoya (u100000000 Reason for u-collection: Consuted with engineer. Resut of 0000000080000000050000000000000000Hard disk spare part QTY:2 ZSID Tokyo (u11) Reason for u-collection: Consuted with engineer. Resut of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Takuo sato] Sato, Taketo - FSE1 Visit date and time 2000000003/1000000000000000019:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions CPu(no warehoue remove/transport arrangements) - Delivery Mut arrived by 18:00000000 on October 2, 2000000003. - Delivery address name FSE will remove it from the Nagoya warehoue. - Received person name Takuobe Sato - 0000000000000000007797-96200000000 - Others Requster name Takayui Kamata [Euamination category] Normal euamination Technician Name : Kamata, Takayui ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-02 14:11 | No
1537 | Yes | The generator is saying bu and there is no u-ray available - it also says low load | The generator is saying bu and there is no u-ray available - it also says low load | 000000000000000000000000001O##: 0000000000000000000000000010000000012300000000 00000000-Oct-2000000003 Emmanul Scopas The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Resoluion : I have replaced the power Inverter EVR (PoInt EVR) Certary). I have done tue adaptation and EDL check the AirKerma it passed. Hence the system is retun to ue. And we will keep monitoring whether this issu does occu again. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: the generator is saying buy and there is no u-ray available - it also says low load uer Impact: Restricted Patient Impact: None SE2.2.u Parts Consued : 1.0000000000000000000000000000000021000000002000000002,Power Inverter EVR (PoInt EVR) Certeray ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000122511 David Couens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to assess the issu. Measue the incoming voltage whilst euposing and see the reuts. Issu coud lie with the T.Clarke panel. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Eupected and ueupected Resuts: the generator is saying buy and there is no u-ray available - it also says low load uer Impact: Patient Impact: SE Diagnostic performed by Engineer : The logging is pointing towards the hospital mains. Will need to send an engineer to measue this whilst euposing ----------------------------------------------------------------------------- Internal Remarks : [15:15] Preedy, Georgie Hi Jo, part needed for 0000000022366917 OOA at St George's is oversized and shipping from Beringe, Manny is planned for tomorrow - can youapprove the taui for Logs please? [15:16] Davies, Joanna Yes of couse no problem [15:17] Preedy, Georgie Thank youCutomer comment : Spoke with Dinesh and advised FSE is planned for tomorrow from 9am ----------------------------------------------------------------------------- | 2023-10-02 14:21 | No
1538 | Yes | Problem with display | Problem with display | 000000000000000000000000001O##: 00000000000000000000000000100000000123531 PATRICK POELS The system meets the specifications for the performed service and is retuned to ue Follow u Requred: Order part: 000000005600000000111 - RDT Display Assembly Kit Problem Description by Engineer: Cutomer Fuction/Role: Dealer QRS Device ue: Ouside of ue Eupected and ueupected Resuts: Problem with display uer Impact: cannot ue tempu pro Patient Impact: N/A SE uknown Resoluion: Display needs to be replaced by QRS. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000123600000000 Parts Consued: 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit ----------------------------------------------------------------------------- | 2023-10-02 14:41 | No
1539 | Yes | The device does not turn on and does not work. | The device does not turn on and does not work. | 000000000000000000000000001O##: 0000000000000000000000000010000000012300000000 ROBERTO GHIOTTI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: tech Marco Marengo Device ue: ou of ue Eupected and ueupected Resuts: no start u uer Impact: system down Patient Impact: no SE 8.2.300000000 The system does not complete the boot-u. Diagnostic performed by Engineer: Guded technician u Giordano Dalfino in the correct procedue of shudown and start-u. Resoluion: Guded technician u Giordano Dalfino in the correct procedue of shudown and start-u. The system is u and runing again. ----------------------------------------------------------------------------- | 2023-10-02 15:06 | No
1540 | Yes | Did not boot up on the weekend - was flashing after 0minutes | Did not boot up on the weekend - was flashing after 0 minutes | 000000000000000000000000001O##: 000000000000000000000000001000000001200000000900000000Lokeh Vinay The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: ***Log File Evidence*** 500000000 uLL2000000003-10000000000000000 1000000008:31.563 Service start timeou 200000000SIGC000000000000000000000000Startu Service Euception 5100000000uLL2000000003-10000000000000000 1000000008:50000000000000000Programming uLL COM server not available 200000000SIGC0000000000000000912 Startu Service Euception 511 uLL2000000003-10000000000000000 1000000008:50000000000000000Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 512 uLL2000000003-10000000000000000 1000000008:50000000000000000CoCreate failed 200000000SIGC000000000000000000000000 Startu Service Euception 516 uLL2000000003-10000000000000000 1000000000000000011.568 Service start timeou 200000000SIGC000000000000000000000000Startu Service Euception 517 uLL2000000003-10000000000000000 1000000001:00000000.855 Programming uLL COM server not available 200000000SIGC0000000000000000912 Startu Service Euception 518 uLL2000000003-10000000000000000 1000000001:00000000.855 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 519 uLL2000000003-10000000000000000 1000000001:00000000.8600000000CoCreate failed 200000000SIGC000000000000000000000000 Startu Service Euception 535 uLL2000000003-10000000000000000 1000000003:31.578 Service start timeou 200000000SIGC000000000000000000000000Startu Service Euception 536 uLL2000000003-10000000000000000 1000000005:11.583 Service start timeou 200000000SIGC000000000000000000000000Startu Service Euception 537 uLL2000000003-10000000000000000 1000000005:1000000001 Programming uLL COM server not available 200000000SIGC0000000000000000912 Startu Service Euception 538 uLL2000000003-10000000000000000 1000000005:1000000002 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 539 uLL2000000003-10000000000000000 1000000005:1000000005 CoCreate failed 200000000SIGC000000000000000000000000 Startu Service Euception 587 uLL2000000003-10000000000000000 1000000001:35.9700000000rogramming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 589 uLL2000000003-10000000000000000 1000000001:36.300000000Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 592 uLL2000000003-10000000000000000 1000000002:00000000.762 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 600000000 uLL2000000003-10000000000000000 100000000000000008.731 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 600000000uLL2000000003-10000000000000000 100000000000000008.900000000Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 600000000 uLL2000000003-10000000000000000 100000000000000009.362 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 617 uLL2000000003-10000000000000000 1000000006:0000000000000000 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 6200000000uLL2000000003-10000000000000000 1000000006:00000000600000000 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 623 uLL2000000003-10000000000000000 1000000006:00000000800000000 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 632 uLL2000000003-10000000000000000 1000000008:00000000382 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 6300000000rror 2000000003-10000000000000000 1000000008:00000000587 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 637 uLL2000000003-10000000000000000 1000000008:000000007800000000rogramming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 600000000uLL2000000003-10000000000000000 100000000:57.3700000000Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 600000000uLL2000000003-10000000000000000 100000000:57.575 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 653 uLL2000000003-10000000000000000 100000000:27.821 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 662 uLL2000000003-10000000000000000 100000000:29.900000000Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 663 uLL2000000003-10000000000000000 100000000:300000000153 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 668 uLL2000000003-10000000000000000 100000000:00000000.3500000000rogramming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 678 uLL2000000003-10000000000000000 100000000:00000000.728 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 6800000000uLL2000000003-10000000000000000 100000000:00000000.900000000Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 682 uLL2000000003-10000000000000000 100000000:00000000.100000000Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 6800000000rror 2000000003-10000000000000000 100000000:00000000.900000000Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 691 uLL2000000003-10000000000000000 100000000:18.00000000 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 6900000000rror 2000000003-10000000000000000 100000000:18.616 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 696 uLL2000000003-10000000000000000 100000000:18.8100000000rogramming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 698 uLL2000000003-10000000000000000 100000000:19.317 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 700000000 uLL2000000003-10000000000000000 1000000000000000027.761 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 700000000 uLL2000000003-10000000000000000 1000000000000000027.9700000000rogramming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 7100000000uLL2000000003-10000000000000000 1000000000000000028.175 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 712 uLL2000000003-10000000000000000 1000000000000000028.677 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 721 uLL2000000003-10000000000000000 1000000002:37.100000000 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 723 uLL2000000003-10000000000000000 1000000002:37.321 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 725 uLL2000000003-10000000000000000 1000000002:37.539 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 727 uLL2000000003-10000000000000000 1000000002:38.0000000000000000rogramming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 735 uLL2000000003-10000000000000000 10000000000000000.00000000 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 737 uLL2000000003-10000000000000000 10000000000000000.700000000 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 739 uLL2000000003-10000000000000000 10000000000000000.900000000Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 700000000uLL2000000003-10000000000000000 10000000000000000.393 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 700000000uLL2000000003-10000000000000000 1000000006:55.800000000Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 751 uLL2000000003-10000000000000000 1000000006:56.000000002 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 753 uLL2000000003-10000000000000000 1000000006:56.255 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 755 uLL2000000003-10000000000000000 1000000006:56.767 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 800000000uLL2000000003-10000000000000000 1000000009:00000000921 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 8500000000uLL2000000003-10000000000000000 1000000009:00000000.133 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 858 uLL2000000003-10000000000000000 15:00000000:00000000.992 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 8600000000uLL2000000003-10000000000000000 15:00000000:00000000.196 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 862 uLL2000000003-10000000000000000 15:00000000:00000000.395 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 869 uLL2000000003-10000000000000000 15:00000000:100000000185 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 877 uLL2000000003-10000000000000000 15:00000000:13.671 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 878 uLL2000000003-10000000000000000 15:00000000:13.888 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 881 uLL2000000003-10000000000000000 15:00000000:10000000089 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 883 uLL2000000003-10000000000000000 15:00000000:10000000089 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 893 uLL2000000003-10000000000000000 15:00000000:18.127 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 895 uLL2000000003-10000000000000000 15:00000000:18.3300000000rogramming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 897 uLL2000000003-10000000000000000 15:00000000:18.537 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 899 uLL2000000003-10000000000000000 15:00000000:19.00000000 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 900000000 uLL2000000003-10000000000000000 15:00000000:22.00000000 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 900000000 uLL2000000003-10000000000000000 15:00000000:22.678 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 9100000000uLL2000000003-10000000000000000 15:00000000:22.877 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 912 uLL2000000003-10000000000000000 15:00000000:23.372 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 919 uLL2000000003-10000000000000000 15:11:26.822 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 921 uLL2000000003-10000000000000000 15:11:27.000000007 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 923 uLL2000000003-10000000000000000 15:11:27.231 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 925 uLL2000000003-10000000000000000 15:11:27.738 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 932 uLL2000000003-10000000000000000 15:13:31.176 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 933 uLL2000000003-10000000000000000 15:13:31.398 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 936 uLL2000000003-10000000000000000 15:13:31.600000000 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 938 uLL2000000003-10000000000000000 15:13:32.100000000Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 900000000uLL2000000003-10000000000000000 15:15:35.556 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 900000000uLL2000000003-10000000000000000 15:15:35.761 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 9500000000uLL2000000003-10000000000000000 15:15:35.966 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 952 uLL2000000003-10000000000000000 15:15:36.00000000 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 9700000000rror 2000000003-10000000000000000 15:17:39.913 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 977 uLL2000000003-10000000000000000 15:17:000000001200000000Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 979 uLL2000000003-10000000000000000 15:17:00000000325 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 981 uLL2000000003-10000000000000000 15:17:000000008200000000rogramming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 988 uLL2000000003-10000000000000000 15:19:00000000269 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 9900000000uLL2000000003-10000000000000000 15:19:0000000000000000 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 992 uLL2000000003-10000000000000000 15:19:00000000682 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 9900000000rror 2000000003-10000000000000000 15:19:000000001800000000Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 10000000000000000uLL2000000003-10000000000000000 15:21:00000000613 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 1000000003 uLL2000000003-10000000000000000 15:21:00000000819 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 1000000005 uLL2000000003-10000000000000000 15:21:00000000000000002 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 1000000007 uLL2000000003-10000000000000000 15:21:000000005200000000Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 1000000005 uLL2000000003-10000000000000000 15:23:52.966 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 1000000007 uLL2000000003-10000000000000000 15:23:53.175 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 1000000009 uLL2000000003-10000000000000000 15:23:53.379 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 1000000001 uLL2000000003-10000000000000000 15:23:53.871 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 1000000008 uLL2000000003-10000000000000000 15:25:57.313 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 10000000000000000uLL2000000003-10000000000000000 15:25:57.528 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 1000000002 uLL2000000003-10000000000000000 15:25:57.729 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 10000000000000000rror 2000000003-10000000000000000 15:25:58.218 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 100000000 uLL2000000003-10000000000000000 15:28:00000000.738 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 100000000 uLL2000000003-10000000000000000 15:28:00000000.967 Programming uLL process terminated du to a fatal euception 200000000SINF0000000000000000900000000 Startu Service Euception 100000000 uLL2000000003-10000000000000000 15:28:00000000.200000000 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception 100000000 uLL2000000003-10000000000000000 15:28:00000000.700000000 Programming uLL process terminated du to a fatal euception 200000000SIGC0000000000000000900000000 Startu Service Euception Repair Action : The case to be Monitored for now. Resoluion : If the issu happens then a SW reload on Sute PC is requred and then if the issu keeps repeating even after that the Sute PC has to be replaced. The case is to be monitored for now. Resoluion : no futher issus case can be closed for now. Diagnostic performed by Engineer : Discused and confirmed the issu with the cutomer A cutomer advised they were called in for an emergency case yesterday aroud 20000000000000000 you time. On boot u system did not progress to the clinical interface and kept going in loops trying to start u. All the time the review screen 000000000000000000000000001ld show “u-ray system starting” and did not advance any futher. I have verified logs and I can relate to what the cutomer is saying. I am not sue abou the time stamps that we get in log files bu as per the log files time, they started u the system at 1000000006: 00000000 and then the system tried to start u bu kept timing u till 15:29 which is more than 00000000minues util the uer has tried to force shu it down. The susequnt boot u is all normal incluing the boot u this morning. I know abou a known issu where du to an incorrect handling of parallel access to data, a crash may occu while starting acqusition, or when stopping acqusition and the system auomatically recovers in 100000000seconds. This issu is present in Azuion 1.2.3 systems and was solved in the 1.2.5 SW release. This system is already on 1.2.5 SW and also happened at start-u. The logs go in this pattern as below where the startu timeous and then there are a lot of Process Crashes, on the sute PC. The cutomer has ued the system this morning with no issus. If the issu happens then a SW reload on Sute PC is requred and then if the issu keeps repeating even after that the Sute PC has to be replaced. The case is to be monitored for now. Problem Description by Engineer : Problem description by engineer : Did not boot u on the weekend - was flashing after 00000000minues When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: system Shoud comeplete the Boot u Sequnce uer Impact: Cant ue the system Patient Impact: No futher information is available as per the cutomer Curent Software Version: 1.2.5.00000000----------------------------------------------------------------------------- | 2023-10-02 15:31 | No
1541 | Yes | device is not getting any power, therefore not turning on | device is not getting any power, therefore not turning on |  | 2023-10-02 15:36 | No
1542 | Yes | M0A - u00- Charge button Issue (Issue did not reproduce) | u- Charge button Issue (Issue did not reproduce) | 000000000000000000000000001O##: 00000000000000000000000000100000000123739 Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Incident occured on 26/00000000/2000000003 at 7:58, the operating test had not passed (uLLwith the charge buton). Since then, the problem has not occured and coud not be reprodued. The cutomer wants to know if the device can be pu back into service. The problem already occured once in 20000000000000000 Resoluion: Hello u BATAILE, As agreed, can youget back to me with the printed logs of the device? Also, can youcheck if the event was stored in the internal memory of the device and possibly euport it to a PC card? Sincerely, Gullaue PANET Resoluion: Hello u BATAILLE, I am coming back to youregarding the uuu First of all, is the device equpped with the stimuator option (option B00000000)? Becaue if it is the case, it is mandatory to do the operating test with the charge resistance. So the charge buton that needs to be pressed is the one coming from the front of the device. In the technical docuentation, it is possible that it is the mechanics of the charge buton (least likely) or rather the Processor board. The Processor board is no longer in stock becaue of the obsolescence of the device (12/31/2000000002) and we will not be able to repair the device. ufortuately, I fear that this uLLis still coming back and that it can no longer be solved by means of an operating test. Sincerely, Gullaue PANET Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the charge buton of the device. ----------------------------------------------------------------------------- | 2023-10-02 16:16 | No
1543 | Yes | CV*Intermittent problem* Power problem | CV*Intermittent problem* Power problem | 000000000000000000000000001O##: 00000000000000000000000000100000000123578 Khaled Megud Fawzy The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem description by engineer: Problem description by engineer: Power is sudenly drop down How often is this issu occuring? : Intermittent Malfuction area: Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer: power is sudenly drop down - mains panel needs to be checked. Internal Comments: uLL# and/or description of uLL Power is sudenly drop down - mains panel needs to be checked. Repair Action: Power is sudenly drop down - mains panel needs to be checked. Resoluion: Power is sudenly drop down - mains panel needs to be checked. Sucontractor 000000000000000000000000001will be open Follow u Requred: uLL# and/or description of uLL Power is sudenly drop down - mains panel needs to be checked. Repair Action: Power is sudenly drop down - mains panel needs to be checked. What are the neut Trouleshooting steps: Power is sudenly drop down - mains panel needs to be checked. Technician Name: Khaled Megud Fawzy ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000013000000007 ----------------------------------------------------------------------------- | 2023-10-02 16:36 | No
1544 | Yes | Audio-no sound/No audio found at Bench-Foulink s/n u0/other HW-Speaker-HW/Replaced | Audio-no sound/No audio found at Bench-Foulink s/n u00/other HW-Speaker-HW/Replaced |  | 2023-10-02 18:57 | No
1545 | Yes | DFM0 not getting power- Duplicate of PR 00 | DFM0 not getting power- Duplicate of PR 0001 |  | 2023-10-02 19:11 | No
1546 | Yes | AED repeatedly gives "Reinsert the battery for a test" msg with good battery | AED repeatedly gives "Reinsert the battery for a test" msg with good battery | 000000000000000000000000001O##: 00000000000000000000000000100000000123272 FALASCA, IVO The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Make NEMO withou collection at the same time. Tu IREDEEM Spa Piazza dei Martiri 19000000001900000000n.1, 0000000021 Bologna - Tel.0000000010000000035879 Ac. u Domenico Giordano Cell.300000000855700000000Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: It reports to "re-insert the battery for a test" after every insertion of the battery. uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u Domenico Giordano confirms that althouh the battery is good, the defibrillator reports to "re-insert the battery for a test" after every insertion of the battery. Accept the shipment of the spare part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000136000000006 Parts Consued: 1.000000000000000000000000000000005600000000000000001, HS1 (C B) Defib, Italy Italian 00000000, Euch ----------------------------------------------------------------------------- | 2023-10-02 19:41 | No
1547 | Yes | 00Y0System down | 00CDY0System down | 000000000000000000000000001O##: 0000000000000000000000000010000000013700000000 00000000-Oct-2000000003 McDowell Preston The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: patient image scans Eupected and ueupected Resuts: able to record images / coud not record- low data disk avail uer Impact: uknown Patient Impact: uknown SE 8.2.3000000005 Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replace ippc image drives -performed re-create image drive. Verified proper recording / playback of images. ￼ ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000013590000000000000000-Oct-2000000003 McDowell Preston The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : ship to Fuloc: 1639200000000 Preston McDowell Deliver 10000000000000000am on 1000000003-23. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: patient image scans Eupected and ueupected Resuts: able to record images / coud not record- low data disk avail uer Impact: uknown Patient Impact: uknown SE 8.2.3000000005 Diagnostic performed by Engineer : was able to view the reported problem. Was able to Recreate Image Drive and ALso Database consistency. All passed and now system is fuctional, bu du to issu happening previouly, I will ordered the Hard Drives and replace tomorrow. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001357200000000George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Attempt a recreate image store as temporarary soluion util IPPC and or disks can be replaced. Frontal IPPC coud be defective, not powered on or net000000000000000000000000001k cabling incorrectly attached. If needed, PN 0000000080000000050000000000000000- Hard disk spare part. Coud not determine which IPPC installed. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Debriefed caller. They have a patient on the table and uLLmessage "Euposue not possible. Image disk ful". Asked caller to save LOD and a review shows system reporting "Free space low: Delete euaminations." ----------------------------------------------------------------------------- Internal Remarks : Case: 00000000223793700000000 FSE: McDowell Preston Low image disk uLL FSE 000000000000000000000000001ld like to know if IPPC HDDs need to be replaced. Instruted FSE to perform Database Consistency Test(Passed) and Re-create Image Store. System is now u. ----------------------------------------------------------------------------- | 2023-10-02 19:46 | No
1548 | Yes | Monitor not turning on | Monitor not turning on | 000000000000000000000000001O##: 00000000000000000000000000100000000120000000067 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to Bench Diagnostic performed by Engineer : Problem Description: "will not power u to a uer state. Light on power buton flashes green and screen jut comes on completely black. Also uable to enter service mode. Monitor completely uuable." Cutomer callback Tempu Pro: "will not power u to a uer state. Light on power buton flashes green and screen jut comes on completely black. Also uable to enter service mode. Monitor completely uuable." Form ued: Bench Trouleshooting Steps Taken: EMS site manager reports Tempu Pro: "will not power u to a uer state. Light on power buton flashes green and screen jut comes on completely black. Also uable to enter service mode. Monitor completely uuable." Created Bench 000000000000000000000000001 and ##d bench repair form. Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS site manager Device ue:Not in Clinical ue Eupected and ueupected Resuts: "will not power u to a uer state. Light on power buton flashes green and screen jut comes on completely black. Also uable to enter service mode. Monitor completely uuable." uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: "will not power u to a uer state. Light on power buton flashes green and screen jut comes on completely black. Also uable to enter service mode. Monitor completely uuable." Follow u Requred : EMS site manager reports Tempu Pro: "will not power u to a uer state. Light on power buton flashes green and screen jut comes on completely black. Also uable to enter service mode. Monitor completely uuable." ----------------------------------------------------------------------------- | 2023-10-02 19:46 | No
1549 | Yes | Cuomer advises the charger pin has sheared off in the port - device will not charge. | Cuomer advises the charger pin has sheared off in the port - device will not charge. | 000000000000000000000000001O##: 00000000000000000000000000100000000136716 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Follow u Requred : 0000000056000000005000000001-Rear Cover Assembly 0000000056000000005000000001-Front Case Assembly 00000000ous for repair Diagnostic performed by Engineer : Confirmed complaint there is physical damage to the charging port, additionally the Rear cover needs to be replaced du to the kickstand being physically damaged. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000120000000098 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Sent to bench Diagnostic performed by Engineer : Problem Description: "Cutomer advises the charger pin has sheared off in the port - device will not charge." Cutomer callback Tempu Pro: "Cutomer advises the charger pin has sheared off in the port - device will not charge." Form ued: Bench Trouleshooting Steps Taken: EMS site manager called to report Tempu Pro "Cutomer advises the charger pin has sheared off in the port - device will not charge." Cutomer requst to send device to bench for repair. Create Bench 000000000000000000000000001 ##d bench form and requsted shipping label for cutomer. Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS site manager Device ue:Not in Clinical ue Eupected and ueupected Resuts: "Cutomer advises the charger pin has sheared off in the port - device will not charge." uer Impact:No impact Patient Impact:No impact SEuuProblem Description/ueupected Resuts: "Cutomer advises the charger pin has sheared off in the port - device will not charge." Follow u Requred : EMS site manager called to report Tempu Pro "Cutomer advises the charger pin has sheared off in the port - device will not charge." Cutomer requst to send device to bench for repair. ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-02 21:21 | No
1550 | Yes | Blower temperature high. (EC 0) | Blower temperature high. (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000136600000000Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : 10000000000000000/2000000003 Biomed has reported testing the uit for 200000000ous after replacing the air inlet filter. The Cutomer is uable to reprodue the reported issu with the blower temperatue uLL Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes uer Impact: V600000000has been removed from service Patient Impact: Biomed is not reporting any adverse oucome to patient care or therapy SE uu Problem Description/ueupected Resuts: V600000000displayed blower temperatue high uLLDiagnostic performed by Engineer : Problem Description: V600000000displayed blower temperatue high uLL Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000000000000/2000000003 Trouleshooting Steps Taken: Advised Biomed to confirm the diagnostic code in the significant event log. Diagnostic code 1122 is the blower temperatue high diagnostic code. The blower temperatue was greater than 95ºC for longer than 600000000seconds. Advised Biomed to change the air inlet filter and test the V600000000for 200000000ous and attempt to reprodue the uLL 1. Verify that the air inlet filter is not dirty or blocked. 2. Verify that the cooling fan is operational. 3. Replace the blower. 00000000Replace the MC PCBA. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Letha Hargraves IP & Issu: 80000000033200000000Vent blower reach too hot uLLmessage. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and qute. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-02 21:41 | No
1551 | Yes | Audio-No Sound / Speaker Malf-No Tone-Keyrin s/n u0 / No Trouble-Working to Spec-Other / Replaced Parts | Audio-No Sound / Speaker Malf-No Tone-Keyrin s/n u00 / No Trouble-Working to Spec-Other / Replaced Parts | 000000000000000000000000001O##: 00000000000000000000000000100000000282100000000 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000137000000003 Roper, Ken The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Problem Description: speaker malf no tone speaker malf no tone Form ued: Bench Trouleshooting Steps Taken: speaker malf no tone Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: speaker malf no tone Follow u Requred : Problem Description: speaker malf no tone speaker malf no tone Form ued: Bench Trouleshooting Steps Taken: speaker malf no tone ----------------------------------------------------------------------------- | 2023-10-02 21:46 | No
1552 | Yes | Prouimital pressure sensor auto zero failed alert. | Prouimital pressure sensor auto zero failed alert. | 000000000000000000000000001O##: 0000000000000000000000000010000000013700000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer stated that 1 of the pins was bent from the solenoids, and he straightened it ou. After installing the DA pcba properly, he stated it ran for a hou, and no issus. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: check vent prouimal pressue auozero failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it has a Prouimal pressue auozero failed. He verified there is a code 1100000000 in the history. He stated that they jut replaced the flow sensor assy. Instruted him to check the pin alignment between the solenoids and the DA pcba. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: John Correira - phone 600000000 65000000007153 IP & Issu: 570000000000000000 - Prouimital pressue sensor auo zero failed alert. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then ufer to remote tech quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-02 23:31 | No
1553 | Yes | 00uuu System is not turning on | 00CFuuu System is not turning on | 000000000000000000000000001O##: 0000000000000000000000000010000000015952000000005-Oct-2000000003 Midgett Pau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System is not tuning on When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System is not tuning on uer Impact: System is not tuning on Patient Impact: No impact to patient SE 7.2.16.00000000Tools & Equpment : DOSE-KVP-LIGHT METER,20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,20000000000000000000000000000000000000000:00000000:00000000 Resoluion : Trouleshot MPDuby removing load and testing. No failue. Reconnected each power connection one by one with no failue. Performed generator calibration and also detector calibration. Tested system repeatedly with no failues. MPDucontrol board is believed to have fiued the issu which was replaced previouly . ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000159267 Clark, Byron The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System is not tuning on When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System is not tuning on uer Impact: System is not tuning on Patient Impact: No impact to patient SE 7.2.16.00000000Diagnostic performed by Engineer : biomed replaced part MPDucontrol uit 000000002300000000160000000000000000. issu is still occuring with fue F100000000blowing. Contact to discus with his facility and determine neut step, if onsite needs to be created... FSE opened a case 0000000022393900000000with LS - 1200000000 / AT - 2227300000000Internal Comments : Alan has requsted on site 000000000000000000000000001to be created. The MPDucontrol did not resolve the sitution and F100000000is still blowing the part for the entire MPDu as requsted Material Nuber 00000000800000000868300000000- MPDureplacement assembly Follow u Requred : ongoing issus with system not powering on MPDucontrol has been replaced and same issu with F100000000blowing has occured. part for entire MPDu Material Nuber 00000000800000000868300000000- MPDureplacement assembly... On site for Pau Midget requsted ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000013700000000 Sean Kiernan The system meets the specification for the performed service and is retuned to ue Resoluion : Recommended replacement of MPDucontrol board which appears to be defective. Internal Comments : uLL# and/or description of uLL: System is not tuning on Repair Action : MPDucontrol uit 000000002300000000160000000000000000 spare fues 00000000230000000000000000952 Diagnostic performed by Engineer : Biomed called and reported that when they try to tun the system on nothing happens. Had Biomed check for L1, L2 and L3 phase indicator lights on the MPDuand foud that all three lights were lit so incoming suply power for the MPDuwas present. Had uer then try to tun system on from the M cabinet power on switch and it still did not tun on. Instruted uer to check fue F100000000in the M cabinet and foud that fue F100000000was blown, uer replaced fue F100000000and the fue blew again. Recommended replacement of MPDucontrol board which appears to be defective. Problem Description by Engineer : Problem description by engineer : System is not tuning on When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : System Control Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System is not tuning on uer Impact: System is not tuning on Patient Impact: No impact to patient SE 7.2.16.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000152379 Parts Consued : 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Alan Billal IP & Issu: 5000000006 Need to speak with the parts Dpt Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: ufr call to parts dept per caller req Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : 000000000000000000000000001ked with FSE Pau to determine why the F100000000fue failed on the MPDu We ued Du376800000000Section 00000000o determine power rouing and disconnected all loads on the F100000000(1A) fue. The fue was then replaced and the system powered on withou blowing the fue. We added each load one at a time and the fue did not fail. It was determined the fue was never replaced before arriving FSE on site. There was a video issu for data monitor after bu power was getting to the monitor. The FSE was to look into it and call back if needed. The FSE also noticed it was DVI monitor and the inpu needed setting on the monitor. Internal Remarks : call back PO 150000000092 for $000000006.00000000 Internal Remarks : Interaction Type: Cutomer Call Back Caller: Alan Billal IP & Issu: " 5000000006 cutomer called to get statu udate on the case " Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Gave cutomer the manager information Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-02 23:51 | No
1554 | Yes | unit is not charging the battery | unit is not charging the battery | 000000000000000000000000001O##: 00000000000000000000000000100000000136631 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit is not charging the battery uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it is not charging the battery. She verified that the power suply is operating fine. 1200000000AC volts are going into the power suply and 200000000C volts are going into the PM pcba. She also euplained that the circuation fan is tuning on for 00000000econds than tuns off for 00000000econds, and repeats again and again. I had her check the battery terminal on both ends for cable. She noticed 1 terminal on the power harness cable was puhed in. She was able to pul it back ou and connected the battery. It seems to be 000000000000000000000000001king fine and the battery is charging properly. Recommended replacing the power harness cable. Part #00000000561511621 cable, power harness. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Sarah Red Issu: 56627862 Part ID Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-03 00:36 | No
1555 | Yes | HS0ED is emitting the triple chirp | HS0AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000137856 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 13 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A00000000A-0000000000000000 Battery Install Before# 12/31/2000000008 Lot#00000000900000000 Pads Eup Date: 2/12/2000000005 Pads Lot # Y000000001222-00000000----------------------------------------------------------------------------- | 2023-10-03 00:46 | No
1556 | Yes | u00quality concerns over u-Ray PC Biplane and Suite PC parts | u0B0quality concerns over u-Ray PC Biplane and Suite PC parts |  | 2023-10-03 00:47 | No
1557 | Yes | u ray unuble | u ray unuble | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000000000000 00000000Oct-2000000003 Hobbs Matthew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: charge nuse Device ue: ep lab Eupected and ueupected Resuts: image shows only digital noise uer Impact: room uuable Patient Impact: delays SE8.1.3 Resoluion : replaced detector and calibrated to system Parts Consued : 1.0000000000000000000000009896000000000000000000000000,Detector Pu0000000000000000High Speed pack Tools & Equpment : DOSE-KVP-LIGHT METER,20000000000000000000000000000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,20000000000000000000000000000000000000000:00000000:00000000 TORQu WRENCH,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000136351 00000000Oct-2000000003 Oduuwe Chijioke The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uray tech Device ue: Diagnostic Eupected and ueupected Resuts: Vertical and Horizontal lines on images uer Impact: system down Patient Impact: patient reschedued SE 8.1 Diagnostic performed by Engineer : Checked and Adjuted FD power suply to 25V and also replaced Flashlite FDC bu issus were not resolves. Flat Detector replacement recommended and ordered ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000135966 00000000-Oct-2000000003 Hobbs Matthew The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: charge nuse Device ue: ep Eupected and ueupected Resuts: improper image on screen uer Impact: system uuable Patient Impact: inconvenience SE 8.1.3 Diagnostic performed by Engineer : ran diagnostics on image detector. frontend test failed right away. ----------------------------------------------------------------------------- Internal Remarks : Matt Hobbs - qustion abou FDC part nr. Internal Remarks : approval for detector replacement fse ran diags on FDC controller all passes and image is scrambled , approved detector replacement Internal Remarks : Spoke to FSE Matthew Hobbs, he said he already 000000000000000000000000001ked on this 2 hous ago, assigned to him. ----------------------------------------------------------------------------- | 2023-10-03 01:31 | No
1558 | Yes | 0 CAN HY000My toothbru caught on fire and melted while it was on charge / SPLIT | 00CAN HY00/0_My toothbru caught on fire and melted while it was on charge / SPLIT |  | 2023-10-03 02:48 | No
1559 | Yes | Backup alarm failed (EC 00 | Backup alarm failed (EC 00) |  | 2023-10-03 03:05 | No
1560 | Yes | usystem will not power up | usystem will not power up | 000000000000000000000000001O##: 00000000000000000000000000100000000139000000008 00000000-Oct-2000000003 Luwig, Gabrielle The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: na Device ue: na Eupected and ueupected Resuts: na uer Impact: none Patient Impact: none SE Resoluion : System 000000000000000000000000001ld not power on. Trouleshoot to defective/damaged fan tray and DCPS. Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000138976 00000000-Oct-2000000003 Luwig, Gabrielle The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: na Device ue: na Eupected and ueupected Resuts: na uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : System down, part has been ordered and follow u 000000000000000000000000001O#created Follow u Requred : Replace PDuFan Tray and DCPS ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022379961 FSE: Luwig, Gabrielle System has no power; Hospital power pass. No LEDs tuning on, on PDu MIM mains coming in and ou of Mains Disconnect Switch; all fues pass. FSE check NY1 (from uS) L1, L2, L3; 2300000000+ measued on all phases, Pass. FSE inspected Fan Power Tray; no visible damage. Instruted FSE to place part order for 0000000080000000000000000900000000 - PDuFanTray and DCPS. Internal Remarks : Planner Initials: GM Action taken: Follow u activity for FSE Gabrielle Luwig for 100000000300:300000000M Oucome:dispatched Internal Remarks : Interaction Type: Internal Caller: Carey Gruzie IP & Issu: Case 0000000022379961 / usystem will not power u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Created qute Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-03 05:01 | No
1561 | Yes | 00 CN00- BNM 00VIII/PCCI/NJ/GI/ the front panel button failure | u- BNM 000VIII/PC0NJ/GI/ the front panel button failure | 000000000000000000000000001O##: 00000000000000000000000000100000000122171 Rahmat Kasyfu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Confirmed reported problem the Operational Check was failed Cutomer Fuction/Role: Biomed engineer Device ue: no Eupected and ueupected Resuts: n/a uer Impact: no Patient Impact: no SE Rev D Resoluion : Need Replace Processor PCBA and u PCA, the device is retuned to fuly fuctionality ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001000000008 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Therapy 1.0000000000000000000000000000000056000000007200000000,DFM1 ASSY u Modue 1.1 ----------------------------------------------------------------------------- | 2023-10-03 05:06 | No
1562 | Yes | THERE are places that don't work on THE CLIENT uNG THE CLIENT's WORTH. | THERE are places that don't work on THE CLIENT uNG THE CLIENT's WORTH. |  | 2023-10-03 06:16 | No
1563 | Yes | Speaker is damaged | Speaker is damaged | 000000000000000000000000001O##: 0000000000000000000000000010000000012100000000 Alhad Apte The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Device ue: Ou of ue Eupected and ueupected Resuts: speaker is damaged uer Impact: No impact on uer Patient Impact: no impact on patient SE Not provided Diagnostic performed by Engineer : Speaker assembly is damaged. Informed cutomer that spare parts are not available separately for repair. Repair to be done by euchange. Qute is sent for the same. ----------------------------------------------------------------------------- | 2023-10-03 06:41 | No
1564 | Yes | does not have enough flo | does not have enough flo | 000000000000000000000000001O##: 000000000000000000000000001000000001399500000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer has a V600000000and uit does not have enouh flow. Respiratory was trying to trouleshoot bu was limited to knowledge of technical skills. Recommended to take uit off patient, and wait for biomed to come in to perform the PVT. The qute was sent to the cutomer and I remindered them via ## Problem Description by Engineer : Cutomer Fuction/Role: respiratory Device ue: in ue Eupected and ueupected Resuts: not enouh flow uer Impact: none Patient Impact: un SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Andrea IP & Issu: 5700000000729 does not have enouh flo Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason For Entitlement Claim N/A Action Taken: created case, sq's, remote technical w/o, informed cutomer that investigation needs to be started du to a patient being harmed while on machine, transferred to on call v600000000RSE RObert Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-03 07:06 | Yes
1565 | Yes | ProAct: Suite PC could not open or lost connection with Certeray generator - Frontal | ProAu Suite PC could not open or lost conneuon with Certeray generator - Frontal | 000000000000000000000000001O##: 0000000000000000000000000010000000012100000000 13-Dec-2000000003 Yuhianto, Marthin The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: device in ue Eupected and ueupected Resuts: - uer Impact: - Patient Impact: - SE 2.2.6 Resoluion : Action take, Follow the soluion step remote service all in good condition and commuication with uer no more issu as long they ue this device Device is fuctionally ----------------------------------------------------------------------------- | 2023-10-03 07:16 | No
1566 | Yes | 0_0Surgery_FD0_System can’t boot up. | 0K_00Surgery_FD00System can’t boot up. | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 00000000K_000000003_Sugery_FD2000000005_System can’t boot u. Safety Qustions : Case: 0000000022366166 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes : This PO is a Reference Nuber only SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-03 08:41 | No
1567 | Yes | u00quality concerns over u-Ray PC Biplane and Suite PC parts | u0B0quality concerns over u-Ray PC Biplane and Suite PC parts |  | 2023-10-03 09:30 | No
1568 | Yes | 00 CN00 Therapy delivery test failed | u Therapy delivery test failed | 000000000000000000000000001O##: 00000000000000000000000000100000000122100000000Nandha, Rahu The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : therapy cap to be replaced Problem Description by Engineer : Cutomer Fuction/Role: Bme Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : therapy cap to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000037 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-03 09:51 | No
1569 | Yes | The Smart Select knob did not respond.. | The Smart Seleuknob did not respond.. |  | 2023-10-03 10:41 | No
1570 | Yes | M0A - u00 Failed test, issue with paddles | u Failed test, issue with paddles |  | 2023-10-03 10:51 | No
1571 | Yes | 00- CN00 - Therapy test failed | 000 - CN000 - Therapy test failed | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000 00000000-Oct-2000000003 Neha, Deshmuh The system meets the specification for the performed service and is retuned to ue Resoluion : Foud dut on the paddles, thorouhly cleaned the paddles , performed operational check , system 000000000000000000000000001king fine. Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Foud uer complaining the therapy test getting failed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001350000000000000000andha, Rahu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Fse visit requred. Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: no Eupected and ueupected Resuts:na uer Impact: no Patient Impact: no SEna Diagnostic performed by Engineer : Fse visit requred. ----------------------------------------------------------------------------- | 2023-10-03 11:01 | No
1572 | Yes | M0A / u00/ Display was not clear. | M00A / u000/ Display was not clear. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_118_HDR_M000000005A_Display was not clear. SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 0000000022366762 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-03 12:01 | No
1573 | Yes | Display was not clear. | Display was not clear. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 0000000022366769 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes : This PO is a Reference Nuber only SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : 00000000R_118_HDR_M000000005A_Display was not clear. ----------------------------------------------------------------------------- | 2023-10-03 12:01 | No
1574 | Yes | Failed Testing | Failed Testing |  | 2023-10-03 12:51 | No
1575 | Yes | Touch position misalignment during calibration | Touch position misalignment during calibration | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000027 October 16, 2000000003 MERLI, Enzo The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Reguatory Qustions: No/No Consued parts: 1.00000000000000000000000000000000561539000000001, RP-Touh Screen, V600000000V600000000Lu/V6800000000Resoluion: Screen replaced and calibrated. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000122800000000Assad MAJEED The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: On-site intervention requred. Follow u Requred: Service on site for replacement of the touh panel. Please look for resouces. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that duing the calibration of the V600000000touh screen, the calibration can only be performed and the V600000000can only be tuned off by removing the battery and the power cable. Duing a calibration of the touh screen, the V600000000can only be tuned off at the end of the calibration. So it seems that the touh panel is defective and that the calibration does not detect the suport of the first target (offset too significant) and that afterwards, the V600000000can only be tuned off by removing the battery and the power cable. The cutomer therefore wants on-site service to replace the touh panel. ----------------------------------------------------------------------------- | 2023-10-03 13:21 | No
1576 | Yes | Anode Problems | Anode Problems | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000096 Oct 000000002000000003 Anesthesia Tusday Nardo The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Boss TRM Device ue: u-ray tue Eupected and ueupected Resuts: Anode problems uer Impact: Problems with tue Patient Impact: No SE 8.1.u Resoluion: Sostituto tuo radiogeno. Eseguta calibration geometrica, calibrazioni u-Ray e check qulità d'immagine. Sistema OK. Controllate anche alimentazioni sugenerators 00000000VAC L1,2,3 -N e fase -fase 2300000000AC. OK Parts Consued: 1.000000000000000000000000989000000000000000085100000000u 200000000 0000000000000000 ROT-GS 1000000003 Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000006-28 00000000:00000000:00000000 MuTIMETER, 2000000003-11-00000000 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 200000000000000000000000031 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000007-10000000031 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000156355 Oct 000000002000000003 Bernni Nicola The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: BioMed Device ue: Yes Eupected and ueupected Resuts: Euchange of tue and OK uer Impact: none Patient Impact: none SE 8.1 Resoluion: Suporto duante la sostituione del Tuo Ru ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000 Oct 000000002000000003 Anesthesia Tusday Nardo The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: TRM Device ue: anode problems Eupected and ueupected Resuts: anode rotation problems uer Impact: Bad qulity images Patient Impact: No SE 8.1.u Resoluion: Problem with roto anodo. Eseguta diagnosi, ordinato nuvo tuo via taui. Inizio mout tuo esistente. Tools & Equpment: MuTIMETER, 2000000003-11-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000 Patrick Ruegg The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: Tue Eupected and ueupected Resuts: Anode uLL uer Impact: restart Patient Impact: no SE na Diagnostic performed by Engineer: Remote suport of the local FSE. uLLfoud on rotor or RoCo. Material will be ordered. ----------------------------------------------------------------------------- Internal Remarks: We have raised a taui to Switzerland with DN# 862100000000000000001 and DN# 862100000000816. The ticket nuber is 2922593. We will keep youposted with futher udate. ----------------------------------------------------------------------------- | 2023-10-03 14:06 | No
1577 | Yes | Failing CO0alibration | Failing CO0calibration | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000021 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : The following part will be issud to the end uer to install: 1u 0000000056000000001711 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Service Suport Coordinator Device ue: not in ue Eupected and ueupected Resuts: Failing co2 calibration duing PVT uer Impact: n/a Patient Impact: n.a SE n/a Resoluion : The following part will be issud to the end uer to install: 1u 0000000056000000001711 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000069 Parts Consued : 1.0000000000000000000000000000000056000000001711,Intrepid CO2 Modue ASSY ----------------------------------------------------------------------------- | 2023-10-03 14:21 | No
1578 | Yes | We have a faulty upro monitor, malfunctioning screen (S/N 00. | We have a faulty tempupro monitor, malfunuoning screen (S/N 00. | 000000000000000000000000001O##: 000000000000000000000000001000000001000000005 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: In ue Eupected and ueupected Resuts: malfuctioning screen uer Impact: n/a Patient Impact: Swapped to known 000000000000000000000000001king device SE n/a Resoluion : Damaged screen protector cauing issus with the touh screen/replaced. Tested OK. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000135173 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u 0000000056000000005000000001 1u 0000000056000000005000000001 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: In ue Eupected and ueupected Resuts: malfuctioning screen uer Impact: n/a Patient Impact: Swapped to known 000000000000000000000000001king device SE n/a Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u 0000000056000000005000000001 1u 0000000056000000005000000001 ----------------------------------------------------------------------------- Internal Remarks : From: MOODY, Gary (LONDON AMBuANCE SERVICE NHS TRuT) <gary.moody100hs.net> Sent: Monday, October 9, 2000000003 11:00000000AM To: PH.PMu.Suport (Fuctional Accout) <PH.PMu.Suportucom> Suject: FW: u Health Systems Bench Repair SN 6000000008700000000Case 0000000022379000000003 Cauion: This e-mail originated from ouside of u, be carefu for phishing. Good morning, In response to Tracey’s ou of office, can youthis device be schedued for collection on Monday 16th October? Many thanks Gary Gary Moody BSc(Hons) Clinical Suport Manager - Motorcycle Response uit GMB u 000000000000000000000000001kplace Organiser ¦ ## uGMBrepresentative00ond-amb.nhs.u ¦ London Ambuance service ¦ Motorcycle Response uit ¦ Ilford Ambuance Station ¦ Aldborouh Road Souh ¦ London IG3 8HQ T: 000000000000000078328900000000nternal: 000000009800000000 Personal Mobile: 0000000080000000057700000000 ¦ ## gary.moody00ond-amb.nhs.u Internal Remarks : Hi Gary , Please can youconfirm a sutable collection date and collection address for SN 6000000008700000000 Tempu Pro, Printer Internal Remarks : Tried calling cutomer Internal Remarks : Hi Gary , Please can youconfirm a sutable collection date and collection address for SN 6000000008700000000 Tempu Pro, Printer Internal Remarks : Tried calling cutomer no response / ##sent yesterday /aware may be on shifts Internal Remarks : tried contacting cutomer went to voic##Internal Remarks : TRIED CALLING CuTOMER WENT STRAIGHT TO VOIC##AGAIN & RESENT ##Internal Remarks : COLLECT MONDAY 16.10000000023 Tracking Nuber:1Z3R6000000006695259829 ----------------------------------------------------------------------------- | 2023-10-03 14:56 | No
1579 | Yes | uPro ¦ device detected an uLLrequiring shut down and restart after ung the video laryngoscopy (Trizeps-0 | TempuPro ¦ device deteud an uLLrequiring shut down and restart after ung the video laryngoscopy (Trizeps-0 | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000099 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Device to be collected from end uer and sent to u Bench do that logs can be collected for Investigation Follow u Requred : Device to be collected from end uer and sent to u Bench do that logs can be collected for Investigation Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: Device notified the uer that the device detected an uLLand needed to be shu down and restarted uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : Issu escalated to RDT Buas possibly related to issu raised on FSN-2000000003-CC-EC-000000005 ----------------------------------------------------------------------------- Internal Remarks : Hi Desmond , You device is showing as not collected on Monday ,please let u know a sutable date for a recollection Internal Remarks : Escalation raised and sent to RDT Tech team - issu possibly related to FSN-2000000003-CC-EC-000000005 Internal Remarks : Hi Sam , Cutomer has requsted we close the case as device still with end uer Thank youInternal Remarks : Hi Rose, Can youplease close this case as the device is still with the uer. Regards, Desmond Internal Remarks : End uer has advised that the Tempu Pro is still ou with the overall end uer and is yet to be retuned to them for it to be collected and shipped. I will follow u with Desmond in a few days to get clarification on progress of Pro being collected. Internal Remarks : Hi Desmond , You device is showing as not collected on Monday ,please let u know a sutable date for a recollection Internal Remarks : Tracking Nuber:1Z3R600000000669900000000319 COLLECT MONDAY COLLECTION & RETuNS ADDRESS London Ambuance Service NHS Trut ¦ uit 7 Consu Avenu ¦ Rainham ¦ RM13 8HY Internal Remarks : Hi Desmond , I have been asked to collect you Tempu Pro, Printer, IBP SN 60000000011 Is this the collection address below London Ambuance Service NHS Trut ¦ uit 7 Consu Avenu ¦ Rainham ¦ RM13 8HY ----------------------------------------------------------------------------- | 2023-10-03 15:16 | No
1580 | Yes | DFM0 low battery (Replaced DFM0 Lithium Ion Battery) | DFM0 low battery (Replaced DFM0 Lithium Ion Battery) | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000082 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Make NEMO withou collection at the same time. Tu Polygon S.p.A. C/O ASST Franciacorta 00000000azzini St. - 25000000002 Chiari (BS) Administration Bulding Ac. En. Anna Morotti Tel.: 0000000000000000710000000000000000 Cell.: 366 6300000000100000000Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: The fuctional test shows that the battery is flat and does not recharge completely. uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Low battery for which replacement is asked. I proceed with the order of the requsted part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000089 Parts Consued: 1.0000000000000000000000000000000098000000000000000000000000DFM100000000 Lithiu Ion Battery ----------------------------------------------------------------------------- File uloaded by CS Portal uer: File name: 20000000031000000003_000000000000000056.jpg. 20000000031000000003_000000000000000039.jpg. Comments: Docuents loaded duing the opening of the requst. ----------------------------------------------------------------------------- | 2023-10-03 15:31 | No
1581 | Yes | uuucopy failed error | uuucopy failed error | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000077 00000000Oct-2000000003 Singh, Dule The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : In logs Converter 1 Short circut is coming Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : Replaced converter Q2, now flouo & cine is fuctional,observation done in few cases. Parts Consued : 1.00000000000000000000000000000000210000000016291,Converter 8E Velara ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: In logs Converter 1 Short circut is coming Repair Action : In logs Converter 1 Short circut is coming Follow u Requred : uLL# and/or description of uLL: In logs Converter 1 Short circut is coming Repair Action : In logs Converter 1 Short circut is coming What are the neut Trouleshooting steps : Need to replace below part 00000000210000000016291 Converter 8E Velara Technician Name : Kuar, Atu Problem Description by Engineer : Problem description by engineer : In logs Converter 1 Short circut is coming Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : In logs Converter 1 Short circut is coming ----------------------------------------------------------------------------- | 2023-10-03 15:46 | No
1582 | Yes | Escalation 000 Quench pipe broke off during quench and caud helium to dissipate in the room (root cau- uuu | Escalation 00000 Quench pipe broke off during quench and caud helium to dissipate in the room (root cau- uuu |  | 2023-10-03 16:07 | No
1583 | Yes | System went out with patient on table, table was floating. Tried restart but no connection to hard disk | System went out with patient on table, table was floating. Tried restart but no conneuon to hard disk | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000050000000000000000Oct-2000000003 Berrig Clau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Cutomer Device ue: y Eupected and ueupected Resuts: See Problem Description uer Impact: Yes, partly Patient Impact: No SE NA Resoluion : Host pc fudet defekt og er blevet skiftet. Parts Consued : 1.00000000000000000000000000000000800000000200000000611,Nehalem Host PC Monoplane Rev_III NO_HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000NAESS BJORNAR The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Dear CCC Please plan a FSE at site to do fautfinding on system not starting u, most likely du to break down on HOST PC and/or disk. HOST PC and disk are ordered as spareparts. Plan 00000000at site + travel Covered by contract Parts: 00000000800000000200000000612 Nehalem Host PC Monoplane Rev_III NO_HDD 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II Problem Description by Engineer : Cutomer Fuction/Role: Cutomer Device ue: y Eupected and ueupected Resuts: See Problem Description uer Impact: Yes, partly Patient Impact: No SE NA Diagnostic performed by Engineer : Contacted cutomer and we coud sort ou that it seems like a total breakdown on HOST PC and/or disk/sw. We need to plan and onsite for FSE, and we order a new HOST PC and OS disk 00000000800000000200000000612 Nehalem Host PC Monoplane Rev_III 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II Case will be followed u by Service Center ----------------------------------------------------------------------------- | 2023-10-03 16:41 | No
1584 | Yes | M0A / u00 Functional test failed. | M00A / u000/ Funuonal test failed. | 000000000000000000000000001O##: 0000000000000000000000000010000000015760000000026-Oct-2000000003 Laezza Francesco The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: co-rd Device ue: def. Eupected and ueupected Resuts: NBP calibration eupired uer Impact: no Patient Impact: no SE no Resoluion: Performed NBP calibration with a positive oucome. Tests and tests were OK. ----------------------------------------------------------------------------- | 2023-10-03 17:21 | No
1585 | Yes | This record has been captured in SAP PRD RA 00for SRC. This record will be closed. | This record has been captured in SAP PRD RA 0000for SRC. This record will be closed. |  | 2023-10-03 17:32 | No
1586 | Yes | Getting hard disk full error | Getting hard disk full error | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000073 00000000-Oct-2000000003 Chauhary, Deepak The system meets the specification for the performed service and is retuned to ue Tools & Equpment : MuTIMETER,200000000000000006-27 00000000:00000000:00000000 Diagnostic performed by Engineer : Checked and foud image disk fauty . Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replaced the hard disks also reloaded the software of IPPC, after that problem has been resolved. Handed over the system in 000000000000000000000000001king condition. Problem Description by Engineer : Problem description by engineer : Getting hard disk ful uLL How often is issu occuring? : Mutiple times per month Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000073 Kuar, Atu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Getting hard disk ful uLL How often is issu occuring? : Mutiple times per month Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Image Disk ful uLLis coming very frequntly Internal Comments : uLL# and/or description of uLL: RMT - IPu Euposue not possible. Image disk ful Repair Action : FSE to visit and replace Image disk and recreate image store Follow u Requred : uLL# and/or description of uLL: RMT - IPu Euposue not possible. Image disk ful Repair Action : FSE to visit and replace Image disk and recreate image store What are the neut Trouleshooting steps : FSE to visit and replace Image disk and recreate image store Technician Name : Kuar, Atu ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-03 17:36 | No
1587 | Yes | Alarm issues | Alarm issues | 000000000000000000000000001O##: 00000000000000000000000000100000000135100000000ABDELBARI MOHAMMED The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: In ue. Eupected and ueupected Resuts: Issus with CareEvent server. uer Impact: Issus with CareEvent server. Patient Impact: None. SE C.00000000 Diagnostic performed by Engineer: Biomed informed that they had lost connection to the CareEvent server. Tuns ou that the issu occured du to a fauty rouer in the net000000000000000000000000001k infrastrutue. Once the IT department had corrected the issu on the rouer, everything was 000000000000000000000000001king as intended again and we agreed to close the case. ----------------------------------------------------------------------------- | 2023-10-03 17:46 | No
1588 | Yes | GDU error | GDU error | 000000000000000000000000001O##: 000000000000000000000000001000000001000000007 00000000Oct-2000000003. Simon Lua The system meet the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: n/a. Device ue: n/a. Eupected and ueupected Resuts: n/a uer Impact: n/a. Patient Impact: n/a. SE n/a. Resoluion: Replacement of the ND, fuctionality was tested. The device is free from defects and it is operational. All tests OK. Parts Consued: 1.00000000000000000000000000000000800000000289731, Fast Ethernet Switch 16-Port. 1.0000000000000000000000000000000080000000050000000000000000, Host PC Monoplane Revised_II no HDD. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000136000000003 Martin Horace The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: n/a. Device ue: n/a. Eupected and ueupected Resuts: n/a uer Impact: n/a. Patient Impact: n/a. SE n/a. Diagnostic performed by Engineer: replace part. Follow u Requred: replace part. ----------------------------------------------------------------------------- | 2023-10-03 18:06 | No
1589 | Yes | Battery failed. (EC 0) | Battery failed. (EC 00) |  | 2023-10-03 18:39 | No
1590 | Yes | post flow sensor cal error | post flow sensor cal error | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000057 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: flow calibration uLL uer Impact: none Patient Impact: none SE 2.300000000Resoluion : Cutomer will order the flow sensor assy and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and it is alarming post flow sensor cal uLL He stated it has software 2.300000000 and average air and o2 is reading 0000000000000000 Once he tun the flow to 100000000 the average air go u the 200000000slpm. The uLLcodes in the history was a 1100000000 and 1100000000. Recommended replacing the flow sensor assy Part #989800000000617100000000 flow sensor assy ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: mike IP & Issu: 7258000000003 post flow sensor cal uLL Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-10-03 18:51 | No
1591 | Yes | Constant tone and not seeing any uLLmessages | Constant tone and not seeing any uLLmessages | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000072 00000000-Oct-2000000003 Reu, Timothy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Constant tone and not sending any uLLmessages. uer Impact: system restricted Patient Impact: none reported SE 2.100000000Diagnostic performed by Engineer : Inspected device and coud not dulicate the reported problem. Verified/reseated internal connections. Verified proper resistance of speakers. Electrical safety and alarms tests passed, device is ready for patient ue. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 MuTIMETER,200000000000000001-31 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000069 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Constant tone and not sending any uLLmessages. uer Impact: system restricted Patient Impact: none reported SE uknown Follow u Requred : Cutomer states there is a constant auible tone from device. device is not showing any uLLcodes in service mode. Device will make tone in service mode and in normal operational mode. Cutomer is requsting on site service for FSE to evaluted uit and determine what needs to be repair and fiu uit. Resoluion : Creating on site service. Diagnostic performed by Engineer : Cutomer states there is a constant auible tone from device. device is not showing any uLLcodes in service mode. Device will make tone in service mode and in normal operational mode. Cutomer is requsting on site service for FSE to evaluted uit and determine what needs to be repair and fiu uit. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Casey Roberson IP & Issu: 59785500000000 Constant tone and not sending any uLLmessages. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-03 19:16 | No
1592 | Yes | The room does not turn on | The room does not turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000189676 October 100000000 2000000003 Jean-lous MAHIER The system meets the specifications for the performed service and is retuned to ue. Resoluion: Reguarization of spare part consuption. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Problem restarting the system following a power ouage duing the grou test. uer Impact: system not available. Patient Impact: no SE u Diagnostic performed by Engineer: Reguarization of spare part consuption. Consued parts: 1.000000000000000000000000000000008000000000000000000000000, PD. ASSY. PDM Tools and Equpment: MuTIMETER, 200000000000000002-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001500000000100000000October 3, 2000000003 Jean-lous MAHIER The system meets the specifications for the performed service and is retuned to ue. Resoluion: Replacement of the power suply modues of the main cabinet and the imaging cabinet, the uS and sabattry. Loaded the software of the replaced modues. Fuctionality tests. The device is fuly operational again. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Problem restarting the system following a power ouage duing the grou test. uer Impact: system not available. Patient Impact: no SE u Diagnostic performed by Engineer: Following test of the generator, no powering u of the equpment. uS and battery, power suply management modues of the main cabinet and imaging cabinet are defective. Follow u Requred: No follow-u. Equpment operational. Consued parts: 1.0000000000000000000000000000000080000000000000000300000000, PD. ASSY. FRONTPANEL BOuED 1.00000000000000000000000000000000800000000117161, uS 1Phase Data Integrity Controller SP 1.00000000000000000000000000000000800000000117151, uS 1Phase Data Integrity Battery SP Tools and Equpment: MuTIMETER, 200000000000000002-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000000000000025 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000088 Lionel GIROu The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: CDP TAuI: 00000000800000000117161 uS 1 phase data integrity controller Quntity: 1 00000000800000000117151 uS 1 phase data integrity battery Quntity: 1 0000000080000000000000000300000000 PD. ASSY. frontpanel boued Quntity: 1 Delivery by TAuI, Biomedical Department FSE: Jean Lous Mahier +33612552700000000 Delivery for 000000001000000002000000003 00000000:00000000 Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Problem restarting the system following a power ouage duing the grou test. uer Impact: system not available. Patient Impact: no SE u Diagnostic performed by Engineer: Problem restarting the system following a power ouage duing the grou test. Checked the electrical power suply of the system with the electricians. Intervention by Jean Lous Mahier to check the power suply. The system is powered electrically bu does not start. It seems that the problem comes from the uS in the electrical cabinet M which is off. Parts ordered by Taui for delivery tomorrow 000000001000000002000000003 ----------------------------------------------------------------------------- Internal Remarks: NO E-MAIL ----------------------------------------------------------------------------- | 2023-10-03 19:21 | No
1593 | Yes | The power cable has more resistance than allowed in electrical tests - Pending repair as of 06 | The power cable has more resistance than allowed in eleuical tests - Pending repair as of 0/6 |  | 2023-10-03 19:22 | No
1594 | Yes | 00 CN00- uCable is Damaged | u- uCable is Damaged | 000000000000000000000000001O##: 00000000000000000000000000100000000185938 October 17, 2000000003 Hidalgo Flores, Lenin Otton The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: ucable is damaged. uer Impact: Impossible to register u Patient Impact: The equpment is ued with limitations. SE Resoluion: Broken part is removed and the new one is installed. Cable test is secued and the person in charge is shown. I close the part. Parts Consued: 1.000000000000000000000000000000005600000000871,DFM100000000 Therapy Collar ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000150000000071 00000000-Oct-2000000003 Hidalgo Flores, Lenin Otton The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Material is requred for cover 000000005600000000871 Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: ucable is damaged. uer Impact: Impossible to register u Patient Impact: The equpment is ued with limitations. SE Diagnostic performed by Engineer: We checked and delivered consuable 98980000000010000000000000000. Ref M1672A or 989800000000170000000071 the consuable is in poor condition, the client is left, signatue will be signed, the person in charge will close the part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000052 Imbernon Masip, Francisco The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine Device ue: Defibrillator Eupected and ueupected Resuts: ucable is damaged. uer Impact: Impossible to register u Patient Impact: The equpment is ued with limitations. SE Diagnostic performed by Engineer: The client (Migul Angel 93227500000000) says that the ucable is damaged and requsts its replacement. Material order is handled and the case is transferred onsite so that the field service assigned to the center performs the euchange. Resoluion: Material order is handled. Follow u Requred: The client (Migul Angel 93227500000000) says that the ucable is damaged and requsts its replacement. Material order is handled and the case is transferred onsite so that the field service assigned to the center performs the euchange. - Equpment located in H-E00000000-00000000-000000001 - CARDIAC HEMODYNAMICS. - Assign FSE Lenin - Order the mentioned material ----------------------------------------------------------------------------- | 2023-10-03 19:31 | No
1595 | Yes | 00T0ystem not booting up no power to system | 00CDT0System not booting up no power to system | 000000000000000000000000001O##: 0000000000000000000000000010000000015000000000000000000000000-Oct-2000000003 Rodriguz Lazaro The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System not booting u, system down When was the first occurence? Within... : 00000000-Oct-2000000003 00000000:53:25 How often is issu occuring? : First occurence Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System not tuning on uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer : The system was already u and there was nothing in the logs to show what caued it. Resoluion : Biomed had reset the wall breaker for the room. System operational. No issu foud. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000150000000036 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System not booting u, system down When was the first occurence? Within... : 00000000-Oct-2000000003 00000000:53:25 How often is issu occuring? : First occurence Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System not tuning on uer Impact: System down Patient Impact: No futher information is available as per the cutomer SE Not provided Diagnostic performed by Engineer : John said this morning they coud not tun on the system. The tried resetting the system, the uS, reset breakers, and the shut trip. John said the only thing that seems to be powering on is the cooling uit. John wants to have FSE onsite to trouleshoot the issu. Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Check incoming power to system, quck check Green LED’s L1, L2, & L3 in bottom of the M cabinet. Check F1, F2, F3, F00000000& F5 if LED’s are off. If LED’s are on then check F100000000incoming power for Control uit Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Check incoming power to system, quck check Green LED’s L1, L2, & L3 in bottom of the M cabinet. Check F1, F2, F3, F00000000& F5 if LED’s are off. If LED’s are on then check F100000000incoming power for Control uit Technician Name : Coffman, Josh ----------------------------------------------------------------------------- | 2023-10-03 19:41 | No
1596 | Yes | M0A / u0 / Broken screen. | M00A / u00 / Broken screen. | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000096 Wauers, Anthony The system does not meet the specifications for the performed service and was not retuned to ue. Follow u Requred : Bench repair : 1*0000000056000000006581 +2 hous Problem Description by Engineer : Cutomer Fuction/Role: biomed. Device ue: Eupected and ueupected Resuts: uer Impact: None Patient Impact: SE Diagnostic performed by Engineer: Bench repair : 1*0000000056000000006581 +2 hous ----------------------------------------------------------------------------- | 2023-10-03 19:51 | No
1597 | Yes | This is a split case of - 00- Safety stop in CS is not installed, mu be installed. | This is a split case of - 000 - Safety stop in CS is not installed, mu be installed. |  | 2023-10-03 20:00 | No
1598 | Yes | Display does not work | Display does not work | Internal Remarks : New Case Caller: John Stasiewicu Issu: 56633300000000 Display does not 000000000000000000000000001k Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: No contract or warranty Action Taken: Created case an requsted qute Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: John Stasiewicu Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-03 20:06 | No
1599 | Yes | It does not turn on. | It does not turn on. | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000029 00000000Oct-2000000003 Jun Manul Lozano Garcia The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The uers say that the equpment does not start. There seems to be a problem with the power suply, J. M. Lozano is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment does not start. Patient Impact: No impact SE R2.2.6 Diagnostic performed by Engineer: Client's notice: "The equpment does not start". uon arriving at the room we saw that the general magnetothermic of the electrical board had come off and the SAI batteries had ru ou. There mut have been a prolonged power ouage in the hospital duing the weekend that affected the power suply of that electrical frame. The electrical frame is reasse re-set and it 000000000000000000000000001ks again. Tools & Equpment: MuTIMETER, 200000000000000007-12 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: The uers say that the equpment does not start. There seems to be a problem with the power suply, J. M. Lozano is in charge of checking it. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: The equpment does not start. uer Impact: The equpment does not start. Patient Impact: No impact SE 2.2 Diagnostic performed by Engineer: The uers say that the equpment does not start. There seems to be a problem with the power suply, J. M. Lozano is in charge of checking it. Internal Comments: uLL# and/or description of uLL N/A Repair Action: The uers say that the equpment does not start. There seems to be a problem with the power suply, J. M. Lozano is in charge of checking it. Follow u Requred: The uers say that the equpment does not start. There seems to be a problem with the power suply, J. M. Lozano is in charge of checking it. ----------------------------------------------------------------------------- | 2023-10-03 20:11 | No
1600 | Yes | Loudspeaker problem; Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker problem; Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 000000000000000000000000001000000001576800000000ctober 5, 2000000003 Frederic LOMBARD The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: No. Eupected and ueupected Resuts: uer Impact: None Patient Impact: None SE N Parts Consued: 1.00000000000000000000000000000000261000000009000000001,IV-MP5 CBL Speaker Assembly Resoluion: Replacement of the louspeaker. Reguatory Qustions: No/No ----------------------------------------------------------------------------- | 2023-10-03 20:11 | No
1601 | Yes | Needing help with calibration / LCD touchscreen/ Part ID | Needing help with calibration / LCD touchscreen/ Part ID | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000085 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Touhscreen issus / parts ID uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 00000000561539000000001 - RP-Touh Screen,V600000000- 990000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states Touhscreen will not calibrate on device. Cutomer attempted to perform touhscreen calibration and device will not pass. Cutomer is requsting replacement part nuber and price for Touhscreen. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Armando Guterez 250000000055000000009 IP & Issu: 10000000000000000600000000 / Needing help with calibration / LCD touhscreen/ Part ID Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-03 20:16 | No
1602 | Yes | AED emits triple chirp and fails self-test | AED emits triple chirp and fails self-test | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000092 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : End uer to be provided with service euchange device 1u 000000005600000000000000001 - HS1 (W B) Defib u Eng 00000000 Ed, Euchg Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Clinical Engineer Device ue: not in ue Eupected and ueupected Resuts: Machine triple chirping and failing self test. End uer tried new batteries and pads. Same resut. Machine failing. uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : End uer to be provided with service euchange device 1u 000000005600000000000000001 - HS1 (W B) Defib u Eng 00000000 Ed, Euchg ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000150000000055 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Defib, u English 00000000, Euch ----------------------------------------------------------------------------- | 2023-10-03 20:26 | No
1603 | Yes | Problem z wykonaniem zdjęcia ze skosu | Problem z wykonaniem zdję0 ze skosu | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000083 00000000-Oct-2000000003 Marek Jeske The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: uknown. Eupected and ueupected Resuts: Problems with C-arc movement. uer Impact: none. Patient Impact: none. SE 8.2.3000000005 Resoluion: The uer reported a problem with the device's stand movement. Locating the damage. Replacement of the motor control modue. A thorouh adjutment of positions and curents for all movements of the stand has been carried ou. Check of operation - correct. The faut was repaired. The device was handed over for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000990000000052,AMC Triple Servo Drive HP-LP-LP ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000022 Roman Kuka The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: Replacement of the drive motor for movement of the C-arm and/or the AMC modue (2SPC). Problem Description by Engineer: Cutomer Fuction/Role: uknown. Device ue: uknown. Eupected and ueupected Resuts: Problems with C-arc movement. uer Impact: none. Patient Impact: none. SE 8.2.3000000005 Diagnostic performed by Engineer: Analysis of the event log registered by the system revealed uLL related to euceeding the power suply voltage of the C-arm movement motor: 5100000000 uLL2000000003-10000000000000000 10000000000000000000000000000000000000000Application uLL Motor assembly uLL[auis=FrtCarc] 000000003000000001126 Normal Operation Service NoDetailInfo. [FREE_TEuT]. Time of occurence : 2000000003-10000000000000000, 1000000000000000000000000[SEVERITY] uLL[EuCEPTION_DESCRIPTION] [detail=eds(DRV=00000000DRVR=00000000 AE=3; AEF=00000000SBR=3; SBRF=3; FBR=00000000 FBRF=00000000 CR=00000000 CRF=00000000 CRIF=00000000 PS=00000000 P=3; V=00000000 I=00000000), MMS=00000000MSS=3; actPM=-10000000000000000958; actPL=-10000000000000000958 setP=-10000000053900000000 actV=-23.86000000009; setV=-23.8100000000; actI=00000000000000007233; SetI=1.0000000000000000 5100000000 Information 2000000003-10000000000000000 10000000000000000000000000000000000000000ATEM slave device event 000000003000000001500000000 Normal Operation Service NoDetailInfo [FREE_TEuT]. [device=AMC_OBJECT_AuIS SAM_CleaCARC]. [event=AMC_EV_FAuT_NOTIFICATION]. [detail=Auis is fauted.] Time of occurence: 2000000003-10000000000000000, 1000000000000000000000000 5100000000 Information 2000000003-10000000000000000 10000000000000000000000000000000000000000ATEM slave device event 000000003000000001500000000 Normal Operation Service NoDetailInfo [FREE_TEuT]. [device=AMC_OBJECT_AuIS SAM_CleaCARC]. [event=AMC_EV_EVENT_NOTIFICATION]. [detail=Over Voltage]. Time of occurence : 2000000003-10000000000000000, 1000000000000000000000000 The drive motor of the C-arm movement and/or the AMC control modue (2SPC) is probably damaged. On-site service visit is necessary in order to diagnose/repair the faut. ----------------------------------------------------------------------------- Cutomer comment: 000000003000000000000000000000000rg9NnQAI,000000003000000000000000000000000200000000gRQAu0000000030000000000000000000000006MFmHQAW. ----------------------------------------------------------------------------- | 2023-10-03 20:26 | No
1604 | Yes | System will not boot into the application software | System will not boot into the application software | 000000000000000000000000001O##: 00000000000000000000000000100000000136576 00000000-Oct-2000000003 Singh Amit The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: fse Device ue: no Eupected and ueupected Resuts: yes uer Impact: no Patient Impact: no SE8u Resoluion : Confirmed the reported failue. Confirmed IPPC issu. To resolve I replaced the IPPC and the OS drive. Configued the IPPC. Confirme dthe system is fuly operantional. Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- Internal Remarks : Case: 0000000022380000000055 FSE: Singh Amit System is not coming u; first HDD on IPPC is not lighting u. FSE reseated OS Disk; issu persisted. Reviewed logs; foud "DatauY uLL(DMuLLIMAGESTORE_NOT_AVAILABLE)". Instruted FSE to place part order for 0000000080000000050000000022 - ClarityIQ_II IP PC no HDD and 0000000080000000050000000000000000- Hard disk spare part. ----------------------------------------------------------------------------- | 2023-10-03 20:41 | No
1605 | Yes | System does not completely start up | System does not completely start up | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000083 00000000-Oct-2000000003 Alvizo, Lus The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: TR Device ue: u Eupected and ueupected Resuts: System does not start. uer Impact: No Patient Impact: No SE 000Resoluion: We replace the FLEu VISION CPuand the hard disk, we install the software and the necessary configuations, we perform operation tests, the equpment is left 000000000000000000000000001king correctly. Parts Consued: 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Tools & Equpment: MuTIMETER, 20000000000000000000000009 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000137656 00000000-Oct-2000000003 Mauimo, Migul The system does not meet the specification for the performed service and is not retuned to ue. Follow u Requred: Fleuvision needs to be changed. pC and OS HDD Problem Description by Engineer: Cutomer Fuction/Role: TR Device ue: u Eupected and ueupected Resuts: System does not start. uer Impact: No Patient Impact: No SE 000Diagnostic performed by Engineer: The equpment is checked and a problem is foud in the Fleuvision drive and this prevents the system from starting, which will be requsted for replacement. ----------------------------------------------------------------------------- | 2023-10-03 20:56 | No
1606 | Yes | touch screen on the reset alarm is not working | touch screen on the reset alarm is not working | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000093 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the touhscreen and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and the touhscreen is not operating right. He did not calibrate the touhscreen and now it is open on his 000000000000000000000000001k bench. The alarm reset was not responding to the touh. Provided some parts #s Part #00000000561539000000001, touhscreen Part #0000000056153100000000, caster lock Part #00000000561530000000051, base assy V600000000stand. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Eddie Do IP & Issu: 77830000000011 & touh screen on the reset alarm is not 000000000000000000000000001king Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-03 21:06 | No
1607 | Yes | Machine and prouimal pressure sensors failed (EC 00 | Machine and prouimal pressure sensors failed (EC 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000098 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uLLcode 1000000007 uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and uit is not reading the MC, PM, DA or the 2 flow sensor on the vent info screen. The FCO66 PM pcba was replaced on 8-30000000023, and soon after that, they encoutered this failue. Cutomer already tried another PM pcba within a past coule week, and it did solve the issu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000160000000079 Barragan, Nick Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Tyler Gresko IP & Issu: 70000000096600000000Not reading the flow sensors or the boards Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-03 21:11 | No
1608 | Yes | Hard drive error | Hard drive error | 000000000000000000000000001O##: 00000000000000000000000000100000000137267 00000000-Oct-2000000003 Henry, Shane The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: recreate image discs uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : recreated image discs for the frontal and lateral stands. fiued the issu for no ready light and not sending to PACS du to low storage space. retuned to service ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: James Morton Issu: 51250000000075 Hard drive uLL Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: No contract or warranty Action Taken: Created case an requsted qute Neut Step Action: Awaiting Qute Verbal PO Nuber: 0000000000000000-15600000000500000000 Verbal PO Approver: James Morton Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-03 21:16 | No
1609 | Yes | 00UNIT WILL SHUT DOWN AND RESTART BY ITSELF | 000 UNIT WILL SHUT DOWN AND RESTART BY ITSELF | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000000000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : 10000000000000000/2000000003 Cutomer has reinstalled software version 2.300000000and will ru the uit for a few hous. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000not shu down. uer Impact: V600000000has been removed from service. Patient Impact: Biomed is not reporting and adverse oucome to patient care or therapy. The uLLwas removed from service and the patient was placed onto another uit. SE 2.300000000 Problem Description/ueupected Resuts: V600000000uLLshu down while on a patient. Diagnostic performed by Engineer : Problem Description: V600000000uLLshu down while on a patient Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: 2.300000000 Date and Time of Issu: 10000000000000000/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version R. Biomed has confirmed V600000000diagnostic code 1100000000 in the significant event log. uLLrestarted du to anomalies detected duing operation. This coud inclue invalid or corruted information, uanticipated situtions, or object allocation uLL. NOTE If diagnostic code 1100000000 occus in conjuction with diagnostic code 3000000007 (software euception), no action is requred. Biomed has confirmed diagnostic code 3000000007 in the log. 1. Reinstall operational software. 2. Replace the CPuPCBA. Biomed is going to reinstall software version 2.300000000 ----------------------------------------------------------------------------- Internal Remarks : **CuT REQuSTED TECH SuPORT Interaction Type: New Case Caller: RICHARD ARIAS IP & Issu: 89617900000000 uIT WILL SHu DOWN AND RESTART BY ITSELF - DEVICE WAS IN CLINICAL uE Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND REMOTE 000000000000000000000000001 RTAC Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-03 21:21 | No
1610 | Yes | WFS not working | WFS not working | 000000000000000000000000001O##: 00000000000000000000000000100000000292800000000 00000000-Nov-2000000003 Anand Pishe The system meets the specification for the performed service and is retuned to ue Resoluion : Repalced table base connection board . Problem Description by Engineer : Cutomer Fuction/Role: technician Device ue: yes. Eupected and ueupected Resuts: fuctionality of the system uer Impact: no Patient Impact: no uuDiagnostic performed by Engineer : table base connection board problem Parts Consued : 1.0000000000000000000000000000000080000000037100000000,TABLE BASE CONNECTION BOu Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000151111 26-Oct-2000000003 Anand Pishe The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: technician Device ue: yes. Eupected and ueupected Resuts: fuctionality of the system uer Impact: no Patient Impact: no uuDiagnostic performed by Engineer : INSTALLED WFS Bu SuPECTING SD1 ISSu. NEED TO REPLACE SD1 Resoluion : NEED TO REPLACE SD1 Follow u Requred : PART REPALCEMENT Parts Consued : 1.00000000000000000000000000000000800000000238252,Wireless Footswitch 000000002 1.00000000000000000000000000000000800000000257862,WFS Base Station Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000006-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000063 00000000-Oct-2000000003 Anand Pishe The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Need to replace WFS Follow u Requred : WFS and base station replacement Problem Description by Engineer : Cutomer Fuction/Role: technician Device ue: yes. Eupected and ueupected Resuts: fuctionality of the system uer Impact: no Patient Impact: no uuDiagnostic performed by Engineer : WFS is not 000000000000000000000000001king ----------------------------------------------------------------------------- | 2023-10-03 21:26 | No
1611 | Yes | Shut down uneupectedly | Shut down uneupeudly |  | 2023-10-03 21:26 | No
1612 | Yes | Prouimal pressure sensor autozero failed. (EC 0B) - Final 00 | Prouimal pressure sensor autozero failed. (EC 00) - Final 0/0 | 000000000000000000000000001O##: 00000000000000000000000000100000000151000000009 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised cutomer to perform eutended ru in to dulicate complaint and perform PVT testing, advised that Sol 3 and 00000000an be replaced as a precauion as well as DA pcb and cable if needed, cutomer acknowledged, no futher follow u needed, issu resolved. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:uit alarms for uLL1100000000 uer Impact: removed from service Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer states uit was in ue and alarmed for uLL1100000000, uit removed from ue and no patient harm reported, in biomed the uLLcode was verified bu not able to be dulicated, no service performed recently, requsts trouleshooting. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Rafael Rod 800000000723-8100000000 IP & Issu: uLL1100000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Was the device in clincal ue? No Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-03 21:31 | No
1613 | Yes | Giving uLLrunning on internal battery | Giving uLLrunning on internal battery | 000000000000000000000000001O##: 00000000000000000000000000100000000151100000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the power suply and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: uit switching from AC to battery duing every change of pressue uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the uit switching from AC to battery at the beginning of higher air flow. When youlower the Ipap to below 100000000the issus goes away. He tried a go PM pcba, and the issu remained. Recommended replacing the power suply Part # 00000000561512591, power suply, list price $890000000000000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Art IP & Issu: 56631119/ giving uLLruning on internal battery. When on AC Power the uLLmsg comes u. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-03 21:36 | No
1614 | Yes | Paddle Tray assembly broken. | Paddle Tray assembly broken. | 000000000000000000000000001O##: 00000000000000000000000000100000000197900000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001000000001 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: The requsted spare part is blocked: please verify if it can be ordered throuh another channel. Thank you Make NEMO withou collection at the same time. Tu c/o ASST Papa Giovanni uuIII Piazza Oms, 1 2000000007 Bergamo (BG) Tel. 0000000052678187 C/O LABORATORY H.C./u BASEMENT FLOOR Ac. Francesco Mazzoleni Tel. 3200000000320000000091 Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Paddle Tray assembly is broken and to be replaced uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Technician Francesco Mazzoleni asks to send Paddle Tray assembly (uu0000000056300000000000000001) in replacement of the fauty part. ----------------------------------------------------------------------------- | 2023-10-03 21:36 | No
1615 | Yes | Touchscreen is nonresponsive. | Touchscreen is nonresponsive. |  | 2023-10-03 21:36 | No
1616 | Yes | Would not power on | Would not power on |  | 2023-10-03 21:41 | No
1617 | Yes | green check mark is not working | green check mark is not working | 000000000000000000000000001O##: 00000000000000000000000000100000000151100000000Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer stated when he opened the u assy, an reseated the cable going to the switch pcba. This solved the issu. He was able to go into diagnostic mode sucessfuly. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and the green check mark is not 000000000000000000000000001king. Instruted him to take apart the u assy, and inspect the cable going into the switch pcba. If needed disconnect it and secuely reconnect the cable. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: green check mark is not 000000000000000000000000001king uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: joel gonzalez IP & Issu: 57000000009500000000 the buton by the uuudoes not 000000000000000000000000001k Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: provided free remote created case r000000000000000000000000001and roued to rse Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-03 21:51 | No
1618 | Yes | Not generating u-Rays | Not generating u-Rays |  | 2023-10-03 22:01 | No
1619 | Yes | SCREENS BLANK WITHOUT PATIENT DATA.-Unable to confirm cau-Unknown reason - insufficient information | SCREENS BLANK WITHOUT PATIENT DATA.-Unable to confirm cau-Unknown reason - insuffi0nt information |  | 2023-10-03 22:29 | Yes
1620 | Yes | Patient u-Paddles created short | Patient u-Paddles created short |  | 2023-10-03 22:37 | Yes
1621 | Yes | The equipment is beeping. | The equipment is beeping. | BuQ&R Feedback: No Reason: Corrective Maintenance. Safety Qustions: Case: 0000000022380000000097 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Bruo Miranda. Last udated by: Bruo Miranda. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown Problem Reported by Cutomer: The equpment's problem: The equpment is beeping. Priority: Stopped Look for: Denise Department: Sugical Center 00 +55110000000056300000000 00 +551190000000039995 Contact ##CentroCirugico00nstituodeolhosleste.com.br uit: Instituo de Olhos Leste Address: Ru Gabriel Machado nuero 00000000in front of praça getuio vargas segudo andar #chamado ----------------------------------------------------------------------------- | 2023-10-03 22:41 | No
1622 | Yes | Split - Speaker malfunction; Visual alarming only at the Mu0. Cu resolved | Split - Speaker malfunuon; Visual alarming only at the Mu0. Cu resolved | 000000000000000000000000001O##: 00000000000000000000000000100000000151600000000Adams, Devon The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : Is Caue Known: No Caue: Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Problem Description: Speaker malfuction uLLmessage alerting at Mu700000000 and PIC iu. PIC iu & Mu700000000 & u2- Speaker malfuction at the central and at the bedside. No injuies or deaths occured. Visul alarming only at the Mu700000000. Form ued: Trouleshooting/Investigation Object: u2 or Mu700000000? Not Object: PIC iu. Software Version: PIC iu C.00000000.00000000/ Mu00000000 uu u2 uu Date and Time of Issu: 1000000003/2000000003 uer Actions: Monitoring patient Where is the Deviation: Mu700000000 SN: DE1250000000026u2 SN: biomed rejected requst to get hardware SN. Trouleshooting Steps Taken: Biomed stated that the bedside monitor uY was an Mu00000000 and then changed information to be Mu700000000. Collected data. Recommended that the patient be disconnected from the bedside monitor so that the biomed coud trouleshoot which device had the speaker malfuction uLL Biomed advised staff to disconnect patient. Recommended that AFTER the patient had been disconnected, the biomed attempt to isolate the speaker malfuction uLLto either the Mu700000000 or u2 by disconnecting the u2. The call dropped.Called back biomed and biomed said that the staff 000000000000000000000000001ld swap the u2 to start. The biomed rejected futher remote trouleshooting. The biomed requsted the case be closed. Was PRS ued: No PRS Permission: Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Get remote trouleshooting help to uderstand how to isolate issu. uer Impact: uable to trut bedside setu du to lacking auible alert and uLLmessage. Patient Impact: None at this time. SE PIC iu C.00000000.00000000/ Mu00000000 uu u2 uu Problem Description/ueupected Resuts: Speaker malfuction uLLmessage alerting at Mu700000000 and PIC iu. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: scott hilton \ 97000000005865/ speaker malfuction Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-03 23:06 | No
1623 | Yes | Relay test failed during weekly test | Relay test failed during weekly test | 000000000000000000000000001O##: 000000000000000000000000001000000001517100000000ent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Equpment manager states they were performing the weekly test on this uit and it failed the relay test; needs to get the device euchanged uder their contract. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Diagnostic performed by Engineer : PT# 000000005600000000181 Tempu LS Man Defibrillator Verified shipping address: Johnson Couty Ambuance District ATTN: Jeremy West E Youg Ave 500000000 Warrensbug, MO 6000000003-1200000000 u Follow u Requred : PT# 000000005600000000181 Tempu LS Man Defibrillator Verified shipping address: Johnson Couty Ambuance District ATTN: Jeremy West E Youg Ave 500000000 Warrensbug, MO 6000000003-1200000000 u Resoluion : Verified cutomers contact and shipping information in the system is correct and placed a parts order. Created a follow u activity for the parts order and attached the part needed for the activity to the 000000000000000000000000001 Attached the log and rescu files to the 000000000000000000000000001agnostic performed by Engineer : Sent follow u ##requsting the cutomer send the log and rescu files for this device. Diagnostic performed by Engineer : Cutomer followed u and sent all the files. I attached all the files to the case and ##d them to the tempu team. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000152168 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Caller: Jeremy West IP & Issu: Relay test failed duing weekly test Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Verbal PO Nuber: ----------------------------------------------------------------------------- | 2023-10-04 00:06 | No
1624 | Yes | “Audible & Visual notifications affected-No product malfunction-software issue-upgrade-FCO implemented | “Audible & Visual notifications affeud-No produumalfunuon-software issue-upgrade-FCO implemented | 000000000000000000000000001O##: 00000000000000000000000000100000000137811 00000000Oct-2000000003 Gud Casado The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: monitor patient vital signs Eupected and ueupected Resuts: vital signs to patient display monitor. uer Impact: Patient Impact: SE Resoluion : Implemented FCO86200000000882. uit 000000000000000000000000001king properly. Site has more uits affected as per FCO. Tools & Equpment : TEST SIMuATOR KIT,200000000000000005-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: JJ Action Taken: FSE-Gud Casado ##d me to create a FCO 000000000000000000000000001from Serial# u000000006C5900000000 and dispatch to him for 1000000003/23 from 11am-12pm with no travel time. I created Case#0000000022381000000005/00000000000000000000000000100000000137811 I dispatched the 000000000000000000000000001per FSE requst I Pu FADB in the Planning Notes Last udated: 2:38PM EDT 1000000002/23 ----------------------------------------------------------------------------- | 2023-10-04 00:11 | No
1625 | Yes | Low leak CO0ebreathing risk alarm | Low leak CO0rebreathing risk alarm | 000000000000000000000000001O##: 00000000000000000000000000100000000152357 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the flow sensor assy and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and they are getting a check vent low leak Co2 rebreathing risk. He also euplained they cannot get the uit to stay in stand by. Recommended relacing the flow sensor Part #989800000000617100000000 flow sensor assy Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: check vent low leak Co2 rebreathing risk alarm uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Harry Weaver IP & Issu: 6000000000000000000000000/The uit is giving a "low leak CO2 rebreathing risk alarm". Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-04 00:51 | No
1626 | Yes | Euposure not possible. Only Fluoroscopy works. | Euposure not possible. Only Fluoroscopy works. | 000000000000000000000000001O##: 0000000000000000000000000010000000015200000000 TSCHuTZ NIKLAS The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: na Device ue: Eupected and ueupected Resuts: uer Impact: Na Patient Impact: na SE Diagnostic performed by Engineer: Checked Euposue and Fluroscopy. Resoluion: No faut foud. Euposue and uuucopy 000000000000000000000000001k. Internal Comments: Poor information from the uer, who called me directly and informed them regarding euposue not possible. ----------------------------------------------------------------------------- | 2023-10-04 00:51 | No
1627 | Yes | Monitor fell off the wall/cracked. Unable to Confirm Cau-No Further Info Available. Provided information. | Monitor fell off the wall/cracked. Unable to Confirm Cau-No Further Info Available. Provided information. | 000000000000000000000000001O##: 0000000000000000000000000010000000015200000000 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states the arm on the wall broke and the MP700000000has a lot of physical damage; needs to check what options are available the have the uit repaired. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Resoluion : Informed cutomer the uit is end of life and we no longer have any parts or offer any repairs. Verified cutomers ##address and sent the end of life letter SB86200000000592A. Diagnostic performed by Engineer : Informed cutomer the uit is end of life and we no longer have any parts or offer any repairs. Verified cutomers ##address and sent the end of life letter. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ryan Baker IP & Issu: monitor fell off the wall/ cracked Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: created 000000000000000000000000001 roued to tech suport ----------------------------------------------------------------------------- | 2023-10-04 01:36 | No
1628 | Yes | switch button not working | switch button not working | 000000000000000000000000001O##: 00000000000000000000000000100000000152532 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: switch buton not 000000000000000000000000001king uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the switch buton is not 000000000000000000000000001king properly and is requsting to have a replacement switch PCBA and cabling shipped for installation. confirmed cutomers ship to address. setting u material 000000000000000000000000001to have parts shipped to cutomer. Resoluion : confirmed cutomers ship to address. setting u material 000000000000000000000000001to have parts shipped to cutomer. Follow u Requred : Please ship replacement parts to: MEMORIAL HOSPITAL WEST Attn: Myra Becerra / BIOMED 700000000 N Flamingo Rd Pembroke Pines, FL 33000000008 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000152800000000 Parts Consued : 1.00000000000000000000000000000000561511600000000RP-Cable, Switch PCBA, V600000000 1.00000000000000000000000000000000561512600000000,RP-PCBA, Switch, V600000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Myra IP & Issu: 59710000000000000000 switch buton not 000000000000000000000000001king Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Entitled Reason for Entitlement Claim: Entitled Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-04 02:21 | No
1629 | Yes | The device was not secure on the universal stand. | The device was not secure on the universal stand. | 000000000000000000000000001O##: 00000000000000000000000000100000000152600000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the needed parts, and repair V600000000 “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000 and they need to secue the uit to the uiversal stand. Provided part #s Part #989800000000630000000061, thubwheels, 2 pk Part #00000000561500000000131, foot, gray 00000000k Part #00000000561531961, cover CPutray Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: secue the uit to the uiversal stand. uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: dillion zellers \ 5662800000000 mouting screws parts id Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created case/r000000000000000000000000001transf to tech Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-04 02:31 | No
1630 | Yes | Battery failure | Battery failure | 000000000000000000000000001O##: 00000000000000000000000000100000000152693 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Advised FSE that non u battery is not compatible with new PM pcb containing PIC 1.300000000software, cutomer will need to replace with a u brand battery to correct issu, FSE acknowledged and no futher suport needed, issu resolved. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: FCO service Eupected and ueupected Resuts: battery failed alarm after FCO66 PM pcb replacement uer Impact: not uable needs repair Patient Impact: none SEPIC 1.300000000Diagnostic performed by Engineer : FSE performed FCO86600000000000000006 PM pcb replacement and uit alarmed Battery failed alarm on screen and verified uLLcode 11200000000n event log, uit has non u brand battery installed, requsts trouleshooting. ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-04 02:36 | No
1631 | Yes | The power cord needs to be replaced due to electrical safety issues | The power cord needs to be replaced due to eleuical safety issues |  | 2023-10-04 02:38 | No
1632 | Yes | 0 (CBIT) Check Vent: CPU PCBA ADC Failed | 00 (CBIT) Check Vent: CPU PCBA ADC Failed | 000000000000000000000000001O##: 00000000000000000000000000100000000152837 Mccleary Barry The system meets the specification for the performed service and is retuned to ue Resoluion : Advised cutomer to perform eutended ru in and testing, if uLLcannot be dulicated then proceed with FCO66, cutomer acknowledged and will continu with FCO testing, no follow u needed, issu resolved. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: FCO service Eupected and ueupected Resuts:111E in event log uer Impact: requsts trouleshooting Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states duing FCO86600000000000000006 PM pcb replacement when reviewing the event log uLL111E was noted from 3/2000000003, no other VI or CV uLLcodes foud and code was not dulicated after that, uit cycles normally withou alarming and no malfuctions, requsts neut step action. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jim Rajala +190000000022875500000000 IP & Issu: 62100000000627 Need help with a FCO and have some qustions Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: R000000000000000000000000001created and roued call to RSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-04 02:46 | No
1633 | Yes | Lack of a red alarm | Lack of a red alarm | 000000000000000000000000001O##: 0000000000000000000000000010000000015000000002 Youg, Maria The system meets the specification for the performed service and is retuned to ue Resoluion : Was able to show the tech why the monitor did not alarm becaue the beat was labels as normal. Monitor keep losing leads and doing relearning. Problem description by Clinical Suport Specialist : Cutomer Fuction/Role: biomed and monitor tech Device ue: monitoring a patient Eupected and ueupected Resuts: Monitor did not alarm for the Vtach alarm uer Impact: none Patient Impact: none SE Rev C Diagnostic performed by Engineer : The monitor did not alarm Vtach at 1000000001/23 at 16:51 Review the strip and it is labeling the PVCs as at normal. Requsted the monitor determine where the wide beat was learned as normal. That was done at 9/300000000at 15:500000000 See attached strips. The monitor tech and biomed uderstands the why it did not alarm. No patient data available. ----------------------------------------------------------------------------- File uloaded by CS Portal uer : File Name: uerou Scan_1000000002200000000310000000000000000.pdf. Comments: File(s) uloaded duing case creation ----------------------------------------------------------------------------- | 2023-10-04 02:51 | No
1634 | Yes | 00 CN00- Defective speaker | u- Defeu speaker |  | 2023-10-04 03:05 | No
1635 | Yes | Audio & Sounds/No sound even after speaker replacement/Cuomer repairing the unit | Audio & Sounds/No sound even after speaker replacement/Cuomer repairing the unit | 000000000000000000000000001O##: 00000000000000000000000000100000000150000000079 Harris, Alvin The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Is Caue Known: Yes Caue: API board appears to be defective and requre replacement, may also requre replacing the NBP assembly Is Resoluion Known: Yes Resoluion Sumary: Malfuction Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Agreed to send ##to curequsting his repair decision when available Custated he 000000000000000000000000001ld trouleshoot futher, speak with his Management, and determine their repair decision Cuagreed to reply by ##with their repair decision Parts/Suplies Information Provided Part ID and Price: API board - PN 00000000261000000005111 - $1,800000000.00000000 NBP assembly - PN 0000000056000000005571 - $950000000000000000 Main board - PN 0000000056000000008621 - $1,200000000.00000000 euchange Prevention: --- Sent ##to curequsting repair decision when available --- Checked ##on 1000000006/23, foud no response from cu Called back to 773-296-500000000 for repair decision, left voic## sent ## --- Checked ##on 10000000010000000023, foud no response from cu Called back to 773-296-500000000 for repair decision, left voic## sent ## --- Checked ##on 10000000011/23, foud no response from cuFollow u Requred : ATT: Joshu Rivera/Biomed/Case 0000000022392231 ##d Irvine CA Bench form to cuProblem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts:Device will give soud throuh its speaker when powered on/Will power on uer Impact:Can't ue device in this state Patient Impact:No impact SEuu cucoud not provide this becaue he coud not get the device to power on Problem Description/ueupected Resuts: Device will not give soud throuh its speaker when powered on even after replacing the speaker/Now will not power on Diagnostic performed by Engineer : Problem Description: Device will not give soud throuh its speaker when powered on even after replacing the speaker/Now will not power on No soud even after speaker replacement Form ued: Parts Trouleshooting Steps Taken: Cucoud not say when the issu was first seen, the device was handed over to him from a co-000000000000000000000000001ker Custated the device came from their Labor & Delivery department Custated a speaker assembly was already puchased from u, installed in the device, bu the issu remains Mentioned the possibility of the speaker being defective, informed cuthat he coud requst another speaker and retun the originally puchased speaker Reviewed the Service gude sugestions with the cu discused the involvement of the API board Asked cuif he had the Suport Tool, custated he believed 1 of his co-000000000000000000000000001kers have the tool Asked cuto power device on so that I coud verify its software version Cuattempted to power the device on, bu the device 000000000000000000000000001ld not respond althouh the AC power LED illuinated green when connected to AC Cuconfirmed the issu was not with the power cord or AC mains Sugested curemove NBP assembly, disconnecting it from the main board and attempting to power the device on, if this resolved the issu, the NBP assembly shoud be replaced Cualso asked for the main board part information Discused repair options incluing Bench repair ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Joshu Rivera IP & Issu: No soud even after speaker replacement Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-04 03:06 | No
1636 | Yes | system will not power up | system will not power up | 000000000000000000000000001O##: 0000000000000000000000000010000000013880000000000000000Oct-2000000003 Alford, Matthew The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: cathlab Device ue: cathlab Eupected and ueupected Resuts: n/a uer Impact: Patient Impact: SE Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS Diagnostic performed by Engineer : part replaced uit is u. ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action taken: DSM Greg Boley requsted call to be dispatched to FSE Matthew Alford for 1000000002000000000 Oucome:dispatched Internal Remarks : Case: 0000000022381752 FSE: Alford, Matthew System not booting u; No LEDs not powering on PDu (From uS) L1, L2, L3 measues 2300000000+; Pass. Instruted FSE to replace 0000000080000000000000000900000000 - PDuFanTray and DCPS. Internal Remarks : Interaction Type: New Case Caller: Peter Roberts IP & Issu: system will not power u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Qute Neut Step Action: Awaiting Qute Verbal PO Nuber: TMuC1788900000000 Verbal PO Approver: Peter Roberts Verbal PO CAP: N/A Internal Remarks : Interaction Type: cut callback Caller: Peter Roberts IP & Issu: system will not power u 8700000000200000000 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created AA Neut Step Action: N/A Verbal PO Nuber: TMuC1788900000000 Verbal PO Approver: Peter Roberts Verbal PO CAP: N/A Internal Remarks : Planner Notes Planner Initials: TB Action Taken: The assigned FSE Matthew Alford confirmed he was onsite for 00000000000000000000000000100000000138800000000on 10000000000000000 and the issu was resolved, and he spoke to the cutomer physically onsite. The engineer stated that he 000000000000000000000000001ld follow u and call cutomer: Peter Roberts at 000000006275159 via the requst listed on action assignment#00000000000000008185600000000 Oucome: Requst completed Neut Steps: N/A Last udated: 10000000016/23 3:00000000M EST ----------------------------------------------------------------------------- | 2023-10-04 03:16 | No
1637 | Yes | Self inspection failure | Self inspeuon failure | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: 黄丽智 18900000000918838 Safety Qustions: Case: 0000000022382200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhang Caiqin Last udated by: Zhang Caiqin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Self inspection failue ----------------------------------------------------------------------------- | 2023-10-04 06:06 | No
1638 | Yes | 0/u FD0Unable to enter the system when turned on | 0R/u FD0Unable to enter the system when turned on | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 00000000.1000000002000000003 15:30000000031 uC (86000000000000000000000000 RDI Accuacy: YES $Q6: (The device can't boot u.) $Q7:(Comfirmed reported problem.The device can't boot u.) $Q8:(Check and find it is the HostPC software problem, need to reinstall the HostPC software.) $Q9:(Reinstall the HostPC software.) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: N/A. Euternal Remarks: 00000000.1000000002000000003 15:30000000029 uC (86000000000000000000000000 Service content: I cannot enter the system when I tun on the compuer. I checked and foud that it was a software problem. I reinstalled the software and it was normal. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022382658 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: uiyuyang Last udated by: uiyuyang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: uable to enter the system when tuned on ----------------------------------------------------------------------------- | 2023-10-04 06:51 | No
1639 | Yes | System does not get past count down | System does not get past count down | 000000000000000000000000001O##: 00000000000000000000000000100000000153867 100000000Oct-2000000003 Nevo, Eyal The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: FSE completing a PM Eupected and ueupected Resuts: Once completed PM, did a system shu down and system started freezing at coutdown. Application never loaded. uer Impact: System is restricted and shoud not be shu down. Patient Impact: No harm SE 8.2 Resoluion : Replaced IPPC. Rebult database. System fuly fuctional Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000015300000000 00000000Oct-2000000003 Nevo, Eyal The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: FSE completing a PM Eupected and ueupected Resuts: Once completed PM, did a system shu down and system started freezing at coutdown. Application never loaded. uer Impact: System is restricted and shoud not be shu down. Patient Impact: No harm SE 8.2 Follow u Requred : IPPC replacement Diagnostic performed by Engineer : System freezing u at start u. Does not start application and coutdown finishes. Spoke with RTAC and IPPC is confirmed to be freezing u duing start u. Replaced for PC is needed in order for system to fuction properly. Cutomer is able to ue system util compuer replacement. ----------------------------------------------------------------------------- Internal Remarks : Eyal, system stuk at 0000000000000000after PM. Log shows Frontal IPPC has all POST failue bu powered on as lateral IPPC. Advised him to check ethernet cable u100000000to THu12. Tried new cable, and powered u sucessfuly at one point. Approved the IPPC, also tried another ethernet cable. Internal Remarks : TPR review and approval ----------------------------------------------------------------------------- | 2023-10-04 07:21 | No
1640 | Yes | Equipment will not boot up completely. | Equipment will not boot up completely. | 000000000000000000000000001O##: 0000000000000000000000000010000000015350000000000000000-Oct-2000000003 Belzer Juie The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: manager Device ue: vascuar lab Eupected and ueupected Resuts: table cables got stuk uder table while moving swivel base uer Impact: uable to ue Patient Impact: uable to ue SE 000Resoluion : Replaced hard drive for FleuVision PC and reconfigued system. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000151770000000000000000-Oct-2000000003 Belzer Juie The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FleuVision PC replacement Diagnostic performed by Engineer : Inspected cables for damage and inspected system. Retrieved uLLlogging to analyze. Attempted to reload software to system FleuVision PC. Problem Description by Engineer : Cutomer Fuction/Role: manager Device ue: vascuar lab Eupected and ueupected Resuts: table cables got stuk uder table while moving swivel base uer Impact: uable to ue Patient Impact: uable to ue SE 000----------------------------------------------------------------------------- | 2023-10-04 07:51 | No
1641 | Yes | The AED beeped once - log analysis requested (Device is replaced) | The AED beeped once - log analysis requested (Device is replaced) | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000096 Suui, Kosue The system meets the specification for the performed service and is retuned to ue Resoluion: When I analyzed the AED log at a factory in the uS., I foud a record of a high voltage (HV) calibration test not being passed. The AED was not diagnosed du to an electrical failue on a high voltage element in the AED, so I replaced the AED. I replaced the AED and confirmed that the AED passed a battery fit test. * In this case, abnormal records were confirmed at low temperatues from the recommended AED temperatue. The AED and accessories are fauty and can caue prematue wear, so please store the AED uder the recommended temperatue environment. [Simutaneou FCO 000000000000000000000000001k] There is no FCO for the device in which 000000000000000000000000001k is requred. October 16, 2000000003 Fuua Denshi Nagano Hanbai I will send the 000000000000000000000000001k report along with the equpment to the person in charge. Internal Comments: [000000000000000000000000001k report (paper) sumission] Did yousumit a 000000000000000000000000001k report on paper? ＜YES＞ Final check: I checked the operation following the 000000000000000000000000001k based on the set of orders. Yes Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Person in charge of Sales at Fuda Electronics Device ue: Not ued Eupected and ueupected Resuts: The AED beeped once. uer Impact: None Patient Impact: None SE None Diagnostic performed by Engineer: When I received the AED, I confirmed that the AED passed a battery fitted test. I did not confirm any abnormalities in the standby monitoring. Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 + (C G) Defib, Japanese, Euch Tools & Equpment : MuTIMETER, 200000000000000003-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks: Failue content 2 The AED beeped once and replacing the battery did not improve the sitution. ----------------------------------------------------------------------------- | 2023-10-04 08:21 | No
1642 | Yes | 0 / u0/ High voltage uuuitor aging. | 00/ u000/ High voltage uuuor aging. | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 0000000022383596 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_1100000000ndoscopy Room_8612900000000High voltage uuuitor aging SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-04 08:46 | No
1643 | Yes | System can't startup application | System can't startup application | 000000000000000000000000001O##: 0000000000000000000000000010000000015000000009 00000000-Oct-2000000003 CHATCHAI SANTANA The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Nuse Device ue: not ue Eupected and ueupected Resuts: system can't fluroscopy uer Impact: no Patient Impact: no SE8.2.17 Diagnostic performed by Engineer : Grabber card in FleuVision PC not initial. Resoluion : Replace FleuVision PC. The system tested and retun to ful fuction. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000150000000000000000SANGBOONSONG MuDGROM The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: Nuse Device ue: not ue Eupected and ueupected Resuts: system can't fluroscopy uer Impact: no Patient Impact: no SE8.2.17 Diagnostic performed by Engineer : System can't startu application Follow u Requred : System can't startu application ----------------------------------------------------------------------------- | 2023-10-04 11:06 | No
1644 | Yes | 0 / u0/ Battery issues (Battery is replaced) | 00/ u000/ Battery issues (Battery is replaced) | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 000000002238000000009 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? N/A 00000000Was it an ou-of-the-bou failue? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP Problem Reported by Cutomer : 1600000000763-Battery issus SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-04 11:51 | No
1645 | Yes | Need a new host PC | Need a new host PC | 000000000000000000000000001O##: 000000000000000000000000001000000001500000000 JALAR JIM The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow-u 000000000000000000000000001k schedued Follow-u Requred : Host PC has issus with starting, also last week there were power issus to the HDDs in the HOST PC. Need to replace it. Order a new host PC for the cutomer as material only. I will suport them against replacing the part. The system code is 7220000000000000000Problem Description by Engineer: Cutomer Fuction/Role: MT Device ue: in ue Eupected and ueupected Resuts: Intermittent issus with HOST PC the wake on power not 000000000000000000000000001king as it shoud, becaue PC has power to it. uer Impact: manul start of PCs Patient Impact: NA SE000Diagnostic performed by Engineer: Host PC has issus with starting, also last week there were power issus to the HDDs in the HOST PC. Need to replace it. Order a new host PC for the cutomer as material only. I will suport them against replacing the part. The system code is 7220000000000000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000015000000009 Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD ----------------------------------------------------------------------------- | 2023-10-04 12:36 | No
1646 | Yes | 00 CN00- Defibrillator does not turn on (Parts are replaced) | u- Defibrillator does not turn on (Parts are replaced) | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000037 PATRICK POELS The system meets the specifications for the performed service and is retuned to ue Resoluion: 0000000056000000009000000001 DFM100000000 SOM PCA ASSY and 0000000056000000009000000001 - DFM100000000 ASSY Processor need to be replaced Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: ouside of ue Eupected and ueupected Resuts: DFM100000000 is not powering on uer Impact: cannot ue DFM100000000 Patient Impact: N/A SE uknown ----------------------------------------------------------------------------- | 2023-10-04 12:51 | No
1647 | Yes | uLLmessage - there is a restriction and the system cannot acquire | uLLmessage - there is a restriuon and the system cannot acquire | 000000000000000000000000001O##: 00000000000000000000000000100000000256261 200000000ct-2000000003 Nigel Huphries The system meets the specification for the performed service and is retuned to ue Resoluion : Following failue to load software onto IP-PC on previou visit, the new disks were transferred to the 2nd replacement PC and software reloaded. This time the software reload was sucessfu. Image disk storage verified ok. Fuctional test of imaging and storage for both cardiac and vascuar is sucessfu. System can be retuned to clinical ue. Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Cardiac/Vasc Eupected and ueupected Resuts: Imaging stops uer Impact: High Patient Impact: High SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000255663 200000000Oct-2000000003 Nigel Huphries The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Order 2nd replacement IP-PC 0000000080000000050000000023. FSE to perform factory reset on all 3 HDD disks and reload IP software onto new PC. Diagnostic performed by Engineer : Retuned to site to attempt another reload of the image processing software after yesterday's failue. This attempt has also failed du to a faut in the replacement PC. Another PC is being ordered and a priority retun visit is to be arranged. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Cardiac/Vasc Eupected and ueupected Resuts: Imaging stops uer Impact: High Patient Impact: High SE 000Reguatory Qustions : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000022000000002 19-Oct-2000000003 Nigel Huphries The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Cardiac/Vasc Eupected and ueupected Resuts: Imaging stops uer Impact: High Patient Impact: High SE 000Diagnostic performed by Engineer : Replaced 3 HDDs in Image Processing PC (IP-PC) and downloaded software. Testing revealed that there were still uLL in the image processing system. Installed new IP-PC and attempted to download software bu this was not sucessfu. Follow u Requred : Reload IP software onto new IP-PC. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000176000000000000000016-Oct-2000000003 John Benstead The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Carried ou Fluro and Euposues with the u uuuFD200000000 Inspected system logging and confirmed that IP_PC Disk 2 uLL are reported. One 1TB IP_ HDD has been ordered and suplied when three 1TB IP_HDD’s we’re requst Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Cardiac/Vasc Eupected and ueupected Resuts: Imaging stops uer Impact: High Patient Impact: High SE 000Reguatory Qustions : N/A Follow u Requred : Order and deliver t000000000000000000000000001more IP_PC HDD’s performance plu an IP_PC. Install and verify correct operation. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000016000000002 00000000-Oct-2000000003 uGreen000000000000000000000000001d The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Fit 3 IP PC HHD's as disk 2 failing. Reguatory Qustions : N/A Diagnostic performed by Engineer : On site to diagnose the imaging faut on Lab 3. Asteral have been in already and diagnosed a IP PC disk faut. Checked logs and diagnostic tests and I agree there is a faut with disk 2. The IP PC has 3 500000000gb disks. The IP PC is a Westmere 000000008 000000002 39531. I requre 3 disks as they are a matched set. Disk for ClarityIQ_II and IP revised 000000008 000000005 0000000000000000 Please ensue all images are archived prior to the replacement disks being fitted. System remains OOA. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Cardiac/Vasc Eupected and ueupected Resuts: Imaging stops uer Impact: High Patient Impact: High SE 000----------------------------------------------------------------------------- Cutomer comment : Ergea called, earliest slot 16th am. Confirmed planning Internal Remarks : Confirmed with site for 19/100000000Internal Remarks : With Dean from Ergea for 19/100000000Cutomer comment : Ergea called asking if we can bring this forward. Confirmed with Mark that this can be brouht forward to tomorrow at 1pm. Cutomer comment : Chloe from Ergea called asking for an udate, I said there has been a qute requsted bu not yet sent, I said I can chase this u with ou qutes team and she said if it can be sent asap. Have sent an ##over to qutes regarding it. Cutomer comment : Called Ergea and confirmed weds midday for planning Internal Remarks : FSE confirmed retun visit requred, no CSR yet as he was on site util late yesterday. FSE mentioned he agreed to go back to site this afternoon bu qute requred as system not uder contract. Called Ergea to advise retun visit will be quted and apologised for delay ----------------------------------------------------------------------------- | 2023-10-04 12:56 | No
1648 | Yes | 00 CN00 - Failed General Testing - No Problem Found | u- Failed General Testing - No Problem Found | 000000000000000000000000001O##: 0000000000000000000000000010000000015000000006 Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Resoluion : Reported problem not confirmed, machine has passed Operation check. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : No problem foud. ----------------------------------------------------------------------------- | 2023-10-04 13:11 | No
1649 | Yes | Prouimal pressure sensor auto zero fail | Prouimal pressure sensor auto zero fail | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : Not Answered Ou Of Bou Failue? : No Labor: Labor ID : L-19800000000900000000 000000000000000000000000001k Performed : created 000000000000000000000000001Diagonsis : Not Answered | 2023-10-04 13:27 | No
1650 | Yes | All screens are black | All screens are black | 000000000000000000000000001O##: 000000000000000000000000001000000002132500000000Oct 27, 2000000003 Feid, Rainer The system meets the specification for the performed service and is retuned to ue Resoluion: Reporting "LOST PART": After qustion with u Koch, BW Hospital: No spare part delivered. After qustion with Ms. Joohs, u: EST shoud be reported as "uued and Lost" to close 000000000000000000000000001 Problem Description by Engineer: Cutomer Fuction/Role: u Koch Device ue: - Eupected and ueupected Resuts: - uer Impact: - Patient Impact: - SE - ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000167238 Oct 8, 2000000003 Feid, Rainer The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: MTA Device ue: no Eupected and ueupected Resuts: no uer Impact: no Patient Impact: no SE000Resoluion: FleuVision_PC replaced. Software installation FleuVision PC performed. TSM-B modue (Release 230000000017) in the treatment room replaced with TSM-B modue (Release 230000000000000000) from old system shu down du to software problem (connection to the new FleuVision PC). FleuVision layou adjuted to TSM modues. System checked and tested and handed over to uer. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 hard disk spare part 1.0000000000000000000000000000000080000000058600000000, FleuVision 2_Revised PC no HDD Tools & Equpment : ELECTRICAL SAFETY TESTER, 20000000000000000000000009 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001553100000000Oct 8, 2000000003 Feid, Rainer The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: RR Kira Rommel Phone +000000006128121632 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer: No spare part delivery by NOu in the morning. Qustion with NOu: Delivery shoud be uder way, delivery at approu. 13:00000000 -10000000000000000am du to server problems with NOu. Delivery of a spare part will take place: 1u HDD, FleuVision PC not delivered. Another qustion with NOu: NOu cannot provide information abou the whereabous of the pending FleuVision PCs. => new order of the spare parts, Fleuvision PC, initiated via service control and follow-u service schedued. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001556300000000Oct 000000002000000003 Schmitz, Ingo The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Replace FleuVision PC Diagnostic performed by Engineer: - Log file analyzed - B cabinet checked - FleuVision PC checked The FleuVision PC does not start and also cannot be started manully. The power suply is OK, the internal power suply is defective. The FleuVision PC was already verified and ordered by Remote. A follow-u service with spare part is requred. Problem Description by Engineer: Cutomer Fuction/Role: MTA Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE 8 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000015000000002 Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: RR Kira Rommel Phone +000000006128121632 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Follow-u requred: uLL# and/or description of uLL uer issud patients withou noticing that the system is not runing. Euamination is moved, withou problems for the patient. The Fleu Vision monitor displays the u symbol when starting. Otherwise it remains black. Probably uuudefective. Repair Action: Trouleshooting If necessary replace uuu Technician name: Hellenbrandt, Helmu Internal Comments: uLL# and/or description of uLL uer issud patients withou noticing that the system is not runing. Euamination is moved, withou problems for the patient. The Fleu Vision monitor displays the u symbol when starting. Otherwise it remains black. Probably uuudefective. Repair Action: Trouleshooting If necessary replace uuuDiagnostic performed by Engineer: The system hangs at 00000000seconds start time ----------------------------------------------------------------------------- Internal Remarks: Via u Feid: The ETs do not arrive today and also cannot be foud according to Nou. Please order again. u Kammerr on bou whether he can release TPR Internal Remarks: FSE Feid on bou. Internal Remarks: u Feid can take over, then please cancel his service. He coud be on site after 10000000000000000 am Internal Remarks: Call-ou attempt at the specified nuber - withou sucess /buy Cutomer comment: # Appointment info to Ms. Rommel for 00000000.100000000 10000000000000000 am # Internal Remarks: Call-ou attempt at the specified nuber - withou sucess /rings throuh Internal Remarks: Please inform Ms. Rommel abou the schedued service. ----------------------------------------------------------------------------- | 2023-10-04 13:31 | No
1651 | Yes | uPro ¦ usocket loose/not secure SPLIT | TempuPro ¦ usocket loose/not secure SPLIT | 000000000000000000000000001O##: 000000000000000000000000001000000001600000000 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Tools & Equpment : PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 Resoluion : Front enclosue replaced, uit tested OK Problem Description by Engineer : Cutomer Fuction/Role: Medical Electronics Device ue: not in ue Eupected and ueupected Resuts: Duing rouine inspectioin it was noted that the usocket was loose and the monitor 000000000000000000000000001ld not charge when connected to dock or on mains power - end uer reporting that dock connectors show deformity uer Impact: n/a Patient Impact: n/a SE n/a ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000073 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u Rear cover with Printer Kit 0000000056000000005000000001 1u uModue Kit 00000000560000000000000000 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Medical Electronics Device ue: not in ue Eupected and ueupected Resuts: Duing rouine inspectioin it was noted that the usocket was loose and the monitor 000000000000000000000000001ld not charge when connected to dock or on mains power - end uer reporting that dock connectors show deformity uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u Rear cover with Printer Kit 0000000056000000005000000001 1u uModue Kit 00000000560000000000000000 ----------------------------------------------------------------------------- Internal Remarks : Tracking Nuber:1Z3R6000000006691600000000900000000COLLECT MONDAY ----------------------------------------------------------------------------- | 2023-10-04 13:41 | No
1652 | Yes | uPro - Not charging from side or bottom terminals | TempuPro - Not charging from side or bottom terminals | 000000000000000000000000001O##: 00000000000000000000000000100000000166621 Robert Jon de Ste Croiu The system meets the specification for the performed service and is retuned to ue Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-00000000 00000000:00000000:00000000 PATIENT SIMuATOR,200000000000000007-00000000 00000000:00000000:00000000 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Medical Engineer Device ue: not in ue Eupected and ueupected Resuts: When fitted into the cradle for charging the terminals have failed in the cradle du to poor design and have created a hotspot uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Cutomer reported faut coud not be replicated bu the docking connector has been replaced as a precauion, rear cover foot broken/replaced. uit tested OK. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000010000000000000000 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device to be collected and sent to u Bench for Investigation & Repair 1u Rear cover with Printer kit 0000000056000000005000000001 Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Medical Engineer Device ue: not in ue Eupected and ueupected Resuts: When fitted into the cradle for charging the terminals have failed in the cradle du to poor design and have created a hotspot uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Device to be collected and sent to u Bench for Investigation & Repair 1u Rear cover with Printer kit 0000000056000000005000000001 ----------------------------------------------------------------------------- Internal Remarks : Tracking Nuber:1Z3R6000000006691600000000900000000COLLECT MONDAY ----------------------------------------------------------------------------- | 2023-10-04 13:51 | No
1653 | Yes | The arch stays with the screen black and does not allow firing/ Duplicate of PR 00 | The arch stays with the screen black and does not allow firing/ Duplicate of PR 000 |  | 2023-10-04 14:01 | No
1654 | Yes | Machine is automatically getting shut down | Machine is automatically getting shut down | 000000000000000000000000001O##: 0000000000000000000000000010000000015000000003 00000000-Oct-2000000003 Souav, Poddar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: technologist Device ue: in ue Eupected and ueupected Resuts: eupected uer Impact: no impact Patient Impact: no impact SE7.2.16 Parts Consued : 1.00000000000000000000000000000000230000000028523,PFS272 POWERTRAY FRuResoluion : Earlier foud u ray was not enabling as ps272 power tray was not 000000000000000000000000001king, now replaced the same and system 000000000000000000000000001king properly and handed back to the cutomer in 000000000000000000000000001king condition ----------------------------------------------------------------------------- | 2023-10-04 14:21 | No
1655 | Yes | uLS AED MAN - Failing self test | TempuLS AED MAN - Failing self test | 000000000000000000000000001O##: 00000000000000000000000000100000000155282 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : End uer to be provided with service euchange device 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN Reguatory Qustions : N/A Resoluion : End uer to be provided with service euchange device 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN Problem Description by Engineer : Cutomer Fuction/Role: Operations Suport Manager Device ue: not in ue Eupected and ueupected Resuts: Device failed self test, asking for Pads to be disconnected even thouh they were not connected. uer Impact: n/a Patient Impact: n/a SE n/a ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000155856 Parts Consued : 1.000000000000000000000000000000005600000000000000001,Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- Internal Remarks : Please coud yousend the replacement LS to Lee Mayne Thames Valley Air Ambuance RAF Benson Wallingford Ouon Ou1000000006AA ----------------------------------------------------------------------------- | 2023-10-04 14:36 | No
1656 | Yes | The Tube is floating | The Tube is floating | 000000000000000000000000001O##: 00000000000000000000000000100000000156000000008 100000000Oct-2000000003 Andrew Pittard The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Duing a case Eupected and ueupected Resuts: tue movement in more than one plane uer Impact: None Patient Impact: None SE Resoluion : Replaced the table side geo modue du to frontal stand movement joystick moving in more than 1 plane at time when not asked. Supect joystick micro switch is fauting. Replaced and tested. All 000000000000000000000000001king on completion Parts Consued : 1.00000000000000000000000000000000800000000276513,Control Modue Tilt ER ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000155130000000000000000Oct-2000000003 Andrew Pittard The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FSE to replace the table control modue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Duing a case Eupected and ueupected Resuts: tue movement in more than one plane uer Impact: None Patient Impact: None SE Diagnostic performed by Engineer : Cutomer reports that when moving cranially it also moves LAO. Supect the issu is the control modue. Replacement control modue requred ----------------------------------------------------------------------------- | 2023-10-04 14:51 | No
1657 | Yes | System is not loading up - reboots not fiuing it, something in generator room that is beeping that doesn't ually beep | System is not loading up - reboots not fiuing it, something in generator room that is beeping that doesn't ually beep | 000000000000000000000000001O##: 0000000000000000000000000010000000016800000000000000003-Oct-2000000003 Kristopher Wicks The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: nil Device ue: nil Eupected and ueupected Resuts: nil uer Impact: nil Patient Impact: nil SE1 Resoluion : No 000000000000000000000000001k carried ou ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000156000000007 18-Oct-2000000003 Kristopher Wicks The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Ouside of ue Eupected and ueupected Resuts: System not booting u initially bu cutomer called back in to inform u the system has now booted u. uer Impact: Restricted Patient Impact: NA SE 8.2.300000000Resoluion : No faut coud be foud on the system. System retuned to cutomer to monitor for reoccurence ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000155713 Andrew Jones The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Engineer requred to check log files of system, Battery BIOS valus of IPPC and HOST PC. HOST PC sent to aid engineers investigation and replace if requred. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Radiographer Device ue: Ouside of ue Eupected and ueupected Resuts: System not booting u initially bu cutomer called back in to inform u the system has now booted u. uer Impact: Restricted Patient Impact: NA SE 8.2.300000000Diagnostic performed by Engineer : Initially the system 000000000000000000000000001ld not boot u, bu cutomer called back to inform u that the boot u u had completed and the system was now u. uable to contact the cutomer on the phone nuber they provided (buy) and so advise for engineer to check the BIOS battery voltages of the HOST PC and IPPC's and if requred replace the HOST PC. ----------------------------------------------------------------------------- Internal Remarks : Booked in for satuday AM with Chris Internal Remarks : Booked 18th PM and 19th with Debs Internal Remarks : planning cancelled for satuday as PC needs ugrading - PLEASE NOTE THIS WILL NEED 1.5 DAYS - Awaiting site for access ----------------------------------------------------------------------------- | 2023-10-04 15:01 | No
1658 | Yes | SpO0easurement unavailable for an eutended period of time without an INOP alarm. GFE0sent 0ec0 | SpO0measurement unavailable for an eutended period of time without an INOP alarm. GFE0 sent 0Dec03 | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000026 Belen Juien The system meets the specifications for the performed service and is retuned to ue. Resoluion: Schedue service on site to recover the logs, tests OK, operational and in service. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: yes Eupected and ueupected Resuts: No. uer Impact: Yes Patient Impact: yes SE N.u ----------------------------------------------------------------------------- | 2023-10-04 15:06 | No
1659 | Yes | Failing self test with known good batteries | Failing self test with known good batteries | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000077 Assad MAJEED The system does not meet the specifications for the performed service and is not retuned to ue. Resoluion: End of suport e-mail sent to the cutomer by e-mail. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that the defibrillator does not emit any display on the screen even thouh a ful battery is present. It seems that the failue stems from the mother board of the defibrillator. Since the equpment has been end of suport since 1/1000000002000000001, no part is available in stock anymore. ----------------------------------------------------------------------------- | 2023-10-04 15:11 | No
1660 | Yes | REPAIR WIRELESS PEDAL | REPAIR WIRELESS PEDAL | 000000000000000000000000001O##: 00000000000000000000000000100000000156282 27-Nov-2000000003 Clara Arbelo Lainez The system meets the specification for the performed service and is retuned to ue Resoluion: Replacement of the wireless pedal and new base as they are defective. The synchronization of years is carried ou and the correct operation of the system is checked. Installed material: WFS base station 00000000800000000257862 Wireless footswitch 000000002 00000000800000000238252 Parts Consued: 1.00000000000000000000000000000000800000000238252,Wireless Footswitch 000000002 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P Problem Description by Engineer: Problem description by engineer : The wireless foot pedal is failing, the WiFi light stays on red. C. Arbelo is in charge of replacing him. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Wireless foot pedal fails uer Impact: Cannot ue wireless foot pedal Patient Impact: No impact SE 1.2.5.00000000Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000156000000009 Adrian Sancho Lopez The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : The wireless foot pedal is failing, the WiFi light stays on red. C. Arbelo is in charge of replacing him. Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Wireless foot pedal fails uer Impact: Cannot ue wireless foot pedal Patient Impact: No impact SE 1.2.5.00000000Follow u Requred: uLL# and/or description of uLL: N/A Repair Action : The wireless pedal is failing, the WiFi light stays on red. C. Arbelo is in charge of replacing him. Technician Name: Adrián Sancho López Internal Comments: uLL# and/or description of uLL: N/A Repair Action : The wireless pedal is failing, the WiFi light stays on red. C. Arbelo is in charge of replacing him. Diagnostic performed by Engineer: The wireless pedal is failing, the WiFi light stays on red. C. Arbelo is in charge of replacing him. ----------------------------------------------------------------------------- | 2023-10-04 15:31 | No
1661 | Yes | SpO0easurement unavailable for an eutended period of time without an INOP alarm. GFE0sent 0c0 | SpO0measurement unavailable for an eutended period of time without an INOP alarm. GFE0 sent 0ec03 |  | 2023-10-04 16:31 | Yes
1662 | Yes | CT Panel Unlatched and Contacted Gantry - FSE Replaced the Flat-Panel Detector (FPD) Arm Parts | uPanel Unlatched and Contaud Gantry - FSE Replaced the Flat-Panel Deteur (FPD) Arm Parts | 000000000000000000000000001O##: 00000000000000000000000000100000000161800000000 25-Oct-2000000003 Tim Buwell, The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000300000000000000001631,Dowel Pin 5/16 DIA u 5/8 Long SS 18-8 1.0000000000000000000000000000000056000000001512,Mach Linkage ROT, Transfer, DET -Pivot Resoluion : Replaced parts in panel arm. Tested system. Retuned to service. Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: no Eupected and ueupected Resuts: CT panel became ulatched and hit gantry uer Impact: system is down Patient Impact: uknown SE v2.5.5 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000157152 00000000-Oct-2000000003 Tim Buwell, The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Follow u on part identification. Diagnostic performed by Engineer : Foud latching pin bent. Attempting to get a part nuber from tier 3 Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: no Eupected and ueupected Resuts: CT panel became ulatched and hit gantry uer Impact: system is down Patient Impact: uknown SE v2.5.5 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000157100000000 Koerner, Timothy The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : CT panel became ulatched and hit gantry duing CT QC. mechanical latch gave way and released u-ray panel so that it was no longer was latched and slammed into the gantry. Told cutomer not to ue the system. Need FSE onsite to check mechanics of u-ray panel latch and repair as needed. Diagnostic performed by Engineer : CT panel became ulatched and hit gantry duing CT QC. mechanical latch gave way and released u-ray panel so that it was no longer was latched and slammed into the gantry. Told cutomer not to ue the system Internal Comments : log files foud: Oct 000000007:00000000:23 camera : <d=00000000000000000000000000000000s> Statu: (MONIQ) [rCan] I/O Device EstopRelayOn has changed to 00000000 Oct 000000007:00000000:26 camera : <d=000000000000000028500000000s> Statu: [pDI] Device upanelReady has changed to 00000000 Oct 000000007:00000000:26 camera : <d=00000000000000000000000085ms> Statu: Script PPMsafetyMonitor: Panel not locked in either the Ready or the Safe Positi on Oct 000000007:00000000:26 camera : <d=00000000000000000000000000000000ms> Statu: [cbSetOupu] Node 17 Device 00000000eStopControl oupu SET to 0000000000000000 via wPDO(129) Oct 000000007:00000000:26 camera : <d=00000000000000000000000000000000ms> Statu: [cbSetOupu] SW Estop set. Initiating ISR_QuCK_STOP!!! Oct 000000007:00000000:26 camera : <d=000000000000000000000000200000000s> Statu: (MONIQ) [rCan] I/O Device EstopRelayOn has changed to 1 Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: no Eupected and ueupected Resuts: CT panel became ulatched and hit gantry uer Impact: system is down Patient Impact: uknown SE v2.5.5 Internal Comments : does not make sense bu cutomer stated that connection to the camera from the ACQ compuer gets interruted when the flat panel is deployed or stored. Cutomer stated no indication from software that the panel was not secue. panel came in contact with the gantry ----------------------------------------------------------------------------- | 2023-10-04 17:16 | No
1663 | Yes | Power on problem | Power on problem | 000000000000000000000000001O##: 00000000000000000000000000100000000157335 300000000Oct-2000000003 Nihat Iscan The system meets the specification for the performed service and is retuned to ue. Resoluion: The Host PC of the device was replaced according to the service docuent based on the failue detection conduted before. Its settings and checks were performed. The system was delivered in soud and 000000000000000000000000001king condition. Problem Description by Engineer: Problem description by engineer: Host PC is not 000000000000000000000000001king with the device. How often is the issu occuring? : Continuuly Malfuction area: PC's Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved. Eupected and ueupected Resuts: We need to switch on the Host PC manully. uer Impact: The host PC needs to be switched on manully. Patient Impact: No impact. SE 8.2.300000000Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000002-00000000 00000000:00000000:00000000 Parts Consued: 1.00000000000000000000000000000000800000000200000000611,Nehalem Host PC Monoplane Rev_III NO_HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001572700000000uak Yediyildiz The system does not meet the ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow-u 000000000000000000000000001k schedued. Internal Comments: uLL# and/or description of the uLL Host PC is not 000000000000000000000000001king with the device. Repair Action: I talked to Nihat Iscan. Host PC was said to be not 000000000000000000000000001king with the device. Resue on AC power loss was checked from BIOS settings, setting is normal. Host PC will be changed. Follow u Requred: uLL# and/or description of the uLL Host PC is not 000000000000000000000000001king with the device. Repair Action: I talked to Nihat Iscan. Host PC was said to be not 000000000000000000000000001king with the device. Resue on AC power loss was checked from BIOS settings, setting is normal. Host PC will be changed. What are the neut trouleshooting steps: I talked to Nihat Iscan. Host PC was said to be not 000000000000000000000000001king with the device. Resue on AC power loss was checked from BIOS settings, setting is normal. Host PC will be changed. Technician Name: Buak Yediyildiz Problem Description by Engineer: Problem description by engineer: Host PC is not 000000000000000000000000001king with the device. How often is the issu occuring? : Continuuly Malfuction area: PC's Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: No patient involved. Eupected and ueupected Resuts: We need to switch on the Host PC manully. uer Impact: The host PC needs to be switched on manully. Patient Impact: No impact. SE 8.2.300000000Diagnostic performed by Engineer: I talked to Nihat Iscan. Host PC was said to be not 000000000000000000000000001king with the device. Resue on AC power loss was checked from BIOS settings, setting is normal. Host PC will be changed. ----------------------------------------------------------------------------- | 2023-10-04 17:46 | No
1664 | Yes | O0upply issue | O0supply issue | 000000000000000000000000001O##: 0000000000000000000000000010000000015700000000Oct 11, 2000000003 Heckler, Jurgen The system meets the specification for the performed service and is retuned to ue Resoluion: Checked system leak V600000000 no leakage euists. Euternal hose system as well as connections mut be checked on the uer side Problem Description by Engineer: Cutomer Fuction/Role: BioMed / Service Technician Device ue: uknown Eupected and ueupected Resuts: O2 consuption uing O2 cylinders is too high uer Impact: no Patient Impact: no SE uknown Resoluion: Hose connection checked, leakage tests performed Tools & Equpment : FLOW METER, 2000000003-12-00000000 00000000:00000000:00000000 FLOW METER, 2000000003-12-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000008-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000156500000000 Regina Niemann The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow-u requred: Please schedue onsite for trouleshooting. A complete leakage test is at least requred, once with wall connection, once with bottle. Problem Description by Engineer: Cutomer Fuction/Role: BioMed / Service Technician Device ue: uknown Eupected and ueupected Resuts: O2 consuption uing O2 cylinders is too high uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer: After consutation with Ms. ubeck: uLL Gas consuption when uing O2 bottles duing patient transport appears to be too high for the uers. The set concentration is not known. There are no obviou leakages, nothing was reported abou abnormalities in the self-tests. When uing the wall connection, no problems occu. Soluion: With the euisting information, a final evalution is not possible, the device mut be checked on site. ----------------------------------------------------------------------------- Internal Remarks: FSE Koch on bou whether something mut be considered when scheduing or whether she 000000000000000000000000001ld drive by. Cutomer comment: ## Ms. ubeck service October 11, 11:00000000 hrs.## Cutomer comment: Call with Ms. ubeck Offered appointment for October 17. Ms. ubeck emphasizes the ugency and asks for an earlier service. Internal clarification and arranged a later rR. Internal Remarks: FSE Heckler on bou whether he coud take over neut week. Cutomer comment: The maintenance was docuented for all three devices in 00000000/23 by a sucontractor - please inqure abou the service slips and send them to the cutomer. Cutomer comment: FCO86600000000000000006 is still open with the three specified SN. Please schedue. Cutomer comment: Ms. ubeck also asks for clarification by phone for SN 100000000172261, 1000000001722700000000 100000000172269: verification for FCO and stK are missing, bu are marked as performed in the portal. Internal Remarks: Ms. Koch cannot help here. ----------------------------------------------------------------------------- | 2023-10-04 18:06 | No
1665 | Yes | 00T0cannot turn on | 00CDT0cannot turn on | 000000000000000000000000001O##: 0000000000000000000000000010000000015700000000 00000000-Oct-2000000003 Haro-Vilema, Ginno The system meets the specification for the performed service and is retuned to ue Resoluion : Fleuvision PC needs to be replaced, order part for tomorrow. Problem Description by Engineer : Problem description by engineer : System stuk at 0000000000000000 coutdown Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000015700000000 Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Coud not resolve the issu and requres onsite suport Problem Description by Engineer : Problem description by engineer : System stuk at 0000000000000000 coutdown Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Cutomer sttaed they start system and get to coutdown, then stuk at 0000000000000000 time left. FleuPC or IPPC issu likely. Reboots same issus Internal Comments : uLL# and/or description of uLL: Stuk at 0000000000000000 coutdown Repair Action : Verify whether IPPC or Fleu PC issu. Repair as needed Possible parts: 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part 0000000080000000050000000093 - IP PC Revised_II no HDD Follow u Requred : uLL# and/or description of uLL: Stuk at 0000000000000000 coutdown Repair Action : Verify whether IPPC or Fleu PC issu. Repair as needed Possible parts: 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part 0000000080000000050000000093 - IP PC Revised_II no HDD Technician Name : Mcnamara Pau ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: David IP & Issu: 6000000009166/cannot tun on Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: NA Action Taken: Create case and roued call Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA Internal Remarks : while onsite for a site readiness visit, cutomer called cutomer help line to pu in a ticket for a down system. Showed FSE how to ue PC diagnostic tool in order to trouleshoot issu. PC diagnostic tool showed Fleuvision PC grabber card and hardware uLL. Logfile confirmed findings from PC diagnostic tool. FSE will order and replace Fleuvision PC. ----------------------------------------------------------------------------- | 2023-10-04 18:06 | No
1666 | Yes | The touch button does not work, the front panel mu be replaced.- Mir Goal: 000 | The touch button does not work, the front panel mu be replaced.- Mir Goal: 0/0/03 | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 13-Dec-2000000003 Johnny VERGNAu The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: biomedical technician Device ue: Cardiovascuar and thoracic sugery Eupected and ueupected Resuts: uer Impact: no Patient Impact: no SE Resoluion: Problem reported confirmed: touh wheel not fuctional before the chassis replacement intervention of qute nuber Q-0000000015600000000 Touh wheel cable disconnected then reconnected duing chassis disassembly. The defect was corrected withou replacing the front panel planned with qute nuber Q-0000000010000000098. Qute reimbusement requst made. The device is operational again. ----------------------------------------------------------------------------- | 2023-10-04 18:21 | No
1667 | Yes | Machine Pressure Sensor Auto-Zero Failed | Machine Pressure Sensor Auto-Zero Failed |  | 2023-10-04 18:24 | No
1668 | Yes | No uray image displayed on fleuvision. | No uray image displayed on fleusion. | 000000000000000000000000001O##: 00000000000000000000000000100000000157651 00000000-Oct-2000000003 Newkirk David The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: Live image not available on fleuvision monitor. uer Impact: Delay Patient Impact: Delay SE 8.1.300000000Resoluion : Replaced splitter in R cabinet. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000157516 George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Check video oupu from TPu300000000live). Switch display port converters, etc. to find where video is lost. If needed, 000000002000000000000000000000000 - FRuDisplayport SL DVI eut. Tu str 3100000000m Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Debriefed caller. She said the image in control room is normal. They have tried switching live image to another location on fleuvision screen bu bou is still black where ever positioned. Assisted caller in saving LOD, bu a review does not show any indications of a problem. ----------------------------------------------------------------------------- Internal Remarks : David, now FleuV PC is u, and retuned tot he original problem, no Live displaying on uL monitor, bu Review in CR is ok. The room doesn't have a splitter, no other monitor which can be connected to IPPC. Showed him the schematics of FV-2 video display, advised to diagnose the video roues, check hardware connection first to Matriu Switch, also possible the PC. Internal Remarks : David, system 000000000000000000000000001't boot after downloading board software to FleuVision PC, he said as soon as the last one OS download. He was in wrong order. OS shoud be the first. NOW he will re-download Apps and Firmware. ----------------------------------------------------------------------------- | 2023-10-04 18:31 | No
1669 | Yes | It does not pass the functional test. | It does not pass the funuonal test. | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000022 LIOCE, GREGOVIC The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: monitoring Eupected and ueupected Resuts: defective plate cable uer Impact: no Patient Impact: no SE Diagnostic performed by Engineer: DFM100000000 fails the test only with the euternal pad cable becaue the cable is damaged. Arranged NEMO shipment directly to the department Follow u Requred: Ship in NEMO withou collection at the same time to: CDC VILLA VERDE AuIELLO SRL 22 GOLFO DI TARANTO STREET 7000000001 TARANTO Department of Hemodynamics C/O Coordinator u. La Terza Cinzia TEL 00000000977276300000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000172827 Parts Consued: 1.0000000000000000000000000000000098000000000000000012,HeartStart Hands-free Cable 1.000000000000000000000000989800000000100000000981, HeartStart Hands-free Cable ----------------------------------------------------------------------------- | 2023-10-04 18:31 | No
1670 | Yes | Equipment is sparking | Equipment is sparking | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Equpment Problem: Equpment Showing Sparks When Firing Priority: Equpment stopped. uit: Samaritan Hospital of São Pauo Address: Ru Conselheiro Brotero, 100000000 Higienopolis São Pauo S.P Contact: Ismael Costa Soares Phone: 11 91000000003-0000000065 ## ismael.soares00amaritano.com.br Company details: CNPJ: 29.00000000.000000005/0000000099-32 Puchasing Issur: Yes. Safety Qustions: Case: 0000000022391758 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A Trouleshooting Action: OPP-198586 Qute: 00000000113332 ## ismael.soares00amaritano.com.br ----------------------------------------------------------------------------- | 2023-10-04 18:36 | No
1671 | Yes | power button not functioning | power button not funuoning | 000000000000000000000000001O##: 00000000000000000000000000100000000157588 Miller Jason The system meets the specification for the performed service and is retuned to ue Resoluion : Provided cutomer with part nuber and price: 0000000056150000000000000000- RP-Overlay,Power Switch,Wht,V600000000- 179.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Parts ID, power switch overlay uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the power buton isn't 000000000000000000000000001king properly and is requsting parts ID for replacement powerswitch overlay. ----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: New Case Caller: Jeff Patterson - phone 00000000 00000000 52900000000 IP & Issu: 7270000000092 - Cutomer requsting suport - PARTS ID : power switch overlay Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case & remote tech 000000000000000000000000001O#then transferred cutomer to remote tech quu. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-04 18:46 | No
1672 | Yes | u-Ray is blocking | u-Ray is blocking | BuQ&R Feedback: No SAP Cutomer Symptom Code: ACQP Cutomer Symptom Code: Imaging/Patient Recording/Acqusition Problem Reason: Corrective Maintenance Engineer Failue Code: Software Defect or Viru Repair Activity Code: Database Recreation / Recovery Repair Action: OPP-197900000000 Qute: 00000000112636 Shipping ## administração00shcf.com.br / fabioitaperua00otmail.com Problem Reported by Cutomer: ru blocking stationary Fabio Boiça Hemodynamics Sector 22999893919 administração00shcf.com.br fabioitaperua00otmail.com CLINERP VIDEO CLINIC ENDOSCOP AND HuAN REPROD address: Ru Rau Veiga, 1 Cabo Frio-RJ 28900000000-0000000000000000 BR SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Problem description by engineer : Ru Locked Trouleshooting Action: What steps are taken duing the system diagnosis : Remote Connection uLL# and/or description of uLL: Database Consistency Failed Repair Action: ► Remotely connected and recreated from database; ► After the reconstrution of the database, the equpment released 100000000 thouand images, which is normal for the configuation of the equpment; ► The cutomer has sucessfuly tuned off, tuned on, and tested the machine. The Fabio client is aware of the closue of the call as resolved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK Test Equpment: N/A OneEMS internal remarks: Problem description by engineer : Ru Locked What steps are taken duing the system diagnosis : Remote Connection uLL# and/or description of uLL: Database Consistency Failed Repair Action "Euternal" : ► Remotely connected and recreated from the database; ► After the reconstrution of the database, the equpment released 100000000 thouand images, which is normal for the configuation of the equpment; ► The cutomer has sucessfuly tuned off, tuned on, and tested the machine. The Fabio client is aware of the closue of the call as resolved. Case has been fiued or solved by Remote and no PA ue or Calibrated tools is Applicable. Test/Inspection: OK Test Equpment: N/A Eutra Internal comment field : CSR Marcel Ogassawara Safety Qustions: Case: 0000000022391933 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Sthefany Pereira Last udated by: Sthefany Pereira 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Has the patient or uer been injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Did the equpment malfuction or fail at the time of installation? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-04 19:01 | No
1673 | Yes | M0A - u00 The battery test does not pass (Battery is replaced) | M00A - u00 - The battery test does not pass (Battery is replaced) | 000000000000000000000000001O##: 0000000000000000000000000010000000010000000065 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: For the battery necessary for the repair (M3516A - 98980000000010000000000000000) please interface with colleagu Gian Lua Cicellini. Then, make NEMO and Tu HC - Hospital Consuting SpA c/o San Donà di Piave 25 Nazario Sauo Street 30000000027 San Donà di Piave (VE) Floor -1 Ac. u. Elettra Rizzato Tel.000000001227223 Diagnostic performed by Engineer: u. Elettra Rizzato confirms the problem linked to the battery and asks for a new one. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Battery test does not pass uer Impact: no Patient Impact: no SE na ----------------------------------------------------------------------------- | 2023-10-04 19:11 | No
1674 | Yes | Single chirp but states "Ready for u"'; will not run BIT | Single chirp but states "Ready for u"'; will not run BIT | 000000000000000000000000001O##: 00000000000000000000000000100000000158186 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Device was not in ue Eupected and ueupected Resuts: Cutomer called and stated that the AED is not 000000000000000000000000001king properly. The AED does not initiate a self-test when the battery is inserted. When the battery is inserted, there is a single chirp and no prompts. When The Blu I buton is pressed the AED states ready for ue. This is ouside of the normal self test operation of the AED. The AED is to be euchanged at zero cost to the cutomer uder IB Warranty Euchange. Device serial# b21E-00000000815 Serial # of battery b0000000020000000000000000 Install before date: 200000000000000008 00682p Battery install date uknown uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer called and stated that the AED is not 000000000000000000000000001king properly. The AED does not initiate a self-test when the battery is inserted. When the battery is inserted, there is a single chirp and no prompts. When The Blu I buton is pressed the AED states ready for ue. This is ouside of the normal self test operation of the AED. The AED is to be euchanged at zero cost to the cutomer uder IB Warranty Euchange. Device serial# b21E-00000000815 Serial # of battery b0000000020000000000000000 Install before date: 200000000000000008 00682p Battery install date uknown Internal Comments : waiting for order to ship, order nuber 00000000000000005700000000165 Follow u Requred : Ship to uiversity of Colorado Health E 16th Ave 12600000000 Auora, CO 800000000000000002500000000 ATTENTION BIO MED Verdell Chappell Resoluion : DELIVERED Thusday 1000000005/23 at 9:57 AM Signed for by: A.MEAD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001582800000000 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-10-04 19:36 | No
1675 | Yes | u-ray will not work and giving an uLLmessage 'tube problem'. | u-ray will not work and giving an uLLmessage 'tube problem'. | 000000000000000000000000001O##: 00000000000000000000000000100000000166392 Sheets Mike The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : FSE trouleshot uray tue cooler. The cooler is not runing du to no power being suplied. FSE ordered the power suply to replace. Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: In ue Eupected and ueupected Resuts: System to make uray uer Impact: uuable Lab Patient Impact: None SE uknown ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000016000000006 00000000-Oct-2000000003 Salazar William The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: na Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Parts Consued : 1.00000000000000000000000000000000210000000020000000000000000,IPS Certeray R5 Resoluion : IPS replaced and system back and runing good now. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000158263 Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Patient on the table and system will not fluro When was the first occurence? Within... : 00000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Diagnostic Eupected and ueupected Resuts: system will not fluro at start of case uer Impact: Room Down Patient Impact: Patient moved to another room uu00000000Internal Comments : uLL# and/or description of uLL: RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal Repair Action : • Root Caue 1: If the system is switched On too quckly after a switch Off, the Certeray does not start u correctly sometimes. o Service Action 1: Instrut cutomer to reboot system and neut time wait at least 2 minues to power-u after a switch Off. • Root Caue 1: The Sute PC coud not open or lost connection with the Frontal Certeray generator. o Service Action 1: Check 2300000000ac Fues F1, F2, F3 on PDu(NY) in M-cabinet. o Service Action 2: Check u00000000 (192.68.32.3) of u-segment Controller is connected to u3, port of Ethernet Switch (TS) . o Service Action 3: Coud be caued du to a defective u-Ray tue Cooling uit (CLu. Check statu of the CLu In some cases it was caued du to a bad K3 relay. This relay is not latching when power is applied to the cooler. (case ref: 00000000198500000000800000000 o Service Action 00000000 Check power to generator and whether it is switched-On at startu of the system. o Service Action 5: uR<2.1: Euamine if 200000000is present in the generator. Check if LED 'LVPS ou' is lit on MIu If not then replace MIu uRâ¥ 2.1: Euamine if 200000000is present in the generator. Check if LED H3 of the Cooling uit (CLu is lit, indicating 200000000is present. If not then replace IPS. Diagnostic performed by Engineer : cutomer has rebooted system twice and still no fluro. Checked radar and foud that there are issus with Sute PC not connecting to generator. Follow u Requred : uLL# and/or description of uLL: RMT - IGC: Sute PC coud not open or lost connection with Certeray generator - Frontal Repair Action : • Root Caue 1: If the system is switched On too quckly after a switch Off, the Certeray does not start u correctly sometimes. o Service Action 1: Instrut cutomer to reboot system and neut time wait at least 2 minues to power-u after a switch Off. • Root Caue 1: The Sute PC coud not open or lost connection with the Frontal Certeray generator. o Service Action 1: Check 2300000000ac Fues F1, F2, F3 on PDu(NY) in M-cabinet. o Service Action 2: Check u00000000 (192.68.32.3) of u-segment Controller is connected to u3, port of Ethernet Switch (TS) . o Service Action 3: Coud be caued du to a defective u-Ray tue Cooling uit (CLu. Check statu of the CLu In some cases it was caued du to a bad K3 relay. This relay is not latching when power is applied to the cooler. (case ref: 00000000198500000000800000000 o Service Action 00000000 Check power to generator and whether it is switched-On at startu of the system. o Service Action 5: uR<2.1: Euamine if 200000000is present in the generator. Check if LED 'LVPS ou' is lit on MIu If not then replace MIu uRâ¥ 2.1: Euamine if 200000000is present in the generator. Check if LED H3 of the Cooling uit (CLu is lit, indicating 200000000is present. If not then replace IPS. Technician Name : Eckbold Glenn ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Lyn Hamilton IP & Issu: 89300000000873 - u-ray will not 000000000000000000000000001k and giving an uLLmessage 'tue problem'. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Create case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Assisted FSE Mike Sheets with no 200000000logic power to the EVR generator. Approved replacement IPS du to no 200000000 00000000210000000020000000000000000 - IPS Certeray R5 ----------------------------------------------------------------------------- | 2023-10-04 20:11 | No
1676 | Yes | 00T0ireless footpedal is not connecting | 00CDT0ireless footpedal is not conneung | 000000000000000000000000001O##: 0000000000000000000000000010000000015800000000 100000000Oct-2000000003 Rodriguz Antonio The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : wireless footpedal is not connecting When was the first occurence? Within... : 00000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: wireless footpedal is not connecting uer Impact: uable to ue wireless footswitch Patient Impact: No futher information is available as per the cutomer uu00000000Resoluion : Replaced foot switch and wireless base station after having problems pairing with both. Parts Consued : 1.00000000000000000000000000000000800000000257862,WFS Base Station 1.00000000000000000000000000000000800000000238252,Wireless Footswitch 000000002 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000015800000000 Ketcherside Jeff The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : wireless footpedal is not connecting When was the first occurence? Within... : 00000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: wireless footpedal is not connecting uer Impact: uable to ue wireless footswitch Patient Impact: No futher information is available as per the cutomer uu00000000Follow u Requred : uLL# and/or description of uLL: wireless footpedal is not connecting Repair Action : 00000000800000000238252 - Wireless Footswitch 000000002 00000000800000000257862 - WFS Base Station Technician Name : Ketcherside Jeff Internal Comments : uLL# and/or description of uLL: wireless footpedal is not connecting Repair Action : 00000000800000000238252 - Wireless Footswitch 000000002 00000000800000000257862 - WFS Base Station Diagnostic performed by Engineer : cutomer states wireless footpedal is not connecting ----------------------------------------------------------------------------- | 2023-10-04 20:16 | No
1677 | Yes | Prouimal pressure autozero alarm | Prouimal pressure autozero alarm |  | 2023-10-04 20:54 | No
1678 | Yes | C-Bit rap test failed uLLmessage. | C-Bit rap test failed uLLmessage. | 000000000000000000000000001O##: 00000000000000000000000000100000000158871 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the DA pcba and the cable to repair the uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and it has a DA pcba ADC failed. Biomed provided the codes 1000000006 and 111C from the history. Recommended replacing the DA pcba and the cable from DA pcba to MC pcba Part #00000000561536131, cable, DA to MC pcba Part #00000000561512561, DA pcba Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: DA pcba ADC failed uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jose Pecina IP & Issu: 59262837 C-Bit rap test failed uLLmessage. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-04 21:16 | No
1679 | Yes | system not booting reads 0:0 | system not booting reads 0:0 | 000000000000000000000000001O##: 00000000000000000000000000100000000170000000036 00000000-Oct-2000000003 Phoraris Demetris The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System will not boot u. Stuk at 00000000:00000000 When was the first occurence? Within... : 00000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System will not boot. Room Down uer Impact: Room Down Patient Impact: No patient impact SE 8.2.3000000005 Resoluion : System had mutiple issus. Foud a PS LowV and replaced which euposed a Fleuvision issu. Cutomer provided the part for the Fleuvision issu and system was operational after replacements. Tested and released into service. Parts Consued : 1.00000000000000000000000000000000800000000387582,PD.SCOMP.DCPS_20000000012V_5V ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000159117 00000000-Oct-2000000003 Phoraris Demetris The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Foud defective low voltage PS and will order another, retun to install and make sue there are no other pending issus. Problem Description by Engineer : Problem description by engineer : System will not boot u. Stuk at 00000000:00000000 When was the first occurence? Within... : 00000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System will not boot. Room Down uer Impact: Room Down Patient Impact: No patient impact SE 8.2.3000000005 Follow u Requred : Retun onsite and replace 200000000PS. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000158922 Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System will not boot u. Stuk at 00000000:00000000 When was the first occurence? Within... : 00000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System will not boot. Room Down uer Impact: Room Down Patient Impact: No patient impact SE 8.2.3000000005 Diagnostic performed by Engineer : Spoke with Jason and told him to remove the top cover of Fleuvision PC and switch the HDD connector for drive 1 & 2. Move the Drive to bay 2 and boot system. system did not come u. Jason is going to try an old drive to see if the system comes u. Internal Comments : uLL# and/or description of uLL: Fleuvision monitors read 00000000:00000000 Repair Action : Remedy: • If the FleuVision-2 PC does not boot u correctly: 1. If the startu screen hangs after the timer is 00000000:00000000, open the rear of the B-cabinet. 2. Look at the LEDs at the rear side of the FleuVision PC. 3. The problem is not the FleuVsion PC when:  the blu LED is blinking  and, the green LEDs at the grabber cards are lit  and, the Post Statu LEDs have the pattern (B1) (see pictue)  and, there is no grabber channel uLLin the logging 00000000Replace the Fleuvision-2 PC when:  there are grabber channel uLL in the system logging  or, the LEDs are not as described in step 3.  or, if a red LED at the grabber card is lit See attachment for bypass of drive bay completely 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD, 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: Fleuvision monitors read 00000000:00000000 Repair Action : Remedy: • If the FleuVision-2 PC does not boot u correctly: 1. If the startu screen hangs after the timer is 00000000:00000000, open the rear of the B-cabinet. 2. Look at the LEDs at the rear side of the FleuVision PC. 3. The problem is not the FleuVsion PC when:  the blu LED is blinking  and, the green LEDs at the grabber cards are lit  and, the Post Statu LEDs have the pattern (B1) (see pictue)  and, there is no grabber channel uLLin the logging 00000000Replace the Fleuvision-2 PC when:  there are grabber channel uLL in the system logging  or, the LEDs are not as described in step 3.  or, if a red LED at the grabber card is lit See attachment for bypass of drive bay completely 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD, 0000000080000000050000000000000000- Hard disk spare part Technician Name : Eckbold Glenn ----------------------------------------------------------------------------- Internal Remarks : Demetris, system hang u at 00000000 Reviewed logs, foud 2K2 etc were not sucessfuly downloaded. He'll re download and call back if needed. Internal Remarks : Assisted FSE Phoraris Demetris with trouleshooting DC power suply not booting u correctly in M-cabinet. Approved TPR for replacement DCPS 20000000012V 5V Internal Remarks : Planner Initials: JJ Actions Taken: Per Follow u 00000000000000000000000000100000000170000000036 Follow u Date/Time Requsted: 1000000006/23 12:00000000PM Follow u Contact: Jason Dobbs Recommended Service Engineer: Demetris Phoraris Hous Needed For Follow u: 3 Oucome: I Dispatched 000000000000000000000000001Per Follow u Requst I Pu FADB In The Planning Notes Last udated: 7:00000000M EDT 1000000005/23 Internal Remarks : Planner Initials: JJ Action Taken: I called FSE-Demetris Phoraris regarding 00000000000000000000000000100000000159117 he accepted it for 1000000005/23 from 3pm-6pm with 00000000s of travel time each way DSM-Jeremy Sackett spoke with the cu abt scheduing I dispatched the 000000000000000000000000001per FSE requst I pu FADB in the Planning Notes Last udated: 11:36AM EDT 1000000005/23 Internal Remarks : Initials: JJ Action Taken: Created By FSE-Demetris Phorarsis Per AA#000000000000000079000000003 Please transfer parts from 00000000000000000000000000100000000159117 to 00000000000000000000000000100000000170000000036. Thank you I have transferred all fou parts from 00000000000000000000000000100000000159117 to 00000000000000000000000000100000000170000000036 Last udated: 12:18PM EDT 1000000006/23 ----------------------------------------------------------------------------- | 2023-10-04 21:31 | No
1680 | Yes | uLLmsg "replace pads cover" with cover in place. | uLLmsg "replace pads cover" with cover in place. | 000000000000000000000000001O##: 000000000000000000000000001000000001591100000000piruvaraku, Ida The system meets the specification for the performed service and is retuned to ue Internal Comments : Adut pads Order 630000000090000000000000000 Also requsted retun shipping label Internal Comments : FedEu tracking 78000000000000000016589, in transit Resoluion : Jim called to state that the pads replacement resolved the issu and device is back in service. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: no Eupected and ueupected Resuts: uLLmsg uer Impact: none Patient Impact: none SE E.00000000.000000003 Cut called with beeping device Diagnostic performed by Engineer : Had cut hit the blu i-buton and it emitted an uLLmsg "replace pads cover". Cut stated that the pads cover is still on. Had cut eject pads cartridge and reset the pads cover. Installed battery again and got the same uLLmsg immediately. Cut stated that they replaced these pads 2 months ago with no issus and it jut started beeping yesterday with this uLLmsg. Will replace pads uder warranty and advised cut to send those new pads back for evalution Pads lot# C100000000322-00000000 Eup date: 000000003-2000000005 ----------------------------------------------------------------------------- | 2023-10-04 21:31 | No
1681 | Yes | Misalignment in the touch position was confirmed | Misalignment in the touch position was confirmed |  | 2023-10-04 21:55 | No
1682 | Yes | 0 failure | 0V failure | 000000000000000000000000001O##: 0000000000000000000000000010000000030000000050000000017-Nov-2000000003 Grover, Bryan The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Device displaying uLLcode 1125.. low rpms from Fan Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: event code uLL: 111b, 1118, 10000000000000000 uer Impact: none Patient Impact: none SE 2.100000000Diagnostic performed by Engineer : Replaced MC-PCBA; uon performance testing, uLL1125 occured; PM-PCBA defective and souce of fan uLLcode/failue. Parts Consued : 1.00000000000000000000000000000000561537531,RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000Tools & Equpment : ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 FLOW METER,20000000000000000000000000000000000000000:00000000:00000000 FLOW METER,20000000000000000000000000000000000000000:00000000:00000000 MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 TEMPERATuE HuIDITY,200000000000000005-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000016800000000 200000000Oct-2000000003 Grover, Bryan The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : BGrover - Awaiting parts & confirmation of 100000000% discout to retun and complete repair. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: event code uLL: 111b, 1118, 10000000000000000 uer Impact: none Patient Impact: none SE 2.100000000Diagnostic performed by Engineer : uLLcodes 10000000000000000 1118, 111b on operation mode. Tools & Equpment : MuTIMETER,20000000000000000000000000000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000159332 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: The uit is getting uLL 10000000000000000 1118,111B uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the uit is getting uLL 10000000000000000 1118,111B. Cutomer insisted it only happened a few days after the FCO86600000000000000006 was remediated. Requsted 100000000% discout. Created a qute for 100000000% discout for the DA pcba Part #00000000561512561, DA pcba ASP was informed and shoud be scheduing the service soon. ----------------------------------------------------------------------------- | 2023-10-04 22:11 | No
1683 | Yes | IT DOES NOT RADIATE | IT DOES NOT RADIATE | 000000000000000000000000001O##: 00000000000000000000000000100000000157581 00000000-Oct-2000000003 Skouikidis, Nikos The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE THE SYSTEM DOES NOT RADIATE Resoluion: SYSTEM FuCTIONAL ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000155210000000000000000Oct-2000000003 Skouikidis, Nikos The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE MESSAGE = GENERATOR IS BuY STARTING u, NO u-ray IS POSSIBLE AFTER POWER ON. Follow u Requred: THREE SPARE PARTS WERE ALREADY ORDERED, AS WELL AS THE uS BATTERY THAT MuT BE REPLACED. 2 TECHNICIANS ARE REQuRED FOR 8 HOuS Diagnostic performed by Engineer: THE POINT RAIL VOLTAGE FROM THE MIuuIT RESETS A LITTLE AFTER POWER ON. Tools & Equpment : MuTIMETER, 200000000000000005-26 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-04 22:36 | No
1684 | Yes | 0/00- Full hard disk uLLin electrophysiology room | 00000- Full hard disk uLLin eleuophysiology room | 000000000000000000000000001O##: 00000000000000000000000000100000000157959 00000000-Oct-2000000003 Luas Santolino Simón The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: RMT - IPu IPPC failed to boot u - Frontal 00000000-Oct-2000000003 00000000:52:28 Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Images cannot be saved. uer Impact: Images cannot be saved. Patient Impact: No impact SE 8.1.3000000005 Resoluion: The database test is passed and a problem is detected. Database repair procedue is performed. The database is reset. Correct operation of the system is confirmed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000157700000000 Adrián Sancho López The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: uLL# and/or description of uLL Root Caue for RMT - IPu IPPC failed to boot u - Front: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power or the OS/AS did not boot u in-time. Repair Action: Service actions for RMT - IPu IPPC failed to boot u - Front: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu on InCenter in the produt manul of the PC. 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IP-PC and perform Cold reboot of the system. 1.3. If LED statu is not Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IP-PC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct! 1.5. Measue/replace the CMOS battery inside the IPPC becaue a bad/weak battery can caue the PC not to start the first time. 1.6. This pattern can also be caued if the BIOS is set wrong. Please check: Advanced->Boot Featue->Restore On AC Power Loss set to 'Last State' or 'Power Off', defaut setting mut be 'Power On'. Follow u Requred: uLL# and/or description of uLL Root Caue for RMT - IPu IPPC failed to boot u - Front: 1. The Host PC is uable to commuicate with the Frontal IPPC. Either the IPPC did not boot u du to lack of (Mains) power or the OS/AS did not boot u in-time. Repair Action: Service actions for RMT - IPu IPPC failed to boot u - Front: 1.1. Check the back of the IP-PC, there are 8 green POST LEDs indicating the boot statu of the PC. Normal boot statu is B1h, youwill find more info for the boot statu on InCenter in the produt manul of the PC. 1.2. If LED statu is Normal then issu is probably net000000000000000000000000001k related (Boot Code B1h is normal boot) Check wiring between Host PC and IP-PC and perform Cold reboot of the system. 1.3. If LED statu is not Normal then issu is Boot problem in IP-PC, check IP-PC power and perform cold reboot of the system and replace IP-PC if problem persists/repeats. 1.00000000Before changing IPPC, check if the Power for the PC is switched off when the system is OFF. When Power is ON, the IPPC will not boot correct! 1.5. Measue/replace the CMOS battery inside the IPPC becaue a bad/weak battery can caue the PC not to start the first time. 1.6. This pattern can also be caued if the BIOS is set wrong. Please check: Advanced->Boot Featue->Restore On AC Power Loss set to 'Last State' or 'Power Off', defaut setting mut be 'Power On'. Technician Name: Adrián Sancho López Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer: RMT - IPu IPPC failed to boot u - Frontal 00000000-Oct-2000000003 00000000:52:28 Information to suport the complaint handling process Cutomer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Eupected and ueupected Resuts: Images cannot be saved. uer Impact: Images cannot be saved. Patient Impact: No impact SE 8.1.3000000005 Diagnostic performed by Engineer: uLL are foud in the image disks. It is recommended to perform the service procedue "Recreate Image Store". Additionally, it might be necessary to replace them, need to go and check it. ----------------------------------------------------------------------------- | 2023-10-04 22:46 | No
1685 | Yes | uLLCODE BLOWER IS NOT HAVING ENOUGH POWER | uLLCODE BLOWER IS NOT HAVING ENOUGH POWER | Internal Remarks : 10000000010000000023 YP 00000000000000000000000000100000000151871 Attempt 3 Details: Cncted the below nos for qutes. Cut not reachable. left vm Contact Joe Stevenson Phone +1(57000000006000000003527 ##jstevenson00eaconhealthsystem.org Internal Remarks : 10000000012/23 YP 00000000000000000000000000100000000151871 Attempt 00000000 Details: Cncted the below nos for qutes. Cut not reachable. left vm Contact Joe Stevenson O: 5700000000000000003527 M: 570000000000000000-1000000003 Internal Remarks : 10000000000000000/23 YP 00000000000000000000000000100000000151871 Attempt2 Details: Cut is on vacation. will be back on Monday. FSE to cnct Joe. Contact Joe Stevenson Phone +1(57000000006000000003527 ##jstevenson00eaconhealthsystem.org Internal Remarks : 10000000012/23 YP 00000000000000000000000000100000000151871 Details: Cutomer called back and stated that he will look into the qutes ##and reply by EOD Contact Joe Stevenson Phone +1(57000000006000000003527 ##jstevenson00eaconhealthsystem.org Internal Remarks : 1000000000000000023 YP 00000000000000000000000000100000000151871 Details: Cncted cut for qutes. Cut not reachable. left vm. ##sent to cut for qutes. Suject: P.O.C is on vacation Contact Joe Stevenson Phone +1(57000000006000000003527 ##jstevenson00eaconhealthsystem.org Internal Remarks : 10000000012/23 YP 00000000000000000000000000100000000151871 Details: Cutomer does not want to proceed with the repair. Cut sent ##confirmation. Escalated to NP to cancel the 000000000000000000000000001----------------------------------------------------------------------------- | 2023-10-04 23:16 | No
1686 | Yes | shut down, will not restart | shut down, will not restart | 000000000000000000000000001O##: 00000000000000000000000000100000000152758 00000000Oct-2000000003 Venters Ricky The system meets the specification for the performed service and is retuned to ue Resoluion : uS and Batteries circut breakers tripped Reset Breakers and test system Problem Description by Engineer : Cutomer Fuction/Role: technology Device ue: system Eupected and ueupected Resuts: resoluion uer Impact: none Patient Impact: None SE2 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Margaret Walotsky IP & Issu: shu down, will not restart 75532686 Event Type: Incident Service Requst: onsite Did Cutomer Decline Remote Suport? yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created O000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-04 23:31 | No
1687 | Yes | Prouimal pressure auto zero failed | Prouimal pressure auto zero failed | 000000000000000000000000001O##: 00000000000000000000000000100000000159769 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer stated he reinstalled the DA pcba on to the flow sensor assy, and it corrected the problem. He believes the pin alignment was off. He ran the uit for over 300000000minues with no issus. uit passed the testing too. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: prouimal pressue auozero failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and they are getting a prouimal pressue auozero failed alarm. Biomed verified there is a code 1100000000 in the history. Cutomer claims they replaced the sol 3 and 00000000nd the DA pcba and problem was not solved. Recommended to check the pin alignment again. 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA. ----------------------------------------------------------------------------- Internal Remarks : Caller: Jacob 300000000890000000093700000000 IP & Issu: 661526900000000 Getting uLLcode prouimal pressue auo zero failed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Verbal PO Nuber: ----------------------------------------------------------------------------- | 2023-10-04 23:31 | No
1688 | Yes | Geometry movements unavailible - system down | Geometry movements unavailible - system down | 000000000000000000000000001O##: 00000000000000000000000000100000000159838 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Geometry problems. When was the first occurence? Within... : 00000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Cutomer states that Geometry movements uavailable - system down Internal Comments : uLL# and/or description of uLL: Cutomer states that Geometry movements uavailable Repair Action : System was in the process of shuting down and I was not able to get log files. Remote service tool did not indicate any geo issus from this morning,. Resoluion : Created on-site 000000000000000000000000001llow u Requred : uLL# and/or description of uLL: Cutomer states that Geometry movements uavailable Repair Action : System was in the process of shuting down and I was not able to get log files. Remote service tool did not indicate any geo issus from this morning,. Technician Name : Johansson Anders ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000160000000000000000 Parts Consued : 1.00000000000000000000000000000000800000000151715,GEO MODuE WP Kit ----------------------------------------------------------------------------- | 2023-10-04 23:31 | No
1689 | Yes | uPro: device is not charging the battery when plugged into the wall. | TempuPro: device is not charging the battery when plugged into the wall. | 000000000000000000000000001O##: 00000000000000000000000000100000000195518 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Resoluion : Device was repaired by u engineer Chongolnee, Roongruchai, with parts suplied by the Bu Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Diagnostic performed by Engineer : Confirmed complaint the uits Charging port is physically damaged. To correct the issu we will need to replace the Front Case Assembly ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000159900000000Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE v00000000.300000000Diagnostic performed by Engineer : Confirmed complaint the uits Charging port is physically damaged. To correct the issu we will need to replace the Front Case Assembly Follow u Requred : 0000000056000000005000000001-Front Case Assembly 00000000ous for repair ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000159679 Wilbu, Phillip The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: Device is not charging the battery when pluged into the wall. uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Device is not charging the battery when pluged into the wall. Follow u Requred : EMS site Manager called to report Tempu Pro: device is not charging the battery when pluged into the wall. Cutomer reports they have tried known good batteries and charger cord with like resuts. Resoluion : Sent to Bench Diagnostic performed by Engineer : Problem Description: Device is not charging the battery when pluged into the wall. Tempu Pro: device is not charging the battery when pluged into the wall. Form ued: Bench Trouleshooting Steps Taken: EMS site Manager called to report Tempu Pro: device is not charging the battery when pluged into the wall. Cutomer reports they have tried known good batteries and charger cord with like resuts. Created Bench 000000000000000000000000001 ##d cutomer bench form, and requsted shipping label. ----------------------------------------------------------------------------- | 2023-10-04 23:36 | No
1690 | Yes | navigation ring was unresponsive | navigation ring was unresponsive |  | 2023-10-04 23:50 | No
1691 | Yes | uLS: DMP Hardware failure, failed to pace during PM | TempuLS: DMP Hardware failure, failed to pace during PM | 000000000000000000000000001O##: 000000000000000000000000001000000001598300000000Brooks-Jordan, Kim The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Replacement uit ordered Follow u Requred : Please Tu Lincoln Couty EMS Attention: CASEY BROOM 1000000006 Global Dr LINCOLNTON GA 30000000017-500000000 Problem Description by Engineer : Cutomer Fuction/Role:Other: EMS site manager Device ue:Not in Clinical ue Eupected and ueupected Resuts: DMP Hardware failue, failed to pace duing PM uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: DMP Hardware failue, failed to pace duing PM Diagnostic performed by Engineer : Problem Description: DMP Hardware failue, failed to pace duing PM Tempu LS: DMP Hardware failue, failed to pace duing PM Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 10000000000000000000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: EMS site manager reports Tempu LS failed to pace duing PM. uit displayed DMP Hardware failue. Cutomer requst replacement Tempu LS. Created Material for replacement uit. The following information was requsted from cutomer: *Log and Rescu Files: requsted *Date of event? 10000000000000000000000003 *Brief description of the event? Tempu LS failed to pace duing PM. uit displayed DMP Hardware failue. *Operator of the produt? FSE *Did the uit display any uLLmessages or warnings before the incident? DMP Hardware failue. *Was the Patient or uer injued or caued any harm? NA *Patient Oucome? NA *Age of Patient? NA *Patient Gender? NA Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000160000000017 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- Internal Remarks : Cutomer called to ask when their replacement Tempu LS is going to arrive. Cutomer was uset becaue they are a small service and only have 3 devices. Will ##Carlos Maldonado to make sue device is approved and ships today. Internal Remarks : Interaction Type: Cutomer Callback Caller: CASEY BROOM IP & Issu: 93297172 System is down (cannot pace) Per the cutomer. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: CASEY BROOM IP & Issu: "93297172 System is down (can not pace) Per the cutomer." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created case and created free remote tech service 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-05 00:06 | No
1692 | Yes | power supply issue with the unit/powering on and off without ur interaction | power supply issue with the unit/powering on and off without ur interauon | 000000000000000000000000001O##: 00000000000000000000000000100000000159800000000Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Onsite service requred, FSE to resolve Follow u Requred : V600000000is powering on and off withou uer interaction Biomed advised of possible power management board problem. FCO86600000000000000006 Biomed is reporting the uit is located at Mercy Hospital 2215 Trutu Ave. Bakersfield, u300000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000to operate as eupected uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is powering on and off withou uer interaction Diagnostic performed by Engineer : Problem Description: V600000000is powering on and off withou uer interaction Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 100000000000000002000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Advised cutomer the u board or power management board as the possible caue. Part # 00000000561528871 RP-PCBA, uer Interface, u 2nd GEN, V600000000----------------------------------------------------------------------------- Internal Remarks : CSA INTERACTION TEMPLATE Interaction Type: Cutomer Call Back Caller: Roy Sands - phone 661 632 5965 IP & Issu: 5660000000098 - CCB becaue they do not need the FSE to come ou onsite bc they don't need the power board installed anymore. The other part fiued the issu. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? N/A Was the device in Clinical ue at the time of the event? N/A Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Canceled onsite 000000000000000000000000001& closed/completed case. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Roy Sands IP & Issu: 5660000000098/There is a power suply issu with the uit. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-05 00:06 | No
1693 | Yes | Backup alarm failed. (EC 00 | Backup alarm failed. (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000152513 200000000Oct-2000000003 Dung, Johnny The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE 3.100000000Diagnostic performed by Engineer : Performed per-op check. Was able to recreate issu. Replaced Motor Controller board, and CPuboard. Issu was resolved. uit passed performance verification and electrical safety tests. Resoluion : Performed per-op check. Was able to recreate issu. Replaced Motor Controller board, and CPuboard. Issu was resolved. uit passed performance verification and electrical safety tests. Tools & Equpment : MuTIMETER,2000000003-12-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,2000000003-12-31 00000000:00000000:00000000 TORQu SCREWDRIVER,2000000003-12-31 00000000:00000000:00000000 TEMPERATuE HuIDITY,2000000003-12-31 00000000:00000000:00000000 FLOW METER,2000000003-12-31 00000000:00000000:00000000 FLOW METER,2000000003-12-31 00000000:00000000:00000000 STATIC CONTROL DEVICE,2000000003-12-31 00000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000561535700000000RP-PCBA,CPuGEN2, (SW 2.100000000, V600000000 1.00000000000000000000000000000000561537531,RP-PCBA, MOTOR CONTR, 3rd Gen, V600000000V6800000000----------------------------------------------------------------------------- Internal Remarks : 10000000000000000/23 YP 00000000000000000000000000100000000152513 S/N- 200000000000000000000000058 Dispatched to FSE Johnny as cut confirmed the P.O.C and S/N. FCO66 needs to be linked as having uLLin SF Internal Remarks : Interaction Type: New Case Caller: Cedrick 5126300000000300000000 IP & Issu: 200000000000000000000000058 / 000000000000000000000000001ld like to have FSE to come onsite to repair. Check vent backu alarm failed. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Created case and Page QA Neut Step Action: Awaiting Qute Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : 1000000000000000023 YP 00000000000000000000000000100000000152513 Details: Incident follow u ##sent to FSE and cutomer. Cncted cut. Cut not reachable. left VM. FSE Johnny Contact Cedric Huter Phone 512 63000000005300000000 ##cedric.huter00rimedu.com Internal Remarks : 10000000000000000/23 YP 00000000000000000000000000100000000152513 Attempt2 Details: Cncted cut. cut not reachable. left vm Contact Cedric Huter Phone 512 63000000005300000000 ##cedric.huter00rimedu.com Facility: (512) 32000000000000000000000000 Transferred to Biomed. spoke to Aaron. cut stated that will pass the message to Cedric ----------------------------------------------------------------------------- | 2023-10-05 00:11 | No
1694 | Yes | system shut down | system shut down | 000000000000000000000000001O##: 00000000000000000000000000100000000160000000062 00000000-Oct-2000000003 Gokey Pau The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : system shu down. no warning. When was the first occurence? Within... : 00000000Oct-2000000003 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: system shu down no uLL noted uer Impact: system shu off as they were prepping case Patient Impact: Not provided SE 000Tools & Equpment : DOSE-KVP-LIGHT METER,2000000008-0000000030000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,2000000005-0000000030000000000000000:00000000:00000000 Resoluion : Replaced Internal Battery in KV/MA PCB (GEN) anad performed new cals - ie; Tue Conditioning, adaptation, Tue Yield, and EDL, Detector Front End Calibration. Checked Adjutment statu for correct date of resuts. Created Ghost on system. Retuned to the cutomer for normal operation. Diagnostic performed by Engineer : Reviewed RSE Notes and recommendations. Reviewed Logfiles from system on site. Determined Possible corruted Data in Generator. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000160000000023 Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : system shu down. no warning. When was the first occurence? Within... : 00000000Oct-2000000003 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: system shu down no uLL noted uer Impact: system shu off as they were prepping case Patient Impact: Not provided SE 000Diagnostic performed by Engineer : euamined log and saw that there was a tue overload that caued the system to shu down. Internal Comments : uLL# and/or description of uLL: system shu down. no warning. Repair Action : Pattern Name: RMT - uGEN: TueOverload buzer activated - Frontal uY Name: 811fa6000000003a6a-000000003-8d6b-69fb1ac3700000000c • Root Caue 1: The u-ray tue overload buzer was activated. The tue load indicator is displayed when the u-ray tue is too hot, uully du to heavy uage by uer. When the tue load level indicator is displayed blinking (red), the tue load levels are defined as follows: A) âOverload levelâ: the tue is hot, bu the requsted tue load can be provided. B) âStop levelâ: the tue is too hot, Fluroscopy is still possible, Euposue is not possible, a curent Euposue ru is aborted. C) âMauimu levelâ: A curent Fluroscopy or Euposue ru is aborted, uconditionally. D) âTue Houe Contact is Openâ: only âLow Loadâ Fluroscopy possible. o Service Action 1: Let the frontal u-ray tue cool down. • Root Caue 2: Something in the generator configuation or adjutment wrt the u-Ray tue are not correct or valid anymore. o Service Action 1: Please verify the correct u-Ray tue type is configued. o Service Action 2: Verify u-Ray tue adaptation is done (recently) withou any overload warnings.Re-do if necessary. o Service Action 3: Verify the EDL calibrations are done recently.Re-do if necessary. Follow u Requred : uLL# and/or description of uLL: system shu down. no warning. Repair Action : Pattern Name: RMT - uGEN: TueOverload buzer activated - Frontal uY Name: 811fa6000000003a6a-000000003-8d6b-69fb1ac3700000000c • Root Caue 1: The u-ray tue overload buzer was activated. The tue load indicator is displayed when the u-ray tue is too hot, uully du to heavy uage by uer. When the tue load level indicator is displayed blinking (red), the tue load levels are defined as follows: A) âOverload levelâ: the tue is hot, bu the requsted tue load can be provided. B) âStop levelâ: the tue is too hot, Fluroscopy is still possible, Euposue is not possible, a curent Euposue ru is aborted. C) âMauimu levelâ: A curent Fluroscopy or Euposue ru is aborted, uconditionally. D) âTue Houe Contact is Openâ: only âLow Loadâ Fluroscopy possible. o Service Action 1: Let the frontal u-ray tue cool down. • Root Caue 2: Something in the generator configuation or adjutment wrt the u-Ray tue are not correct or valid anymore. o Service Action 1: Please verify the correct u-Ray tue type is configued. o Service Action 2: Verify u-Ray tue adaptation is done (recently) withou any overload warnings.Re-do if necessary. o Service Action 3: Verify the EDL calibrations are done recently.Re-do if necessary. Technician Name : Eckbold Glenn ----------------------------------------------------------------------------- | 2023-10-05 00:56 | No
1695 | Yes | Triple chirp, will not pass the self-test, BIT result is "Remove and reinsert battery to test" | Triple chirp, will not pass the self-test, BIT result is "Remove and reinsert battery to test" | 000000000000000000000000001O##: 000000000000000000000000001000000001600000000600000000impson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below uiversity of Illinois Houing 511 W Hazel000000000000000000000000001d Dr ubana, IL 61800000000 Attn: Charles Edmison Resoluion : Closing Case Tracking #78000000002861900000000, DELIVERED Thusday 1000000005/23 at 9:00000000 AM Signed for by: T.NOuIN ##cutomer to confirm that the AED was 000000000000000000000000001king as it shoud Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported the AED kept chirping and will not pass the self-test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000007 Cutomer reported the AED kept chirping and will not pass the self-test. Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and will not pass the self-test. Cutomer is uaware how long it was chirping and removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Remove and reinsert battery to test. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Advise cutomer to remove device ou of service. util he received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# B16F-00000000900000000 Battery Install Before# 2000000009/00000000/300000000 Lot#000000000000000000000000 Pads Eup Date: 1/1000000002000000006 Pads Lot # 23000000001000000000000000058 Warranty 6/27/20000000000000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000160000000083 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-10-05 01:01 | No
1696 | Yes | Speaker Malfunction-IPCR Location change/no audio found at Bench-Foulinks s/n u00other HW-speaker-HW/replaced | Speaker Malfunuon-IPCR Location change/no audio found at Bench-Foulinks s/n u000 /other HW-speaker-HW/replaced | 000000000000000000000000001O##: 00000000000000000000000000100000000287693 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- Internal Remarks : Left a vmail for the cutomer advising him to ##a pictue of the device 000and the accout name and address Internal Remarks : Interaction Type: cut callback Caller: Adrian Cortez 9162978169 IP & Issu: "u000000007d0000000000000000/IPCR Location change. Located at: Cpmc 1100000000 Van Ness Ave San Fransisco u00000009 Speaker malfuction." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE/ Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : Called and left detailed vmail advising Adrian to ##or call me back Final attempt Internal Remarks : Interaction Type: New Case Caller: Adrian Cortez 9162978169 IP & Issu: "u000000007d0000000000000000/IPCR Location change. Located at: Cpmc 1100000000 Van Ness Ave San Fransisco u00000009 Speaker malfuction." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE/ Sumitted IPCR relocation Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-05 01:11 | No
1697 | Yes | THE SOFTWARE DOES NOT START | THE SOFTWARE DOES NOT START | 000000000000000000000000001O##: 00000000000000000000000000100000000162000000006 October 27, 2000000003 Hernandez, Jose The system meets the specification for the performed service and is retuned to ue. Resoluion: The reported problem is the same one evidenced in the field visit, the equpment does not start the application and fleuvision PC does not start the operating system. Problem Description by Engineer: Problem description by engineer: THE SOFTWARE DOES NOT START When was the first occurence? Within... : 00000000Oct-2000000003 00000000:30000000000000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: THE SOFTWARE DOES NOT START uer Impact: No Patient Impact: No SE 8.1.3000000005 Resoluion: Fleuvision PC format is changed and software is loaded to fleuvision PC disk. The equpment is able to start, 3D image tests are performed and fleuvision fuction is checked to change layou. OK. Equpment is delivered in good operating order. Ordered rouer is installed proactively. Parts Consued: 1.0000000000000000000000003000000000000000070000000000000000,Cisco CBS350000000020000000000000000Managed Switch DVI 1.0000000000000000000000000000000080000000058600000000, FleuVision-2_Revised PC not HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000159386 00000000-Oct-2000000003 Ortega, Roberto The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: THE SOFTWARE DOES NOT START When was the first occurence? Within... : 00000000Oct-2000000003 00000000:30000000000000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: THE SOFTWARE DOES NOT START uer Impact: No Patient Impact: No SE 8.1.3000000005 Follow u Requred: Failues are foud in the FleuVision hardware and Switch GB is proactively requsted. Diagnostic performed by Engineer: Equpment with initialization problems is received. We proceed to physically check the connections in switch GB, host, IPPC and FleuVision PC. While checking the FleuVision PC, hardware failues were foud, so it shoud be replaced. Switch GB is also requsted proactively. Equpment ou of order. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000158100000000Koc, Migul The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: THE SOFTWARE DOES NOT START When was the first occurence? Within... : 00000000Oct-2000000003 00000000:30000000000000000 How often is issu occuring? : Continuuly Malfuction area: PCs Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: THE SOFTWARE DOES NOT START uer Impact: No futher information is available as per the cutomer. Patient Impact: No futher information is available as per the cutomer. SE 8.1.3000000005 Follow u Requred: uLL# and/or description of uLL THE SOFTWARE DOES NOT START Repair Action: It finishes the software start-u cout, bu the software does not start, GeoPC is restarted, it responds well, the brakes soud releasing, Fleuvision is restarted, on the touh screen it allows viewing selection bu it does not complete the software start-u completely, ipPC is restarted, the faut persists, it seems to be a HD and/or Host PC failue, they requst to derive case to field. A field service engineer mut go to the site to test the IPPC, HostPC and HDs, try to reload the software, otherwise evalute hd and/or Host PC replacement, also ru Fleuvision test, reload software, euchange grabers or evalute its replacement. What are the neut Trouleshooting steps: A field service engineer mut go to the site to test the IPPC, HostPC and HDs, try to reload the software, otherwise evalute hd and/or Host PC replacement, also ru Fleuvision test, reload software, euchange grabers or evalute its replacement. Technician Name: Koc, Migul Internal Comments: uLL# and/or description of uLL THE SOFTWARE DOES NOT START Repair Action: It finishes the software start-u cout, bu the software does not start, GeoPC is restarted, it responds well, the brakes soud releasing, Fleuvision is restarted, on the touh screen it allows viewing selection bu it does not complete the software start-u completely, ipPC is restarted, the faut persists, it seems to be a HD and/or Host PC failue, they requst to derive case to field. A field service engineer mut go to the site to test the IPPC, HostPC and HDs, try to reload the software, otherwise evalute hd and/or Host PC replacement, also ru Fleuvision test, reload software, euchange grabers or evalute its replacement. Diagnostic performed by Engineer: The software start-u cout is finished, bu the software does not start, GeoPC is restarted, it responds well, the brakes soud and release, Fleuvision is restarted, on the touh screen it allows viewing selection bu it does not complete the software start-u completely, IPPC is restarted, failue persists, apparently hd and/or Host PC failue, they requst to derive case to field. A field service engineer mut go to the site to test the IPPC, HostPC and HDs, try to reload the software, otherwise evalute hd and/or Host PC replacement, also ru Fleuvision test, reload software, euchange grabers or evalute its replacement. ----------------------------------------------------------------------------- | 2023-10-05 01:51 | No
1698 | Yes | Seeking RCA of why the MP0did not alarm. Unable to Confirm Cau-No Further Information Available. | Seeking RCA of why the MP0did not alarm. Unable to Confirm Cau-No Further Information Available. | 000000000000000000000000001O##: 00000000000000000000000000100000000152860000000000000000Oct-2000000003 Edwin Tandoh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion : Patient incident report has been placed with u. Photos and logs taken for escalation. Patient incident team will be investigating to the matter. ----------------------------------------------------------------------------- Internal Remarks : Cutomer wanted to know why she has not received a qute and no on site has been schedued. Informed cutomer that I will need to requst qute and ##planners to schedue on site. ----------------------------------------------------------------------------- | 2023-10-05 02:16 | No
1699 | Yes | uCVu - Freeze up issues - Unknown TEE procedure | uCVu - Freeze up issues - Unknown TEE procedure | 000000000000000000000000001O##: 00000000000000000000000000100000000152791 300000000Oct-2000000003 Hueini, Salma The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: lead Tech Device ue: yes Eupected and ueupected Resuts: system freezes u when uing TEE probe. uer Impact: restarting system several times to recognize probe Patient Impact: uknown SE9.000000005 Resoluion : Downloaded log files, ru diagnostics and reseated all boards at the back of the system ----------------------------------------------------------------------------- | 2023-10-05 02:16 | No
1700 | Yes | lost c-arm rotation | lost c-arm rotation | 000000000000000000000000001O##: 00000000000000000000000000100000000160000000023 Chhlam, Sy The system meets the specification for the performed service and is retuned to ue Resoluion : Follow u with Clint, system is still operational after part replace. no futher action needed, ok to close case. Problem Description by Engineer : Problem description by engineer : WARNING: Motorized movement uavailable When was the first occurence? Within... : 00000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: lost geometry movment and had to reboot uer Impact: lost geo movement Patient Impact: No futher information is available as per the cutomer SE 8.2.16.1 Diagnostic performed by Engineer : Clint reported that the techs lost geometry movements and after reboot they lost the sid movement. Want log checked. Noticed last movement command before the message, "WARNING: Motorized movement uavailable," was the detector being commanded to move. Internal Comments : uLL# and/or description of uLL: 5800000000000000009200000000WARNING: Motorized movement uavailable Repair Action : Clint to check movements and cables and he also have an spc 7 mvs power suply he can ue. Biomed to also calibrate the SID movement and agreed to keep case open util Friday 1000000006/23 and close if no call backs. Resoluion : Follow u with Clint before closing. ----------------------------------------------------------------------------- | 2023-10-05 02:31 | No
1701 | Yes | IC no resp¦uCVui - freeze up issues - Unknown Procedure | IC no resp¦uCVui - freeze up issues - Unknown Procedure | 000000000000000000000000001O##: 00000000000000000000000000100000000295891 300000000Oct-2000000003 Hueini, Salma The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000561970000000000000000CBA, Harmony ACB 1.00000000000000000000000000000000562000000000000000000000000EPIQ B-Cart Acqusition Modue Problem Description by Engineer : Cutomer Fuction/Role: lead Tech Device ue: yes Eupected and ueupected Resuts: lock us when uing TEE uer Impact: rebooting several times Patient Impact: none SE9.000000005 Resoluion : Replaced acqusition modue and ACB board. Did a ful software reload ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000152900000000 25-Oct-2000000003 Hueini, Salma The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: lead Tech Device ue: yes Eupected and ueupected Resuts: device locks u when uing TEE probe uer Impact: locks u when uing TEE and probe Patient Impact: none SE9.000000005 Resoluion : This issu is intermittent and affects 7 systems. So acqusition modue and ACB of one system 000000000000000000000000001ld be replaced. Ful software reload 000000000000000000000000001ld be done. Whereas another system 000000000000000000000000001ld. Jut get a ful software reload to see which 000000000000000000000000001ks better ----------------------------------------------------------------------------- | 2023-10-05 02:36 | No
1702 | Yes | when running 0leads showing stemi not in stemi | when running 0 leads showing stemi not in stemi | 000000000000000000000000001O##: 00000000000000000000000000100000000160000000068 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Resoluion : RDT determined no faut with device. Calibrated NBP, CO2 and Touh Screen. All fuctions are 000000000000000000000000001king correctly. Device software ugraded to v00000000.300000000No parts requred calibration only. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Coud not confirm uer Impact: uable to ue Patient Impact: none SE v00000000.300000000iagnostic performed by Engineer : Cannot confirm complaint the uit is 000000000000000000000000001king correctly at the bench. the logs will be evaluted by RDT. Cutomer requsts device to be ugraded to Rev. 300000000---------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000160000000072 Brooks-Jordan, Kim The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE .300000000 Problem Description/ueupected Resuts: Accout manager reports false alarm from the device on 9/8/200000000 0000000000000000, sending to bench for repair; no loaner needed, cut already has loaner Diagnostic performed by Engineer : Problem Description: Accout manager reports false alarm from the device on 9/8/200000000 0000000000000000, sending to bench for repair; no loaner needed, cut already has a loaner; no patient or uer harm Delivery Address: Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repairConfirmed delivery address: 259 Belcher Street, CENTREVILLE, AL 3500000000##d bench repair form Follow u Requred : Problem Description: Accout manager reports false alarm from the device on 9/8/200000000 0000000000000000, sending to bench for repair; no loaner needed, cut already has loaner Delivery Address: Form ued: Bench Trouleshooting Steps Taken: Sending to bench for repair Confirmed delivery address: 259 Belcher Street, CENTREVILLE, AL 3500000000 ##d bench repair form ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Lewie Sandel IP & Issu: when runing 12 leads showing stemi not in stemi 95000000005971 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Lewie Sandel IP & Issu: 95000000005971/ calling in for udate on bench and has qustions for bench FSE Event Type: Cutomer Suport Action Taken: Advise in Decontaminiation per tracker and sent AA for udate and callback to Lewie as they have more qustions ----------------------------------------------------------------------------- | 2023-10-05 02:41 | No
1703 | Yes | Defective touchscreen | Defeu touchscreen | 000000000000000000000000001O##: 00000000000000000000000000100000000189179 17-Oct-2000000003 Ramirez, Eletise The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus and crack on the top cover uer Impact: none Patient Impact: none SE un Tools & Equpment : TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 Parts Consued : 1.00000000000000000000000000000000561511811,RP-Cover, Top, V600000000Diagnostic performed by Engineer : Top cover had crack near a screw houing . Top cover replace with new part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000016000000000000000000000000-Oct-2000000003 Ramirez, Eletise The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus and crack on the top cover uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Original lcd was non responsive. Bu after installing new one it is non responsive. Follow u Requred : uit is now fuly fuctional and retun for service. Parts Consued : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000Tools & Equpment : ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 TORQu SCREWDRIVER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000160000000067 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: touhscreen issus and crack on the top cover uer Impact: none Patient Impact: none SE un Follow u Requred : Cutomer has a V600000000and the touhscreen is defective and the top cover is damaged Part #00000000561539000000001 touhscreen Part #00000000561511811, top cover Created an onsite 000000000000000000000000001Diagnostic performed by Engineer : Cutomer has a V600000000and the touhscreen is defective and the top cover is damaged Part #00000000561539000000001 touhscreen Part #00000000561511811, top cover Created an onsite 000000000000000000000000001 ASP will be onsite on 10000000012-23 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Isaiah Davenport +12132800000000876 IP & Issu: "defective touh screen 60000000053855" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: Created R000000000000000000000000001and roued to RSE ----------------------------------------------------------------------------- | 2023-10-05 03:01 | No
1704 | Yes | The cuomer needs support to install the Air O0low sensor assembly. | The cuomer needs support to install the Air O0flow sensor assembly. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000155 GALLARDO, DANIEL The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005615000000001 PCBA, Power Management (PM) 989800000000617100000000Flow Sensor Assembly 00000000561512681 PM Kit, Annul 00000000561512881 Bou 5 hous. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm service Eupected and ueupected Resuts:Flow sensor replacement uer Impact: installation instrutions Patient Impact: none SEun Diagnostic performed by Engineer : Confirmed uit is getting uLL 1000000007, 1118, 111A, 1127, 111B, and 1115 on boot u, need to replace Power Management PCBA to fiu problem. uit is also failing O2 Miu Accuacy, measuing 23.5% O2 when set to 300000000. Will order Flow Sensor Assembly to fiu problem. Tools & Equpment : FLOW METER,200000000000000002-29 00000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000005 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm service Eupected and ueupected Resuts:Flow sensor replacement uer Impact: installation instrutions Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer states he attempted to replace the flow sensor assembly and is getting all kinds of uLLcodes and has been dealing with this device for some times. Cutomer states he doesn't have the time to repair device and is requsting bench repair for device. Confirmed cutomers ##address. Cutomer is requsting bench repair. euplained bench process. ##d cutomer bench requst form. Resoluion : Sent to bench Follow u Requred : Cutomer states he replaced flow sensor assembly and is getting all kinds of uLLcodes and doesn't have bandwidth to 000000000000000000000000001k on device. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000019000000002 Newcomer, Keith The system meets the specification for the performed service and is retuned to ue Follow u Requred : Bench repair requst Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Post flow sensor replacement. at tun on uit goes vent inop w/a 1000000007 diagnostic code. Also goes into vent inop when in diagnostic mode and the "service" buton is pressed. removed, inspected and reseated the MC to DA cable, cutomer has tried one from another uit also, no help. Advised to remove the GDS and verify pin alignment from the auozero solenoids to the DA pcba. If no issu foud advised to install the original flow sensor assy. If problem is resolved to call parts and requst replacement uder warranty. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm service Eupected and ueupected Resuts:Flow sensor replacement uer Impact: installation instrutions Patient Impact: none SEun Diagnostic performed by Engineer : cutomer called back to verify rouing of the pneuatic lined from the patient oulet port. Verified the prouimal and machine lines are connected correctly at the flow sensor and the euhaut line is connected to the internal euhaut port on the bottom of the uit. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000160000000073 Mccleary Barry The system meets the specification for the performed service and is retuned to ue Follow u Requred : Cutomer calling for follow u on the gas delivery system and the hose configuation. Resoluion : Advised cutomer on correct installation, cutomer acknowledged and will complete pm and testing, issu resolved. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: pm service Eupected and ueupected Resuts:Flow sensor replacement uer Impact: installation instrutions Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer is replacing flow sensor assembly duing pm service and requsts installation instrution. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mike 3600000000628-7000000007 IP & Issu: case 00000000220000000000000000 statu udate Event Type: Bench Repair Service Requst: N/A Did Cutomer Decline Remote Suport? N/A Was the device in Clinical ue at the time of the event? Y/A Preferred Service Time: Entitlement Options: N/A Reason for Entitlement Claim: Action Taken: "created AA 0000000000000000953621 Callback-Qust Bench Please contact Mike with statu udate by phone and ##at mtuman00eacehealth.org and 3600000000628-7000000007 thanks" Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mike 3600000000287000000007 IP & Issu: case 00000000220000000000000000 cutomer had additional qustions abou qute and stated that he did receive a ##from the qute team after I pu in the requst for them to contact him by phone and ## Event Type: Bench Repair Service Requst: N/A Did Cutomer Decline Remote Suport? N/A Was the device in Clinical ue at the time of the event? N/A Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: I advised for him to respond to the ##with any qustions in regards to the qute and ##they sent him Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mike Tuman IP & Issu: "75955177 The cutomer needs suport to install the Air O2 flow sensor assembly." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Provide statu udate on bench repair Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : Interaction Type: New Case Caller: Mike Tuman IP & Issu: 75955177/The cutomer needs suport to install the Air O2 flow sensor assembly. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created a case/remote 000000000000000000000000001Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mike Tuman IP & Issu: " 75955177 CCB regarding case#00000000220000000000000000" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: created requst follow u 000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: Cutomer Call Back Caller: Mike Tuman IP & Issu: 75955177 The cutomer needs suport to install the Air O2 flow sensor assembly Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Create follow u R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-05 03:01 | No
1705 | Yes | [Q] LS - 00 AT - 00 | [Q] LS - 000/ AT - 000 | 000000000000000000000000001O##: 0000000000000000000000000010000000015300000000 11-Oct-2000000003 Shula Chintan The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced Frontal and Lateral IP PC’s. Also replaced OS disk and both Image storage disks. Loaded OS and application SW. Recreated image storage. Tested system to confirm proper operation. Retuned system back to clinical ue. Parts Consued : 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Problem Description by Engineer : Cutomer Fuction/Role: lead tech Device ue: store images Eupected and ueupected Resuts: Image space low, delete images popping u uer Impact: inconvenience Patient Impact: uknown SE 000----------------------------------------------------------------------------- | 2023-10-05 04:56 | No
1706 | Yes | Q/O Speaker is falling from TV BOOM | Q/O Speaker is falling from TV BOOM | 000000000000000000000000001O##: 0000000000000000000000000010000000015900000000 00000000Oct-2000000003 Davis Barry The system meets the specification for the performed service and is retuned to ue Resoluion : Speaker fell off of the TV boom. Problem Description by Engineer : Cutomer Fuction/Role: u Ray technologist Device ue: system not in ue Eupected and ueupected Resuts: The speaker was knocked off of the TV boom. uer Impact: no uer impact Patient Impact: no, patient impact SE 8 ----------------------------------------------------------------------------- | 2023-10-05 05:11 | No
1707 | Yes | [Q] system suddenly rebooting | [Q] system suddenly rebooting | 000000000000000000000000001O##: 000000000000000000000000001000000001530000000000000000-Oct-2000000003 Hereford Rick The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Sute PC started continuuly rebooting Eupected and ueupected Resuts: Sute PC started continuuly rebooting uer Impact: Patient Impact: Sute PC started continuuly rebooting uuResoluion : restored software duing PM, retuned parts 1000000006/2000000003 as no futher issus after 2 days of monitoring ----------------------------------------------------------------------------- Internal Remarks : Rick, system keep rebooting after powered on. Log shows Process Crashed. He's not sue the symptom only happened after PM this afternoon or even before. Advised to Reinstall Sute PC software from TSM. ----------------------------------------------------------------------------- | 2023-10-05 05:21 | No
1708 | Yes | HV uuuitor issue observed during product analysis (Split from 00 | HV uuuor issue observed during produuanalysis (Split from 000) |  | 2023-10-05 06:12 | No
1709 | Yes | SA:0u FD0cannot turn on | SA:0/u FD0cannot turn on | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes: This PO is a Reference Nuber only SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 18.1000000002000000003 00000000:53:26 uC (8600000000000000009) $Q1:(Y)$Q2:(N)$Q3:()$Q00000000)$Q5:() $Q6:( Can't boot u ) $Q7: (Confirmation & Description of problem: Confirmed, HOST PC can't power u.) $Q8: (Trouleshooting & Resuts: Check the HOST PC, the Power suply of the HOSTPC is defective.) $Q9: (Repair steps: Replace the HOST PC, renew the License. ) $Q100000000 (Test and Inspection data&Device statu: Device retrued to ful fuctionality. Test and Inspection data refer to PA. ) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Neut Calibration date (N/A) For Mainland China CV/uR, CT engineers only. Please fill in: Curent Tue: 3983600000000 Installed data: 00000000/00000000/2000000002 Euternal Remarks: 17.1000000002000000003 00000000:55:00000000uC (8600000000000000009) Service content: uable to boot, the HOST PC power suply was damaged. I puchased a local power suply and replaced it and the faut was solved.Finally, after the spare parts arrived, I replaced it with a new HOST PC. Service resut: The faut is repaired and the machine can be ued normally. SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002239000000003 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: uable to boot Neut Step Action: Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-05 06:36 | No
1710 | Yes | Wired foot switch not working in the euamination room for fluoros | Wired foot switch not working in the euamination room for fluoros | 000000000000000000000000001O##: 00000000000000000000000000100000000161700000000 00000000-Oct-2000000003 Shimizu Shuei The system meets the specification for the performed service and is retuned to ue Resoluion: Part ued: Footswitch CV 3p 8m Part No.: 000000002700000000000000003800000000sed: 1 Reguatory Qustions: N/A Problem Description by Engineer: He was not able to ue the wired foot switch in the euamination room for fluros Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: He was not able to ue the wired footswitch in the euamination room for fluros uer Impact: No effect Patient Impact: No effect SE 8.1.30000000015 Resoluion: I confirmed from the uLLlog that it was not a control system failue. I juged that the foot switch cable was cu bu I was not able to reprodue the problem, so I supect that there is a faut inside the foot switch. I replaced both the foot switch and the cable with new ones. Following the replacement, I was able to ue the foot switch in the euamination room withou a problem, and I confirmed that there is no problem with operation. Follow u Requred: N/A Parts Consued: 1.000000000000000000000000000000002700000000000000003800000000ootswitch CV 3p 8m ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000161600000000Noda, Kotaro The system meets the specification for the performed service and is retuned to ue Internal Comments: Simpl physical [Suject] 00000000DS:KN:u Ishichigawa: Coud not ue the wired footswitch in the euamination room for fluroscathing. [Parts from the equpment] 000000002700000000000000003800000000ootswitch CV 3p 8m QTY:1 ZRFA Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Yui Matsu] Matsu, Satoshi - FSE1 Visit date and time 2000000003/1000000000000000017:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Yui Matsu - 00000000000000000069000000006159 - Others Call mobile uon arrival Requster name Kotaro Noda [Euamination category] Normal euamination Follow u Requred: * Problem Description by Engineer: He was not able to ue the wired foot switch in the euamination room for fluros Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: He was not able to ue the wired footswitch in the euamination room for fluros uer Impact: No effect Patient Impact: No effect SE 8.1.30000000015 Diagnostic performed by Engineer: There was a one-time problem of not being able to do fluroswitch with jut the FootSwitch in the Euamination Room.It coud be ued on the investigation room side.There was no message displayed. ----------------------------------------------------------------------------- | 2023-10-05 06:46 | No
1711 | Yes | 0 KOR Hu0/0harger cup fell off the back of the foot and was sutured. | 00KOR Hu00/0charger cup fell off the back of the foot and was sutured. |  | 2023-10-05 06:51 | Yes
1712 | Yes | device can`t be discharged / high voltage uuuitor aging. | device can`t be discharged / high voltage uuuor aging. | BuQ&R Feedback: No Reason: Corrective Maintenance FSE Notes : This PO is a Reference Nuber only Problem Reported by Cutomer : 00000000R_110000000086_8612900000000High voltage uuuitor aging SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 00000000223950000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: kiki chao Last udated by: kiki chao 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-05 06:56 | No
1713 | Yes | [Q] - u-ray are not working - SPLIT | [Q] - u-ray are not working - SPLIT |  | 2023-10-05 08:04 | No
1714 | Yes | 0 / u0/ Therapy delivery test failed. | 00/ u000/ Therapy delivery test failed. | 000000000000000000000000001O##: 0000000000000000000000000010000000015000000002 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : EOL Produt part cannot be ordered informed cutomer. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : therapy delivery test failed ----------------------------------------------------------------------------- | 2023-10-05 09:51 | No
1715 | Yes | RR-Fluo & Cine short fail | RR-Fluo & 0e short fail | 000000000000000000000000001O##: 00000000000000000000000000100000000162300000000SANGBOONSONG MuDGROM The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Intermittent uray not possible. Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer : The uSC received an uknown message (120000000000000000from the PoInt The uSC received an uknown message (120000000000000000from the PoInt SYMPTOM 1: Intermittent uray not possible. CAuE 1: The uSC received a message ID from a generator part which is not known to it. Therefore a blank line is displayed in the description field. SOLuION 1: The euception ID in the Additional Info is: 120000000000000000Curently (Juy 2000000003) this problem is uder investigation. Best advise is to replace the PoInt. Follow u Requred : uLL# and/or description of uLL: The uSC received an uknown message (120000000000000000from the PoInt The uSC received an uknown message (120000000000000000from the PoInt SYMPTOM 1: Intermittent uray not possible. CAuE 1: The uSC received a message ID from a generator part which is not known to it. Therefore a blank line is displayed in the description field. SOLuION 1: The euception ID in the Additional Info is: 120000000000000000Curently (Juy 2000000003) this problem is uder investigation. Best advise is to replace the PoInt. Repair Action : Order Power Inverter uit to replace Technician Name : SANGBOONSONG MuDGROM Internal Comments : uLL# and/or description of uLL: The uSC received an uknown message (120000000000000000from the PoInt The uSC received an uknown message (120000000000000000from the PoInt SYMPTOM 1: Intermittent uray not possible. CAuE 1: The uSC received a message ID from a generator part which is not known to it. Therefore a blank line is displayed in the description field. SOLuION 1: The euception ID in the Additional Info is: 120000000000000000Curently (Juy 2000000003) this problem is uder investigation. Best advise is to replace the PoInt. Repair Action : Order Power Inverter uit to replace ----------------------------------------------------------------------------- | 2023-10-05 12:01 | No
1716 | Yes | System does not go up | System does not go up | 000000000000000000000000001O##: 00000000000000000000000000100000000168372 Oct 11, 2000000003 Beer, Dietmar The system meets the specification for the performed service and is retuned to ue Resoluion: Du to sporadic change of the boot sequnce in bios, the backu drive W is jammed, Fuction check, system is ready for clinical ue Problem Description by Engineer: Problem description by engineer: Called Gabriele Neuebaur-Kässmann Phone +00000000300000000351600000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001625500000000ct 5, 2000000003 Beer, Dietmar The system meets the specification for the performed service and is retuned to ue Follow-u requred: Futher repair measues after consutation with TSS Problem Description by Engineer: Problem description by engineer: Called Gabriele Neuebaur-Kässmann Phone +00000000300000000351600000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Resoluion: Trouleshooting, boot sequnce in the CPuBios had adjuted itself, corrected boot sequnce, fuction check, system is ready for clinical service ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000016200000000Hellenbrandt, Helmu The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Problem description by engineer: Called Gabriele Neuebaur-Kässmann Phone +00000000300000000351600000000 The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer: System does not start Internal Comments: uLL# and/or description of uLL Duing start-u, the system is asking for a boot disk Maybe it 000000000000000000000000001ld be suficient to take off the backu disk here. Otherwise trouleshooting/replace compuer Repair Action: Trouleshooting system start Hard drive was already replaced. Replace compuer board if necessary Follow-u requred: uLL# and/or description of uLL Duing start-u, the system is asking for a boot disk Maybe it 000000000000000000000000001ld be suficient to take off the backu disk here. Otherwise trouleshooting/replace compuer Repair Action: Trouleshooting system start Hard drive was already replaced. Replace compuer board if necessary Technician name: Hellenbrandt, Helmu ----------------------------------------------------------------------------- Cutomer comment: Software Gabi has appointment info today 12:30000000013:00000000 hrs. Internal Remarks: Asked u Beer on voice bou for consutation regarding scheduing (today still withou compuer?) Internal Remarks: Service withou parts coud make sense. Maybe youjut have to take off the backu disk so that the system is runing again. ----------------------------------------------------------------------------- | 2023-10-05 12:01 | No
1717 | Yes | 0 / u0/ u/ uLLwhen boot. | 00/ u00 / u/ uLLwhen boot. | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: Linzhi Bayi District People’s Hospital Neut Step Action: ##ASP Hong Lejia to follow u, SMS ZhouShidi MATC/u West District/North District 1550000000060000000088 Safety Qustions: Case: 0000000022396177 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Pan Chuhu Last udated by: Pan Chuhu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Transfer to sales handling, This case is considered invalid and withou any futher steps if qutation is not accepted, Will be closed withou final soluion after cutomer refues ou service. Problem Reported by Cutomer: ECR:00000000/HeartStart uL+/uLLwhen boot Neut Step Action: Contacted u Wufor follow-u call Neut Step Action: System uLLprompts when booting, NNY RSE Forestry is following u ----------------------------------------------------------------------------- | 2023-10-05 13:11 | No
1718 | Yes | Euposure not possible | Euposure not possible | 000000000000000000000000001O##: 00000000000000000000000000100000000155268 12-Oct-2000000003 WITTAWAT PATSARA The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: Cv Eupected and ueupected Resuts: Euposue not possible uer Impact: no Patient Impact: no SE Resoluion : Recheck and clean flash light reseat connector ----------------------------------------------------------------------------- | 2023-10-05 13:16 | No
1719 | Yes | Cannot get the unit to stay in stand by | Cannot get the unit to stay in stand by |  | 2023-10-05 14:09 | No
1720 | Yes | Machine is not booting up | Machine is not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000162600000000Kuar, Atu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : As per uer/FSE the system is 000000000000000000000000001king fine. No issu in system Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : As per uer/FSE the system is 000000000000000000000000001king fine. No issu in system Internal Comments : uLL# and/or description of uLL: As per uer/FSE the system is 000000000000000000000000001king fine. No issu in system Repair Action : As per uer/FSE the system is 000000000000000000000000001king fine. No issu in system Resoluion : As per uer/FSE the system is 000000000000000000000000001king fine. No issu in system ----------------------------------------------------------------------------- | 2023-10-05 14:21 | No
1721 | Yes | Audio&Sounds: Speaker uLLInop: Sound Confirmed: Speaker/HW/Other: Work Around Provided. | Audio&Sounds: Speaker uLLInop: Sound Confirmed: Speaker/HW/Other: Work Around Provided. | 000000000000000000000000001O##: 000000000000000000000000001000000001558300000000ABDELBARI MOHAMMED The system meets the specifications for the performed service and is retuned to ue Diagnostic performed by Engineer: Biomed informed that he had looked into the issu and noticed that the alarm was coming from the u3. Gutav had informed to udock the u3 and reboot it and if the issu is cleared and the soud is 000000000000000000000000001king they can pu it back in clinical ue, if not the monitor needs to be repaired. Biomed confirmed that reboot solved the issu. The INOP is gone and the speaker is 000000000000000000000000001king as intended. We agreed to close the case. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: In ue. Eupected and ueupected Resuts: Speaker technical faut INOP. uer Impact: Speaker technical faut INOP. Patient Impact: None. SE NA ----------------------------------------------------------------------------- | 2023-10-05 14:46 | No
1722 | Yes | 00 CN00 Getting umalfunctioning uLL(Replaced DFM0 Measurement Module u | u Getting umalfunuoning uLL(Replaced DFM0 Measurement Module u | 000000000000000000000000001O##: 00000000000000000000000000100000000163000000006 Gnanasekar, D The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : 0000000056000000009131 DFM100000000 Measuement Modue u Ordered. Follow u Requred : 0000000056000000009131 DFM100000000 Measuement Modue u Ordered. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Getting umalfuctioning uLL----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000163500000000 Parts Consued : 1.0000000000000000000000000000000056000000009131,DFM100000000 Measuement Modue u----------------------------------------------------------------------------- | 2023-10-05 14:46 | No
1723 | Yes | Abnormal wear and tear of the green defibrillation connector. | Abnormal wear and tear of the green defibrillation conneur. | 000000000000000000000000001O##: 0000000000000000000000000010000000015600000000 Assad MAJEED The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: Recommendation: 0000000056000000008971 Requst for qute for suply of part sent to the qutation department. Follow u Requred: Qute Diagnostic performed by Engineer: The cutomer wants to receive a qute to order a therapy connector. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Reguatory Qustions: No/No ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001636500000000Parts Consued: 1.0000000000000000000000000000000056000000008971, DFM100000000 ASSY Therapy Receptacle ----------------------------------------------------------------------------- | 2023-10-05 14:51 | No
1724 | Yes | uLS AED MAN does not connect to uPro device | TempuLS AED MAN does not conneuto TempuPro device | 000000000000000000000000001O##: 00000000000000000000000000100000000163600000000 Marc Chambers The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : End uer to be provided with service euchange device 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Medical Equpment Technician Device ue: not in ue Eupected and ueupected Resuts: End uer was uable to pair LS to PRO - end uer was given phone suport bu issu remained - tested with another LS which paired with the same PRO no issu uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : End uer to be provided with service euchange device 1u 000000005600000000000000001 RDT Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000163681 Parts Consued : 1.000000000000000000000000000000005600000000000000001,Tempu LS Defibrillator - AED, MAN ----------------------------------------------------------------------------- | 2023-10-05 14:51 | No
1725 | Yes | 00 CN00- Device Failed Testing | u- Device Failed Testing |  | 2023-10-05 14:51 | No
1726 | Yes | Buzzer/alarm speaker no longer works - Cau unknown - replacement sent. | Buzzer/alarm speaker no longer works - Cau unknown - replacement sent. | 000000000000000000000000001O##: 00000000000000000000000000100000000163700000000HuB LANGE The system meets the specifications for the performed service and is retuned to ue Follow u Requred: 0000000056000000005351 TELE PWM, 800000000.11a/b/g,uonly, eu,non u IntelliVu Mu00000000800000000.11a/b/g/n 865352 Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Ouside of ue Eupected and ueupected Resuts: Mu00000000speaker doesn't 000000000000000000000000001k uer Impact: Ouside of ue Patient Impact: N/A SE B.00000000.59 Resoluion: 0000000056000000005351 TELE PWM, 800000000.11a/b/g,uonly, eu,non u IntelliVu Mu00000000800000000.11a/b/g/n 865352 order ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001637500000000Parts Consued: 1.0000000000000000000000000000000056000000005351,TELE PWM, 800000000.11a/b/g,uonly, eu,non u ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-05 15:01 | No
1727 | Yes | Wireless footswitch requires replacement | Wireless footswitch requires replacement | 000000000000000000000000001O##: 00000000000000000000000000100000000156000000007 15-Oct-2000000003 Ashley Goodwin The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station Resoluion : Wireless Footswitch and base station fitted and fuctionally tested. Parts already with latest firmware loaded. Passes testing ok. Problem Description by Engineer : Cutomer Fuction/Role: radiographer Device ue: cardiac Eupected and ueupected Resuts: wireless Footswitch lost connection uer Impact: na Patient Impact: na SE 2.2 Reguatory Qustions : N/A ----------------------------------------------------------------------------- Internal Remarks : ##from Ashley Goodwin dated the 28th Sep 2000000003 Yasmin confirmed on Tusday/Wednesday that she had ‘binned’ the wireless footswitch – she had logged a case with a faut previouly. It was disregarded when the Field Safety Notice came ou. Futher to ou discusion, please see attached parts nubers. Wireless Footswitch 3P – 000000008-000000002-38232 WFS Base Station 000000008-000000002-57682 Wireless Footswitch Charger (WFSC) – 000000008-00000000000000005552 Andy Pittard sugests to order the base station as a precauion, jut in case we cannot pair the new footswitch. Also order the charger.. that may have been lost/thrown away. Internal Remarks : ##response from David Watson to myself dated the 3rd Oct 2000000003 Hi Su Charge against the contract please. Sent from my iPhone Best Regards Dave Watson Service Manager 00000000717787600000000 Internal Remarks : ##from myself to David Watson dated the 3rd Oct 2000000003 abou the parts on this case Am I puting against their contract or creating a fiued price form for £0000000000000000? ----------------------------------------------------------------------------- | 2023-10-05 15:06 | No
1728 | Yes | Alarms or Messages/Device defective / SN DE00Speaker | Alarms or Messages/Device defeu / SN DE000Speaker | 000000000000000000000000001O##: 00000000000000000000000000100000000156181 Problem Description by Engineer: Louspeaker uLL malfuction Parts Consued: 1.0000000000000000000000000000000056000000000000000081.Mu_100000000/u3 Main board 1.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly Resoluion: Speaker, main board replaced ----------------------------------------------------------------------------- | 2023-10-05 15:16 | No
1729 | Yes | Telemetry (Med. Internal): No acouic alarm- insufficient information-Cau unknown-provided infomation | Telemetry (Med. Internal): No acouic alarm- insuffi0nt information-Cau unknown-provided infomation |  | 2023-10-05 16:06 | No
1730 | Yes | loan - turns self on and off when in u | loan - turns self on and off when in u | 000000000000000000000000001O##: 000000000000000000000000001000000001600000000 00000000-Oct-2000000003 Andrew Hammond The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Patient transport coordinator Device ue: Ouside of ue Eupected and ueupected Resuts: loan - tuns self on and off when in ue uer Impact: None Patient Impact: None SE NA Resoluion : Basic performance assuance completed, Software udated to 0000000000000000to resolve any issus that coud occu du to the software being on 000000002. Device retuned to ue Tools & Equpment : PATIENT SIMuATOR,200000000000000005-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000163538 James Gillham The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Patient transport coordinator Device ue: Ouside of ue Eupected and ueupected Resuts: loan - tuns self on and off when in ue uer Impact: None Patient Impact: None SE NA Reguatory Qustions : N/A Follow u Requred : loan - tuns self on and off when in ue. Loan device was suplied to the cutomer and when they went to ue it it was tuning itself off and back on again. Andy Hammond is going to site to check it over. Onsite follow u requred. Diagnostic performed by Engineer : loan - tuns self on and off when in ue. Loan device was suplied to the cutomer and when they went to ue it it was tuning itself off and back on again. Andy Hammond is going to site to check it over. Onsite follow u requred. ----------------------------------------------------------------------------- | 2023-10-05 16:11 | No
1731 | Yes | Therapy test failed | Therapy test failed | 000000000000000000000000001O##: 0000000000000000000000000010000000016000000003 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : To replace 0000000056000000009000000001, DFM100000000 uuuitance ASSY" Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001, DFM100000000 uuuitance ASSY" ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000166856 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-10-05 18:41 | No
1732 | Yes | System keeps resetting after boot up | System keeps resetting after boot up | 000000000000000000000000001O##: 0000000000000000000000000010000000021813000000007-Oct-2000000003 Boncqut Sterling The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000000000000700000000,TouhScreenModue B with Win7 License Problem Description by Engineer : Cutomer Fuction/Role: GE Device ue: in ue Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Replaced TSM-b and loaded SW. System is 000000000000000000000000001king as it shoud. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000218217 13-Oct-2000000003 Boncqut Sterling The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: GE Device ue: in ue Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : Cutomer called me and said the TSMb in the euam room didn’t have the fleu vision buton. The fleu vision app was red so I tried to download it bu it crashed the TSM. Ordered a new TSM for replacement. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000020000000022 13-Oct-2000000003 Boncqut Trent The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: not in ue Eupected and ueupected Resuts: system commuication resetting every 200000000300000000seconds. Host PC replacement resolved issu uer Impact: system down Patient Impact: none SE 8.2.3000000005 Resoluion : Sterling Boncqut went onsite today and replaced HOST PC. Problem has been resolved and system retuned to service. Send uued SIB back for credit. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000005000000002-Oct-2000000003 Boncqut Sterling The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: GE Device ue: in ue Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Resoluion : Replaced the host PC a second time and the system started 000000000000000000000000001king properly. The PC that was ordered by the cutomer was also bad. Once system was on I had to reload the euam room TSM and recreate image store with both IPPCS. I created a ghost and config backus. Tested and system is 000000000000000000000000001king as it shoud. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.00000000000000000000000000000000800000000200000000581,Nehalem Host PC Biplane Rev_III NO_HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000187322 11-Oct-2000000003 Boncqut Trent The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: ouside ue Eupected and ueupected Resuts: system freezing and resetting commuication every 200000000300000000seconds uer Impact: system down Patient Impact: none SE 8.2.3000000005 Diagnostic performed by Engineer : Replaced TS switch and TG switch. Still same issu. Replaced SIB and no change. Reinstalled original SIB. Noticed that the HOST PC that was installed is not compatible with WIN 7. Ordered correct HOST PC for early am delivery. Follow u Requred : system keeps resetting after boot u Follow u Requred : Replace host PC Parts Consued : 1.00000000000000000000000000000000800000000239552,REPL. KIT ETHERNET SWITCH R/M/K CAB 1.00000000000000000000000000000000800000000239552,REPL. KIT ETHERNET SWITCH R/M/K CAB ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000171591 11-Oct-2000000003 Boncqut Trent The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: ouside ue Eupected and ueupected Resuts: system freezes u and commuications reset uer Impact: system down Patient Impact: none SE 8.2.3000000005 Diagnostic performed by Engineer : Assisted Sterling Boncqut with trouleshooting. Still no resoluion. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000157983 00000000-Oct-2000000003 Boncqut Sterling The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Assist Follow u Requred : Install switch and SIB Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: in ue Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a Diagnostic performed by Engineer : Reloaded SW with a new HDD. After it was done the same issu was happening where the system looks fine bu then loses connection with everything. Nothing is powering off, jut jut dropping ou. NSS recommended a new host PC. The GE gus said they will attempt to install it and let u know if they need u again. ----------------------------------------------------------------------------- Internal Remarks : Additional research into faut. Possible bad host PC. Internal Remarks : Assisted with diagnosis and parts. Internal Remarks : Interaction Type: New Case Caller: William Walker IP & Issu: 69336858 & system keeps resetting after boot u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Qute Reason for Entitlement Claim: n/a Action Taken: created qute and set eupectations Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a Internal Remarks : 000000000000000000000000001ked with fse trying to diagnose problem with system cycling , had him check suplies all were good he sent video of problem , possible Host pc issu fse to qute cutomer pc replacement Internal Remarks : FSE Sterling Boncqut. FSE needed license file swap after Host PC replacement. Performed swap and ##d udated lic file to FSE. Internal Remarks : Part looku host Host PC Win 7 00000000800000000200000000582 Internal Remarks : Sterling Boncqut - qustion abou installing system SW. Internal Remarks : Swapped license file for new host. Host ordered by cutomer was for Windows uP not Windows 7. ----------------------------------------------------------------------------- | 2023-10-05 18:51 | No
1733 | Yes | Patient u - FR0ailed to deliver shock | Patient u - FR0failed to deliver shock | 000000000000000000000000001O##: 00000000000000000000000000100000000157911 Thierry LE GONIDEC The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: opening a CFF nuber: PR # is 3928557 Recovery of dae files, eutraction of the utrace that was sent to the cutomer. Diagnostic performed by Engineer: On October 00000000roud 10000000000000000 the defibrillator did not indicate any shock after analysis of an utrace. The staff was eupecting a shock. The staff took a lifepack 15 defibrillator to shock the patient and retrieve a QRS. The patient was able to be recovered. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: patient monitoring Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Reguatory Qustions: No/No ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-05 18:56 | Yes
1734 | Yes | The pressure line sensor needs checking. | The pressure line sensor needs checking. | 000000000000000000000000001O##: 000000000000000000000000001000000001669200000000Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: The pressue line sensor needs checking. uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states it was reported that the device kept giving prouimal pressue sensor disconnect. Troule shot with cutomer. confirmed uLLcode 1200000000 was present in event log. cutomer attached test adapter to device powered on in normal operational mode. adjuted alarm setting. uit operating properly. cutomer removed prou line. device triggered alarm. cutomer reattached prou line. alarm went away. uable to repeat reported problem cutomer states he will continu to test device and call back if futher assistance is needed. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kyle Buger IP & Issu: 58000000000000000000000000- The pressue line sensor needs checking. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-05 19:11 | No
1735 | Yes | İnternal paddle deformed | İnternal paddle deformed | 000000000000000000000000001O##: 0000000000000000000000000010000000016200000000 00000000-Oct-2000000003 KARAKAYA, Mehmet The system meets the specification for the performed service and is retuned to ue. Resoluion: Daily control test was performed on the system and it was tested with an analysis device. It was detected that the internal part and the internal part 000000000000000000000000001ked within the normal valus. The system was delivered in 000000000000000000000000001king condition. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: In ue Eupected and ueupected Resuts: The inter-spoon is defective. uer Impact: No Patient Impact: No SE 2.00000000 Diagnostic performed by Engineer: It was reported that the inter-ading spoon gave an uLL Tools & Equpment : DEFIBRILLATOR ANALYZER, 2000000003-11-18 00000000:00000000:00000000 ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-05 19:16 | No
1736 | Yes | System is not powering on | System is not powering on | 000000000000000000000000001O##: 00000000000000000000000000100000000158311 00000000-Oct-2000000003 Chase, Brian The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: GE service Device ue: clinical Eupected and ueupected Resuts: will not power up uer Impact: same Patient Impact: none SE Resoluion : Dirty contacts cauing voltage drop across ABB AL16-300000000100000000relay SBB.PRS.300000000. Cleaned contacts. Powers properly now. Tools & Equpment : MuTIMETER,200000000000000008-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Sturt HARR IP & Issu: 600000000660000000000000000uS table stop blocking. Event Type: PM Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Qute Reason for Entitlement Claim: Action Taken: Created case and qute Neut Step Action: Awaiting Qute Verbal PO Nuber: 00000000-11000000007 Verbal PO Approver: GE Health Verbal PO CAP: 5k Internal Remarks : 000000002239800000000 Assisted FSE Chase with trouleshooting GSS PDuand power to the system. Foud Relay SBB PRS 300000000 relay that feeds 00000000VAC to the M-cabinet TB00000000as sticking. Cleaned contacts with Deouit D5 and they are 000000000000000000000000001king correctly. System now powers u; however, cutomer is also 000000000000000000000000001king on replacing GEO PC that appears to have a power suply issu. Cutomer will be 000000000000000000000000001king on that. u portion of the repair is complete. ----------------------------------------------------------------------------- | 2023-10-05 19:21 | No
1737 | Yes | [Q] Geometry failure C-arm | [Q] Geometry failure C-arm | 000000000000000000000000001O##: 00000000000000000000000000100000000271100000000 Oct 31, 2000000003 Adov, Aleuander The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE 1.2.5 Resoluion: Reported problem confirmed: C-arc does not 000000000000000000000000001k. Cable set encoder and potentiometer for C drive replaced and fuction test performed. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000700000000100000000,CABLE SET POLY G FLOOR FD12 Tools & Equpment : ELECTRICAL SAFETY TESTER, 200000000000000009-1000000000000000000000000:00000000 MuTIMETER, 200000000000000009-13 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000271000000009 Oct 31, 2000000003 Malhas, Ali The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Resoluion: The reported faut is confirmed. The cable between Roll motor uit and control board replaced. Fuction check IN O. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000232339 Oct 23, 2000000003 Adov, Aleuander The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE 1.2.5 Diagnostic performed by Engineer: Defective spare parts packed and registered for retun. Follow-u service annouced in order to replace the cable set in the C-arm. The problem with the C-arm movement persists. Parts Consued: 1.0000000000000000000000000000000023000000002300000000,Lu BOARD V00000000 1.0000000000000000000000000000000023000000002300000000, MVR HIGH POWER BOARD 1.0000000000000000000000000000000023000000002300000000,Lu BOARD V00000000 1.0000000000000000000000000000000023000000002300000000,CLEA EuTENSION BOARD 2 1.0000000000000000000000000000000023000000002300000000,CONTROL RACK CV 1.0000000000000000000000000000000023000000002300000000CLEA EuTENSION BOARD 1 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018000000007 Oct 11, 2000000003 Adov, Aleuander The system meets the specification for the performed service and is retuned to ue Parts Consued: 1.00000000000000000000000000000000800000000265772,GEO IO BOu 1.00000000000000000000000000000000800000000660000000083,Sute PC 1.00000000000000000000000000000000800000000000000001572,POWER SuPLY u00000000000000001B 1.000000000000000000000000000000002300000000223500000000OTOR C-ARC RED. POLY-G2 CLEA Tools & Equpment: MuTIMETER, 200000000000000009-13 00000000:00000000:00000000 Follow u requred: Schedue time for retun of defective parts Resoluion: Reported problem confirmed: C-arm movement not possible. Trouleshooting performed, GEO IO bou, control modue, power suply for motor and sute PC replaced one after the other, uLLpersists. Internal cabling in the C-arm checked, a contact is high impedance. A consutation is being held as to whether the cable set for the C-arm shoud be ordered. System reassembled, uLLdoes not occu anymore for the time being. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. After consutation with Application, u Lieberknecht, the system also 000000000000000000000000001ked withou problems in the following day. Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE 1.2.5 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000158166 Oct 5, 2000000003 Adov, Aleuander The system meets the specification for the performed service and is retuned to ue Resoluion: Reported problem confirmed: C-arm movements not possible. Performed trouleshooting, replaced boards and rack in the C-arm stand, performed all geometry adjutments of the C-arm and stand. Fuction test OK. The system corresponds to the manuactuer's specifications in the tests performed and is ready for clinical ue. Problem Description by Engineer: Cutomer Fuction/Role: uknown Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE 1.2.5 ----------------------------------------------------------------------------- Internal Remarks: uu 00000000800000000265772 - GEO IO BOu, Line 81 Cutomer ETA: OCT 100000000 2000000003 Internal Remarks: u Adov on bou with the With the order to be in a place. Can he say which spare parts were needed? Then we can take that into accout in the qute and do not need to offer parts that were not ued. Internal Remarks: SPS LF answered by ## Internal Remarks: According to u Adov: Follow-u service, 2 FSE 7 hous each 1u 00000000800000000700000000100000000 Internal Remarks: Via u Bertelmann: The parts from last week were not all needed. u Adov will get in contact later abou how to continu. Internal Remarks: 00OG: What is with this part: LN-1100000000771 GEO IO BOu 00000000800000000265772 1,0000000000000000 Ordered Cutomer comment: ### Arranged appointment with the uer for 10000000031. Internal Remarks: Via u Bertelmann, by ## Hello, please absoluely pass on the costs to Agito in the couse of my other ## See also attachments. Best regards, Florian Bertelmann Internal Remarks: Schedue time for retun of defective parts Internal Remarks: u Adov will clarify by the beginning of neut week which part shoud be offered and will report back. Internal Remarks: According to u Bertelmann, the service can be schedued on 10000000031. ----------------------------------------------------------------------------- | 2023-10-05 19:36 | No
1738 | Yes | Primary alarm issue | Primary alarm issue | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000016 00000000-Nov-2000000003 BASHIR, MouamadouSharif The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The motherboard as well as the t000000000000000000000000001speakers have been replaced. Parts Consued: 1.00000000000000000000000000000000561537200000000RP-PCBA,CPuGEN2, (SW 2.300000000, V6000000001.00000000000000000000000000000000561511921,RP-Speaker Assy, V600000000Tools & Equpment: FLOW METER, 200000000000000007-12 00000000:00000000:00000000 FLOW METER, 200000000000000006-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 2000000003-12-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001670000000017-Oct-2000000003 BASHIR, MouamadouSharif The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred: Hello, Please order the following parts and have them delivered to you site. Please also schedue me on a new 000000000000000000000000001incident for 10000000025/2000000003 as this will be closed. PM board: 000000005615000000001 RP-PCBA,PWR MGMT, GEN 000000001.300000000PIC motherboard: 00000000561537200000000PCBA, CPu(software 2.300000000 Speaker: 00000000561511921 Speaker assembly Parts Consued: 1.00000000000000000000000000000000561511921,RP-Speaker Assy, V600000000Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: Replacing a single speaker did not resolve the problem. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001626900000000ssad MAJEED The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion: On-site intervention requred. Follow u Requred: On-site intervention to replace the Speaker. Please search for resouces. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Diagnostic performed by Engineer: The cutomer reports that the V600000000displays an uLLmessage: “main alarm defective”. Concerning the analysis of the logs, we note the messages: "5000000001, 00000000:50000000000000000M, 10000000000000000-2000000003, POST Power Speaker: FAIL, 959 at 8, 629 at 00000000"5000000001, 00000000:50000000000000000M, 10000000000000000-2000000003, POST Power Speaker: FAIL, 912 at 8, 11700000000t 00000000It seems that the faut comes from the speaker. The cutomer wishes an on-site intervention to replace it. ----------------------------------------------------------------------------- | 2023-10-05 20:01 | No
1739 | Yes | we failed to operate the alarm reset tab and information message tab on the touch screen. | we failed to operate the alarm reset tab and information message tab on the touch screen. |  | 2023-10-05 21:10 | No
1740 | Yes | Touchscreen issue | Touchscreen issue |  | 2023-10-05 21:25 | No
1741 | Yes | The device would not stay in standby. | The device would not stay in standby. | 000000000000000000000000001O##: 00000000000000000000000000100000000168678 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: staying in standby uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and they are uable to keep the uit in standby. While trouleshooting, we noticed the average air was reading -1.600000000slpm. Provided him with the latest service manul rev T Recommended replacing the flow sensor assy Part #989800000000617100000000 flow sensor assy, list price $1217.00000000 Resoluion : Cutomer will order the flow sensor assy and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” ----------------------------------------------------------------------------- Internal Remarks : Caller: Brian Brownell 300000000 577 2800000000 IP & Issu: 69175200000000 - When uit is powered on it goes immediately to ventilation when placed on stand by it will go for a few seconds and back to ventilation Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: ----------------------------------------------------------------------------- | 2023-10-05 22:11 | No
1742 | Yes | Cooling fan issue | Cooling fan issue | 000000000000000000000000001O##: 00000000000000000000000000100000000168500000000 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : No service will be done on this case. Onsite 000000000000000000000000001was requsted on case #0000000022399533 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: cooling fan tuns on and off every 00000000econds uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and the cooling fan is tuning on and off every 00000000econds. Respiratory verified this was the issu. ##d Dennis in biomed to call back when uit is in front of them. It coud be related to the battery connection. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Brooke Power IP & Issu: 58500000000688 Device is stuk in a startu loop.. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-05 22:16 | No
1743 | Yes | The battery does not hold the charge at all | The battery does not hold the charge at all | 000000000000000000000000001O##: 000000000000000000000000001000000001500000000600000000Raffenne, Virginie The system meets the specifications for the performed service and is retuned to ue. Resoluion: From: RAFFENNE, Virginie Sent: Thusday, October 5 2000000003 18:200000000 To: 'interventions.biomed00h-ajaccio.fr' <interventions.biomed00h-ajaccio.fr> Cc: 'commandes00vf-biomedical.com' <commandes00vf-biomedical.com>; "u. Christelle TERTRAIS" <christelle.tertrais00vf-biomedical.com> Suject: Case 0000000022392300000000 - Requst for batteries - V600000000 Hello u Bufa Regarding you requst for V600000000batteries, I invite youto contact the AVF Biomedical company directly at the contact details below to order them becaue the batteries are not inclued in the contract. - By phone at 00000000 99 32 23 76 - By e-mail to commandes00vf-biomedical.com Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: PART ORDER uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: Requsts suply of batteries. ----------------------------------------------------------------------------- | 2023-10-05 22:16 | No
1744 | Yes | Ongoing monitor issue | Ongoing monitor issue | 000000000000000000000000001O##: 00000000000000000000000000100000000535388 00000000-Dec-2000000003 Kocher, Josh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: in ue Eupected and ueupected Resuts: monitor looses live image uer Impact: delay Patient Impact: delay uuResoluion : Replaced DP/CAT7 converter at u Ray PC TP u300000000replaced CAT7 cable from TP u300000000o splitter, replaced Y cable at TZ u00000000 replaced u Ray PC Loaded software to u Ray PC and reloaded system database ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000527836 00000000-Dec-2000000003 Munerlyn, Freddy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Diagnostic Eupected and ueupected Resuts: Flicker image uer Impact: None Patient Impact: None SE 2.2 Resoluion : Replaced DP/CAT7 converter at u Ray PC TP-u300000000Replaced CAT7 cable from TP-u300000000o R Cab Replaced Y-DVI cable TZ-u00000000on back of FleuVision PC Replaced gender changer for Y Cable Replaced u Ray PC Reloaded SW to u Ray PC and re installed database We did not have enouh display single link transceivers for the CAT7 that comes from TP-u300000000n the M Cab so we swapped with a known good one in the B Cab off a splitter; TZ-u00000000 Freddy is aware of which one we swapped and will replace it when the new one is delivered System re booted with no issus and we had image on both live and reference On arrival system had collision detected on the ID Shift cover Replaced Tilt AMC Drive (SAT12) and Tilt NIuBoard (SAT2) Reloaded SW to the AMC Drive Performed all table calibrations Performed frontal stand calibrations Beam prop potentiometer was ou by .00000000V Beam C Arm potentiometer was floating du to loose belt on potentiometer; tightened u and voltage stabilized ID detector shift potentiometer was ou of spec by .25V. Adjuted potentiometer to spec Performed body gurd sensor calibration After completion made u ray; verified live/ref image on fleuvision monitor and performed geometry movements with no uLL. Follow u Requred : N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000337883 28-Nov-2000000003 Munerlyn, Freddy The system meets the specification for the performed service and is retuned to ue Follow u Requred : Need to order additional parts and come back onsite Problem Description by Engineer : Cutomer Fuction/Role: TECH Device ue: DIAGNOSTIC Eupected and ueupected Resuts: Image Flicker uer Impact: None Patient Impact: None SEuknown Resoluion : Ran diagnostics for the monitor, puled Log files and sent to NSS for suport. Issu was escalated and parts ordered for repairs on another 000000000000000000000000001 Recommended to replace u-Ray PC and cable connectors. See 00000000000000000000000000100000000527836 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000159525 200000000Oct-2000000003 Munerlyn, Freddy The system meets the specification for the performed service and is retuned to ue Parts Consued : 2.0000000000000000000000000000000080000000000000000561,Single Link DVI-eutender Cable 13m 1.00000000000000000000000000000000800000000660000000013,FleuViewing PC 00000000 Problem Description by Engineer : Cutomer Fuction/Role: TECH Device ue: LIVE IMAGE ON FLEuVISION BLINKS Ou Eupected and ueupected Resuts: IMAGE BLINKS Ou uer Impact: NONE Patient Impact: NONE SE8.2 Resoluion : REPLACED THE CAT 7 CABLE FROM THE u-RAY PC TO THE DVI CONNECTOR, REPLACED DuL DVI CONNECTOR AT THE SPLITTER IN THE R CABINET. CHECKED ALL CONNECTION ON u-RAY PC. REPLACED FLEuVISION PC AND VERIFIED OPERATIONS. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000268352 28-Nov-2000000003 Munerlyn, Freddy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Diagnostic Eupected and ueupected Resuts: Image flickers uer Impact: None Patient Impact: None SE uknown Resoluion : assisted with the install of the Fleuviewing PC with other engineer. Verified operations. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000170000000089 00000000-Oct-2000000003 Kocher, Josh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: End uer Device ue: In ue Eupected and ueupected Resuts: Screen blanks ou duing cases on fleuvision monitor uer Impact: no display Patient Impact: delay uuResoluion : Assisted primary FSE installing FleuViewing PC; re installed software and back u. Primary FSE will need to retun to install License file; however system is curently operational and able to be ued. usue if license file will need to be uloaded bu recommended. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000059910000000000000000-Dec-2000000003 Munerlyn, Freddy Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Diagnostic Eupected and ueupected Resuts: Flicker image uer Impact: None Patient Impact: None SE 2.2 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Internal Remarks : Assisted FSE Kocher, Josh with trouleshooting live video on fleuvision not showing correctly at time. Approved TPR for replacement fleuvision PC. Internal Remarks : Fred, intermittent problem of Live disappearing at any view port of FV monitor, yet the rabbit ear or CR stays, many parts replaced by different FSEs, bu still happening. Walked thruwith the video roue (non FS), advised to trace/euchange from r-cab splitter, Tu/Ru, CAT 7 to B-cab cables. Internal Remarks : FSE Josh Kocher. FSE called in looking for parts ID assistance. He is looking for TZu00000000that Y's to TZu000000001 and TZu000000002. uable to locate cable. Created action requst to L3 for assistance with cable. Internal Remarks : Josh - requred new license key. Internal Remarks : Planner Initials: GM Action Taken: FSE Freddy Munerlyn accepted call for 12/1500AM Oucome: dispatched Internal Remarks : Josh, 1). uB stick with system backu greyed ou in Restore command. Walked thrua few folders, no sucess. Tried a backu file in the system, 000000000000000000000000001ked. So re-tried uB, 000000000000000000000000001ked. 2). Needs a new license for FV PC. Wiskey server is down. advised to call in weekend, or early Monday. ----------------------------------------------------------------------------- | 2023-10-05 22:31 | No
1745 | Yes | Not allowing to enable fluoro / disc low on space / Patient on the table | Not allowing to enable fluoro / disc low on space / Patient on the table | 000000000000000000000000001O##: 00000000000000000000000000100000000169135 100000000Oct-2000000003 Hobbs Matthew The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.000000000000000000000000989710000000000000000573,Flashlite High End KIT 1.0000000000000000000000009896000000000000000000000000,Detector Pu0000000000000000High Speed pack Resoluion : replaced defoa detector recently installed. Second replacement detector is 000000000000000000000000001king well Problem Description by Engineer : Problem description by engineer : uable to aqure images du to IPPC uLL. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Tools & Equpment : DOSE-KVP-LIGHT METER,20000000000000000000000000000000000000000:00000000:00000000 DOSE-KVP-LIGHT METER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000168986 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : uable to aqure images du to IPPC uLL. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Cutomer states that the systems is down and that they are able to do uuu They are also getting a disc low on space / Patient on the table Follow u Requred : uLL# and/or description of uLL: After 000000000000000000000000001king with Biomed on-site and looking throuh the logs it seems as if the frontal IPPC is having issus. Repair Action : Pattern Name: RMT - DTC: No net000000000000000000000000001k connection Image Detection - IPPCuY Name: 000000003e39bc-5b95-c000000006-0000000090000000000000000d1300000000ffae Root Caue 1: Defective cable between Image detection and IP PC. Defective net000000000000000000000000001k connection in Flat detector controller. Defective net000000000000000000000000001k card in IP PC. Service Action 1: Ru field service test âVerify Image Processing & Hostâ to isolate the problem. Technician Name : Johansson Anders Resoluion : Created on-site 000000000000000000000000001ternal Comments : uLL# and/or description of uLL: After 000000000000000000000000001king with Biomed on-site and looking throuh the logs it seems as if the frontal IPPC is having issus. Repair Action : Pattern Name: RMT - DTC: No net000000000000000000000000001k connection Image Detection - IPPCuY Name: 000000003e39bc-5b95-c000000006-0000000090000000000000000d1300000000ffae Root Caue 1: Defective cable between Image detection and IP PC. Defective net000000000000000000000000001k connection in Flat detector controller. Defective net000000000000000000000000001k card in IP PC. Service Action 1: Ru field service test âVerify Image Processing & Hostâ to isolate the problem. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Negi, Anand 0000000062700000000 IP & Issu: Alua - 6000000000000000000000000 Not letting enabling to do uuu/ disc low on space / Patient on the table Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: Internal Remarks : FSE Matt Hobbs called the quu to place a NFO for 2 parts, one of the part is in backorder. As per FSE Matt, this part was ordered reguarly and escalated to SPS ASAP. The other part was ordered ordered by NFO. ----------------------------------------------------------------------------- | 2023-10-05 23:06 | No
1746 | Yes | Audio-No Sound / No Audio found at Bench s/n u0 / Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at Bench s/n u000 / Other HW-Speaker-HW / Replaced Parts |  | 2023-10-05 23:45 | No
1747 | Yes | Has a display problem | Has a display problem | 000000000000000000000000001O##: 0000000000000000000000000010000000016900000000 11-Oct-2000000003 Hamman, Craig The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: onsite service uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Downloaded and reviewed the event logs (DRPTA). No significant uLL present. Saved DRPTA and attached to the 000000000000000000000000001 Confirmed that the touh screen coud not be properly calibrated. Dismantled the u and euamined for wear or breakages. Replaced Touhscreen, reassembled and tested. Performed and passed the following verification procedues: 1. Preliminary Procedues (section 9.3) 2. Electrical Safety (Test 1) (section 9.00000000 3. Touh Screen Calibration 00000000uLLControls [Test 3] (section 9.6) - 5. Preparing the uLLfor ue (section 9.17) All tests passed and uits retuned to service. Parts Consued : 1.00000000000000000000000000000000561539000000001,RP-Touh Screen, V600000000V600000000Lu/V6800000000Tools & Equpment : TORQu SCREWDRIVER,200000000000000005-31 00000000:00000000:00000000 MuTIMETER,200000000000000005-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000169000000009 Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: onsite service uer Impact: system restricted Patient Impact: none reported SE uknown Follow u Requred : Cutomer states touhscreen will not respond to touh and will not calibrate Resoluion : Creating on site service. Diagnostic performed by Engineer : Cutomer states touhscreen will not respond to touh and will not calibrate. cutomer is requsting on site service to have FSE repair device. ----------------------------------------------------------------------------- Internal Remarks : New Case Caller: Everette Duk Issu: 79185637 Has a display problem Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case an remote 000000000000000000000000001roued to tech suport Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-05 23:51 | No
1748 | Yes | Blower does not ventilate; doesn't blow at all - Final 00 | Blower does not ventilate; doesn't blow at all - Final 0/0 | 000000000000000000000000001O##: 000000000000000000000000001000000001692800000000ccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Provided part nuber and cutomer will order for repair, no futher follow u needed, issu resolved- V600000000Blower assy pn 00000000561512761 $1000000001.00000000 Case follow u ##sent, no response from cutomer. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts:blower stopped while in ue uer Impact: removed from service Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer states while in ue on patient, the blower stopped 000000000000000000000000001king and no ventilation occured, uit alarmed on screen bu there was no blower oupu, uit removed from patient withou futher incident and no harm reported, in biomed the complaint is dulicated and in Diagnostics the blower rpm remains at 30000000000000000when flow is increased, no oupu from blower, requsts replacement part to correct the issu. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000177000000001 Parts Consued : 1.00000000000000000000000000000000561512761,RP-Blower Assy,V600000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Andrew Dulap IP & Issu: 5900000000300000000 Blower does not ventilate; doesn't blow at all Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-05 23:51 | No
1749 | Yes | V0is displaying low minute volume message | V0is displaying low minute volume message | 000000000000000000000000001O##: 0000000000000000000000000010000000016900000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000not display low minue volue message. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying low minue volue message. Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000is displaying low minue volue message. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000000000000/2000000003 uer Trouleshooting Steps Taken: Advised cutomer to reset the V600000000to defaut settings and then reprogram the uit with the following settings. Mode: S/T IPAP: 15 cmH2O EPAP: 00000000mH2O Biomed is reporting the V600000000has stopped alarming. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Wilbert 31357610000000000000000u 65800000000 IP & Issu: SN 100000000000000007255 . Getting uLLthat can not get rid of, low ventilation uLL Euhalation was going back and fouth real fast. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created case and Roued to RTECH Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-06 00:16 | No
1750 | Yes | Jeanine from GE , problem with the imaging computer | Jeanine from GE , problem with the imaging computer | 000000000000000000000000001O##: 0000000000000000000000000010000000017000000000000000011-Oct-2000000003 Serrano Jessica The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: Memory ful uLL - can not uuuor eupose uer Impact: cannot store any images Patient Impact: none SE 8.1.17 Resoluion : did a re create image storage - monitored issu for a week and had no issus Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000160000000096 00000000-Oct-2000000003 Serrano Jessica The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: Memory ful uLL - can not uuuor eupose uer Impact: cannot store any images Patient Impact: none SE 8.1.17 Diagnostic performed by Engineer : Re created image store in FSF - rebooted system and no uLL. Checked logfiles, POST resuts, ran diagnostics on Image Processor - no issus. If issu occus again HD will be replaced and board software will be loaded Follow u Requred : Monitor issu- if issu retuns replace Hardrives ----------------------------------------------------------------------------- | 2023-10-06 00:21 | No
1751 | Yes | DA PCBA ADC failure | DA PCBA ADC failure |  | 2023-10-06 02:44 | No
1752 | Yes | 00POST) Check Vent: Primary Alarm Failed | 00 (POST) Check Vent: Primary Alarm Failed |  | 2023-10-06 03:06 | No
1753 | Yes | The system does not start.(FleuvisionPC) | The system does not start.(FleusionPC) | 000000000000000000000000001O##: 0000000000000000000000000010000000017100000000 27-Oct-2000000003 Toyama, Takafui The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer: I investigated the problem that occured the other day on October 6, in which the equpment cout was 00000000:00000000 and did not 000000000000000000000000001k. It was restarted that day and it was able to be restarted, and the problem was not reprodued util today, bu I confirmed from the equpment log that an uLLhad occured on a pc uit internal controlling the large-screen monitor. Resoluion: [Parts ued] Part name: u Haswell Fleuvision 2 PC no HDD/Part Nuber: 0000000080000000012200000000/シリアル番号 722300000000217882/個数: 1 Resoluion: I replaced the HDD in line with the compatibility between the large screen control PC uit and the PC in the Machine Room. I will make a separate report. I udated the software on the equpment, so I have replaced it with a new PC that is more reliable. Following the replacement, I installed the software and I was able to verify that images were displayed normally on the large-screen monitor. I also confirmed that it is possible to set/change the layou on the touh screen. Internal Comments: November 2, 2000000003 / Following the completion of the sitution sitution, I reported to u Korai in Radiology that there was no problem, and he agreed. Reguatory Qustions: N/A Problem Description by Engineer: Problem description by engineer : The system does not start.(FleuvisionPC) u Frihata - After "00000000:00000000" was displayed, it stopped at the screen of a person's face, and it did not retun to the normal screen. He shu it down, bu the power does not shu off. ＝＝＝ ・I had him wait and it was shudown and improvement was made with a restart. ・He said, "Please take a look at it at the same time as this month's inspection." Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system does not start.(FleuvisionPC) u Frihata - After "00000000:00000000" was displayed, it stopped at the screen of a person's face, and it did not retun to the normal screen. He shu it down, bu the power does not shu off. ＝＝＝ ・I had him wait and it was shudown and improvement was made with a restart. ・He said, "Please take a look at it at the same time as this month's inspection." uer Impact: No effect Patient Impact: No effect SE 8.2.300000000 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.300000000Resoluion: I replaced the HDD in line with the compatibility between the large screen control PC uit and the PC in the Machine Room. I will make a separate report. I udated the software on the equpment, so I have replaced it with a new PC that is more reliable. Following the replacement, I installed the software and confirmed that images were displayed normally on the large-screen monitor and that the layou coud be set/changed on the touh screen. Parts Consued: 1.0000000000000000000000000000000080000000012200000000,u Haswell Fleuvision 2 PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000161368 Kamata, Takayui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : The system does not start.(FleuvisionPC) u Frihata - After "00000000:00000000" was displayed, it stopped at the screen of a person's face, and it did not retun to the normal screen. He shu it down, bu the power does not shu off. ＝＝＝ ・I had him wait and it was shudown and improvement was made with a restart. ・He said, "Please take a look at it at the same time as this month's inspection." Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system does not start.(FleuvisionPC) u Frihata - After "00000000:00000000" was displayed, it stopped at the screen of a person's face, and it did not retun to the normal screen. He shu it down, bu the power does not shu off. ＝＝＝ ・I had him wait and it was shudown and improvement was made with a restart. ・He said, "Please take a look at it at the same time as this month's inspection." uer Impact: No effect Patient Impact: No effect SE 8.2.17 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.2.17.00000000Diagnostic performed by Engineer: ・The neut day (1000000006) in the DailyLog, I confirmed that the Fleuvision PC had a startu uLLdu to a "grabber card" uLL ・Dispatched at the same time as the 10000000027 inspection (FSE Toyama) Internal Comments: uLL# and/or description of uLL: [Suject] 00000000DS:TK:u Frehata: The system does not start.(FleuvisionPC) [Internal Comments] ＝＝＝ 1000000002000000003-10000000000000000 00000000:18:13.0000000000000000uLL Fatal: Failed to initialize all grabber cards. Details: Failed to open grabber channel 00000000 Failed to open grabber channel 1, Failed to open grabber channel 2, Failed to open grabber channel 3, [SEVERITY] Fatal uLL uLLStartu FleuVision 600000000 2000000003-10000000000000000 00000000:22:25.0000000000000000uLL Fatal: u.PmsComp.Au.Cockpit.Private.TsmControl.Component.FieldService.Miscellaneou: FleuVision Monitoring (MediaWall : Software Operational) : FAILED [SEVERITY] Fatal uLL uLLStartu FleuVision ＝＝＝ [u-ed parts] 0000000080000000058600000000:FleuVision-2_Revised PC no HDD QTY:1 ZREP/Traceable Tokyo (u11) Reason for u-collection: Consuted with engineer. Resut of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Keiji Toyama] Toyama, Takafui - FSE1 Visit date and time 2000000003/100000000279:00000000 - FSE1 000000000000000000000000001k time 2 [Euamination category] Other: Repair Action : [Suject] 00000000DS:TK:u Frehata: The system does not start.(FleuvisionPC) [Internal Comments] ＝＝＝ 1000000002000000003-10000000000000000 00000000:18:13.0000000000000000uLL Fatal: Failed to initialize all grabber cards. Details: Failed to open grabber channel 00000000 Failed to open grabber channel 1, Failed to open grabber channel 2, Failed to open grabber channel 3, [SEVERITY] Fatal uLL uLLStartu FleuVision 600000000 2000000003-10000000000000000 00000000:22:25.0000000000000000uLL Fatal: u.PmsComp.Au.Cockpit.Private.TsmControl.Component.FieldService.Miscellaneou: FleuVision Monitoring (MediaWall : Software Operational) : FAILED [SEVERITY] Fatal uLL uLLStartu FleuVision ＝＝＝ [u-ed parts] 0000000080000000058600000000:FleuVision-2_Revised PC no HDD QTY:1 ZREP/Traceable Tokyo (u11) Reason for u-collection: Consuted with engineer. Resut of [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Keiji Toyama] Toyama, Takafui - FSE1 Visit date and time 2000000003/100000000279:00000000 - FSE1 000000000000000000000000001k time 2 [Euamination category] Other: Follow u Requred: * ----------------------------------------------------------------------------- | 2023-10-06 05:11 | No
1754 | Yes | Euternal paddle problem | Euternal paddle problem | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000912 Parts Consued : 1.0000000000000000000000000000000056300000000991,Paddle Repl. Kit Water Resistant ----------------------------------------------------------------------------- | 2023-10-06 05:31 | No
1755 | Yes | SA:0/u FD0Unable to power on | SA:0R/u FD0Unable to power on | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 11.1000000002000000003 13:1000000009 uC (8600000000000000003) Service content: No response after booting, uable to enter the system normally, replace the power suply and FDC board, and the device will 000000000000000000000000001k normally. Service Resut: Equpment is fuly fuctional. Other important matters: None yet. Internal Remarks: 11.1000000002000000003 13:16:00000000uC (8600000000000000003) Ower lost RDI Accuacy: Yes $Q7: (Confirm reported problem, system can't start u) $Q8: (on site check power and led of fdc and MDY, foud power to pc bou lost ) $Q9: (replace power tray and FDC. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name ( NA ) SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: uable to boot SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 0000000022000000000000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-06 07:11 | No
1756 | Yes | Power issue | Power issue | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : 1600000000800000000-Power issu SAP Transfer : Dispatched for Onsite Suport SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 000000002200000000179 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable ----------------------------------------------------------------------------- | 2023-10-06 07:51 | No
1757 | Yes | u0- Control panel cannot be locked - in motion / Resolved - Adju the solenoid valve & spring | u0 - Control panel cannot be locked - in motion / Resolved - Adju the solenoid valve & spring | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Control Panel occasionally fails to lock FSE LiuYao 000000008600000000911700000000 called Euternal Remarks: 00000000.1000000002000000003 12:52:300000000TC (860000000063900000000 Service content: The control panel cannot be locked, the control panel locking fuction is invalid, adjut the control panel solenoid valve and spring. Service resuts: After repair, the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.1000000002000000003 12:52:37 uC (860000000063900000000 $Q1: (Y) $Q2: (N) $Q3: () $Q00000000() $Q5: () $Q6: (Control panel problem) "RDI Accuacy:Yes " $Q7: Confirmed reported problem, the Control panel can't lock when ued. $Q8: Ru control panel lock fuction test foud the Control panel can't lock when ued. $Q9: Adjut control panel Lock. $Q100000000Diagnostic System Test passed; System Fuctional Test passed, Device retuned to ful fuctionality. $Q11: Calibrated Tools name Flue S/N 269900000000800000000Neut Calibration date 6/21/20000000000000000Safety Qustions: Case: 000000002200000000273 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Retun visit sucessfu SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-06 08:36 | No
1758 | Yes | The live monitor in the euamination room does not display. | The live monitor in the euamination room does not display. | 000000000000000000000000001O##: 00000000000000000000000000100000000300000000896 300000000Oct-2000000003 ua, Koji The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Duing inspection Eupected and ueupected Resuts: Lateral live images are not displayed. uer Impact: None Patient Impact: None SE8.1.3000000005 Resoluion: 000000000000000000000000001k details: The lateral live images for the polygraph in the euamination room were not displayed. I confirmed that the signal cable was cu, so I replaced the cable. Operation is fine. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000172000000003 ----------------------------------------------------------------------------- Internal Remarks: IGT: LAN cable 300000000m (ShikokuChildren's and Ononono Medical Center) 2000000003A-000000006359 000000009765929 ----------------------------------------------------------------------------- | 2023-10-06 09:06 | No
1759 | Yes | 0 u Hu0/0- TB charger shorted and caught fire | 00u Hu000- TB charger shorted and caught fire | 2000000003-11-10000000019:21:500000000**claims***No response from cu. Closing case.;2000000003-11-00000000 18:300000000:***Claims*** ##d cu 3rd reminder. Awaiting response.;2000000003-10000000018 19:2000000006:***claims*** ##d cu to follow u with safety qustions. Awaiting response.;2000000003-10000000000000000 17:00000000:300000000***claims*** ##d cu with safety qustions and FedEu retun label. The tracking # is 773677000000006288. Awaiting response. | 2023-10-06 10:31 | No
1760 | Yes | Could not perform fluoro | Could not perform fluoro | 000000000000000000000000001O##: 00000000000000000000000000100000000171831 100000000Oct-2000000003 Kua, Masayui The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Fluros clear did not 000000000000000000000000001k. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Coud not do fluros uer Impact: No effect Patient Impact: No effect SE uknow Resoluion: I was not able to confirm the problem when I visited the site, bu the radiation signal was cu off in the middle of the uLLstatu. Today I replaced the power suply for the related parts. After I replaced it, I confirmed that it was 000000000000000000000000001king withou a problem. Parts Consued: 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000171769 Ohgua, Yui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Fluros clear did not 000000000000000000000000001k. Cutomer Fuction/Role: Radiology Department Device ue: Duing euamination Eupected and ueupected Resuts: Coud not do fluros uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: RC Response: I euplained how to identify the failue. We plan to respond on site. Internal Comments: [Suject] 00000000DS:YO2:u Nakatsuasa: Fluros [Internal Comments] Programming uLL CAN connection with generator Programming uLL CAN connection with generator [Parts from the equpment] 00000000230000000016000000007:MPD CONTROL uIT QTY:1[GATP] ZRFA Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name Kua, Masayui Kua - FSE1 Visit date and time 2000000003/1000000001000000007:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions NBD (miued loading, after the neut day) - Delivery Mut be 2000000003/100000000100000000AM - Delivery address name Yamaguhi Office - Received person name Masayui Kaji - 0000000000000000001600000000-000000006 - Others Requster name Yui Ogua [Euamination category] Normal euamination Follow u Requred: * ----------------------------------------------------------------------------- | 2023-10-06 11:11 | No
1761 | Yes | M0A / u00/ Failed Ops check. - | M00A / u000/ Failed Ops check. - | Test Equpment Info : Nul Device Alarm/Alert as intended? : Yes Problem Observed : failed Ops check Ou Of Bou Failue? : No Labor: Labor ID : L-198100000000 000000000000000000000000001k Performed : Tested uit. Passes Ops check every time. I swapped it ou of this location with another uuufor precauion. Diagonsis : problem not dulicated | 2023-10-06 11:26 | No
1762 | Yes | SPLIT_The loudspeaker output wasn't normal as no sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | SPLIT_The loudspeaker output wasn't normal as no sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. |  | 2023-10-06 11:36 | No
1763 | Yes | Boot up problem | Boot up problem | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 0000000022000000007300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Wendy Cheug Last udated by: Wendy Cheug 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer : Boot u problem SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport FSE Notes : This PO is a Reference Nuber only ----------------------------------------------------------------------------- | 2023-10-06 11:41 | No
1764 | Yes | Service request 00/ SN: A0-00ailed Self Test | Service request 000/ SN: A0F-000Failed Self Test | 000000000000000000000000001O##: 0000000000000000000000000010000000016200000000 PATRICK POELS The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Dealer AED Partner Device ue: ouside of ue Eupected and ueupected Resuts: AED not reliable, does not detect the battery. Dealer tried different batteries uer Impact: cannot ue HS1, not reliable Patient Impact: N/A SE N/A Follow u Requred: This concerns a Bench Repair. Please order the following part for HTC52 attn Patrick Poels: 000000005600000000000000001 - HS1 (C B) Defib, Duch 00000000, Euch Resoluion: AED needs to be replaced ----------------------------------------------------------------------------- | 2023-10-06 11:46 | No
1765 | Yes | Flow sensor out of range | Flow sensor out of range | 000000000000000000000000001O##: 00000000000000000000000000100000000162763 12-Oct-2000000003 SAMAPORN ROMPHOREE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: uLLV600000000 Eupected and ueupected Resuts: Flow sensor ou of range refer preventive maintenanceuer Impact: No impact Patient Impact: No impact SE 2.1 Resoluion : - Replace RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 - Performance Verification test :Pass Parts Consued : 1.000000000000000000000000989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-27 00000000:00000000:00000000 FLOW METER,200000000000000002-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- | 2023-10-06 12:36 | No
1766 | Yes | NEMO: Audio &Sounds/No Sound: The device is reporting speaker uLL, no sound: Speaker/HW: Replaced 000 | NEMO: Audio &Sounds/No Sound: The device is reporting speaker uLL, no sound: Speaker/HW: Replaced 00000 | 000000000000000000000000001O##: 00000000000000000000000000100000000162918 Regina Niemann The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u requred: offer 1u 0000000056000000003831, cutomer will then replace it himself Problem Description by Engineer: Cutomer Fuction/Role: BioMed / Service Technician Device ue: uknown Eupected and ueupected Resuts: speaker failue uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer: uLL Louspeaker is electrically defective, in companion mode both monitors display a louspeaker uLL stand alone, the uLLmessage only remains on the u3, the u3 louspeaker no longer gives off any souds Soluion: Part 0000000056000000003831 mut be replaced System is ou of ue u to the installation of the spare parts(s) A patient was not harmed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000189869 Parts Consued: 1.0000000000000000000000000000000056000000003831.Mu_100000000/u3 Speaker assembly ----------------------------------------------------------------------------- Cutomer comment: Louspeaker uLL are potentially reportable - list of qustions sent by ##to cutomer. Cutomer comment: Cutomer was informed by ##that there is probably no incident to be reported here and a replacement part qute is coming. Cutomer comment: Sent the kVa to MT again Cutomer comment: Today the uLLmessage "louspeaker defective" occured when starting the monitor in companion mode. No soud can be heard anymore. The monitor was replaced. Only on the small monitor did the message remain. No patient was harmed. Kind regards, Stefan Kaase Medical technician Internal Remarks: Auomatic response: Requst for futher information - processing nuber: 0000000022000000000000000000000000 FROM stefan.kaase00vkb.de I will not be in the bulding from 1000000006 to 10000000016/23 incluively. In ugent cases, please contact Medical Technology. ## medizintechnik00vkb.de ----------------------------------------------------------------------------- | 2023-10-06 13:01 | No
1767 | Yes | 0/V0uLLUnable to start the machine | 0K/V0uLLUnable to start the machine | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Neut Step Action: 000000003586000000006351 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 13.1000000002000000003 00000000:25:33 uC (86000000000000000063) $ ASP engineer: [Pangbo Li Hongbin] $Q1:（No） $Q2:（No） $Q3:（Yes） $Q00000000No） $Q7: (Confirmed report problem & the device can't tuned on, and cannot 000000000000000000000000001k normally.) $Q8: (After arriving at the site, it will be confirmed that the machine is can't tuned on and cannot 000000000000000000000000001k normally. & 1. After reconnecting the power cord, the machine is not tuned on, and the faut remains. 2. According to the scene to confirm, the power suply cannot 000000000000000000000000001k.) $Q9: (check the device and replace Power suply (1st generation) ( 00000000561512591)) $Q100000000 (Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality..) $Q11: Calibrated Tools name (Flue Safety Analyzer ESA6200000000 S/N (2886000000006) Neut Calibration date (17/1000000002000000003) Euternal Remarks: 13.1000000002000000003 00000000:25:300000000uC (86000000000000000063) *Department information: [Department: Emergency Ward Contact: Wang Tao Tel: 13586000000006351 ]* $ ASP engineer: [Pangbo Li Hongbin] $Q1:（No） $Q2:（No） $Q3:（Yes） $Q00000000No） $Q7: (Confirmed report problem & the device can't tuned on, and cannot 000000000000000000000000001k normally.) $Q8: (After arriving at the site, it will be confirmed that the machine is can't tuned on and cannot 000000000000000000000000001k normally. & 1. After reconnecting the power cord, the machine is not tuned on, and the faut remains. 2. According to the scene to confirm, the power suply cannot 000000000000000000000000001k.) $Q9: (check the device and replace Power suply (1st generation) ( 00000000561512591)) $Q100000000 (Test and inspection data refer to: PA checklist. Device retuned to ful fuctionality..) $Q11: Calibrated Tools name (Flue Safety Analyzer ESA6200000000 S/N (2886000000006) Neut Calibration date (17/1000000002000000003) Problem Reported by Cutomer: TaizhouEnze Medical Center Can't tun on the machine FSE Notes: This PO is a Reference Nuber only Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: uiyuyang Last udated by: uiyuyang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-06 14:06 | No
1768 | Yes | M0A / u00 Request for batteries (Device is EoL and no work performed by uYy) | M00A / u00 / Request for batteries (Device is EoL and no work performed by uYy) | 000000000000000000000000001O##: 000000000000000000000000001000000001635100000000DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: The part is EOL, therefore try to order it. If it is not possible to order the parts, it is necessary to commuicate the ou of suport. If it is possible, order the parts and ship them to: San Vito al Tagliamento Hospital 2 SAVORGNANO STREET City SAN VITO AL TAGLIAMENTO Was PN Zip code 33000000008 At Clinical E.M., Althea Italia. +390000000000000000683, Framcesco Lapenna. Problem Description by Engineer: Cutomer Fuction/Role: Biomed dept Device ue: no Eupected and ueupected Resuts: The system has low battery. uer Impact: no Patient Impact: no SE na The batteries are euhauted. Diagnostic performed by Engineer: Spoke to u Lapenna who confirmed problems on both the batteries that do not pass the calibrations. ----------------------------------------------------------------------------- | 2023-10-06 14:26 | No
1769 | Yes | LCD brightness is low and cannot be adjued | LCD brightness is low and cannot be adjued | 000000000000000000000000001O##: 0000000000000000000000000010000000021000000003 23-Oct-2000000003 DAMIANI, MAuIZIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: ventilation Eupected and ueupected Resuts: display with low brightness, cannot be adjuted uer Impact: no Patient Impact: no SE na Resoluion: The display is not visible Internal Comments: Carried ou antistatic pad measuement valu 1.1Mhom Parts Consued: 1.00000000000000000000000000000000561528900000000,RP-Screw,M2u3 Socket hd, PKG/100000000V600000000 1.00000000000000000000000000000000561512000000001, RP-Washer, M3 ID u 7mm OD, Pkg/100000000 V600000000 1.0000000000000000000000000000000056153600000000, RP-Cable, LCD, NEC, 2nd Gen, V600000000V6800000000 1.00000000000000000000000000000000561528911, RP-Cable, u PCBA to LCD, V600000000 1.00000000000000000000000000000000561511751, RP-Clamp, Cable, 3/16 Dia, V600000000 1.00000000000000000000000000000000561513131, RP-Tape, Kapton, 3/00000000Wide 1.00000000000000000000000000000000561528911, RP-Cable, u PCBA to LCD, V600000000 1.00000000000000000000000000000000561528891, RP-LCD, Assy, 2nd GEN, V600000000 1.00000000000000000000000000000000561528871, RP-PCBA, uer Interface, u 2nd GEN, V600000000Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000003-15 00000000:00000000:00000000 Resoluion: Replaced spare parts. Carried ou check of the case withou detecting futher anomalies ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000163631 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: The cutomer reports display with low brightness, it cannot be adjuted. Colleagu Mauizio Damiani is aware of the problem and asks to order spare parts. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: ventilation Eupected and ueupected Resuts: display with low brightness, cannot be adjuted uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: The cutomer reports display with low brightness, it cannot be adjuted. Colleagu Mauizio Damiani is aware of the problem and asks to order spare parts. ----------------------------------------------------------------------------- | 2023-10-06 14:41 | No
1770 | Yes | It does not turn on in the morning. | It does not turn on in the morning. | 000000000000000000000000001O##: 0000000000000000000000000010000000017300000000 00000000-Oct-2000000003 ROBERTO BRAN The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: The problem with the system that does not tun on in the morning re-occus. An intervention is needed on site to replace the start-u console. Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: ou of ue Eupected and ueupected Resuts: No uer Impact: No Patient Impact: No SE 000Resoluion: Replaced the system start-u console and ran fuctional tests with a favorable oucome. The problem was solved. The system is released for clinical ue. Parts Consued: 1.00000000000000000000000000000000800000000000000005100000000 Service part REVIEW MODuE ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000017300000000 ROBERTO GHIOTTI The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: The problem with the system that does not tun on in the morning re-occus. An intervention is needed on site to replace the start-u console. Information to suport the complaint handling process Cutomer Fuction/Role: u Device ue: ou of ue Eupected and ueupected Resuts: No uer Impact: No Patient Impact: No SE 000Diagnostic performed by Engineer: Bran's information Internal Comments: uLL# and/or description of uLL The problem with the system that does not tun on in the morning re-occus. An intervention is needed on site to replace the start-u console. Repair Action: Onsite with parts already at the cutomer's Follow u Requred: uLL# and/or description of uLL The problem with the system that does not tun on in the morning re-occus. An intervention is needed on site to replace the start-u console. Repair Action: Onsite with parts already at the cutomer's What are the neut Trouleshooting steps: Onsite with parts already at the cutomer's, swapted from another case, Bran asks for plan Technician Name: ROBERTO GHIOTTI ----------------------------------------------------------------------------- | 2023-10-06 14:41 | No
1771 | Yes | System Failure | System Failure | 000000000000000000000000001O##: 000000000000000000000000001000000001636600000000 Parts Consued : 1.0000000000000000000000000000000056000000001600000000,Intrepid PCCRD System on Modue (SOM) ----------------------------------------------------------------------------- | 2023-10-06 14:41 | No
1772 | Yes | A problem with operation of the AP arm. | A problem with operation of the AP arm. | 000000000000000000000000001O##: 00000000000000000000000000100000000167000000009 Szymon Konarski The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: A remote verification of the uLLlog indicated that an on-site service visit is necessary in order to perform conditioning and adaptation of the u-ray lapmpy in the AP plane and a verification of the euposue foots suject to operation. The device was left with no changes. Problem Description by Engineer: Cutomer Fuction/Role: A 000000000000000000000000001ker in the medical apparatu department. Device ue: neuoradiology laboratory. Eupected and ueupected Resuts: The euposue foots trigger does not 000000000000000000000000001k. uer Impact: not foud. Patient Impact: not foud. SE8.2.17 ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-06 14:51 | No
1773 | Yes | 0/V0uLLDetect ouygen alarm not connected | 0K/V0uLLDeteuouygen alarm not conneud | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: Detect ouygen alarm not connected Internal Remarks: 13.1000000002000000003 00000000:32:00000000uC (86000000001000000005) $ASP engineer: [qiangyi zhang ziyang] $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7: (Confirmed the reported problem, Detect ouygen alarm not connected $Q8:（Restart the machine, Adjuting the ouygen souce pressue, and foud the device normally.So I can ensue that the problem is the ouygen souce pressue. ） $Q9:（Adjuting the ouygen souce pressue.） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11:Calibrated Tools name (NA) S/N (NA) Neut Calibration date (NA) FSE Notes: This PO is a Reference Nuber only Euternal Remarks: 13.1000000002000000003 00000000:32:00000000uC (86000000001000000005) *[Department: Equpment Department Contact Person: Teacher LiuTel: 00000000980000000000000000600000000* $ASP engineer: [qiangyi zhang ziyang] $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7: (Confirmed the reported problem, Detect ouygen alarm not connected $Q8:（Restart the machine, Adjuting the ouygen souce pressue, and foud the device normally.So I can ensue that the problem is the ouygen souce pressue. ） $Q9:（Adjuting the ouygen souce pressue.） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11:Calibrated Tools name (NA) S/N (NA) Neut Calibration date (NA) SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000000000007 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: han suu Last udated by: han suu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-06 15:56 | No
1774 | No | 0 JPN AIR_FLOSS_SU The nozzle was broken off, and the part of the nozzle remained on the handle | 00JPN AIR_FLOSS_SU The nozzle was broken off, and the part of the nozzle remained on the handle | 2000000003-10000000011 00000000:2000000005:A_ occurence date Several years ago.The euact date is uknown. B_ produt puchase date Aroud 2000000003 C_ manuactuer's code uknown D_ I actully heard abou the part nuber Hu8100000000 E_ code change date N/a; 2000000003-10000000000000000 10000000015:300000000PSI500000000; 2000000003-10000000000000000 00000000:10000000056:1_ハンドルにノズル the battery installed, did younotice before cleaning the teeth? uknown 2. Did younotice it in the oral cavity (did youfeel it)? uknown 3. Injuy? uknown 00000000Concern? uknown 5_ Comment (Safety or Qulity jugement) Safety jugement.; 2000000003-10000000000000000 00000000:00000000000000001_症状概要 Is broken 2. Please tell me the details of the sitution when this happened.(Please also tell me abou the sitution if there is any injuy or injuy to the wheren) It seems that a few years ago, it broke when something was being ued. 3. When did the produt start to be ued? 2000000003 Puchase 00000000What is the condition of the produt after the problem occured?Can the appearance of the produt in particuar be confirmed to be fauty or damaged?Please attach the images.(If the problem occus with an accessory, please also tell me the statu of the accessory.) Was broken and there was a part remaining in the main uit 5. Please specifically tell me the frequncy of ue of the produt in qustion (if the problem occus with the accessory, please also tell me the frequncy of ue of that accessory). Random ue from once a month to several times a week. 6. When the problem occured, where (in what room) was the produt ued?(If the problem occus with an accessory, please also tell me abou that accessory.) Wash wash 7. Where was the produt stored? Wash wash 8_ (if applicable) Where did the produt charge? Wash wash When did youpuchase 33_Manor? 20000000000000000.15 has not been replaced since the repair Did the 36-engine system break down duing ue? In ue. Did youcontact 38_euternal (other than the puchased store)?Please provide the details. uknown | 2023-10-06 16:36 | No
1775 | Yes | Device reported not coming ON | Device reported not coming ON | 000000000000000000000000001O##: 00000000000000000000000000100000000198832 11-Oct-2000000003 Ndizihiwe Balinda, Gilbert The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technical/Sonographer Device ue: NO Eupected and ueupected Resuts: System to boot when ON buton is pressed. uer Impact: not able to ue the machine Patient Impact: non SE2.2.3 Resoluion : System is not booting. We foud the power modue damaged on the Power distribuion uit (M-Cabinet). Part to be ordered and replaced before futher trouleshooting. Tools & Equpment : MuTIMETER,200000000000000003-0000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-06 16:41 | No
1776 | Yes | Lights flickering (ball lights) one section of the light bulb has blown. (0FE SEND 00 | Lights flickering (ball lights) one seuon of the light bulb has blown. (0GFE SEND 0-0 | 000000000000000000000000001O##: 00000000000000000000000000100000000175189 Christopher Pimbley The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Diagnostic performed by Engineer : Previou case raised, advised planning to arrange site visit asap. Problem Description by Engineer : Cutomer Fuction/Role: RAD Device ue: IN uE Eupected and ueupected Resuts: DEFECTIVE PART uer Impact: NONE Patient Impact: REPORTED CAuE OF EPILEPTIC SEIZuE. SE 5.6.1 ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-06 16:51 | Yes
1777 | Yes | Defibrillator beeps from time to time. | Defibrillator beeps from time to time. | 000000000000000000000000001O##: 000000000000000000000000001000000001600000000 Kristopher Schardin The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion: Info that specialty dealer will provide the service by ##to the cutomer Problem Description by Engineer: Cutomer Fuction/Role: - Device ue: - Eupected and ueupected Resuts: - uer Impact: Defibrillator beeps from time to time. Patient Impact: - SE - ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-06 16:51 | No
1778 | Yes | The system power does not come on | The system power does not come on | 000000000000000000000000001O##: 00000000000000000000000000100000000189200000000 100000000Oct-2000000003 Tachinami, Makoto The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Engineer Device ue: No Eupected and ueupected Resuts: The equpment power did not come on. uer Impact: No Patient Impact: No SE000Resoluion: I was not able to reprodue the problem, bu I replaced, jut in case, the uit that controls the equpment's power suply. After the 000000000000000000000000001k, there was no problem with operation. Parts Consued: 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000175000000008 00000000-Oct-2000000003 Toyama, Takafui The system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: The system power does not come on.(No power to the M-cabinet PC) u Inano: uhin: 90000000091 ・Other than the DataMonitor pilot lamp flashing, nothing is displayed and the startu screen doesn't display. ・All of the pilot lamps on all other monitors are off. ・Each modue's LED is not flashing and it's off. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system power does not come on.(No power to the M-cabinet PC) u Inano: uhin: 90000000091 ・Other than the DataMonitor pilot lamp flashing, nothing is displayed and the startu screen doesn't display. ・All of the pilot lamps on all other monitors are off. ・Each modue's LED is not flashing and it's off. uer Impact: No effect Patient Impact: No effect SE 8.1.17 Resoluion: Du to the cutomer's circustances, the 000000000000000000000000001k was not carried ou, only Travel was recorded.Not applicable to CSR creation. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000017000000008 Kamata, Takayui The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: The system power does not come on.(No power to the M-cabinet PC) u Inano: uhin: 90000000091 ・Other than the DataMonitor pilot lamp flashing, nothing is displayed and the startu screen doesn't display. ・All of the pilot lamps on all other monitors are off. ・Each modue's LED is not flashing and it's off. Cutomer Fuction/Role: Radiology Department Device ue: Not ued Eupected and ueupected Resuts: The system power does not come on.(No power to the M-cabinet PC) u Inano: uhin: 90000000091 ・Other than the DataMonitor pilot lamp flashing, nothing is displayed and the startu screen doesn't display. ・All of the pilot lamps on all other monitors are off. ・Each modue's LED is not flashing and it's off. uer Impact: No effect Patient Impact: No effect SE 8.1.17 Diagnostic performed by Engineer: ・When the power transformer in the Machine Room was tuned off, the soud of the chir operation also disappeared and power to the main power suply was tuned on. ・Power to the HOST-PC and IPPC inside the M-Cabinet is not on and there is no response when the power switch of the same PC is pressed. ・The MPD main breaker and F2 ~ F5 were tuned off (at this time, there were no breakers that were tripped), and when they were tuned back on, the L1 ~ L3 lamps tuned on. ・After that, it started u normally by restarting. He said he wanted an engineer to come jut to be on the job, so I asked FSE Toyama to respond. Internal Comments: [Suject] 00000000DS:TK:u Inano: The system power does not come on.(No power to the M-cabinet PC) [Internal Comments] Comp [Parts from the equpment] 00000000230000000016000000007:MPD CONTROL uIT QTY:1 ZRFA Tokyo (u11) [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name [Keiji Toyama] Toyama, Takafui - FSE1 Visit date and time 2000000003/1000000000000000022:00000000 - FSE1 000000000000000000000000001k time 2 [Shipping instrutions] ----------------- shipping instrutions #1----------------- - Delivery conditions ASAP (transport arrangement) - Delivery Asap - Delivery address name Send directly to the site - Received person name Keiji Toyama - 0000000000000000007797-9631 - Others Call mobile uon arrival Requster name Takayui Kamata [Euamination category] Other: Follow u Requred: * ----------------------------------------------------------------------------- | 2023-10-06 17:11 | No
1779 | Yes | [Q] Intermittently no euposure when pressing the foot pedal | [Q] Intermittently no euposure when pressing the foot pedal | 000000000000000000000000001O##: 00000000000000000000000000100000000166600000000 Barry White The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Diagnostic performed by Engineer : Monitored , no uLL in logging or futher reports Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Intermittently no euposue when pressing the foot pedal uer Impact: NA Patient Impact: NA SE8.2 ----------------------------------------------------------------------------- Internal Remarks : Have called mutiple times bu not able to speak to Matt Internal Remarks : Spoke to Matt, it is not clear whether on frontal or lateral. There is a new IP PC du for lateral plane. Will monitor for a week ----------------------------------------------------------------------------- | 2023-10-06 18:06 | No
1780 | Yes | No power to the bedside controls, control panel and one of the pumps | No power to the bedside controls, control panel and one of the pumps | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000296 13-Oct-2000000003 Mark Tuker The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: No power to the bedside controls, control panel and one of the pups uer Impact: NA Patient Impact: NA SE7.2 Resoluion : Proscribe failed as resut of reported issu. Proscribe replaced by TSM-B. Software loaded and tested OK Parts Consued : 1.0000000000000000000000000000000080000000000000000393,Proscribe to TSM repl.kit Patient Table 1.00000000000000000000000000000000800000000239552,REPL. KIT ETHERNET SWITCH R/M/K CAB Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000003-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000192670000000012-Oct-2000000003 Mark Tuker The system does not meet the specification for the performed service and is not retuned to ue Reguatory Qustions : N/A Follow u Requred : Replace Touh Screen Modue and/net000000000000000000000000001k switch in R Cab Diagnostic performed by Engineer : First issu resolved….power now restored to geometry control uits and injector. Futher issu now identified- no booting of TSM. TSM and Ethernet switch ordered. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: No power to the bedside controls, control panel and one of the pups uer Impact: NA Patient Impact: NA SE7.2 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000175580000000000000000Oct-2000000003 Mark Tuker The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: No power to the bedside controls, control panel and one of the pups uer Impact: NA Patient Impact: NA SE7.2 Diagnostic performed by Engineer : Cameleon power suply had blown fue. 200000000olt oupu now eratic. Suply has to be replaced Follow u Requred : Replace cameleon power suply in M cabinet Parts Consued : 1.00000000000000000000000020000000000000000600000000262,FuE SLOW 6,25A 1/000000001/00000000 1.00000000000000000000000000000000230000000000000000952,SPARE FuES Tools & Equpment : MuTIMETER,200000000000000003-29 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000175232 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to futher investigate , check F15 in MPDuand F21 in NC Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: No power to the bedside controls, control panel and one of the pups uer Impact: NA Patient Impact: NA SE7.2 Reguatory Qustions : N/A Diagnostic performed by Engineer : Logging is not concluive , supect PSuFue issu in M Cab ----------------------------------------------------------------------------- Internal Remarks : Nuber provided not in service. Internal Remarks : ##sent : Hello u, I have been trying to call youregarding a call logged on you FD200000000. Please can yousave a log file on the system and call me on 00000000867 500000000913 once done. Regards Barry Internal Remarks : Hello, Parts ordered via TAuI – ETA shoud be 9am tomorrow 13/100000000 Any issus we will let youknow. Kind regards, Rachael Stevens Logistics Coordinator u ELECTRONICS u LIMITED Registered Office: Ascent 1, Aerospace Bouevard, Farnborouh, Gu00000000uW. Registered in England No 00000000897 Tel: +00000000(00000000 8000000000000000060000000000000000– u Engineering Line Tel: +00000000(00000000 8000000000000000060000000000000000 – Cutomer Line (Service Parts and Pricing) www.u.co.u The Cutomer Services Portal is 200000000. Create and manage cases, manage produts, view reports and view contracts. Find ou more abou the portal here ________________________________________ The information contained in this message may be confidential and legally protected uder applicable law. The message is intended solely for the addressee(s). If youare not the intended recipient, youare hereby notified that any ue, forwarding, dissemination, or reprodution of this message is strictly prohibited and may be ulawfu. If youare not the intended recipient, please contact the sender by retun e-mail and destroy all copies of the original message. From: Tuker, Mark <mark.tukerucom> Sent: Thusday, October 12, 2000000003 5:100000000PM To: dl_100000000H_CSLogistics <dl_100000000H_CSLogisticsucom>; Planning uI <planning.uiucom>; Parker, Aiden <Aiden.Parkerucom> Cc: Pittard, Andy <Andy.Pittarducom>; Cutis, Scott <scott.cutisucom>; Jones, Aiden <aiden.jonesucom> Suject: RE: **uGENT TECH INFO REQuRED** 00000000000000000000000000100000000200000000296 Importance: High Hi All, Scott has helped and confirmed 0000000080000000000000000393 (TSM-B) and 00000000800000000239552 (ethernet switch) are the parts to order. Forget all others please! Kind Regards Mark From: Tuker, Mark Sent: Thusday, October 12, 2000000003 0000000000000000PM To: dl_100000000H_CSLogistics <dl_100000000H_CSLogisticsucom>; Planning uI <planning.uiucom>; Parker, Aiden <Aiden.Parkerucom> Cc: Pittard, Andy <Andy.Pittarducom>; Cutis, Scott <scott.cutisucom> Suject: RE: **uGENT TECH INFO REQuRED** 00000000000000000000000000100000000200000000296 Hi Again, It is, indeed very confuing. Please order all …..00000000230000000035161 + 00000000200000000000000000000000000000000+0000000080000000000000000300000000 Thanks. Mark. Internal Remarks : logs advised part sumitted not valid, they are chasing FSE to advise asap Internal Remarks : follow u requred, cutomer called to chase/ Advised parts and FSE are being looked into to get someone back asap Cutomer comment : planning added for tomorrow, tried calling on both nubers on 000000000000000000000000001 no answer Cutomer comment : Retun visit with part planned Thusday am. Tried calling no answer Cutomer comment : planned for 100000000100000000am tried to call cutomer bu line is engaged Cutomer comment : called and confirmed for thusday morning with phillip Cutomer comment : Richard called asking for an udate, I told him an engineer will be onsite tomorrow at 9am. ----------------------------------------------------------------------------- | 2023-10-06 18:11 | No
1781 | Yes | C arm is not moving | C arm is not moving | Internal Remarks : Interaction Type: Call New Case Caller: Daniel Browne IP & Issu: 93000000008700000000/ C arm is not moving can ue service. Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Qute Reason for Entitlement Claim: na Action Taken: Created qute Neut Step Action: Awaiting Qute Verbal PO Nuber: na Verbal PO Approver: Daniel Browne/ Aleu Walker Verbal PO CAP: na ----------------------------------------------------------------------------- | 2023-10-06 18:16 | No
1782 | Yes | System not switching on | System not switching on | 000000000000000000000000001O##: 00000000000000000000000000100000000175277 00000000-Oct-2000000003 Moolman, Juie The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: radiographer Device ue: u Eupected and ueupected Resuts: System not swthing on uer Impact: cannot ue the machine Patient Impact: cannot do patients SE8.1 Resoluion : Foud host compuer not booting u. Need to replace host PC. Resoluion : Will provide cutomer with ugrade qutation ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001751900000000oolman, Juie The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Foud host PC Fauty. Have to be replaced. Will provide qutation to cutomer Problem Description by Engineer : Cutomer Fuction/Role: Radiogrpher Device ue: u Eupected and ueupected Resuts: System not swithing on uer Impact: Cannot ue system Patient Impact: No patients can be treated SE8.1 Follow u Requred : CHECK SYSTEM ----------------------------------------------------------------------------- | 2023-10-06 18:16 | No
1783 | Yes | System will not boot, hangs at 0:0 | System will not boot, hangs at 0:0 | 000000000000000000000000001O##: 0000000000000000000000000010000000019500000000 12-Oct-2000000003 Chen,Ying The system meets the specification for the performed service and is retuned to ue Resoluion : The new hard drive was installed and software reloaded. The FleuVision is displaying properly and the system can make euposues and fluro properly. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000175766 100000000Oct-2000000003 Chen,Ying The system meets the specification for the performed service and is retuned to ue Follow u Requred : Order/Replace Hard drive on FleuVision PC Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Resoluion : The new FleuVision PC has been installed, software reloaded and settings udated to 000000000000000000000000001k. Verified the system now displays on the FleuVision monitor and shoots u-rays. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000175678 George Owens The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Call in houe engineer Bill for udate as to what he has foud and address as needed. 900000000 of the time, its the fleuvision PC not commuicating. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: uknown Patient Impact: uknown SE uknown Diagnostic performed by Engineer : Debriefed caller. He has checked all the PC's and they all seem to be powered. Recommended he start FSF uing alt + f and review POST resuts, download board software, and event log to determine caue. He will check TS net000000000000000000000000001k switch and call back if needed. He 000000000000000000000000001ld like FSE dispatched while he does the checks. ----------------------------------------------------------------------------- Internal Remarks : 000000002200000000500000000 FSE Chen is replacing Fleuvision PC and was asking abou the software load. Recommend FSE go into FSF Open profile and download OS/APP for fleuvision there....if needed, download grabber cards after OS/APP load. Internal Remarks : Bill called back and got into FSF and saw collimator uLL, informed him that that 000000000000000000000000001ldnt prevent the system form booting. He will look for fleuvision uLL and wait for the FSE to be dispatched. No futher suport was needed. ----------------------------------------------------------------------------- | 2023-10-06 18:41 | No
1784 | Yes | Wall mount is broken connecting to the CARM making it restricted on the unit. | Wall mount is broken conneung to the CARM making it restriud on the unit. | 000000000000000000000000001O##: 00000000000000000000000000100000000169897 00000000-Oct-2000000003 Mckeon Joseph The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: Diagnostic Eupected and ueupected Resuts: Bent stand stop uer Impact: Bent stand stop Patient Impact: None SE1 Diagnostic performed by Engineer : Removed stop Resoluion : Realigned stop and reinstalled. Verified operation ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Call New Case Caller: Beatriz Patino IP & Issu: 72711200000000all mout is broken connecting to the CARM making it restricted on the uit. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: created R000000000000000000000000001roued Neut Step Action: Awaiting Qute Verbal PO Nuber: na Verbal PO Approver: na Verbal PO CAP: na ----------------------------------------------------------------------------- | 2023-10-06 18:46 | No
1785 | Yes | Had a recent field safety notice about the foot pedal as they need a wired one but they can only u a wireless one | Had a recent field safety notice about the foot pedal as they need a wired one but they can only u a wireless one | 000000000000000000000000001O##: 00000000000000000000000000100000000175820000000019-Oct-2000000003 John Benstead The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000772193,Footswitch CV 3p 00000000Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Wired pedal defective uer Impact: NA Patient Impact: NA uuResoluion : Fitted new wired footswitch as instrutions in Du00000000879 D ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001757500000000Barry White The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : FSE to site to replace wired pedal and test Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Wired pedal defective uer Impact: NA Patient Impact: NA SE1 Diagnostic performed by Engineer : New pedal requred ----------------------------------------------------------------------------- Cutomer comment : 5pm on 19th Oct confirmed with Emma (time agreed with FSE as well) Internal Remarks : Confirmed with site for 20000000000000000----------------------------------------------------------------------------- | 2023-10-06 18:46 | No
1786 | Yes | Presenting problems with the touch. | Presenting problems with the touch. | BuQ&A Feedback: No Reason: Corrective Maintenance Euternal Remarks: 12.12.2000000003 16:23:38 POLARIS (POLARIS) ========= Rafael Simielli - 000000000000000057000000000000000000000000100000000000000000000000051 ===================== == EVALuTION OF EQuPMENT CARRIED Ou IN THE LABORATORY. THE DEFECT REPORTED WHERE THE EQuPMENT DOES NOT 000000000000000000000000001K TOuHSCRE EN. TO CORRECTE THE DEFECT, IT WILL BE NECESSARY TO REPLACE THE TOuH PANEL. ALARM LED AND FILTERS KIT WILL ALSO BE ESTIMATED. IF THE QuTE IS APPROVED, THE PARTS WILL BE REPLACED AND THE EQuPMENT WILL BE SuJECTED TO TESTS, AND THERE MAY BE COMPLEMENTARY PARTS IN THE QuTE. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Test Equpment: Directed to Representative: State/City: RS Representative: GLOBALMED E-mail: m.azzi00lobalmed.com.br Service deadline: 100000000100000000Internal Remarks: 12.12.2000000003 16:23:38 POLARIS (POLARIS) ================= Rafael Simielli - 000000000000000057000000000000000000000000100000000000000000000000051 ===================== == EVALuTION OF EQuPMENT CARRIED Ou IN THE LABORATORY. THE DEFECT REPORTED WHERE THE EQuPMENT DOES NOT 000000000000000000000000001K TOuHSCRE EN. TO CORRECTE THE DEFECT, IT WILL BE NECESSARY TO REPLACE THE TOuH PANEL. ALARM LED AND FILTERS KIT WILL ALSO BE ESTIMATED. IF THE QuTE IS APPROVED, THE PARTS WILL BE REPLACED AND THE EQuPMENT WILL BE SuJECTED TO TESTS, AND THERE MAY BE COMPLEMENTARY PARTS IN THE QuTE. Problem Reported by Cutomer: Problem: Having problems with the touh Priority: Stopped uit: uiversidade Santa Maria Address: Avenida Roraima 10000000000000000Santa Maria-RS 97100000000-900000000 BR Responsible: Mauicio Nogara Sector: Clinical Engineering 001: +555532131671 002: E- mail: mauicio.nogara00bserh.gov.br Safety Qustions: Case: 000000002200000000551 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Adriano Santos Last udated by: Adriano Santos 1. Was the equpment in clinical ue at the time the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Was the equpment defective or failed at the time of installation? uknown ----------------------------------------------------------------------------- | 2023-10-06 18:56 | No
1787 | Yes | Power supply shows no power | Power supply shows no power | 000000000000000000000000001O##: 000000000000000000000000001000000001671600000000DI FOGGIA, FRANCESCO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Nuse in caharge Device ue: no Eupected and ueupected Resuts: The AC light was off. uer Impact: no Patient Impact: no SE na The AC light was off. Resoluion: 00000000\10000000023, 2:300000000pm. Heard from the Rosanna Capozza: last Friday the device had the power suply light off, the problem was re-foud when moving the emergency cart. At the moment the system does not show any problem, eucept for the NBP calibration that will be done with another intervention. I am waiting for information from Peconio. Friday 20000000010000000023: heard from Peconio the problem was attribuable to connection problems with theabatta. The problem has been re-foud. The call is closed. ----------------------------------------------------------------------------- | 2023-10-06 19:06 | No
1788 | Yes | The equipment's functions locked, it was restarted and did not turn on again | The equipment's funuons locked, it was restarted and did not turn on again | BuQ&A Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Euternal Remarks: 3000000001000000002000000003 18:0000000000000000POLARIS (POLARIS) ================= Gulherme Do Rio Costa - 000000000000000057000000001000000000000000000000000000000000000000071 =================== ==== Parts were applied to the uDFM100000000 Defibrillator, 000CN3263000000002, the u Interface Modue, Flat Cable and I/O PCA modue were replaced. The parts above did not resolve the equpment failue, du to incompatibility of the interface modue hardware. Collection of the above parts requsted to retun to stock. Youwill be asked to send new parts for application. NF39637, NF396300000000 NF2200000000 Repair Action: schedued for 10000000016 - FSE Dawson Euternal Remarks: 1000000001.2000000003 00000000:00000000:00000000POLARIS (POLARIS) ================= Gulherme Do Rio Costa - 000000000000000057000000001000000000000000000000000000000000000000071 =================== ==== Parts were applied to the uDFM100000000 Defibrillator, 000CN3263000000002, after applying the parts, the equpment was sujected to Test and Operational where it was possible to detect that it continud to show failue in the Sync Buton and Load Buton. New parts have been removed from the equpment and collection will be requsted. NF0000000081, NF0000000091, 0000000085. SAP Transfer: Dispatched for Onsite Suport Problem Reported by Cutomer: The equpment's fuctions stopped, it was restarted and it did not tun on again SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000659 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Nycolle Silva Last udated by: Nycolle Silva 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? AT Repair Action: schedued for 11/00000000 - fse Gulherme SAP Transfer: Case Transferred to SAP Internal Remarks: 1000000001.2000000003 00000000:00000000:00000000POLARIS (POLARIS) ========== Gulherme Do Rio Costa - 000000000000000057000000001000000000000000000000000000000000000000071 =================== ==== Service not yet completed. Tests and Calibrated Tools in PA Tool. Internal Remarks: 3000000001000000002000000003 18:0000000000000000POLARIS (POLARIS) ======================== Gulherme Do Rio Costa - 000000000000000057000000001000000000000000000000000000000000000000071 =================== ==== Service not yet completed. Tests and Calibrated Tools in PA Tool. ----------------------------------------------------------------------------- | 2023-10-06 19:06 | No
1789 | Yes | Intermittent ureading failure. Tested with different leads and electrodes | Intermittent ureading failure. Tested with different leads and eleuodes | 000000000000000000000000001O##: 0000000000000000000000000010000000017600000000 October 31, 2000000003 Diaz, Aaron The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: I ru tests with the patient simuator, and duing the tests I did not detect loss of usignal. I checked the configuation of the equpment regarding the umeasuement, I modified 2 parameters that were not properly configued and I tested it again with the patient simuator. I still don't get any signal loss duing the tests. After corrective maintenance, the equpment is operational and meets the manuactuer's specifications. The necessary verifications have been performed to certify that the equpment is back to the operating conditions prior to the failue. The equpment has been managed following all the necessary secuity and disinfection measues to avoid the cross contamination. To gurantee the correct operation of the equpment, we recommend them to ue only the accessories and consuables approved by u. Tools & Equpment: PATIENT SIMuATOR, 200000000000000001-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001752900000000Ghazwan Hueidi Sordo The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: uDFM100000000 Defibrillator/Monitor CN32635185 Intermittent ureading failue. Tested with different leads and electrodes Assign fSE udate software and check uconnection. uCARDIAC HEMODYNAMICS # 23/2696 Follow u Requred: uDFM100000000 Defibrillator/Monitor CN32635185 Intermittent ureading failue. Tested with different leads and electrodes Assign fSE udate software and check uconnection. uCARDIAC HEMODYNAMICS # 23/2696 ----------------------------------------------------------------------------- | 2023-10-06 19:06 | No
1790 | Yes | uis not functioning (Problem solved by replacing the ucable from cuomer stock) | uis not funuoning (Problem solved by repla0g the ucable from cuomer stock) | 000000000000000000000000001O##: 00000000000000000000000000100000000172825 Gnanasekar, D The system meets the specification for the performed service and is retuned to ue Resoluion : Problem solved by replacing the ucable from cutomer stock. Diagnostic performed by Engineer : uis not fuctioning Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-06 19:11 | No
1791 | Yes | Aquisition_u-ray problem | Aquisition_u-ray problem | 000000000000000000000000001O##: 00000000000000000000000000100000000176137 Buak Yediyildiz The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE We have CV Centron, SW: 1.000000002.1, SN: 272 and there is problem with acqusition, the following messages are appeared on the screen: Diagnostic performed by Engineer : Please ru Image Processing and Host test from FSF . Please perform recreate image store . It looks image disks or IPPC are defective . There are also other uLL in the logfile . Please euamine the statu of the Ethernet switch (TS). Is the powered u? If not then check Fue F12 located in the NP uit of the M-cabinet. Also euamine the brick power suply (2300000000c/12Vdc) for proper operation. If Fue F12 is blown please euamine the GeoIPC. A defective internal power suply of the GeoIPC (RC) coud be the reason of the blown fue. Replace GeoIPC (RC) when necessary. If all power suplies are Ok, then euamine the control net000000000000000000000000001k cables from HostPC to EthernetSwitch (TS) and SIBBOu. PLease follow manuls below ; ----------------------------------------------------------------------------- | 2023-10-06 19:51 | No
1792 | Yes | The unit does not turn off via power switch | The unit does not turn off via power switch |  | 2023-10-06 19:56 | No
1793 | Yes | AED emits triple chirp, repeatedly gives "Remove and reinsert battery" msg | AED emits triple chirp, repeatedly gives "Remove and reinsert battery" msg | 000000000000000000000000001O##: 00000000000000000000000000100000000176632 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 8 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test Resoluion : Closing case AED no longer has warranty and will be removed from service. ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-06 20:16 | No
1794 | Yes | AED emits triple chirp, repeatedly gives "Remove and reinsert battery" msg | AED emits triple chirp, repeatedly gives "Remove and reinsert battery" msg | 000000000000000000000000001O##: 00000000000000000000000000100000000176773 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 9 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A10000000000000000656 Battery Install Before# 1/2000000007 Lot# Pads Eup Date: 6/29/2000000005 Pads Lot # C122922-00000000----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-06 20:26 | No
1795 | Yes | uLLrestarted due to anomalies detected during operation. (EC 00 | uLLrestarted due to anomalies deteud during operation. (EC 00) | 000000000000000000000000001O##: 00000000000000000000000000100000000176729 Carey, Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue. Eupected and ueupected Resuts: V600000000not log uLLcodes uer Impact: NA Patient Impact: NA SE uknown Problem Description/ueupected Resuts: Agiliti FSE has located a 1100000000 in the V600000000significant event log while completing FCO86600000000000000006 Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: Agiliti FSE has identified diagnostic code 1100000000 in the V600000000significant event log while completing FCO86600000000000000006. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uknown Date and Time of Issu: 10000000000000000/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Biomed has confirmed V600000000diagnostic code 1100000000 in the significant event log. uLLrestarted du to anomalies detected duing operation. This coud inclue invalid or corruted information, uanticipated situtions, or object allocation uLL. NOTE If diagnostic code 1100000000 occus in conjuction with diagnostic code 3000000007 (software euception), no action is requred. Biomed has confirmed diagnostic code 3000000007 in the log. 1. Reinstall operational software. 2. Replace the CPuPCBA. Advised cutomer to proceed with udating the V600000000to the 00000000 generation power management board. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: James Rajala IP & Issu: 70000000052873 FCO power management board. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roured to RSE. Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-06 20:41 | No
1796 | Yes | 0/u FD0Unable to power on | 0R/u FD0Unable to power on | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: uable to power on SAP Transfer: An uLLoccured whilst processing the notification - Euternal Remarks: 1000000001000000002000000003 00000000:56:32 uC (8600000000600000000) Service content (requred): The device cannot be tuned on; replace the MPD Control uit insuance. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 1000000001000000002000000003 00000000:56:38 uC (8600000000600000000) ''RDI Accuacy:YES'' $Q7: Comfirmed the reported problem, uable to power on. $Q8: Field inspection revealed a blown fue in the MPD Control uit. $Q9: Replace the fue of the MPD Control uit. $Q100000000 Test and Inspection data refer to:PA.Device retuned to ful fuctionality. $Q11: Calibrated Tools refer to: PA. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000358 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: han suu Last udated by: han suu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-06 20:41 | No
1797 | Yes | AED emits triple chirp with "Remove and reinsert battery" msg, does not resolve with new battery | AED emits triple chirp with "Remove and reinsert battery" msg, does not resolve with new battery | 000000000000000000000000001O##: 00000000000000000000000000100000000176858 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test uer Impact: NONE Patient Impact: NONE SE N/A Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test Resoluion : Closing case AED no longer has warranty and will be removed from service. Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test.. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 100000000ears old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# B00000000F-00000000383 Battery Install Before# 2000000007/2 Lot#98158P Pads Eup Date: 5/3/20000000000000000 Pads Lot # 211100000000-00000000500000000---------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-06 20:41 | No
1798 | Yes | cannot calibrate touchscreen. | cannot calibrate touchscreen. | 000000000000000000000000001O##: 000000000000000000000000001000000001769900000000Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000pass the touhscreen calibration. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000000000000/2000000003 Trouleshooting Steps Taken: Biomed has attempted mutiple touhscreen calibrations. Cutomer advised to replace the V600000000touhscreen. Advised cutomer the latest version of the service manul is version T. Part # 00000000561539000000001 RP-Touh Screen, V600000000List price is $9900000000age 181 for referencing the repair Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000pass the touhscreen calibration and diagnostics. uer Impact: V600000000has been removed from service. Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000touhscreen is uresponsive when calibrating touhscreen. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Marqus Baker IP & Issu: " 7780000000000000000 cannot calibrate touhscreen." Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-06 21:01 | No
1799 | Yes | Systems stops at 00 | Systems stops at 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000187551 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : SW download issus. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Diagnostic performed by Engineer : Cutomer states that they are in the middle of replacing the FleuVision PC and is having issus getting SW to load. Internal Comments : uLL# and/or description of uLL: Cutomer states that they are in the middle of replacing the FleuVision PC and is having issus getting SW to load. Repair Action : Advised cutomer to replace the HDD of FleuVision PC and reboot system. SW shoud start to load to the new drive which shoud take abou an hou to load. After the SW has finished loading youhave to go into the download board sw and download any "red" items. FleuVision layous will manully configued. Resoluion : Advised cutomer to replace the HDD of FleuVision PC and reboot system. SW shoud start to load to the new drive which shoud take abou an hou to load. After the SW has finished loading youhave to go into the download board sw and download any "red" items. FleuVision layous will manully configued. Will keep 000000000000000000000000001open util 1000000001000000002000000003 in case there are any call backs. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000186822 Coffman, Josh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : RMT - FLV: Fleuvision PC grabber cards are not initialized. 00000000-Oct-2000000003 00000000:0000000052 When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System to boot u normally/System is stuk at 0000000000000000 uer Impact: System got stuk at 0000000000000000 at reboot and uable to come u Patient Impact: No patient involved SE 8.2.3000000005 Diagnostic performed by Engineer : Kenneth called in after installing Fleu PC, wanted to know if he coud ue the previou OS drive. I said he shoud try to ue the drive and if it 000000000000000000000000001ks that great. If it does not 000000000000000000000000001k then I advised him to ue new drive and download board sw. He said he 000000000000000000000000001ld call back if he needed any futher assistance. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000176000000003 Vilkaukas Viktorija The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : RMT - FLV: Fleuvision PC grabber cards are not initialized. 00000000-Oct-2000000003 00000000:0000000052 When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: System to boot u normally/System is stuk at 0000000000000000 uer Impact: System got stuk at 0000000000000000 at reboot and uable to come u Patient Impact: No patient involved SE 8.2.3000000005 Resoluion : Called and left a voic##for Ken. At this time the system is ok with the replacement of the Fleuvision PC. Closing case. Diagnostic performed by Engineer : Ken is reporting that the system is stuk at 0000000000000000 at morning boot u and takes 7 minues to shu down. He booted the system back u in service mode and sent the logfile which is showing that the Fleuvision PC is having an issu with the Grabber Cards. Advised Ken to replace the Fleuvision PC and the hard drive. Ordered parts for Ken. Internal Comments : uLL# and/or description of uLL: Root Caue for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1. Fleuvision PC (TZ) grabber card initialization uLL The grabbercard is malfuctioning. Repair Action : Service actions for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1.1. Please replace the Fleuvision PC (TZ). Diagnostic performed by Engineer : Followed u with Ken and he said that he had some troules loading the new PC, bu it is curently 000000000000000000000000001king. He is missing the Ref. screen. He said he may have a bent pin in the cable. ##d him the setting of the fleuvision before it was changed. He will check those and make sue that it is set correctly. Also asked him to check and make sue that his download board software for the APP is matching eupected and curent valus. He will check the above mentioned items and will let me know if futher assistance is needed. Follow u Requred : uLL# and/or description of uLL: Root Caue for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1. Fleuvision PC (TZ) grabber card initialization uLL The grabbercard is malfuctioning. Repair Action : Service actions for RMT - FLV: Fleuvision PC grabber cards are not initialized.: 1.1. Please replace the Fleuvision PC (TZ). Technician Name : Vilkaukas Viktorija ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000177000000009 Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- | 2023-10-06 21:06 | No
1800 | Yes | Area dose too high at the last euamination | Area dose too high at the last euamination |  | 2023-10-06 21:19 | Yes
1801 | Yes | No prep or euposure while patient was on the table | No prep or euposure while patient was on the table | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000002 00000000-Nov-2000000003 Ferrer Anthony The system meets the specification for the performed service and is retuned to ue Resoluion : Had follow u 000000000000000000000000001O#from Koby Hicks. uS for this room jut needed a PM. uS is now p.m. closing 000000000000000000000000001oblem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: Yes Eupected and ueupected Resuts: None uer Impact: ￼ uS needed a PM no uer impact Patient Impact: ￼ uS needed a PM no patient impact SE N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000185700000000100000000Oct-2000000003 Hicks Koby The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : System 000000000000000000000000001ld not eupose du to uS alarms. Wired aroud uS connector and informed cutomer system is fuctional withou the uS back u. Scheduing battery replacement and PM for uS. Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Follow u Requred : Assist on connecting us back u to system after uS has battery replacement and PM from Toshiba. Tools & Equpment : MuTIMETER,2000000005-00000000-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000180000000080000000000000000-Oct-2000000003 Ferrer Anthony The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer : Went to site and nobody in the OR area to let me in to trouleshoot. Today is Colubu day and saw no staff at site. Cutomer Juian said to call him and no picku from cutomer either at site. Problem Description by Engineer : Problem description by engineer : no prep or euposue Patient on table How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: no prep or euposue Patient on table uer Impact: no prep or euposue Patient on table Patient Impact: patient on table SE 8.1.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000177195 00000000-Oct-2000000003 Small, John The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Trouleshoot CY bou, System interface uit, Host pc, and EZcue Send follow-u 000000000000000000000000001k order to FSE Anthony Ferrer. He is aware that the 000000000000000000000000001O#is going on his schedue for Monday, October 9th 2000000003. Diagnostic performed by Engineer : Was not able to physically trouleshoot system du to cutomer uing the room. Looked at the schematics to find ou possible issus which are the CY bou, System interface uit, host pc, or EZcue Problem Description by Engineer : Problem description by engineer : no prep or euposue Patient on table How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: no prep or euposue Patient on table uer Impact: no prep or euposue Patient on table Patient Impact: patient on table SE 8.1.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000177155 Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : no prep or euposue Patient on table How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: no prep or euposue Patient on table uer Impact: no prep or euposue Patient on table Patient Impact: patient on table SE 8.1.3000000005 Follow u Requred : uLL# and/or description of uLL: no prep or euposue Patient on table Repair Action : Chack attached logfile to see if anything stands ou for issus of no fluro Technician Name : Eckbold Glenn Internal Comments : uLL# and/or description of uLL: no prep or euposue Patient on table Repair Action : Chack attached logfile to see if anything stands ou for issus of no fluro Diagnostic performed by Engineer : Checked uLLlog and all i can see is that u-ray is not possible. System bootu appears normal. ----------------------------------------------------------------------------- Internal Remarks : Koby - uable to uray - en_uray missing - uS transfer switch defective. FSE removed MDY u82 uS juper cable and tested okay. Internal Remarks : John Small - Enable u missing messages. Sugested checking footswitch, SIB and Cue. ----------------------------------------------------------------------------- | 2023-10-06 21:21 | No
1802 | Yes | Alarm, check vent backup alarm failed | Alarm, check vent backup alarm failed | 000000000000000000000000001O##: 00000000000000000000000000100000000176919 Cerchio Robert The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer has a V600000000and it has a check vent back u alarm failed. Biomed verified there is a 1100000000in the significant log. She requsted onsite. Part # 00000000561537200000000 CPupcba, software 2.300000000 ASP 000000000000000000000000001ld need to installed 3.200000000software and reprogram the serial #, add the power on hous, and enable the software options Avaps, C-fleu, ramp, auo trak + and HFT -600000000 Commuicate with the HFT-600000000abou the FCO72. Resoluion : Cutomer decided to replace the CPupcba themselves. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in ue Eupected and ueupected Resuts: check vent backu alarm failed uer Impact: none Patient Impact: none SE un Diagnostic performed by Engineer : Cutomer has a V600000000and it has a check vent back u alarm failed. Biomed verified there is a 1100000000in the significant log. She requsted onsite. Part # 00000000561537200000000 CPupcba, software 2.300000000 ASP 000000000000000000000000001ld need to installed 3.200000000software and reprogram the serial #, add the power on hous, and enable the software options Avaps, C-fleu, ramp, auo trak + and HFT -600000000 Commuicate with the HFT-600000000abou the FCO72. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Alicia 00000000000000003628 IP & Issu: 10000000000000000282/Alarm Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-06 21:21 | No
1803 | Yes | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00/ Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at Bench-Keyrin s/n u000 / Other HW-Speaker-HW / Replaced Parts |  | 2023-10-06 21:30 | No
1804 | Yes | Low leak CO0ebreathing risk uLLcode | Low leak CO0rebreathing risk uLLcode | 000000000000000000000000001O##: 000000000000000000000000001000000001772600000000erchio Robert The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer will order the flow sensor assy and repair uit. “Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event.” Diagnostic performed by Engineer : Cutomer has a V600000000and they have a check vent low leak Co2 rebreathing risk, and uable to keep the uit in standby. While trouleshooting, we noticed the average air was reading -2.300000000slpm. Recommended replacing the flow sensor assy Part #989800000000617100000000 flow sensor assy, list price $1217.00000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: check vent low leak Co2 rebreathing risk uer Impact: none Patient Impact: none SE un ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: George Hall IP & Issu: 566300000000100000000low leak co2 rebreathing risk Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: free remote Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-06 21:36 | No
1805 | Yes | System not booting up | System not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000177793 00000000-Oct-2000000003 Andreoni, Mark The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: No Eupected and ueupected Resuts: The system didn't boot u uer Impact: Yes Patient Impact: Yes SE 8.1 Resoluion : Replaced the FleuVision PC and the system came u and is fuctioning properly. I notified the cutomer that the system is ready to ue for patients. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000017610000000000000000-Oct-2000000003 Andreoni, Mark The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: No Eupected and ueupected Resuts: The system didn't boot u uer Impact: Yes Patient Impact: Yes SE 8.1 Diagnostic performed by Engineer : Foud that there was no power to the B cabinet. Replaced a fue (F100000000 and the B cabinet had power, bu the fleuVision PC 000000000000000000000000001ldn't tun on. Ordering a PC. Follow u Requred : N/A Tools & Equpment : MuTIMETER,200000000000000002-28 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : Mark Andreoni - system does not bootu. Fue F100000000efective. After replacing fue, FleuVision-2 PC does not bootu. Sugested replacing FV2PC. ----------------------------------------------------------------------------- | 2023-10-06 21:36 | No
1806 | Yes | HS0ED is emitting the triple chirp | HS0AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000177373 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Closing case AED no longer has warranty and will be removed from service Diagnostic performed by Engineer : Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. Instruted the Cutomer to do du another self-test to remove battery and reinsert, AED had the same uLLmessage Not ready to ue "Remove and reinsert battery to test. Advise cutomer that the AED has a 8 to 100000000year life span and you curent machine is 19 years old, the issu is with the device. AED no longer has warranty, Instruting the cutomer they will need to remove from service and not ue the device on anyone. AED Serial# A00000000-00000000152 Battery Install Before# 10000000020000000000000000 Lot#82562P Pads Eup Date: 2000000005/000000003 Pads Lot # C100000000322-00000000Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test uer Impact: NONE Patient Impact: NONE SEE.00000000.000000003 Cutomer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test ----------------------------------------------------------------------------- ----------------------------------------------------------------------------- | 2023-10-06 21:41 | No
1807 | Yes | Unable to move and the TABLE is FREE FLOATING | Unable to move and the TABLE is FREE FLOATING | Internal Remarks : Interaction Type: New Case Caller: Jermell HINNANT on behalf of Renee BRIGGS IP & Issu: " 5118000000006 This device is DOWN becaue we're NOT able to move the CAMERA, and the TABLE are FREE FLOATING" Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes, bu this is N/A (IuR) Was the device in Clinical ue at the time of the event? N/A becaue the cutomer stated nothing was reported to him Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: Qute Action Taken: I created a case, qute, and the cutomer gave me this PO#: 00000000112236200000000CAP of $5000000000000000000000000) Neut Step Action: Awaiting Qute Verbal PO Nuber: PO#: 00000000112236200000000CAP of $5000000000000000000000000) Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-06 22:11 | No
1808 | Yes | No stand movement available when booting up system | No stand movement available when booting up system | 000000000000000000000000001O##: 0000000000000000000000000010000000017700000000 00000000-Oct-2000000003 Gashaw, Kunegere The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: No stand movements when system is booted u. Eupected and ueupected Resuts: Enable movements uLLwhen system is booted u. uer Impact: None Patient Impact: NOne SE 8.2.3000000005 Resoluion : Replaced the Geo modue the equpment performed, and checked all the movement of stand fuctionality, specification passed the system retuned to cutomer for clinical ue. Follow u Requred : N/A Parts Consued : 1.00000000000000000000000000000000800000000151715,GEO MODuE WP Kit ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000177173 00000000-Oct-2000000003 Beall Jordan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace Geo TSO Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: No stand movements when system is booted u. Eupected and ueupected Resuts: Enable movements uLLwhen system is booted u. uer Impact: None Patient Impact: NOne SE 8.2.3000000005 Diagnostic performed by Engineer : Tested and verified issu, foud that geo TSO is not 000000000000000000000000001king sometimes. The Cran/Cau movement is sticking all the time. Shoed staff 000000000000000000000000001k aroud and will order and replace GEO TSO. ----------------------------------------------------------------------------- | 2023-10-06 22:16 | No
1809 | Yes | Triple chirp OOB | Triple chirp OOB | 000000000000000000000000001O##: 00000000000000000000000000100000000177700000000Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer called and stated that the AED has been received and is in service. Retun Tracking information 7910000000077000000009 Internal Comments : waiting for order to ship order nuber 63000000009000000009100000000Problem Description by Engineer : Cutomer Fuction/Role: End uer Device ue: Device was not in ue Eupected and ueupected Resuts: Cutomer called and informed that the aed is new ou of the bou and alarming. The AED is Chirping a triple chirp and is uder warranty and to be replaced at zero cost to the cutomer uder IB Warranty Euchange. Device serial# A23g-10000000076 Serial # of battery B00000000200000000900000000 Install before date: 2000000009-00000000-28 0000000059P uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer called and informed that the aed is new ou of the bou and alarming. The AED is Chirping a triple chirp and is uder warranty and to be replaced at zero cost to the cutomer uder IB Warranty Euchange. Device serial# A23g-10000000076 Serial # of battery B00000000200000000900000000 Install before date: 2000000009-00000000-28 0000000059P ----------------------------------------------------------------------------- | 2023-10-06 23:06 | No
1810 | Yes | M0A - u00 Damaged battery | u Damaged battery |  | 2023-10-06 23:35 | No
1811 | Yes | M0A - u00 Damaged battery | u Damaged battery |  | 2023-10-06 23:43 | No
1812 | Yes | M0A - u00- Damaged battery | u- Damaged battery |  | 2023-10-06 23:48 | No
1813 | Yes | System not booting up | System not booting up | 000000000000000000000000001O##: 0000000000000000000000000010000000022791000000003-Oct-2000000003 Rodriguz Lazaro The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced the uS on the system. Tested system is operational. Tools & Equpment : MuTIMETER,200000000000000003-31 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Cathlab Eupected and ueupected Resuts: the compuer only uS intermittently fails. uer Impact: None Patient Impact: Not available SE 000Diagnostic performed by Engineer : Done on previou 000000000000000000000000001by other FSE. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000169533 100000000Oct-2000000003 Henry,Delvin The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000230000000016000000007,MPD CONTROL uIT Problem Description by Engineer : Cutomer Fuction/Role: fse Device ue: na Eupected and ueupected Resuts: uit will not power u uer Impact: none Patient Impact: SE Resoluion : Replace pducontrol uit. uit has issus with teal power conditioner recommended to replace battery back u ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000228520000000000000000Oct-2000000003 ----------------------------------------------------------------------------- Internal Remarks : 00000000220000000000000000 FSE Henry is reporting that system will not start. Had FSE check TB2 in PDuzero volts...FSE wiggled the PDuuS bypass knob on the PDuand was able to get the uS to light u and provide 2300000000AC to the M-cabinet....then FSE coud start the system. FSE will need to contact MTE or APM to determine if the rotary knob is available. If knob not available, can bypass the uS by jupering NP u100000000/100000000. Internal Remarks : Delvin Henry - system does not bootu. FSE checking NP100000000board fues, which are good. Sugested replacing NP100000000board. Internal Remarks : Interaction Type: New Case Caller: John M IP & Issu: same issu as case 000000002239200000000 System not booting u no power to system Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Remote Bypass Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created 000000000000000000000000001 advised FSE will be in contact Internal Remarks : Delvin Henry - system does not bootu. FSE checking NP100000000board fues, which are good. Sugested replacing NP100000000board. ----------------------------------------------------------------------------- | 2023-10-07 00:11 | No
1814 | Yes | Alarming battery malfunction even while being plugged in | Alarming battery malfunuon even while being plugged in |  | 2023-10-07 00:53 | No
1815 | Yes | Audio-No Sound / No Audio found at bench-Keyrin s/n u00/ Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at bench-Keyrin s/n u000/ Other HW-Speaker-HW / Replaced Parts |  | 2023-10-07 00:55 | No
1816 | Yes | Split Speaker Malfunction INOP - no audible alert; u0Cu. Resolved | Split Speaker Malfunuon INOP - no audible alert; u0 Cu. Resolved |  | 2023-10-07 01:00 | No
1817 | Yes | Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 Other HW-Speaker-HW / Replaced Parts | Audio-No Sound / No Audio found at Bench-Keyrin s/n u000/ Other HW-Speaker-HW / Replaced Parts |  | 2023-10-07 01:03 | No
1818 | Yes | Audio-no sound/No audio found at Bench-Keyrin s/n u0other HW-Speaker-HW/Replace | Audio-no sound/No audio found at Bench-Keyrin s/n u00other HW-Speaker-HW/Replace |  | 2023-10-07 01:20 | No
1819 | Yes | Audio-no sound/No audio found at Bench- Keyrin/n u0/other HW-Speaker-HW/Replaced Parts | Audio-no sound/No audio found at Bench- Keyrin/n u00/other HW-Speaker-HW/Replaced Parts |  | 2023-10-07 01:40 | No
1820 | Yes | Device is not blowing enough air | Device is not blowing enough air |  | 2023-10-07 02:29 | No
1821 | Yes | Electrical safety test failed | Eleuical safety test failed | 000000000000000000000000001O##: 00000000000000000000000000100000000178800000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Provided part # and cutomer will order, issu resolved- Material Nuber 00000000561535731 - RP-Plate,Retention,O2 Fitting,V600000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: part id Eupected and ueupected Resuts: Electrical safety test failed uer Impact: PVT failed Patient Impact: none SEun Diagnostic performed by Engineer : Cutomer requsts part id to replace O2 retention plate failing electrical safety test after Flow sensor assembly replaced for failing Air flow test, no patient involvement. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Michael GuIERREZ IP & Issu: " 56628200000000This device is DOWN after replacing the FLOW SENSOR ASSEMBLY" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: I created a case, CSP, R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-07 02:46 | No
1822 | Yes | Speaker Malfunction- No audible sound found at Bench Foulinks-s/nu0other HW /Replaced parts. | Speaker Malfunuon- No audible sound found at Bench Foulinks-s/nu000other HW /Replaced parts. | 000000000000000000000000001O##: 00000000000000000000000000100000000000000008500000000he system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000178800000000 Sanchez, Karbin The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: The uit is reading speaker malfuction. Biomed requsted bench repair. Device ue: Mu00000000 Eupected and ueupected Resuts: The uit is reading speaker malfuction. Biomed requsted bench repair. uer Impact: no Patient Impact: no SE uknown Diagnostic performed by Engineer : The uit is reading speaker malfuction. Biomed requsted bench repair. Follow u Requred : The uit is reading speaker malfuction. Biomed requsted bench repair. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: Ralph Lizza IP & Issu: Cutomer called back to add PO nuber to case Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken:Created AA for Bench Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Ralph Lizza IP & Issu: 6695000000005 - The uit is reading speaker malfuction. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Created case and R000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-07 03:11 | No
1823 | Yes | Unit freezing - will not turn off | Unit freezing - will not turn off | 000000000000000000000000001O##: 0000000000000000000000000010000000017900000000000000005-Oct-2000000003 Alshwaiat, Mohamed The system meets the specification for the performed service and is retuned to ue Resoluion : problem solved by FCO66- 000000000000000000000000001O#reference 00000000000000000000000000100000000312100000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:no operation after start u uer Impact: not uable needs repair Patient Impact: none reported SEun ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000179000000006 Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: uit set u Eupected and ueupected Resuts:no operation after start u uer Impact: not uable needs repair Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer states duing set u after 5 seconds from power on the screen freezes and uit does not cycle, mutiple high priority alarm messages on screen, no patient involvement reported, in biomed the complaint is dulicated and verified PM pcb is pn 1000000005900000000 and uLLcodes in event log are 11300000000115,111B, cutomer requsts FCO86600000000000000006 to replace PM pcb to correct issu, created onsite and AA for follow u scheduing. Follow u Requred : FCO86600000000000000006 PM pcb replacement- uLLcode 11300000000115,111B ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jilbert Hakhverdian IP & Issu: 56629000000002/ uit freezing - will not tun off Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-07 03:36 | No
1824 | Yes | When turning on the equipment I tried to turn it on but it turned off. Diagnosis in description | When turning on the equipment I tried to turn it on but it turned off. Diagnosis in description |  | 2023-10-07 04:36 | No
1825 | Yes | Footswitch not functioning | Footswitch not funuoning | 000000000000000000000000001O##: 0000000000000000000000000010000000017850000000000000000-Oct-2000000003 Maua George The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: AD7 Eupected and ueupected Resuts: footswitch uer Impact: intermittently 000000000000000000000000001king Patient Impact: n/a SE2.25 Resoluion : footswitch not 000000000000000000000000001king/ loose connection to AD7 Table/ adjuted ----------------------------------------------------------------------------- | 2023-10-07 05:11 | No
1826 | Yes | Failure | Failure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 000000002200000000932 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Karen Lee Last udated by: Karen Lee 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer : Failue ----------------------------------------------------------------------------- | 2023-10-08 06:45 | No
1827 | Yes | AED emitted the triple chirp | AED emitted the triple chirp | Internal Remarks: Failue content 2 Alarm souded 3 times.Please inspect. LOT Adut pad C110000000022-00000000 Pediatric pad C000000000000000023-00000000 Battery 80000000000000000----------------------------------------------------------------------------- | 2023-10-08 06:50 | No
1828 | Yes | 0/u00No u-ary | uo u-ary | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000R/u7 M200000000No u-ary Internal Remarks: 00000000.1000000002000000003 00000000:00000000:00000000 uC (860000000013300000000 "RDI Accuacy: Yes" $Q7: Confirmation reported problem, uLLreported in ue, no u ray, patient information lost. $Q8:uLLreported in ue, no u ray, patient information lost. uLLlog See patient image disk ful and image disk failue uLLlog information. $Q9:Inform the uer of the importance of cleaning the patient image in time and bypass the hard disk adapter bou of the u-ray PC. $Q100000000Test and Inspection data refer to PA, Do a few patient, Observation of a certain period of time. The system is ok. Device retuned to ful fuctionality. $Q11:Calibrated tools name NA SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.1000000002000000003 00000000:00000000:00000000 uC (860000000013300000000Department: Interventional Department Contact person: ZhuJu Contact nuber: 138970000000031 Faut phenomenon: uLLreported duing ue, no connection, and patient information is lost. Soluion/maintenance measues: An uLLis reported duing ue, the system does not go offline, and patient information is lost. uLLLog See patient image disk ful uLLlog message for image disk failue.Inform uers of the importance of cleaning patient images in time and bypass the u-ray PC's hard drive adapter bou. The equpment 000000000000000000000000001ks normally after service. SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000000000002 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Pan Chuhu Last udated by: Pan Chuhu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: RSE Gu Qingfeng: Bill to site1 uuYonghu 133276613300000000eut Step Action: An uLLwas reported duing ue and the line did not go ou. Now it is restored, bu the patient information is lost YNN ----------------------------------------------------------------------------- | 2023-10-08 06:55 | No
1829 | Yes | Test Load bad contact | Test Load bad contact | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer : Test Load bad contact SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 000000002200000000763 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Karen Lee Last udated by: Karen Lee 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-08 07:05 | No
1830 | Yes | Quality inspection ufailed | Quality inspeuon ufailed | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: First Affiliated Hospital of Zhejiang uiversity School of Medicine, Yuang Campu The qulity inspection ufailed. The printer door was urecognizable and very loose. Neut Step Action: ASP Honglejia follow u Safety Qustions: Case: 000000002200000000992 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: zhounan Last udated by: zhounan 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Governor lihu has been contacted to verify the specific warranty of this device ----------------------------------------------------------------------------- | 2023-10-08 07:10 | No
1831 | Yes | AED is emitting the triple chirp | AED is emitting the triple chirp | BuQ&R Feedback: No Reason: Corrective Maintenance. Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: Not at all. Created By: Viviane Ferreira. Last udated by: Viviane Ferreira. 1. Was the equpment in clinical ue when the problem was detected? No 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some kind of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? uknown 00000000Did the equpment show a defect or failue when it was installed? uknown Problem Reported by Cutomer: The component is in the u São Carlos 000000000000000000000000001kshop and is ued in ou own offices. These are the requsted information: Latam Linhas Aéreas S.A., Cnpj: 00000000.000000002.862/0000000000000000-600000000 LATAM u -São Carlos - address Rodovia SP-318, Nº km 200000000no bairro Águ Vermelha, São Carlos em São Carlos - SP, CEP 13578-0000000000000000 Daniele of Avila PSE Component Engineering Produt Suport Engineering - Stores (+55) 16 3362-1700000000 VOIP 0000000000000000700000000----------------------------------------------------------------------------- | 2023-10-08 07:45 | No
1832 | Yes | uL battery is insufficient, fails testing.. | uL battery is insuffi0nt, fails testing.. | 000000000000000000000000001O##: 00000000000000000000000000100000000172891 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Cicellini is aware. For a similar problem omniacel 110000000000000000 Last time it was installed Fracastoro Hospital 1 Fontanelle Street 3700000000 San Bonifacio HC c/o Fracastoro Hospital warehoue San Bonifacio Michela Giarolo +3900000000000000000000000000000000 Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: no Eupected and ueupected Resuts: battery is euhauted. uer Impact: no Patient Impact: no SE na The battery is flat. Diagnostic performed by Engineer: Following a PM carried ou by the Sidem technician, battery is foud to be flat. ----------------------------------------------------------------------------- Internal Remarks: Good morning, Please find below the problems foud duing the preventive maintenance at the presidio of San Bonifacio AuSS 9 Scaligera. TeleMon s.n.:u510000000000000000fauty screen HeartStart uL s.n.:u0000000022900000000: impossible to carry ou test for insuficient battery. The cutomer informs that the device is abou to be disposed of. Thanks and have a nice day. En. Matteo Carraro FIELD SERVICE ENGINEER HEALTHCARE IT CONSuTANT Via Bergamo, 900000000 20000000082 Belluco (MB) [tel://+39000000009600000000771/]Tel +39 000000009 60000000067 71 Fau +39 000000009 68 83 289 E-mail matteo.carraro00idemspa.it Website: http://www.sidemspa.com The information contained in this message may be confidential and legally protected uder applicable law. The message is intended intended for the addressee (s). If youare not the intended recipient, youare hereby notified that any ue, forwarding, reprodution of this message is strictly prohibited and may be ufufu. If youare not the intended recipient, please contact the sender by retuning the e-mail and destroying all copies of the original message. ----------------------------------------------------------------------------- | 2023-10-08 08:20 | No
1833 | Yes | Speaker operation failure-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Speaker operation failure-No sound. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000179919 13-Oct-2000000003 Yamada, Sho The system meets the specification for the performed service and is retuned to ue Internal Comments: [Cu related] Coud the person in charge of the equpment agree on how to complete the report?Yes On that day, normal report completed [000000000000000000000000001k report (paper) sumission] Did the 000000000000000000000000001k report be sumitted in paper? YES [Check system operation ouside of failue locations] Implemented Parts Consued: 1.0000000000000000000000000000000056000000008621,MS_u2 ASSY CBL u2/MP2 Speaker Assembly Tools & Equpment : PATIENT SIMuATOR, 200000000000000007-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 2000000003-11-17 00000000:00000000:00000000 PRESSuE CALIBRATION TOOL, 200000000000000005-23 00000000:00000000:00000000 Resoluion: I replaced the speaker. After I replaced it, I confirmed normal operation. Reguatory Qustions: N/A Problem Description by Engineer: ICuSY:u2 (A0000000000000000): Bed 6: Speaker operation problem. Cutomer Fuction/Role: Clinical Engineer Device ue: Duing euamination Eupected and ueupected Resuts: ICuSY:u2 (A0000000000000000): Bed 6: Speaker operation failue. uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: I confirmed the problem. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000179918 Endo, Masaki The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: [Suject] 00000000DS:ME:ME u Watanabe:ICuSY:u2 (A0000000000000000): Bed 6: Speaker operation failue. [Parts from the equpment] 0000000056000000008621:MS_u2 ASSY CBL u2/MP2 Speaker Assembly QTY:1 ZSID --None-- 0000000056000000007761:MS_u2 u2/MP2 Mainboard ver2 QTY:1 ZREP --None-- [Dispatch] 1 RC ----------------- FSE registration information----------------- - Visit FSE1 Name --None-- - FSE1 Visit date and time 2000000003/10000000000000000--None-- - FSE1 000000000000000000000000001k time 2 Follow u Requred: * Problem Description by Engineer: ICuSY:u2 (A0000000000000000): Bed 6: Speaker operation problem. Cutomer Fuction/Role: Clinical Engineer Device ue: Duing euamination Eupected and ueupected Resuts: ICuSY:u2 (A0000000000000000): Bed 6: Speaker operation failue. uer Impact: No effect Patient Impact: No effect SE uknow Diagnostic performed by Engineer: Cold booting the system did not restore the recovery.No soud from the speaker, too. ----------------------------------------------------------------------------- | 2023-10-08 08:25 | No
1834 | Yes | Death:Alarms/No Sound: No sound on the monitor in the nurses station:No Malfunction/ur uLLReplaced Euternal Speaker | Death:Alarms/No Sound: No sound on the monitor in the nurses station:No Malfunuon/ur uLLReplaced Euternal Speaker | 000000000000000000000000001O##: 0000000000000000000000000010000000018000000007 Walker, Maria The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: Yes Caue: Euternal speaker was removed Is Resoluion Known: Yes Resoluion Sumary: uer uLL Resoluion Detail: Connected the euternal speaker to the PC Diagnostic performed by Engineer : Problem Description: Not getting auio alarm tone Not getting auio alarm tone Form ued: Trouleshooting/Investigation Object: N2OBSSuV00000000C Not Object: The other suveillance stations in other floors Software Version: C.00000000.00000000 Date and Time of Issu: 1000000006/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: - Cutomer says the patient in room 200000000was in alarm condition aroud 9PM last night bu there was no soud at the central station and the patient eupired - Cutomer says the central station showed the visul alarm bu there was no soud - Cutomer does not know what the alarm was - Patient was on an Mu00000000 - Cutomer did not have the Mu00000000in qustion - I had cutomer simuate an red and yellow alarm on another Mu00000000and there was no soud - The alarm volue is set to 100000000on the suveillance station - The PC is at the nusing station desk - I had cutomer check the euternal speaker and there is no speaker present - Cutomer got a spare speaker and connected it to the PC and can now hear the soud Was PRS ued: No PRS Permission: Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: Hear alarm tone uer Impact: Not getting auio alarm tone Patient Impact: No impact SE C.00000000.00000000 Problem Description/ueupected Resuts: Not getting auio alarm tone ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ross Brouseau IP & Issu: 71000000009275/No soud on the monitor in the nuses station /Answered Yes to patient harmed Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001roued to RSE/advised RP chat of call in quu Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-08 08:40 | Yes
1835 | Yes | Problem with ventricular fibrillation alarm (GFE states case 000 still in progress 00 | Problem with ventricular fibrillation alarm (GFE states case 0000 still in progress 00 | 000000000000000000000000001O##: 00000000000000000000000000100000000173213 Lämsä, Marku The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Wait for CAS response Problem Description by Engineer: Cutomer Fuction/As: Biomed Device ue: In ue Eupected and ueupected Resuts: Alarm shall appear in time uer Impact: None Patient Impact: None SE C.00000000.00000000 Diagnostic performed by Engineer: Contacted cutomer. I checked the log files. I requsted for ucuves to check with clinical specialist. ----------------------------------------------------------------------------- | 2023-10-08 08:45 | No
1836 | Yes | 0 - u0 - Failed the self-test | 00- u000- Failed the self-test | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: ECR：00000000/HeartStart uL+ Defibrillator/Monitor OneEMS internal remarks: The device failed the self-test and repeated attempts failed. Buiness follow-u is requred. Safety Qustions: Case: 00000000220000000012 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-08 08:55 | No
1837 | Yes | Loudspeaker uLLSound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | Loudspeaker uLLSound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000262788 The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- Internal Remarks: PMM CuTOMER!! u GmbH Market DACH c/o Albertinen Hospital Hambug Medical Süntelstraße 11a 2200000000 Hambug ----------------------------------------------------------------------------- | 2023-10-08 09:20 | No
1838 | Yes | Problem with discharge value. | Problem with discharge value. | 000000000000000000000000001O##: 00000000000000000000000000100000000191300000000 17-Oct-2000000003 Marco Tonazzo The system meets the specification for the performed service and is retuned to ue. Resoluion: Check faut. Replaced condenser pn 0000000056000000006351. The system is 000000000000000000000000001king perfectly again. Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: no Eupected and ueupected Resuts: Suplied energy valus close to the low limit. uer Impact: no Patient Impact: no SE na Suplied energy valus close to the low limit. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000173500000000DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: The part is EOS. We try to order it. Intervention is needed on site. Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: no Eupected and ueupected Resuts: Suplied energy valus close to the low limit. uer Impact: no Patient Impact: no SE na Suplied energy valus close to the low limit. Diagnostic performed by Engineer: Heard from technician Carraro: following pm foud energy suplied valu at low limit. I ordered the part. Intervention is needed on site. ----------------------------------------------------------------------------- | 2023-10-08 09:25 | No
1839 | Yes | M0A - u00- Problem with discharge value | u Problem with discharge value | 000000000000000000000000001O##: 00000000000000000000000000100000000191276 18-Oct-2000000003 Marco Tonazzo The system meets the specification for the performed service and is retuned to ue. Resoluion: Check faut. Replaced condenser pn 00000000563338331. The system is u and runing again. Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: no Eupected and ueupected Resuts: Suplied energy valus close to the low limit. uer Impact: no Patient Impact: no SE na Suplied energy valus close to the low limit. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000173575 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: The part is EOS. We try to order it. Intervention is needed on site. Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: no Eupected and ueupected Resuts: Suplied energy valus close to the low limit. uer Impact: no Patient Impact: no SE na Suplied energy valus close to the low limit. Diagnostic performed by Engineer: Heard from technician Carraro: following pm foud energy suplied valu at low limit. I ordered the part. Intervention is needed on site. ----------------------------------------------------------------------------- | 2023-10-08 09:25 | No
1840 | Yes | Problem with discharge value. | Problem with discharge value. | 000000000000000000000000001O##: 00000000000000000000000000100000000191250000000017-Oct-2000000003 Marco Tonazzo The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: no Eupected and ueupected Resuts: Suplied energy valus close to the low limit. uer Impact: no Patient Impact: no SE na Suplied energy valus close to the low limit. Resoluion: Check faut. Replaced condenser pn 0000000056000000006351. The system is 000000000000000000000000001king perfectly again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000173715 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: The part is EOS. We try to order it. Intervention is needed on site. Problem Description by Engineer: Cutomer Fuction/Role: biomed dept Device ue: no Eupected and ueupected Resuts: Suplied energy valus close to the low limit. uer Impact: no Patient Impact: no SE na Suplied energy valus close to the low limit. Diagnostic performed by Engineer: Heard from technician Carraro: following pm foud energy suplied valu at low limit. I ordered the part. Intervention is needed on site. ----------------------------------------------------------------------------- | 2023-10-08 09:35 | No
1841 | Yes | M0A - u00 Defective Screen | M00A - u00 - Defeu Screen | 000000000000000000000000001O##: 0000000000000000000000000010000000018000000009 Gu DE GREVE The system meets the specifications for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Eupected and ueupected Resuts: uer Impact: None Patient Impact: SE Tools and Equpment: PATIENT SIMuATOR, 200000000000000006-00000000 00000000:00000000:00000000 Resoluion: The described problem was observed: the LCD screen is broken. Corrective maintenance: replacement of the screen (brst effort) For the benefit of the dout, u will take care of this repair. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000017500000000Gu DE GREVE The system does not meet the specifications for the performed service and is not retuned to ue. Follow u Requred: Do 000000000000000000000000001Bench repair 3:00000000 for GDG Repaired by u uder the dout. Problem Description by Engineer: Cutomer Fuction/Role: Biomed Device ue: Eupected and ueupected Resuts: uer Impact: None Patient Impact: SE Resoluion: Do 000000000000000000000000001Bench repair 3:00000000 for GDG Repaired by u uder the dout. ----------------------------------------------------------------------------- | 2023-10-08 09:40 | No
1842 | Yes | Fluoroscopy uLL | Fluoroscopy uLL | 000000000000000000000000001O##: 00000000000000000000000000100000000183500000000 ANDREA COLOMBO The system meets the specification for the performed service and is retuned to ue. Resoluion: reported uLLof fluroscopy, log file of October 9 there is a sporadic episode of "fluroscopy failed, please retry", which lasted some time, followed by a reguar fuctioning of the system. a fuctional check will be carried ou duing the activity in couse of management of case 00000000220000000018 (activity to check dose measuements with Health Physics) Internal Comments: uLL# and/or description of uLL A fluroscopy uLLwas reported. In the log file of October 9 there is a sporadic episode of "fluroscopy failed, please retry". It lasted some time, followed by reguar fuctioning of the system. Repair Action: A fluroscopy uLLwas reported. In the log file of October 9 there is a sporadic episode of "fluroscopy failed, please retry". It lasted some time, followed by reguar fuctioning of the system. Problem Description by Engineer: Problem description by engineer: A fluroscopy uLLwas reported. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: fluroscopy failed, please retry. Issu uer Impact: no Patient Impact: no SE 000Diagnostic performed by Engineer: reported fluroscopy uLL----------------------------------------------------------------------------- | 2023-10-08 09:40 | No
1843 | Yes | The ucable does not work. | The ucable does not work. | 000000000000000000000000001O##: 0000000000000000000000000010000000019790000000027-Oct-2000000003 Laezza Francesco The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000056000000001800000000,Intrepid uSpO2+NBP Frame ASSY Resoluion: Replaced the connector card of the parts applied with a positive oucome, tests and tests were OK. The NBP calibration was carried ou with the positive oucome. Problem Description by Engineer: Cutomer Fuction/Role: Inf Lotito Device ue: no Eupected and ueupected Resuts: Difficuty acquring the u uer Impact: no Patient Impact: no SE na. Problems with the uaqusition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000175000000007 DI FOGGIA, FRANCESCO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: From what was reported by the cutomer the problem arose after fco. From the tests and the videos done, it seems to be a problem of false contacts of the uconnector on the main. I still order the parts for eucess of zeal. It is necessary to calibrate the NBP. Problem Description by Engineer: Cutomer Fuction/Role: Inf Lotito Device ue: no Eupected and ueupected Resuts: Difficuty acquring the u uer Impact: no Patient Impact: no SE na Problems with the uaqusition. Diagnostic performed by Engineer: Heard from engineer Lotito.IL defibrillator does not acqure the uIL problem persists even after changing cable. Moreover, the cutomer reported that the NBP calibration eupired. ----------------------------------------------------------------------------- | 2023-10-08 09:45 | No
1844 | Yes | 0 - u00- Cable is worn out | 00- u0000- Cable is worn out | 000000000000000000000000001O##: 0000000000000000000000000010000000017000000001 FALASCA, IVO The system meets the specification for the performed service and is retuned to ue. Follow u Requred: Make NEMO withou collection at the same time. Tu "F. Miuli" Hospital Strada Prov. 127 Acquviva - Santeramo Km. 0000000000000000 70000000021 Acquviva delle Fonti (BA) Department of Cardiology on 1 floor Ac. Head nuse u. Domenica Vesia Tel. 000000000000000000000000000000008 Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: Hands-Free cable is 000000000000000000000000001n ou and can no longer be ued uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: u. Francesca Losito f.f euplained to me that they need a cable for disposable plates with bayonet connector (not roud) and relative test load. He accepts the shipment of the spare parts. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000190000000025 Parts Consued: 1.000000000000000000000000989800000000100000000831,HeartStart 500000000ohm Test Load 1.0000000000000000000000000000000098000000000000000012, HeartStart Hands-free Cable ----------------------------------------------------------------------------- | 2023-10-08 09:45 | No
1845 | Yes | uuunot working - FCO000 | uuunot working - FCO0000 | 000000000000000000000000001O##: 00000000000000000000000000100000000175332 Vongchavana, Eakchai The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uuuNOT 000000000000000000000000001KING uer Impact: Patient Impact: SE Resoluion : Replaced parts and performed verification test. Parts Consued : 1.00000000000000000000000000000000561530000000051,RP-Bezel, Front, w/NAV, Gray, CTD, V600000000Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000001-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-08 10:05 | No
1846 | Yes | System no power | System no power | 000000000000000000000000001O##: 00000000000000000000000000100000000175337 TSCHuTZ NIKLAS The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: na Device ue: Eupected and ueupected Resuts: uer Impact: Na Patient Impact: na SE Diagnostic performed by Engineer: System has one planned restart and, duing power on nothing happens. Checking power in chain to system and fautfinding. Resoluion: Line 1 to the control modue is not 000000000000000000000000001king. Fue replacement (200000000) and fan DC suply rack replaced. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000175300000000 Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS ----------------------------------------------------------------------------- | 2023-10-08 10:10 | No
1847 | Yes | 0 / SN UNK / Treatment Disabled. | 00/ SN UNK / Treatment Disabled. | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action : Level2工程师 朱睿 185331000000003 Safety Qustions : Case: 00000000220000000031 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: nannan wang Last udated by: nannan wang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer : Treatment Disabled Neut Step Action : 已邮件ASP honglejia跟进 Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-08 10:20 | No
1848 | Yes | Problem with the scopy, they cannot close patients, they cannot get it completely out of the system | Problem with the scopy, they cannot close patients, they cannot get it completely out of the system | 000000000000000000000000001O##: 00000000000000000000000000100000000175997 PATRICK ATTN. The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: No remote soluion possible, an on-site visit has been schedued Instruted the onsite FSE to try to delete the patient list, if that fails, initiate a "recreate image store" Problem Description by Engineer: Cutomer Fuction/Role: n/a Device ue: n/a Eupected and ueupected Resuts: n/a uer Impact: n/a Patient Impact: n/a SE n/a ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000176171 00000000-Oct-2000000003 Verckens, Maarten The system meets the specifications for the performed service and is retuned to ue Resoluion: No Scopy possible, memory ful Problem Description by Engineer: Cutomer Fuction/Role: nuse Device ue: yes Eupected and ueupected Resuts: No. uer Impact: yes Patient Impact: no SE 8.1 ----------------------------------------------------------------------------- | 2023-10-08 10:55 | No
1849 | No | Duplicate to 00 GFE confirmed Sound. | Duplicate to 0000 GFE confirmed Sound. |  | 2023-10-08 11:00 | No
1850 | Yes | Speaker operation uLLmessage-Sound not confirmed. Reboot/Restart resolved issue-Cau Unknown. | Speaker operation uLLmessage-Sound not confirmed. Reboot/Restart resolved issue-Cau Unknown. | 000000000000000000000000001O##: 0000000000000000000000000010000000018000000007 Kitaua, Toshihisa The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: In ue Eupected and ueupected Resuts:HCuSN: The message "Speaker operation uLL appears when the Mu5500000000monitor is ued with the u3 (B00000000SP1) monitor. uer Impact: No effect Patient Impact: No effect SE uknown Resoluion: I had him tun the u3 off/on and re-seat the battery after the patient had left the bed, and he confirmed that the message disappeared.He agreed to ue the AED as it is for now and to wait and see. Internal Comments: [Suject] 00000000CS:TK:Medical Equpment Section, u Shin:HCuSY: The message "Speaker operation problem" appears when the Mu5500000000monitor is ued with the u3 (B00000000SP1) monitor in room 9800000000----------------------------------------------------------------------------- | 2023-10-08 11:40 | No
1851 | Yes | Machine is not booting | Machine is not booting | 000000000000000000000000001O##: 00000000000000000000000000100000000180000000055 100000000Oct-2000000003 Pandey, Rajesh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not booting Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 Resoluion : As cutomer reported the issu, checked and foud some time host pc is not booting,reseated all host pc connection and changed cmos battery. Observed the .foud ok. Handed over the system in good 000000000000000000000000001king condition Tools & Equpment : MuTIMETER,200000000000000002-17 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000180000000015 Kuar, Atu The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: no uLL Repair Action : FSE to visit and check the issu Follow u Requred : uLL# and/or description of uLL: no uLL Repair Action : FSE to visit and check the issu Technician Name : Kuar, Atu Diagnostic performed by Engineer : Machine is not booting. Host PC not switching ON Problem Description by Engineer : Problem description by engineer : Machine is not booting Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.3000000005 ----------------------------------------------------------------------------- | 2023-10-08 12:05 | No
1852 | Yes | System stuck during boot up sequence | System stuck during boot up sequence | 000000000000000000000000001O##: 00000000000000000000000000100000000188232 12-Oct-2000000003 Johnson Dylan The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: system hangs u on boot u. Low disk uLL uer Impact: mut restart system mutiple times. Patient Impact: none SE8.1 Resoluion : Replaced OS HDD and downloaded software. Also replaced PC Bios battery. Tested system and confirmed fuctionality. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018000000001 00000000-Oct-2000000003 Johnson Dylan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace IPPC Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: system hangs u on boot u. Low disk uLL uer Impact: mut restart system mutiple times. Patient Impact: none SE8.1 Diagnostic performed by Engineer : Replaced image drives. IPPC also needs replaced. Replace on follow u. Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000177892 00000000-Oct-2000000003 Johnson Dylan The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Replace IPPC HDD Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: system hangs u on boot u. Low disk uLL uer Impact: mut restart system mutiple times. Patient Impact: none SE8.1 Diagnostic performed by Engineer : Replace IPPC HDD disks per NSS. ----------------------------------------------------------------------------- Internal Remarks : 000000002200000000100000000FSE Johnson reporting that IPPC is not coming u and has to be restarted manully...likely a power suply issu. Approve: 0000000080000000050000000092 - IP PC Revised_II no HDD 0000000080000000050000000000000000- Hard disk spare part If system at Windows 7, advised to go with haswell ippc and host if on contract. ----------------------------------------------------------------------------- | 2023-10-08 12:25 | No
1853 | Yes | 0/uDFM0 Defibrillator/Monitor/Do not recognize the handle and ucable | 0K/Effi0 DFM0 Defibrillator/Monitor/Do not recognize the handle and ucable | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Do not recognize the handle and ucable SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Handle electrodes and ucable not recognized The cutomer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Restarted the machine bu it didn't 000000000000000000000000001k Repair Action: T2 Recommended on-site repair steps: It is recommended to check the handle electrodes and ucable Recommended spare part uufor T2 with description: #M3500000000/M16700000000M1669A Notes on cutomer requrements: NA Test/Inspection: NA Test Equpment: NA Safety Qustions: Case: 00000000220000000062 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Internal Remarks: 12.1000000002000000003 00000000:53:100000000TC (860000000000000000100000000 $ASP engineer: [honglejia uuchungfeng] $Q1:NO $Q2:NO $Q3:N/A $Q00000000/A $Q7:Confirmed reported problem: uable to identify electrode plate and ucable. $Q8:Detected that the electrode plate and ucable cannot be identified. $Q9:After replacing the processing board, the instruent retuned to normal ue. $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11: Calibrated Tools: #ESA 6200000000#S/N 3000000002865 #Neut Calibration date#00000000/00000000/20000000000000000 Calibrated Tools: #Prosim8 #S/N #2600000000000000002 #Neut Calibration date #100000000200000000000000003 #. Calibrated Tools: #70000000000000000P # S/N # 3753000000006 # Neut Calibration date # 10000000029/2000000003 #. Euternal Remarks: 12.1000000002000000003 00000000:53:11 uC (860000000000000000100000000 *Department information: [Department: Equpment Department Contact: Teacher Pang Tel: 151777925800000000* $ASP engineer: [honglejia uuchungfeng] $Q1:NO $Q2:NO $Q3:N/A $Q00000000/A $Q7:Confirmed reported problem: uable to identify electrode plate and ucable. $Q8:Detected that the electrode plate and ucable cannot be identified. $Q9:After replacing the processing board, the instruent retuned to normal ue. $Q100000000Test and inspection data refer to:PA checklist. Device retuned to ful fuctionality $Q11:Calibrated Tools: #ESA 6200000000#S/N #3000000002865 #Neut Calibration date#00000000/00000000/20000000000000000 Calibrated Tools: #Prosim8 #S/N #2600000000000000002 #Neut Calibration date #100000000200000000000000003 #. Calibrated Tools: #70000000000000000P # S/N # 3753000000006 # Neut Calibration date # 10000000029/2000000003 #. ----------------------------------------------------------------------------- | 2023-10-08 13:40 | No
1854 | Yes | uCentron computer is off | uCentron computer is off |  | 2023-10-08 13:50 | No
1855 | Yes | Foot switch intermittently working | Foot switch intermittently working | 000000000000000000000000001O##: 00000000000000000000000000100000000187000000006 12-Oct-2000000003 Riley Earl The system meets the specification for the performed service and is retuned to ue Resoluion : Resoluion: I ordered and replaced footswitch and checked operation. AllParts Medical will do the invoicing for the approuimate price of $0000000000000000.75 via qute#11000000001600000000 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Tech Device ue: In ue Eupected and ueupected Resuts: Sometimes when pressing footswitch nothing happens. uer Impact: Inconvenienced Patient Impact: None SE R2.2.3 Diagnostic performed by Engineer : Diagnostics: Checked logfiles and foud many En-u uLL when footswitch was pressed. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Allen IP & Issu: 92336262- Foot switch intermittently 000000000000000000000000001king Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created qute Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-08 14:00 | No
1856 | Yes | Tube defect | Tube defect | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: PPD 989000000000000000085100000000 1 00000000800000000651161 1 Arrangement after TSE+MM approval FSE Yu Jiangmin-86-18630000000000000000181 SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.1000000002000000003 11:11:58 uC (86000000008200000000 Service content: The equpment cannot be euposed, replace the tue and calibrate it. Service resuts: After maintenance, it was confirmed that all fuctions were operating normally and the equpment retuned to normal operation. SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.1000000002000000003 11:0000000037 uC (86000000008217)"RDI Accuacy:Yes" $Q7:(Confirmed reported problem, the system fluroscopy and euposue failed, and uLLno u-ray,tue defected" display on data monitor) $Q8:(test the small and large focu filament with LCR meter, they are 7.5mH, the small and large focu filament are open ) $Q9:(replaced the tue and performed the system adjutments) $Q100000000 (Test and inspection data refer to PA, performing the cardiac scanning, perfectedimages display on monitors, Device retuned to ful fuctionality) $Q11:Calibrated Tools refer toPA For Mainland China CV / uR, CT engineers only.Please fill in: ^Curent Tue installed date: ^Curent Tue uage: scan second_ Euternal Remarks: 00000000.1000000002000000003 11:00000000300000000TC (86000000008217) Service content: The equpment cannot be euposed, replace the tue and calibrate. Service resuts: After maintenance, it was confirmed that all fuctions were operating normally and the equpment retuned to normal operation. SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Tue failue Safety Qustions: Case: 00000000220000000059 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: Tue failue Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: n/a Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Caue 1 Small focu is defect of Frontal u-ray Tue or kV/mA modue is malfuctioning. Soluion 1 Check the filament indutance uing LCR Meter at Frontal u-Ray Tue side. When the filament is open: Youwill measue abou 7.5 mH. 2. If filament indutances are ok, check the HT cables and its contact pins. If the gap at HT pins are not ok, ue pin gage tool to achieve the requred gap. 3. If HT cables are OK bu problem persists, replace kV/mA modue. Soluion 2 HT Cable fauty If filament indutances are ok, check the HT cables and its contact pins. If the gap at HT pins are not ok, ue pin gaue tool to achieve the requred gap. Caue 2 Doserate limit cannot be kept, Q below Qmin. To keep to the dose rate limit, EDL 000000000000000000000000001ld have to go below the Qmin valu. NOTE: This can be caued as collateral of a malfuctioning u-ray tue. (e.g Gridswitch uLL). Then look at that CAT pattern. Soluion 1 Perform Tue adaptation and EDL calibration. Soluion 2 Check EPu data for correct valus. Soluion 3 Check for any filter in u-ray beam. Soluion 00000000 If problem still euists measue Dose oupu of the u-ray tue. Soluion 5 If to less dose oupu try new u-ray tue. Recommended spare parts uuand designation: uu989000000000000000085100000000 u200000000 000000007 rot-GS 10000000000000000nternal Remarks: 00000000.1000000002000000003 11:12:00000000 uC (86000000008200000000 "RDI Accuacy:Yes" $Q7:(Confirmed reported problem, the system fluroscopy and euposue failed, and uLLno u-ray,tue defected" display on data monitor) $Q8:(test the small and large focu filament with LCR meter, they are 7.5mH, the small and large focu filament are open ) $Q9:(replaced the tue and performed the system adjutments) $Q100000000 (Test and inspection data refer to PA, performing the cardiac scanning, perfectedimages display on monitors, Device retuned to ful fuctionality) $Q11:Calibrated Tools refer toPA For Mainland China CV / uR, CT engineers only.Please fill in: ^Curent Tue installed date: ^Curent Tue uage: scan second_ Neut Step Action: SS gun jianbin 18600000000299918 Call to inqure abou repairs SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-08 14:10 | No
1857 | Yes | Fleuvision PC cannot power on | Fleusion PC cannot power on | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 00000000220000000083 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problem: Screen disk read / write uLLDetailed description of cutomer problem or faut phenomenon: Large screen displays hard disk read uLLIs it a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Large-screen PC shows hard disk read uLLwhen booting Remote diagnosis and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Clean, plu and uplu the hard disk of FleuVision-2_Revised PC Recommended spare parts uuand name: uu 0000000080000000012200000000 u Haswell Fleuvision 2 PC no HDD Cutomer demand notes: onsite SAP Transfer: Case Transferred to SAP Euternal Remarks: 00000000.1000000002000000003 17:13:31 uC (8600000000000000001) Service content: Large screen system PC cannot start, replace the compuer and retun to normal. Service Resuts: Equpment leaves site fuly fuctional. Other important matters: None yet. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 00000000.1000000002000000003 17:13:33 uC (8600000000000000001) RDI Accuacy: Yes $Q7: (Confirmation & Description of problem) Confirmed reported problem, The large screen system PC cannot start. $Q8: (Trouleshooting & Resuts) The large screen system PC cannot start. Replace the compuer and restore it to normal. $Q9: (Repair steps) The large screen system PC cannot start. Replace the compuer and restore it to normal. $Q100000000 (Test and Inspection data & Device statu) Test and inspection data refer to PA, Device retuned to ful fuctionality. $Q11:N/A; SAP Transfer: Case Transferred to SAP Neut Step Action: PPD u Haswell Fleuvision 2 PC no HDD 0000000080000000012200000000 1 12-hou shipment has been arranged, please send 000000000000000000000000001k FSE: Cheng Jianping -86-18820000000030000000000000000Problem Reported by Cutomer: uable to boot normally ----------------------------------------------------------------------------- | 2023-10-08 14:35 | No
1858 | Yes | device not charging | device not charging | 000000000000000000000000001O##: 00000000000000000000000000100000000180000000000000000 16-Oct-2000000003 Adale Mengistu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: nuse Device ue: device in ue Eupected and ueupected Resuts: cutomer claims device 000000000000000000000000001't charge and is beeping uer Impact: none Patient Impact: none SE uknown Diagnostic performed by Engineer : Determined device beeped intermittently uon tuning on Resoluion : Confirmed rear AC cable is not fuly secue in port cauing power to be intermittently provided. Performed operational check and shock test with all parameters passing. System operational, device retuned to clinical ue ----------------------------------------------------------------------------- | 2023-10-08 15:20 | No
1859 | Yes | Live monitor problem | Live monitor problem | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: PPD replacement parts 00000000300000000000000006551 1. Shipped 200000000 Cao Peng-86-1810000000062800000000 SAP Transfer: Case Transferred to SAP Internal Remarks: 18.1000000002000000003 00000000:18:500000000uC (8600000000900000000) $Q1: (Y) $Q2: (N) $Q3: (NA) $Q00000000(NA) $Q5: (NA) $Q6:（Occasionally no display of images) $Q7: (Confirmed reported problem, Occasionally no display of images) $Q8: ( Check IPPC, graphics card may be abnormal ) $Q9: (Replace the 0000000080000000090000000052*1 NVIDIA Qudro K600000000-2GB, tun normal ) $Q100000000 (Test and Inspection data refer to PA. Device retuned to ful fuctionality.) SAP Transfer: Case Transferred to SAP T2 Activities:Repair Action: 00000000300000000000000006551 Graphics Card, NVIDIA Qudro K600000000 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the real-time screen of the hanger was intermittently withou signal, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The real-time screen of the hanger has intermittent no signal. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC graphics card is normal and try to re-insert and uplu the IPPC internal graphics card. 2. Check whether the IPPC is normal. 3. Check the machine uLLand perform faut repairs. Recommended spare part uuand name: 0000000080000000090000000052 NVIDIA Qudro K00000000-2GB &0000000080000000050000000092 IP PC Revised_II no HDD Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Case: 000000002200000000117 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: The real-time screen in the operating room is black Neut Step Action: PPD spare parts: 0000000080000000090000000052*1, aging: 00000000 FSE: Cao Peng-86-1810000000062800000000 Euternal Remarks: 18.1000000002000000003 00000000:18:00000000uC (8600000000900000000) Service content: 1. The image occasionally does not display 2. Replace the image PC graphics card and solve the problem Machine retuns to normal Service resuts: The device powers on normally, perspective and euposue are normal, and the image displays normally. ----------------------------------------------------------------------------- | 2023-10-08 15:35 | No
1860 | Yes | Remote Alert: RMT - PWR: Defective Fan and Power Tray unit | Remote Alert: RMT - PWR: Defeu Fan and Power Tray unit | 000000000000000000000000001O##: 00000000000000000000000000100000000181000000005 Lokeh Vinay The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: RADAR Device ue: yes Eupected and ueupected Resuts: no information uer Impact: no information Patient Impact: no information uuDiagnostic performed by Engineer : the system was stuk at the start-u loop. FSE has to reload the SW duing which he forgot to add the systemconfig.ini file. in the root directory hence the issu. ----------------------------------------------------------------------------- | 2023-10-08 15:40 | No
1861 | Yes | Machine is restarting again and again (Replaced DFM0 ASSY Processor) | Machine is restarting again and again (Replaced DFM0 ASSY Processor) | 000000000000000000000000001O##: 00000000000000000000000000100000000180000000071 Madhan, kuar The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : To replace 0000000056000000009000000001, DFM100000000 Processor PCA ASSY Install Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001, DFM100000000 Processor PCA ASSY Install ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000182675 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY Processor ----------------------------------------------------------------------------- | 2023-10-08 16:30 | No
1862 | Yes | No sound for the alarms - No product malfunction - software issue-system restart required | No sound for the alarms - No produumalfunuon - software issue-system restart required | 000000000000000000000000001O##: 00000000000000000000000000100000000185826 11-Oct-2000000003 Dean Lilya The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: PHMIuHMuuV3- central station. no soud for the alarms Resoluion : The PC I got to 000000000000000000000000001k again throuh runing setu bu had some weird behavior when coming back online. When youtried to click on a sector for a patient window it went completely white. instead of letting youpick a piece of equpment it went white and blank with no choices and only came back after clicking inside the patient window. I decided to make sue there were no lingering corrution so I did a software reload jut to be sue. After reload there were no issus. Everything 000000000000000000000000001ked as intended. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000181337 00000000-Oct-2000000003 Dean Lilya The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Reload software trade Bio spare and reload other PC software Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: PHMIuHMuuV3- central station. no soud for the alarms Diagnostic performed by Engineer : Came in saw that patient window wasn’t coming u coudn’t restart withou server tried runing setu again and it 000000000000000000000000001ked for the most part bu something in the software seems corruted. I conclue needs a reload bu coud not find cutomers pic software need to revisit. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000181315 Pirela, Manul The system meets the specification for the performed service and is retuned to ue Resoluion : Is Caue Known: No Caue: Is Resoluion Known: No Resoluion Detail: Resolved Remotely Prevention: Follow u Requred : Problem Description: PHMIuHMuuV3- central station. There is no soud from the alarms PHMIuHMuuV3- central station. There is no soud from the alarms Form ued: Trouleshooting/Investigation Object: PHMIuHMuuV3 Not Object: Software Version: uu Date and Time of Issu: 1000000008/2000000003 uer Actions: Where is the Deviation: PHMIuHMuuV3- central station. There is no soud from the alarms Trouleshooting Steps Taken: Speakers are connected to the PC. Volue is on 2 and it was 000000000000000000000000001king before. no spare speakers. It was 000000000000000000000000001king before. Disconnected and reconnected all speakers. I had the cutomer restart the PC one more time. Was PRS ued: No PRS Permission: Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: uer Impact: No impact Patient Impact: No impact SE uu Problem Description/ueupected Resuts: PHMIuHMuuV3- central station. no soud for the alarms Diagnostic performed by Engineer : Problem Description: PHMIuHMuuV3- central station. There is no soud from the alarms PHMIuHMuuV3- central station. There is no soud from the alarms Form ued: Trouleshooting/Investigation Object: PHMIuHMuuV3 Not Object: Software Version: uu Date and Time of Issu: 1000000008/2000000003 uer Actions: Where is the Deviation: PHMIuHMuuV3- central station. There is no soud from the alarms Trouleshooting Steps Taken: Speakers are connected to the PC. Volue is on 2 and it was 000000000000000000000000001king before. no spare speakers. It was 000000000000000000000000001king before. Disconnected and reconnected all speakers. I had the cutomer restart the PC one more time. Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: David Beal IP & Issu: no soud for the alarms 75513500000000 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created R000000000000000000000000001, notified rp Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-08 18:05 | No
1863 | Yes | Main display isn't showing_SPLIT | Main display isn't showing_SPLIT | 000000000000000000000000001O##: 00000000000000000000000000100000000181750000000000000000-Oct-2000000003 Scott, Bennett The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: cardiac cath lab Device ue: system down Eupected and ueupected Resuts: system 000000000000000000000000001ld power on bu not fuly boot u. uer Impact: uers can not ue the system Patient Impact: high SE8.1 Resoluion : Replaced the fleuvision pc. Reconfigued the system for the correct Fleuvision configuation on this EP system. Checked all video souces on all monitors. Foud the video is not 000000000000000000000000001king for the t000000000000000000000000001operators desk monitors. System is ready for ue bu 2 ou of 6 operators monitors are not 000000000000000000000000001king correctly. I am opening a new case and cutomer is aware and accepted the system as is. Follow u Requred : The t000000000000000000000000001right operators monitors are not showing any video. Internal Comments : Please charge neut flight ou shipping charge. Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment : ELECTRICAL SAFETY TESTER,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001810000000000000000-Oct-2000000003 Scott, Bennett The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: cardiac cath la Device ue: system down Eupected and ueupected Resuts: system does not fuly boot u, stick on 0000000000000000. uer Impact: uers cannot ue Patient Impact: high SE8.2.17 Diagnostic performed by Engineer : Trouleshot fleuvision issu.Checked logs and foud fleuvision image grabber uLL. Follow u Requred : Please create follow u 000000000000000000000000001--------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Edwin Tanchez IP & Issu: 62991186 main display isn't showing anything and bed isn't moving and control panel is stuk and when attempted to reboot, it doesn't come on fuly, they're getting flashing lights. FSE: Edwin Tanchez, Biomed, has left for the day bu if youhave to contact him please call his cell phone on the case in SMAu. The device is located at the Inova heart & vascuar bldg on groud floor in the EP lab in room 100000000 The POC for that dept is be Aaron Agutin and his cell phone nuber is 300000000-222-3238 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? No Preferred Service Time: 1000000008/2000000003 Entitlement Options: Qute Reason for Entitlement Claim: not entitled Action Taken: created qute uing cu's LIFECYCLE ACCOuT Neut Step Action: Awaiting Qute Verbal PO Nuber: LIFECYCLE ACCOuT Verbal PO Approver: Edwin Tanchez Verbal PO CAP: N/A Internal Remarks : Scott - frame grabber card uLL - Approved fleuvision PC part order. ----------------------------------------------------------------------------- | 2023-10-09 04:45 | No
1864 | Yes | 0V0uLLunable to startup | 0/V0uLLunable to startup | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000755 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: 00000000/V600000000uLLNeut Step Action: Contacted ASP shangyi for buiness follow-u Neut Step Action: Baoshan People’s Hospital Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service ----------------------------------------------------------------------------- | 2023-10-09 07:05 | No
1865 | Yes | 0/u FD0HostPC cannot power on | 0R/u FD0HostPC cannot power on | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problem: Hos tPC device cannot be tuned on. Detailed description of cutomer problem or faut phenomenon: Hos tPC device cannot be tuned on. Is it a faut of the device itself? Yes Trouleshooting Action: Key faut codes or phenomena: The Hos tPC device cannot be powered on. Remote diagnosis and test resuts: The Hos tPC device cannot be powered on. Check that the power suply is normal. It is supected that the hostpc is fauty. PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check the HostPC power suply, 2. Reformat the HostPC OS disk and reinstall the system, 3. Replace the host PC. Recommended spare parts uuand name: " 0000000080000000050000000000000000 Host PC Monoplane Revised_II no HDD" Cutomer demand Remarks: Onsite Euternal Remarks: 1000000001000000002000000003 00000000:25:58 uC (86000000008238) "SuT. Ready for Clinical ue local time" : " 2000000003-10000000000000000 12:00000000" Service content (requred): Cutomers reported that the system coud not be tuned on. There was a problem with the battery on the main chassis and motherboard. After replacement, it retuned to normal and everything on the device was normal. Service resut (requred): The system statu is normal and can be ued Other important matters: None Internal Remarks: 1000000001000000002000000003 12:0000000000000000uC (86000000008238) $Q1: (N/A) $Q2: (N/A) $Q3: (N/A) $Q00000000(N/A) $Q5: (N/A) $Q6: (system can't power on) $Q7: (Confirmed reported problem, system can't power on) $Q8: (check the log, buton cell of host pc problem, need replace buton cell of host pc) $Q9: (replace buton cell of host pc) $Q100000000: (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (N/A)S/N (N/A) Neut Calibration date (N/A) ‘ SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Neut Step Action: PPD Host PC Monoplane Revised_II no HDD 0000000080000000050000000000000000 1 200000000arrival at the venu has been arranged Feng Lei -86-19815665565 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Repair Action: PPD/ 0000000080000000050000000000000000 Hard disk spare parts Problem Reported by Cutomer: uable to boot SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-09 07:55 | No
1866 | Yes | ulead is malfunctioning (Replaced u0ead Grabber Limb). | ulead is malfunuoning (Replacedu0Lead Grabber Limb). | 000000000000000000000000001O##: 000000000000000000000000001000000001800000000300000000Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Defective ugrabber Lot No.2000000000000000000000000 to be replaced with same part no. 989800000000160000000000000000 u5-Lead Grabber Limb, IEC DFM100000000 Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : Defective ugrabber Lot No.2000000000000000000000000 to be replaced with same part no. 989800000000160000000000000000 u5-Lead Grabber Limb, IEC DFM100000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000182296 Parts Consued : 1.000000000000000000000000989800000000160000000000000000,u5-Lead Grabber Limb, IEC ----------------------------------------------------------------------------- | 2023-10-09 08:10 | No
1867 | Yes | 0/V0uLLGas leak alarm | 0K/V0uLLGas leak alarm | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 13.11.2000000003 000000000000000053 uC (860000000000000000100000000 *Department information [Department: Respiratory Department Contact Person: Teacher uuContact Information: 1552137100000000]* $ASP engineer:[ HONGLEJIA Liuengjie] $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: Confirmed reported problem: Leakage. $Q8: Check the equpment foud the GDS modue is fauty. $Q9: Replace the GDS modue. $Q100000000Test and inspection data refer to: PA checklist. Device retuned To ful fuctionality $Q11: Calibrated Tools: #ESA 6200000000 S/N #30000000028600000000 Neut Calibration date #00000000/00000000/20000000000000000Problem Reported by Cutomer: Air leakage Neut Step Action: uinhu Traditional Chinese Medicine Hospital, Gungdong Province Internal Remarks: 13.11.2000000003 000000000000000056 uC (860000000000000000100000000 $ASP engineer:[ HONGLEJIA Liuengjie] $Q1: (N) $Q2: (N) $Q3: (N/A) $Q00000000(N/A) $Q7: Confirmed reported problem: Leakage. $Q8: Check the equpment foud the GDS modue is fauty. $Q9: Replace the GDS modue. $Q100000000Test and inspection data refer to: PA checklist. Device retuned To ful fuctionality $Q11: Calibrated Tools: #ESA 6200000000 S/N #30000000028600000000 Neut Calibration date #00000000/00000000/20000000000000000SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Cutomer’s euplanation of the problem: Air leakage alarm The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: Replacing the trachea is ineffective Repair Action: T2 Recommended on-site repair steps: It is recommended to go to the site to check the GDS, sensor, and DA board. Recommended spare part uufor T2 with description: 989800000000617100000000 RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 Notes on cutomer requrements: na Test/Inspection: N/A Test Equpment: N/A Safety Qustions: Case: 000000002200000000927 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A FSE Notes: This PO is a Reference Nuber only ----------------------------------------------------------------------------- | 2023-10-09 09:35 | No
1868 | Yes | 0 CAN HY000Toothbru was on fire | 00CAN HY00/0 Toothbru was on fire | 2000000003-11-27 18:0000000033:***claims***##d cu retun label reminder. Awaiting response.;2000000003-11-15 19:00000000:55:***claims*** ##d cu uS label to retun defective uit. Tracking # is 1ZY73300000000190000000093000000002000000003-10000000010000000019:22:00000000***claims*** ##d cu with safety qustions. Waiting response. | 2023-10-09 10:30 | No
1869 | Yes | Prompt that the device is disabled | Prompt that the device is disabled | BuQ&R Feedback: No SAP Cutomer Symptom Code: OTHR Cutomer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failue Code: uknown Caue Repair Activity Code: Instruted Cutomer SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Display is disabled and treatment cannot be provided Safety Qustions: Case: 000000002200000000100000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju Yu Last udated by: Guju Yu 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: ---------PD---PROBLEM DESCRIPTION--------------- DESCRIPTION OF PROBLEM: Prompt that the device is disabled WHEN WAS THE FIRST OCCuRENCE? WITHIN. . . Same Day HOW OFTEN IS ISSu OCCuRING? First Occurence WHERE IN THE END-uER 000000000000000000000000001KFLOW IS THE ISSu OCCuRING? Boot u POTENTIAL RISK TO THE PATIENT? None WHAT RECENT EVENTS HAVE OCCuRED? Chiller Service/3rd Party Trouleshooting Action: ---------TA-----TROuLESHOOTING ACTION--------------- NOTE uLL# AND / OR DESCRIPTION OF uLL Prompt that the device is disabled WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? N\A WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Gude cutomers to ru operational checks Repair Action: ---------RA-----REPAIR ACTION--------------- RECOMMENDED REPAIR ACTION: Gude cutomers to ru operational checks Service Content: Prompts that the device is disabled and instruts the cutomer to ru an operational check. Service resuts: After remote service, the equpment resues normal operation Test/Inspection: NA Test Equpment: NA ----------------------------------------------------------------------------- | 2023-10-09 12:42 | No
1870 | Yes | CONTACTEL 000 000 | CONTAuL 0/00000:0 | 000000000000000000000000001O##: 000000000000000000000000001000000001827900000000October 16, 2000000003 ue SERRA The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Problem at start-u of the system, the application restarts by itself. uer Impact: No Patient Impact: no SE V2.2.3 Diagnostic performed by Engineer: Deletion of the logs and the database. Tests. Ok. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000182186 Lionel GIROu The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: Cleaned database and log. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Nuse Device ue: no Eupected and ueupected Resuts: Problem at start-u of the system, the application restarts by itself. uer Impact: No Patient Impact: no SE V2.2.3 Diagnostic performed by Engineer: Problem at start-u of the system, the application restarts by itself. The system restarted by itself for 300000000minues. Then the system stabilized and was ued. The cutomer does not tun off the system anymore as a safety measue. Requres on-site intervention by ue Serra to clean the patient database and the log files. ----------------------------------------------------------------------------- Internal Remarks: NO E-MAIL ----------------------------------------------------------------------------- | 2023-10-09 13:17 | No
1871 | Yes | 0/u FD0geometry movement fault(SPLIT) | 0R/u FD0geometry movement fault(SPLIT) | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: PPD 00000000800000000000000007813 1. Shipped in 72H, Song Mingliang -86-18987296523 SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 15.1000000002000000003 00000000:38:00000000uC (86000000006200000000 Service Content: Reconfigue IPC Service resuts: Equpment is normal SAP Transfer: Case Transferred to SAP Internal Remarks: 15.1000000002000000003 00000000:38:00000000uC (86000000006200000000 "RDI Accuacy: yes " $Q7: (Confirmation & geometry movement problems) $Q8: (checked log, IPC commuication time ou, IPC failed on POST) $Q9: (reinstalled IPC software) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problems: Mechanical movement uLL Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Mechanical movement uLL Remote diagnostics and test resuts: It 000000000000000000000000001ks normally after restarting several times. Consider IPC abnormality. PRS TR statu: Doesn't make sense Repair Action: Recommended on-site repair steps: Check whether there are abnormal uLL in the post statu, whether ipc starts normally, whether there are any abnormalities in synqnet commuication and PCu and futher trouleshoot based on on-site uLLreports. Recommended spare part uuand name: PPD, 00000000800000000000000007813 IPC, 200000000to hosp, repair requster will collect Notes on cutomer requrements: onsite Problem Reported by Cutomer: The machine head cannot be tuned on. The uper left corner of the screen is not displayed and the screen is rotated 900000000degrees. Safety Qustions: Case: 000000002200000000200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Weiqiang Zhu Last udated by: Weiqiang Zhu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A ----------------------------------------------------------------------------- | 2023-10-09 13:27 | No
1872 | Yes | 0 DEU SCH0/0 the fever thermometer does not work anymore | 00DEU SCH0/0 the fever thermometer does not work anymore |  | 2023-10-09 13:29 | No
1873 | Yes | Request for technical intervention | Request for technical intervention | 000000000000000000000000001O##: 00000000000000000000000000100000000183327 12-Oct-2000000003 CARDILE, GIANLuA The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Problem description by engineer: The system does not start becaue of a software uLL Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: no Patient Impact: no SE 8.1.3000000005 Resoluion: Performed faut search, foud problem duing the start-u of the system. Replaced spare part previouly ordered and performed necessary configuations. Parts Consued: 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD Tools & Equpment: MuTIMETER, 200000000000000003-27 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000182678 Samule, Villa The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: The system does not start becaue of a software uLL Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: no Patient Impact: no SE 8.1.3000000005 Diagnostic performed by Engineer: The system does not start by software uLLInternal Comments: uLL# and/or description of uLL Start-u of IP-PC not completed Repair Action: An on-site visit with parts is needy. Follow u Requred: uLL# and/or description of uLL Start-u of IP-PC not completed Repair Action: An on-site visit with parts is needy. What are the neut Trouleshooting steps: An on-site visit is necessary with parts in order to continu the diagnosis 1FSE 00000000 2u 0000000080000000050000000000000000- Hard disk spare part 1u 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD Technician Name: Samule, Villa ----------------------------------------------------------------------------- | 2023-10-09 14:47 | No
1874 | Yes | It does not turn on. | It does not turn on. | 000000000000000000000000001O##: 000000000000000000000000001000000001830000000019-Oct-2000000003 MIGLIACCIO, NICOLA The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Tools & Equpment: MuTIMETER, 2000000005-00000000-00000000 00000000:00000000:00000000 Diagnostic performed by Engineer: Fleuvision PC is defective Resoluion: Replacement of Fleuvision PC, reinstallation of operating system and application, reconfiguation of fleuvision and oupu video, fuctional tests were completed with sucess. Problem Description by Engineer: Problem description by engineer: It does not tun on. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000182800000000TESTOLIN, DARIO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Problem description by engineer: It does not tun on. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: Not provided uer Impact: No futher information is available as for the cutomer Patient Impact: No futher information is available as for the cutomer SE 000Diagnostic performed by Engineer: Performed 00contact with u 000000003900000000275 - Salvatore PIRA - they had spoken this morning with some of ou colleagus - at the moment the system 000000000000000000000000001ks reguarly It is assued that cmos Battery of fleuVision PC is partially discharged - the system is kept in monitoring Internal Comments: uLL# and/or description of uLL . Repair Action: Made 00contact with u 000000003900000000275 - Salvatore PIRA - they had spoken this morning with some of ou colleagus - at the moment the system 000000000000000000000000001ks reguarly It is assued that cmos Battery of fleuVision PC is partially discharged - the system is kept in monitoring Follow u Requred: uLL# and/or description of uLL . Repair Action: Made 00contact with u 000000003900000000275 - Salvatore PIRA - they had spoken this morning with some of ou colleagus - at the moment the system 000000000000000000000000001ks reguarly It is assued that cmos Battery of fleuVision PC is partially discharged - the system is kept in monitoring What are the neut Trouleshooting steps: An on-site visit is needed to replace the part to continu the diagnosis - ue the PC diagnostic tool to verify the battery level of the new PC - otherwise send back as DOA. Parts: 0000000080000000058600000000 - FleuVision-2_Revised PC 00CC: 1 FSE 2 hous of room downtime + see plan hous + travel Technician Name: TESTOLIN, DARIO ----------------------------------------------------------------------------- | 2023-10-09 14:57 | No
1875 | Yes | Without Sound-No product malfunction - configuration settings issue -u uLL- work around provided. | Without Sound-No produumalfunuon - configuration settings issue -u uLL- work around provided. | 000000000000000000000000001O##: 00000000000000000000000000100000000182200000000Hendrik Gritzka The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: Order uer Impact: uknown Patient Impact: uknown SE uknown Resoluion: After the replacement of the monitor, the control center has no soud anymore since the new monitor was set u as an oupu device. In the settings of the soud, set u the louspeaker again as the standard device, then restarted the control center so that the settings are also kept duing a restart. ----------------------------------------------------------------------------- Cutomer comment: u Wiesner called and the processing nuber Get. Please call the phone nuber: +000000001735853663 Internal Remarks: Cases were sent together by fau from u Jürgen Restle (no contact data): 000000002200000000589 + 000000002200000000598 ----------------------------------------------------------------------------- | 2023-10-09 15:07 | No
1876 | Yes | Triple Chirp | Triple Chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000183775 PATRICK POELS The system meets the specifications for the performed service and is retuned to ue Follow u Requred: Order part: 000000005600000000000000001 - HS1 (C B) Defib, u English 3ed, Euch Problem Description by Engineer: Cutomer Fuction/Role: Dealer Procardio Device ue: ouside of ue Eupected and ueupected Resuts: Triple Chirp uer Impact: cannot ue AED Patient Impact: N/A SE uknown Resoluion: AED needs to be replaced ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001838300000000Parts Consued: 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Defib, u English 3ed, Euch ----------------------------------------------------------------------------- | 2023-10-09 16:02 | No
1877 | Yes | 0/u FD0Image disk problem | 0R/u FD0Image disk problem | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 00000000.1000000002000000003 15:0000000000000000 uC (8600000000000000009) Service content: The hard disk space is ful and cannot be euposed. Replace t000000000000000000000000001image hard disks and format the hard disk images. Service resut: The faut is repaired and the machine can be ued normally. Problem Reported by Cutomer: 00000000R/u uuuFD200000000 SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the image storage space is low and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Free disk space is too low. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the hard drive is normal and try to bypass the hard drive. 2. If it has no effect, ue PC Diagnostic Tool to perform diagnostic tests and replace the corresponding spare parts. 3. Check the machine uLLreport and perform faut repairs. Recommended spare part uuand name: 0000000080000000050000000000000000Hard disk spare part Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: PPD Hard disk spare part 00000000800000000500000000000000002 Arranged to go to the hospital on 36H, FSE Lu Ji and -86-150000000000000000117700000000SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000565 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: 100000000000000002000000003 12:50000000016 uC (8600000000000000009) $Q1:(Y)$Q2:(N)$Q3:()$Q00000000)$Q5:() $Q6:( Can't euposue.) $Q7: (Confirmation & Description of problem: Confirmed, Warning: Image disk ful, disk write uLL $Q8: (Trouleshooting & Resuts: Perform Image disk data base verify, resut is failed. ) $Q9: (Repair steps: Replace the t000000000000000000000000001image disks, perform image recreate. ) $Q100000000 (Test and Inspection data&Device statu: Device retrued to ful fuctionality. Test and Inspection data refer to PA. ) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Neut Calibration date (N/A) For Mainland China CV/uR, CT engineers only. Please fill in: Curent Tue: 1983600000000 Installed data: 10000000000000000/2000000002 ----------------------------------------------------------------------------- | 2023-10-09 16:48 | No
1878 | Yes | PSA_QA000000uissue - GFE SENT 0 (0DEC 0) - F/U (0DEC 0) | PSA_QA00000000 uissue - GFE SENT 0- (0 DEC 00 - F/U (0 DEC 00 |  | 2023-10-09 17:31 | Unknown
1879 | Yes | The device does not start. | The device does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000183800000000 00000000-Oct-2000000003 Michal Sadowski The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: electrophysiology. Eupected and ueupected Resuts: the device does not start. uer Impact: no. Patient Impact: yes. SE 8.2.3000000005 Resoluion: The Host compuer did not start. Normal fuctionality of the main Host compuer was restored. Testing, mutiple reset of the device. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000182900000000 Michal Sadowski The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Technician. Device ue: electrophysiology. Eupected and ueupected Resuts: the device does not start. uer Impact: no. Patient Impact: yes. SE 8.2.3000000005 Diagnostic performed by Engineer: Consutation with the uer over the phone. On-site visit is necessary. ----------------------------------------------------------------------------- | 2023-10-09 17:42 | No
1880 | Yes | Battery failure | Battery failure |  | 2023-10-09 17:50 | No
1881 | Yes | Damaged battery | Damaged battery |  | 2023-10-09 18:04 | No
1882 | Yes | M0A - u00- Failed self test due to udevice alarm | u- Failed self test due to udevice alarm |  | 2023-10-09 18:06 | No
1883 | Yes | Speaker Problem Loudspeaker problem, it does not work. =CASE IN PROGRESS =- GFE0ENT DEC0 F/U DEC0 | Speaker Problem Loudspeaker problem, it does not work. =CASE IN PROGRESS =- GFE0SENT DEC0. F/U DEC0 | 000000000000000000000000001O##: 00000000000000000000000000100000000182775 Raffenne, Virginie The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: SPEAKER DEFAuT uer Impact: No. Patient Impact: No. SE uknown ********************************* Follow u Requred: ISSu ONLY A QuTE. Despite the contract DO NOT SEND THE PART *************************** Resoluion: ONLY MAKE A QuTE. Despite the contract DO NOT SEND THE PART 0000000056000000003700000000 Card - Main Board, Vsi Diagnostic performed by Engineer: uLLmessages: "Auio failue" at the top in red. "Louspeaker faut" at the bottom, technical alarm uLLlog: uLLCODEEu 881 The auio test passes with the 3 beeps. The louspeaker was already replaced. No symptoms with the soud of the monitor or the alarms. The monitor was restarted several times, the config and the SN were also reloaded. Recommendation for replacement of the mother board: 0000000056000000003700000000 Card - Main Board, Vsi ----------------------------------------------------------------------------- | 2023-10-09 18:12 | No
1884 | Yes | QACASE for PCI imported product#000000The C-arm is faulty during u | QACASE for P0imported produu00000000 The C-arm is faulty during u |  | 2023-10-09 18:18 | No
1885 | Yes | The machine has no u-rays. | The machine has no u-rays. | 000000000000000000000000001O##: 00000000000000000000000000100000000191286 11-Oct-2000000003 LuACCINI, ALESSIO The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Tools & Equpment: MuTIMETER, 200000000000000007-00000000 00000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000080000000000000000900000000,PDuMains Interface Modue Resoluion: Performed intervention for problem with generator power suply. Replaced part as per factory instrutions, reguar tests and checks, system can be ued for clinical ue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018000000006 100000000Oct-2000000003 LuACCINI, ALESSIO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: The machine has no u-rays. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: The machine has no u-rays. uer Impact: yes Patient Impact: yes SE 2.2.6.00000000Diagnostic performed by Engineer: Performed troule shooting, system is not 000000000000000000000000001king. Foud problems with M Cabinet in the block of power distribuion uit. Foud absent L2 phase du to contactor problem. Ordered spare part, corrective action will follow. The system cannot be ued. Tools & Equpment: MuTIMETER, 200000000000000007-00000000 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018300000000 TESTOLIN, DARIO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: The machine has no u-rays. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: The machine has no u-rays. uer Impact: yes Patient Impact: yes SE 2.2.6.00000000Diagnostic performed by Engineer: Performed 00contact with u Lorenzo Malatesti - enable remote connection - check RMT - IGT: Low load flu active - RMT - IPS: Phase faut - Frontal Internal Comments: uLL# and/or description of uLL . Repair Action: Made 00contact with u Lorenzo Malatesti - enable remote connection - check RMT - IGT: Low load flu active - RMT - IPS: Phase faut - Frontal Follow u Requred: uLL# and/or description of uLL . Repair Action: Made 00contact with u Lorenzo Malatesti - enable remote connection - check RMT - IGT: Low load flu active - RMT - IPS: Phase faut - Frontal What are the neut Trouleshooting steps: An on-site visit is needed to replace the part to continu the diagnosis. Parts: 00000000210000000020000000000000000 IPS 00CC: 1 FSE 2 hous of room downtime + see plan hous + travel Technician Name: TESTOLIN, DARIO ----------------------------------------------------------------------------- | 2023-10-09 18:27 | No
1886 | Yes | QACASE for PCI imported product#00000 uLLshowed an overheating alarm of Peng fan | QACASE for P0imported produu00000000 uLLshowed an overheating alarm of Peng fan |  | 2023-10-09 18:28 | No
1887 | Yes | Without Sound - No product malfunction Configuration issue u error | Without Sound - No produumalfunuon Configuration issue u error | 000000000000000000000000001O##: 00000000000000000000000000100000000182257 Hendrik Gritzka The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: biomed Device ue: uknown Eupected and ueupected Resuts: Control center withou soud uer Impact: uknown Patient Impact: uknown SE uknown Resoluion: Processing will take place in case 000000002200000000589, since this case was reported twice. ----------------------------------------------------------------------------- Internal Remarks: Cases were sent together by fau from u Jürgen Restle (no contact data): 000000002200000000589 + 000000002200000000598 ----------------------------------------------------------------------------- | 2023-10-09 18:43 | No
1888 | Yes | uLLwould automatically turn on and could not be turned off | uLLwould automatically turn on and could not be turned off |  | 2023-10-09 18:43 | No
1889 | Yes | Smoke is coming out of the technical room | Smoke is coming out of the technical room | 000000000000000000000000001O##: 000000000000000000000000001000000001800000000 Neagu Cristian The system meets the specification for the performed service and is retuned to ue Resoluion: The angiography system is in 000000000000000000000000001king condition with samples. Diagnosis performed by Engineer: Following some maintenance 000000000000000000000000001k on the electrical panel inside the hospital, components related to the angiography system were affected. It is necessary to go to the site to make a detailed diagnosis. Problem Description by Engineer: Cutomer Fuction/Role: dr Device ue: yes Eupected and ueupected Resuts: no power for IW and CV system uer Impact: no Patient Impact: no 00----------------------------------------------------------------------------- | 2023-10-09 18:47 | No
1890 | Yes | The system will not start up. It has been rebooted 0imes | The system will not start up. It has been rebooted 0times | 000000000000000000000000001O##: 0000000000000000000000000010000000018500000000 11-Oct-2000000003 Nigel Huphries The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Parts Consued : 1.0000000000000000000000000000000080000000058600000000,FleuVision-2_Revised PC no HDD Resoluion : Replaced defective Fleu PC du to failue of grabber cards. The euisting hard disk was reued, hence no loss of cutomer defined setus on the fleuvision screen. Have also replaced the BIOS backu battery in the Host PC to avoid issus with loss of correct date/time. Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Sitting at 00000000 :00000000 uer Impact: NA Patient Impact: NA SE8.1 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018000000002 Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to replace Fleu PC and configue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Sitting at 00000000 :00000000 uer Impact: NA Patient Impact: NA SE8.1 Diagnostic performed by Engineer : Logging sent by FSE shows grabber card uLL ----------------------------------------------------------------------------- Cutomer comment : Faith rang to ask for an udate. I said that the qute has been requsted bu not yet sent as they are not covered uder contract. She said if this coud be sent ugently, so I have ##d qutes to ask this to be sent asap. Internal Remarks : Hi, The service contract has eupired for this case. Iain Rider sorting renewal bu can we provide service in the meantime uing the commercial discouting procedue – Override Cost Center GB90000000015.uR. Many thanks, Aiden Aiden Jones Service Delivery Cutomer comment : confirmed planning for 11/100000000with Karen K ----------------------------------------------------------------------------- | 2023-10-09 18:52 | No
1891 | Yes | M0A - u00 Device fails to turn on | u Device fails to turn on |  | 2023-10-09 19:07 | No
1892 | Yes | Battery failure | Battery failure |  | 2023-10-09 19:24 | No
1893 | Yes | PSE_ QACASE for PCI imported product#000000lack screen | PSE_ QACASE for P0imported produu00000000black screen |  | 2023-10-09 19:26 | No
1894 | Yes | M0A - u00- QACASE for PCI imported product#000000 the defibrillator could not be self-tested | u- QACASE for P0imported produu00000000 the defibrillator could not be self-tested |  | 2023-10-09 19:27 | No
1895 | Yes | EC/00ackup Alarm Failed - Pending FSE site visit 00 | EC/00 Backup Alarm Failed - Pending FSE site visit 0/0 | 000000000000000000000000001O##: 0000000000000000000000000010000000018000000009 Carey, Scott The system does not meet the specification for the performed service and is not retuned to ue Resoluion : FSE to resolve Follow u Requred : Diagnostic code 11200000000 Battery failed. When the PM PCBA detects a battery uLL the uLLcontinus to ventilate and the battery charger tuns off. Advised Biomed of the 00000000 generation power management board that is being installed into all V600000000s as part of a recall. Biomed is reporting swapping the battery with a new u Respironics battery. The uit continus to display uLL11200000000 1. Verify that the PM PCBA is equpped with PIC firmware version 1.300000000 2. Replace the internal battery. 3. Replace the PM PCBA. Down V600000000du to possible power management board - FCO86600000000000000006 Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000is displaying diagnostic code 1100000000and 11200000000iagnostic performed by Engineer : Problem Description: V600000000is displaying diagnostic code 1100000000and 11200000000 Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000000000000/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Biomed has confirmed the uLLcodes in the V600000000significant event log. Cutomer advised of possible bad CPuboard cauing diagnostic code 1100000000- 1100000000(POST) Backu alarm failed - When the Backu Auible Alarm Test fails, alarms normally annuciated by the backu auio device ue the primary speakers. 1. Replace the MC PCBA. 2. Replace the CPuPCBA. 3. Replace the PM PCBA. Advised Biomed of 00000000 generation power management board that is being installed into all V600000000s as part of a recall. Diagnostic code 11200000000 Battery failed. When the PM PCBA detects a battery uLL the uLLcontinus to ventilate and the battery charger tuns off. 1. Verify that the PM PCBA is equpped with PIC firmware version 1.300000000 2. Replace the internal battery. 3. Replace the PM PCBA. Biomed is reporting swapping the battery with a new u Respironics battery. The uit continus to display uLL11200000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Donald pales 70000000090000000000000000 IP & Issu: 8312000000002/ giving red alarm. Says check vent backu Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-09 19:28 | No
1896 | Yes | Connection problems with the wireless u-ray pedal | Conneuon problems with the wireless u-ray pedal | 000000000000000000000000001O##: 00000000000000000000000000100000000185300000000 00000000-Nov-2000000003 Shahbazi, Artikas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer : RR Christoph Eichel Contact Phone 00000000000000002181000000003772 The faut report is confirmed. How often is issu occuring? : Intermittent Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Parts Consued: 1.00000000000000000000000000000000800000000238232,Wireless Footswitch 3P 1.00000000000000000000000000000000800000000257862,WFS Base Station Resoluion: The reported faut is confirmed. Wireless foot switch does not 000000000000000000000000001k sporadically. Foot switch and receiving station in the patient table replaced. System fuly fuctional. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000185139 Hellenbrandt, Helmu The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Diagnostic performed by Engineer: ufoot switch failue The cable foot switch is curently being ued Problem Description by Engineer: Problem description by engineer : RR Christoph Eichel Contact Phone 00000000000000002181000000003772 The faut report is confirmed. How often is issu occuring? : Intermittent Malfuction area : Others Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: Not provided Patient Impact: Not provided SE 8.2.3000000005 Follow u Requred: uLL# and/or description of uLL: The ufoot switch fails sporadically. The battery is charged - the radio signal indicator then lights u red. The installed base station is no longer available Repair Action: Check ufoot switch / replace foot switch and base station if necessary Technician name: Hellenbrandt, Helmu Internal Comments: uLL# and/or description of uLL: The ufoot switch fails sporadically. The battery is charged - the radio signal indicator then lights u red. The installed base station is no longer available Repair Action: Check ufoot switch / replace foot switch and base station if necessary ----------------------------------------------------------------------------- Cutomer comment: ## u Dietrich was schedued to attend the PM on 2.11.2000000003 uon requst ----------------------------------------------------------------------------- | 2023-10-09 19:28 | No
1897 | Yes | QACASE for PCI imported product#000000the C-arm and the table could not move | QACASE for P0imported produu00000000 the C-arm and the table could not move |  | 2023-10-09 19:38 | No
1898 | Yes | QACASE for PCI imported product#000000 the device gave an alarm of high inspiratory volume | QACASE for P0imported produu00000000 the device gave an alarm of high inspiratory volume |  | 2023-10-09 19:52 | No
1899 | Yes | Geo module not working | Geo module not working | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000050000000019-Oct-2000000003 uer, Shafi The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Geo modue not 000000000000000000000000001king When was the first occurence? Within... : 00000000-Oct-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Geo modue not 000000000000000000000000001king uer Impact: No uer impact Patient Impact: No patient impact SE Not provided Tools & Equpment : MuTIMETER,200000000000000002-15 00000000:00000000:00000000 Resoluion : Issu resolved by replacing the geo modue and reloading the software of the system Handed over the system to the uer in proper 000000000000000000000000001king condition ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000239930000000018-Oct-2000000003 uer, Shafi The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Geo modue not 000000000000000000000000001king When was the first occurence? Within... : 00000000-Oct-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Geo modue not 000000000000000000000000001king uer Impact: No uer impact Patient Impact: No patient impact SE Not provided Diagnostic performed by Engineer : Replaced the geo modue however futher issus were foud in SIB and table wiring. Rectified the same by trouleshooting one by one in patient table and technical room cabinets. Geometry movements have been restored however futher 000000000000000000000000001k is in progress for touh screen modue Follow u Requred : TSM Repairs Parts Consued : 1.00000000000000000000000000000000800000000151756,GEO TILT MODuE BIPL. KIT WP Tools & Equpment : MuTIMETER,200000000000000002-15 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000185789 16-Oct-2000000003 uer, Shafi The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Table repair in progress Tools & Equpment : MuTIMETER,200000000000000002-15 00000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : Geo modue not 000000000000000000000000001king When was the first occurence? Within... : 00000000-Oct-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Geo modue not 000000000000000000000000001king uer Impact: No uer impact Patient Impact: No patient impact SE Not provided Diagnostic performed by Engineer : Geo modue not getting detected, 000000000000000000000000001k in progress ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018500000000 Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Geo modue not 000000000000000000000000001king When was the first occurence? Within... : 00000000-Oct-2000000003 20000000000000000:00000000 How often is issu occuring? : First occurence Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Geo modue not 000000000000000000000000001king uer Impact: No uer impact Patient Impact: No patient impact SE Not provided Diagnostic performed by Engineer : Some of the joysticks in geo modue not 000000000000000000000000001king bu keys are 000000000000000000000000001king. Internal Comments : uLL# and/or description of uLL: Geo modue not 000000000000000000000000001king Repair Action : Some of the joysticks in geo modue not 000000000000000000000000001king bu keys are 000000000000000000000000001king. Resoluion : Some of the joysticks in geo modue not 000000000000000000000000001king bu keys are 000000000000000000000000001king. Follow u Requred : uLL# and/or description of uLL: Geo modue not 000000000000000000000000001king Repair Action : Some of the joysticks in geo modue not 000000000000000000000000001king bu keys are 000000000000000000000000001king. Technician Name : Vijayakuar, K.S ----------------------------------------------------------------------------- | 2023-10-09 20:18 | No
1900 | Yes | Low Leak, CO0ebreathing Risk alarm | Low Leak, CO0Rebreathing Risk alarm |  | 2023-10-09 20:20 | No
1901 | Yes | Device could not be turned off - Final 00 | Device could not be turned off - Final 0/0 |  | 2023-10-09 20:21 | No
1902 | Yes | No power to unit | No power to unit | 000000000000000000000000001O##: 0000000000000000000000000010000000018500000000 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Not in Clinical ue Eupected and ueupected Resuts: V600000000power on. uer Impact: V600000000has been removed from service Patient Impact: No impact SE uu Problem Description/ueupected Resuts: V600000000will not power on. Diagnostic performed by Engineer : Problem Description: V600000000will not power on. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000000000000/2000000003 Trouleshooting Steps Taken: Advised cutomer the latest version of the service manul is version T. Advised cutomer of possible power suply or power management board problem. Biomed is reporting they are not curently with the uit. Advised Biomed to reference page 112 of the service manul to confirm the power suply is 000000000000000000000000001king. Biomed advised to confirm 200000000olts DC across the brown and orange wires. Cutomer advised to order power suply if the 200000000olts DC is not present. Biomed to call back if it is determined the power management board is defective. Part # 00000000561512591 Rp - Power Suply, V600000000List price is $8900000000 Page 1600000000for referencing the repair. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: hilario IP & Issu: 7130000000000000000no power Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Entitled Reason for Entitlement Claim: na Action Taken: Created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-10-09 20:32 | No
1903 | Yes | System is not booting up. | System is not booting up. | 000000000000000000000000001O##: 00000000000000000000000000100000000212253 18-Oct-2000000003 Holiday Scott The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part Resoluion : Replaced PC. Reinstalled software. Configued system. Performed calibrations. Tested. Handed over to cutomer. Problem Description by Engineer : Problem description by engineer : System HOST PC is defective and need to be replaced. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: System is not booting u. uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000196383 12-Oct-2000000003 Holiday Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System HOST PC is defective and need to be replaced. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: System is not booting u. uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Parts Consued : 1.0000000000000000000000000000000080000000000000000300000000u8.u.3000000005/uIQ+Centr1.000000003000000005 sw pck 1.0000000000000000000000000000000080000000050000000000000000Hard disk spare part 1.00000000000000000000000000000000800000000200000000581,Nehalem Host PC Biplane Rev_III NO_HDD Diagnostic performed by Engineer : Attempted sw install Ordered Haswell PCs and hard drives. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000188855 100000000Oct-2000000003 Holiday Scott The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System HOST PC is defective and need to be replaced. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: System is not booting u. uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Follow u Requred : Same as above Follow u Requred : Install PC. Install software. Resoluion : Attempted Sw install. Installed PC. Ordered new PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000185615 00000000-Oct-2000000003 Holiday Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : System HOST PC is defective and need to be replaced. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: System is not booting u. uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Follow u Requred : Install new pc Diagnostic performed by Engineer : Ordered host pc for neut day ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000185327 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : uLL# and/or description of uLL: It appears that the system HOST PC is not booting u. Repair Action : Replace HOST PC. On-site BIOMED checked the backplane of the HOST PC which is red and shoud be replaced for HDD connection issus. Technician Name : Johansson Anders Resoluion : Created on-site 000000000000000000000000001oblem Description by Engineer : Problem description by engineer : System HOST PC is defective and need to be replaced. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: System is not booting u. uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Cutomer states that the system is not booting u even after several attempts. Internal Comments : uLL# and/or description of uLL: It appears that the system HOST PC is not booting u. Repair Action : Replace HOST PC. On-site BIOMED checked the backplane of the HOST PC which is red and shoud be replaced for HDD connection issus. ----------------------------------------------------------------------------- Internal Remarks : 000000002200000000200000000 FSE Holiday is inquring if the 8.u.3000000005 software stick in the room over will 000000000000000000000000001k since the software he ordered is on backorder. The software is all the same...shoud 000000000000000000000000001k fine. 0000000080000000000000000351. Internal Remarks : 000000002200000000200000000 FSE Holiday is reporting that he ordered a Biplane Nehalem Revised III PC 00000000800000000200000000581 and received 00000000800000000200000000571 with a disk already in it. The odd thing is that 00000000800000000200000000571 is actully the commercial part nuber that inclues a hard disk...bu it is the same PC as 00000000800000000200000000581...so I believe the part is correct. System appears to be at 8.1.3000000005 according to PRS. FSE does not have an 8.1.3000000005 uB stick. Recommend ordering 0000000080000000000000000300000000- u8.u.3000000005/uIQ+Centr1.000000003000000005 sw pck. FSE had qustions abou ethernet ports for 10000000011,12,13 as these have changed from his last PC. uing 8.u parts trouleshooting docuent dep155672 RevF as a gude if looking left to right by the 8 diagnostic LEDs, u100000000top/u11 bottom...moving to the right, first ethernet net000000000000000000000000001k card is u12 and the futhest right net000000000000000000000000001k card is u13 by the power suply fan. FSE will order software and retun to load when it arrives. Internal Remarks : FSE needed Bios PW Internal Remarks : approved TPR for Host PC no data signal Internal Remarks : Planner Initials: GM Action Taken FSE Scott Holiday accepted call for 100000000130000000000M-00000000 Oucome: dispatched Neut Step Action: Internal Remarks : 000000002200000000200000000 FSE Holiday has loaded 8.1.3000000005; however, the system comes u and when entering FSF, it says " system not installed correctly " and FSE is uable to do anything in FSF. The PC arrived with a hard drive already in it...when it shoudn't have...and when attempting a license swap, the MAC 00000000C00000000E00000000F00000000 was still registered to site 0000000070000000000000000edical city of Fort 000000000000000000000000001th and on case 0000000021893818 cutomer was asking for parts ID for that specific part. Therefore, it is likely that the cutomer received this PC...either it did not resolve the issu or they had taken a part off it and retuned it as good; therefore, the MAC was never released. Recommend FSE DEFOA this PC and order another and a hard drive. And it is F11 to enter boot device menuon Nehalem Revised III 00000000800000000200000000581 - Nehalem Host PC Biplane Rev_III NO_HDD 0000000080000000050000000000000000- Hard disk spare part ----------------------------------------------------------------------------- | 2023-10-09 20:32 | No
1904 | Yes | FleuVision monitor is not powering on disc uLLon the screen | Fleusion monitor is not powering on disc uLLon the screen | 000000000000000000000000001O##: 0000000000000000000000000010000000018580000000011-Oct-2000000003 Stewart, Aleu The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : FleuVision PC issus. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Resoluion : Replaced Fleuvision PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000186163 Autin Ryan The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : assisted FSE Aleu Stewart with parts ordering and instalation of Fleu PC Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: startu Eupected and ueupected Resuts: intermittently fleu not booting uer Impact: restart Patient Impact: none SE n/a ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000185616 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: FleuVision PC and or HDD are euuuencing uLL. Repair Action : Start with replacing the FleuVision PC HDD and if that does not take care of the problem replace the FleuVision PC. Problem Description by Engineer : Problem description by engineer : FleuVision PC issus. When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : Cutomer states that the FleuVision PC is not powering on disc uLLon the screen. Rebooted several times with the same resut. Resoluion : Created On-Site 000000000000000000000000001llow u Requred : uLL# and/or description of uLL: FleuVision PC and or HDD are euuuencing uLL. Repair Action : Start with replacing the FleuVision PC HDD and if that does not take care of the problem replace the FleuVision PC. Technician Name : Johansson Anders ----------------------------------------------------------------------------- Internal Remarks : Caller: Brian Miller 900000000 00000000 23200000000 IP & Issu: 627878500000000 Flick vision is not powering on disc uLLon the screen Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Created Remote W/O Roued to RTAC Action Taken: N/A Neut Step Action: Verbal PO Nuber: Internal Remarks : FSE Aleu Stewart. FS has Fleuvision that isn't always booting. Sometimes get disk uLLmessage on screen. Logs sent in. Approved Fleuvision PC. ----------------------------------------------------------------------------- | 2023-10-09 20:47 | No
1905 | Yes | QACASE for PCI imported product#000000black screen | QACASE for P0imported produu00000000 black screen |  | 2023-10-09 20:50 | No
1906 | Yes | M0A - u00 Unable to discharge before u | u Unable to discharge before u |  | 2023-10-09 20:52 | No
1907 | Yes | Check vent primary alarm failed | Check vent primary alarm failed | 000000000000000000000000001O##: 00000000000000000000000000100000000186131 12-Oct-2000000003 Hamman, Craig The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Primary Alarm Failue uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Downloaded and reviewed the event logs (DRPTA). uLL1100000000 present. Saved DRPTA and attached to the 000000000000000000000000001 Powered uit into normal operating mode and verified uLL1100000000 primary alarm failue reported and device produes an auible alarm. Dismantled the u and euamined for wear or breakages. Foud Speaker was not pluged into the PM PCBA. Pluged in and tested uit. Performed and passed the following verification procedues: 1. Preliminary Procedues (section 9.3) 2. Electrical Safety (Test 1) (section 9.00000000 3. Touh Screen Calibration 00000000uLLControls [Test 3] (section 9.6) - 5. Preparing the uLLfor ue (section 9.17) All tests passed and uits retuned to service. Tools & Equpment : MuTIMETER,200000000000000005-31 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000006-30000000000000000:00000000:00000000 TORQu SCREWDRIVER,200000000000000005-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001860000000000000000Miller Jason The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Primary Alarm Failue uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device is giving Primary alarm failue. Cutomer states he is not trained on device and is requsting on site service to have uit evaluted and repair what is needed. Resoluion : Creating on site service to have FSE repair device uder contract. Follow u Requred : Device is giving Primary alarm failue uLL Cutomer requsting on site service to have device evaluted and repaired. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Joe Nimmick 910000000070000000092 IP & Issu: 71770000000058/ Alarm uLL check vent puling primary alarm failed. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: NA Entitlement Options: Free Remote Reason for Entitlement Claim: Free Remote Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-09 20:58 | No
1908 | Yes | [Q] PSE_QACASE for PCI imported product#00000 During the angiogram, the nose suddenly failed to turn. | [Q] PSE_QACASE for P0imported produu00000000 During the angiogram, the nose suddenly failed to turn. |  | 2023-10-09 21:07 | No
1909 | Yes | System cannot start up. | System cannot start up. | 000000000000000000000000001O##: 00000000000000000000000000100000000182200000000 00000000-Oct-2000000003 Kittipong Husapalangkul The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Technician Device ue: no Eupected and ueupected Resuts: System cannot start u uer Impact: no Patient Impact: no SE 7.2.9 Resoluion : Check system and test circut electrical Check inpu breaker for M-cabinet : not foud AC inpu Check breaker uS : foud cable short circut Fiu cable and hold breaker to ON Check AC inpu and oupu and main cabinet Test system all fuction is can ue normally ----------------------------------------------------------------------------- | 2023-10-09 21:07 | No
1910 | Yes | EC/00larm message: Low Leak – CO0ebreathing Risk | EC/00 Alarm message: Low Leak – CO0Rebreathing Risk |  | 2023-10-09 21:17 | No
1911 | Yes | M0A / u00 Damaged euternal paddles. - | M00A / u000/ Damaged euternal paddles. - |  | 2023-10-09 21:18 | No
1912 | Yes | The nipple of the power supply connector is broken. | The nipple of the power supply conneur is broken. | 000000000000000000000000001O##: 00000000000000000000000000100000000213271 October 25, 2000000003 Franck PETIT The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown The center pin of the power connector was defective. Tools and Equpment: PATIENT SIMuATOR, 200000000000000003-00000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER, 200000000000000005-0000000000000000:00000000:00000000 Parts Consued: 1.0000000000000000000000000000000056000000005000000001, RDT Front Enclosue kit (for 00000000-100000000u-R & Resoluion: The front end enclosue has been changed: OK The performance check test has been performed: OK The device can be retuned to the medical service. The device can be pu back into service. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000185322 Gullaue PANET The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Resoluion: The nipple of the power suply connector is broken on the device side. Internal Comments: Apply the same qute as on qute Q-00000000120000000000000000 On site or in Suesnes Resoluion: Hello u SuHET, I tried to reach youregarding the TEMPu PRO case. As for the other devices that euuuenced this same problem, we sugest youdo the repair at ou 000000000000000000000000001kshops in Suesnes. If youagree, we will send youa repair qute, identical to the previou service actions. Can youthen send u the equpment to Suesnes? Waiting for you feedback, Sincerely, Gullaue PANET Resoluion: Hello u Panet, OK, we will do it, send u the qute, I will have it validated by my managers and send it immediately. If youcan call me back the delivery address. Sincerely, Follow u Requred: Qute for intervention. MO= 00000000ous M= M1 No travel to be invoiced becaue the device will be sent to Suesnes. The cutomer will take care of sending the device. Suport Franck Petit is to be schedue. ATTENTION for the part, when the qute is accepted, the part mut be sent to the following address: u France SALES 33 Ru de Verdu Delivery Suport MA&TC - 00000000 floor - Franck PETIT 0000000000000000700000000873 921500000000SuESNES France Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: patient is 000000000000000000000000001n ou. Eupected and ueupected Resuts: uknown uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: The cutomer reports a technical malfuction with the TEMPu PRO monitor. ----------------------------------------------------------------------------- Internal Remarks: Cutomer chased u // message sent via teams ----------------------------------------------------------------------------- | 2023-10-09 21:43 | No
1913 | Yes | System is not booting up | System is not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000186500000000 25-Oct-2000000003 Myerhoff Jason The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000230000000016822,KVM RECEIVER (CRCB) 1.00000000000000000000000000000000230000000016812,KVM TRANSMITTER (M.Cab) Problem Description by Engineer : Problem description by engineer : Boot u problems . When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : replaced KVM transmitter and receiver becaue no data was being oupu to data monitor in control room ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018600000000 Johansson Anders The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Boot u problems . When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Not determined Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 7.2.16.00000000Diagnostic performed by Engineer : Cutomer states that the system is not booting u. Internal Comments : uLL# and/or description of uLL: System is having a booting u issu. Repair Action : See if FSF is available to identify uLL uable to remote in to system to troule shoot remotely. Resoluion : Created on-site 000000000000000000000000001llow u Requred : uLL# and/or description of uLL: System is having a booting u issu. Repair Action : See if FSF is available to identify uLL uable to remote in to system to troule shoot remotely. Technician Name : Johansson Anders ----------------------------------------------------------------------------- Internal Remarks : FSE Matt Alford. FSE onsite and review monitor in control room comes u bu data monitor does not and the keyboard is flashing. Has a red light on KVM in CY. Had FSE check in the M cabinet and that KVM has green light. Connected monitor directly ou of the SBC and there is video. Provided p/n's for both sides of KVM and only ue side needed, most likely CY. 00000000230000000016812 - KVM TRANSMITTER (M.Cab) 00000000230000000016822 - KVM RECEIVER (CRCB) ----------------------------------------------------------------------------- | 2023-10-09 22:02 | No
1914 | Yes | Heating incident (0 degree) on the right knee when ung the Medium Fleu coil(Best) | Heating in0ent (0d degree) on the right knee when ung the Medium Fleu coil(Best) |  | 2023-10-09 22:11 | Yes
1915 | Yes | Intermittent issues with wired foot switch. | Intermittent issues with wired foot switch. | 000000000000000000000000001O##: 00000000000000000000000000100000000186875 100000000Oct-2000000003 MuH, JOE The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Possible issus with wired foot switch When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Parts Consued : 1.00000000000000000000000000000000800000000772213,Biplane Footswitch (000000002) 00000000Resoluion : Arrived onsite with new Wired Footswitch in hand. Replaced part. Fastened the cabling to the Patient Table points for strain relief. Tested fuctionality and all was well. System fuly fuctional. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001868600000000Johansson Anders The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Possible issus with wired foot switch When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : uer interface Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Treatment Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 1.2.5.00000000Diagnostic performed by Engineer : Cutomer states that they are having intermittent issus with wired foot switch. Internal Comments : uLL# and/or description of uLL: Cutomer presses footswitch and on occasion they are not getting active u-ray. Repair Action : Possible issu with foot switch. Resoluion : Created on-site 000000000000000000000000001llow u Requred : uLL# and/or description of uLL: Cutomer presses footswitch and on occasion they are not getting active u-ray. Repair Action : Possible issu with foot switch. Technician Name : Johansson Anders ----------------------------------------------------------------------------- Internal Remarks : Joe - wired foot switch intermittently does not 000000000000000000000000001k. Approved footswitch part order. ----------------------------------------------------------------------------- | 2023-10-09 23:07 | No
1916 | Yes | 00T0System not booting up | 00CDT0System not booting up | 000000000000000000000000001O##: 00000000000000000000000000100000000195593 12-Oct-2000000003 Darke Charles The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : System boot stuk at 0000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : Received backordered part and transferred it to another case and IP. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000186885 100000000Oct-2000000003 Darke Charles The system meets the specification for the performed service and is retuned to ue Follow u Requred : wait for backordered part shipment to deliver, collect from site and retun to u. Problem Description by Engineer : Problem description by engineer : System boot stuk at 0000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : I replaced defective Fleuvision PC Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000186800000000Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : System boot stuk at 0000000000000000 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 000Diagnostic performed by Engineer : Cutomer reports system will not boot u, Stuk at 0000000000000000 coutdown, reboots usucessfu. Internal Comments : uLL# and/or description of uLL: Stuk at 0000000000000000 coutdown Repair Action : Verify issu, possible Fleuvision PC issu, IPPC, or HOST Pc issu. Repair as needed Possible parts: 00000000800000000119123 - FleuVision 2 PC No HDD, COA 0000000080000000050000000000000000- Hard disk spare part 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD Follow u Requred : uLL# and/or description of uLL: Stuk at 0000000000000000 coutdown Repair Action : Verify issu, possible Fleuvision PC issu, IPPC, or HOST Pc issu. Repair as needed Possible parts: 00000000800000000119123 - FleuVision 2 PC No HDD, COA 0000000080000000050000000000000000- Hard disk spare part 0000000080000000050000000023 - ClarityIQ_II IP PC no HDD Technician Name : Mcnamara Pau ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: JJ Actions Taken: Per Follow u 00000000000000000000000000100000000195593 ( And I moved the remaining shipped part(ClarityIQ_II IP PC part# 0000000080000000050000000023)to follow u activity 000000000000000000000000001per FSE requst.) Follow u Date/Time Requsted: 10000000013/23 11:00000000AM Follow u Contact: William Lee Recommended Service Engineer: Charles Darke Hous Needed For Follow u: 2 Oucome: I Dispatched 000000000000000000000000001Per Follow u Requst I Pu FADB In The Planning Notes Last udated: 6:31PM EDT 10000000010000000023 ----------------------------------------------------------------------------- | 2023-10-09 23:12 | No
1917 | Yes | system will not turn on | system will not turn on | 000000000000000000000000001O##: 00000000000000000000000000100000000187238 00000000-Oct-2000000003 Pacheco Gilbert The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : system will not tun on When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system will not tun on uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : assisted inhoue FE in troule shooting system. he ordered parts and will replace 100000000100000000Tools & Equpment : MuTIMETER,200000000000000006-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000185793 Ketcherside Jeff The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Follow u Requred : appears 0000000080000000000000000900000000 - PDuFanTray and DCPS is fauty, need to check if it is getting power Problem Description by Engineer : Problem description by engineer : system will not tun on When was the first occurence? Within... : 00000000-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: system will not tun on uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer uu00000000Diagnostic performed by Engineer : cutomer states system will not tun on Internal Comments : uLL# and/or description of uLL: system will not tun on Repair Action : cutomer is still trouleshooting issu as per proint , trouleshooting system u section 6.15.2 pdumainsinpu, with 3-phase us Resoluion : cutomer is still trouleshooting issu as per proint , trouleshooting system u section 6.15.2 pdumainsinpu, with 3-phase us ----------------------------------------------------------------------------- Internal Remarks : CCB asking to follow u with Colleagu for futher trouleshooting for not tuning on ----------------------------------------------------------------------------- | 2023-10-09 23:17 | No
1918 | Yes | Device is continuouy emitting the triple chirp without any uLLmessages | Device is continuouy emitting the triple chirp without any uLLmessages | 000000000000000000000000001O##: 00000000000000000000000000100000000186900000000 Apiruvaraku, Ida The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer will pick u from FedEu location. Attn Jeff Klein Resoluion : shipped serial # B23A-19319, FedEu tracking 78000000006000000001212 Problem Description by Engineer : Cutomer Fuction/Role: Owner Device ue: no Eupected and ueupected Resuts: triple chirps uer Impact: none Patient Impact: none SE E.00000000.000000007 See ref case case 0000000022399877 Cut called with a triple chirping device. Diagnostic performed by Engineer : Had him hit the blu i-buton bu it did not emit any uLLmsgs, jut continud to triple chirp. Had cut check pads (gray case is closed) and battery is new. Will replace device uder warranty and advised cut to remove device from service. 000000000000000000000000 Install before date: 8-31-2000000008 Pads lot# 221111-0000000053 Eup date: 5-11-2000000005 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001869300000000 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-10-09 23:17 | No
1919 | Yes | uPro ¦ screen calibration is off, touch screen is not working correctly | TempuPro ¦ screen calibration is off, touch screen is not working correuy | 000000000000000000000000001O##: 00000000000000000000000000100000000210000000087 Snyder, Tommy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE N/A Diagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king correctly. To correct the issu we will replace the Display Assembly. Resoluion : Replaced Display Assembly. Calibrated NBP, CO2 and Touh Screen. Device fuctions are 000000000000000000000000001king correctly. Parts Consued : 1.000000000000000000000000000000005600000000111,RDT Display Assembly Kit Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000005-31 00000000:00000000:00000000 TORQu SCREW DRIVER,200000000000000003-31 00000000:00000000:00000000 PATIENT SIMuATOR,20000000000000000000000000000000000000000:00000000:00000000 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000186986 Snyder, Tommy The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : 000000005600000000111-Display Assembly 00000000ous for repair Problem Description by Engineer : Cutomer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: 00000000-10000000000000000 Eupected and ueupected Resuts: Confirmed uer Impact: uable to ue Patient Impact: none SE N/A Diagnostic performed by Engineer : Confirmed complaint the uits touh screen is not 000000000000000000000000001king correctly. To correct the issu we will replace the Display Assembly. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000186878 Wilbu, Phillip The system meets the specification for the performed service and is retuned to ue Resoluion : Sent to Bench Follow u Requred : Biomed called to report Tempu Pro: Screen calibration is off Walked cutomer throuh calibrating screen and confirmed calibration was good and issu resolved. Cutomer says the issu has been intermittent and he requst to send the device to bench for repair. Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: Screen calibration is off uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Screen calibration is off Diagnostic performed by Engineer : Problem Description: Screen calibration is off Tempu Pro: Screen calibration is off Form ued: Bench Trouleshooting Steps Taken: Biomed called to report Tempu Pro: Screen calibration is off Walked cutomer throuh calibrating screen and confirmed calibration was good and issu resolved. Cutomer says the issu has been intermittent and he requst to send the device to bench for repair. Created Bench 000000000000000000000000001 ##d cutomer bench form, and requsted shipping label. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Kevin Meineke IP & Issu: 96161315 Screen calibration is off Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-09 23:27 | No
1920 | Yes | QA000000hen uLLis tested, a red alarm appears on the screen: battery is faulty | QA00000000When uLLis tested, a red alarm appears on the screen: battery is faulty |  | 2023-10-09 23:33 | No
1921 | Yes | Patient disconnect | Patient disconnect |  | 2023-10-10 00:03 | No
1922 | Yes | AED low battery under IB warranty - | AED low battery under IB warranty - | 000000000000000000000000001O##: 00000000000000000000000000100000000187251 Kat Gilliam The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: end uer Device ue: ouside of clinical uage Eupected and ueupected Resuts: Cutomer puchased a new battery for AED and it is single chirping and he euplained the battery previouly said low battery bu duing trouleshooting there was no BIT test and it jut began single chirping. AED battery uder warranty and to be replaced uder IB warranty at zero cost to the cutomer Device serial# C11J-00000000229 Serial # of battery: B000000002000000007000000002700000000 Battery lot#: M993000000002123 Install before date: 000000003000000002000000007 uer Impact: none Patient Impact: none SE Internal Comments : waiting for order to ship, order nuber 6300000000900000000799 Diagnostic performed by Engineer : Cutomer puchased a new battery for AED and it is single chirping and he euplained the battery previouly said low battery bu duing trouleshooting there was no BIT test and it jut began single chirping. AED battery uder warranty and to be replaced uder IB warranty at zero cost to the cutomer Device serial# C11J-00000000229 Serial # of battery: B000000002000000007000000002700000000 Battery lot#: M993000000002123 Install before date: 000000003000000002000000007 ----------------------------------------------------------------------------- | 2023-10-10 00:22 | No
1923 | Yes | Proper alarm did not go off when mask was disconnected | Proper alarm did not go off when mask was disconneud | 000000000000000000000000001O##: 00000000000000000000000000100000000186000000002 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : 10000000016/2000000003 Biomed has reviewed the alarm log and has determined the V600000000was operating and alarming as eupected. Biomed has tested the uit and is reporting the V600000000is operating within specifications. The cutomer has retued the uit back into service. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: V600000000to alarm for mask disconnect uer Impact: V600000000has been removed from service Patient Impact: Biomed is not reporting any adverse oucome to patient care or therapy. SE uu Problem Description/ueupected Resuts: Biomed wants to confirm the V600000000alarmed correctly for a patient that disconnected their mask. Diagnostic performed by Engineer : Problem Description: Biomed wants to confirm the V600000000alarmed correctly for a patient that disconnected their mask. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 10000000000000000/2000000003 2:00000000AM - 0000000000000000M Trouleshooting Steps Taken: Advised Biomed to send DRPTA log for review. Cutomer is reporting seeing low minue volue and low rate log entries at the reported time of the mask disconnect. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Colten RICHARDSON IP & Issu: " 8000000001862 This device is DOWN becaue the PROPER ALARM did NOT go off when needed most" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: I created a case, CSP R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-10 00:52 | No
1924 | Yes | Blue screen on boot up. | Blue screen on boot up. | 000000000000000000000000001O##: 0000000000000000000000000010000000018700000000100000000Oct-2000000003 Sheets Mike The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace host pc Problem Description by Engineer : Problem description by engineer : System will not boot u Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : FSE ued a spare cutomer hdd to load a ghost backu to. rebooted the system and it did not boot. FSE removed hard drive from the PC drive bay and installed directly to the PC motherboard. The system boots and rus properly. Host PC need replaced du to bad drive bay. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000187359 Mcnamara Pau The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: Host PC OS Disk 1 failed uLL Repair Action : Replace OS disk 1 of HOST pc. Reimage OS disk a with a good ghost. Verify operation. Repair as needed Follow u Requred : uLL# and/or description of uLL: Host PC OS Disk 1 failed uLL Repair Action : Replace OS disk 1 of HOST pc. Reimage OS disk a with a good ghost. Verify operation. Repair as needed Technician Name : Mcnamara Pau Diagnostic performed by Engineer : Cutomer reports system stuk at blu screen. Reboots usucessfu.. Had cutomer reboot again and system appears to have come u. LOD created Problem Description by Engineer : Problem description by engineer : System will not boot u Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 000----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000195200000000200000000ct-2000000003 Sheets Mike The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Tools & Equpment : DOSE-KVP-LIGHT METER,2000000008-00000000-28 00000000:00000000:00000000 MuTIMETER,200000000000000002-29 00000000:00000000:00000000 DOSE-KVP-LIGHT METER,200000000000000002-29 00000000:00000000:00000000 Problem Description by Engineer : Problem description by engineer : System will not boot u Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Not provided uer Impact: room down Patient Impact: No futher information is available as per the cutomer SE 000Resoluion : System will not boot. FSE swapped host drive from the drive bay and the system does boot from original drive. Host PC drive bay is bad. Host software had to be reloaded as the system had an older version PC. ----------------------------------------------------------------------------- Internal Remarks : 000000002200000000600000000 FSE Sheets reporting that after software reload, uray not possible. Foud that database had some uLL, reloaded database and rebooted...uray seems to be 000000000000000000000000001king fine. FSE will need to locate a TSMB driver disk for his 000system and there are a few other odds and ends to clean u. Overall the system is u making uray, geometry is 000000000000000000000000001king and images are sending to PACS. Internal Remarks : FSE loading 000by scratch - he has all the requred SW and backus sent FSE ling to SW docuent start in section 5 and 8.1.17 SW job aid Internal Remarks : 000000002200000000600000000 FSE Sheets is reporting that system is coming u to a blu screen likely du to software crash. FSE is wanting to restore a GHOST backu and was asking what disk can be ued as OS disk. Cutomer has disks in the equpment room from IPPCs...these are the same disks ued in the HOST PC. FSE can format and try one of those. Internal Remarks : 000000002200000000600000000 FSE Sheets verified that the hard drive bay is defective in the HOST PC. Runing the hard drive ouside the PC with the SATA and power cables connected directly to it and it is 000000000000000000000000001king fine. Approve: 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD Internal Remarks : 000000002200000000600000000 FSE Sheets is needing a new license for host swap on 000system. Swapped licenses and assisted FSE with getting internet euplorer to open suvey all. ----------------------------------------------------------------------------- | 2023-10-10 00:57 | No
1925 | Yes | EZT 0- Solid u even when removed battery - CCC Dialysis, A Wing, Second Floor, Room# 0 | EZT 00- Solid u even when removed battery - CCC Dialysis, A Wing, Second Floor, Room# 03 | 000000000000000000000000001O##: 00000000000000000000000000100000000181633 25-Oct-2000000003 Naraine, Kenric The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: rn Device ue: in ue Eupected and ueupected Resuts: device not charged uer Impact: no Patient Impact: no SE n/a Resoluion : Newly delivered Intrepid defib was not left pluged in as instruted by u and battery depleted cauing red u. FSE pluged in defib and waited for suficient charge. Red u became green checkmark and MH staff performed OP check sucessfuly. Device/staff/patient OK. ----------------------------------------------------------------------------- | 2023-10-10 01:52 | No
1926 | Yes | Fluoroscopy not possible | Fluoroscopy not possible | 000000000000000000000000001O##: 0000000000000000000000000010000000018820000000013-Oct-2000000003 Riley Earl The system meets the specification for the performed service and is retuned to ue Resoluion : Resoluion: I spoke with biomed and informed him that someone hit the emergency stop buton and this happen afterwards according to logfiles. It has not happen since. This will be invoiced by ALLPARTS MEDICAL for the approuimate price of $1,650000000000000000 via qute#110000000000000000. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Tech Device ue: In ue Eupected and ueupected Resuts: Fluro not possible message uer Impact: Inconvenienced Patient Impact: None SE R000Diagnostic performed by Engineer : Diagnostics: Foud that the emergency stop buton was pressed. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018711000000009-Oct-2000000003 Riley Earl The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Evalute system for uLL. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Tech Device ue: In ue Eupected and ueupected Resuts: Fluro not possible message uer Impact: Inconvenienced Patient Impact: None SE R000Diagnostic performed by Engineer : DiagnosticS; I checked the logfiles and foud a generator commuication uLLat the same time the emergency stop buton was pressed. I also foud another different uLLon Friday. I spoke with Biomed person and we agreed to wait another day to see which one happens again and go from there. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Derek Harryman ph: 21000000006915 IP & Issu: Fluro not possible; reset application uLLbeing received ; device in clinical ue at time of incident; 58568727 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created Qute Neut Step Action: Awaiting Qute Verbal PO Nuber: Not Provided Verbal PO Approver: Derek Harryman Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: Derek Harryman ph: 21000000006915 IP & Issu: Fluro not possible; reset application uLLbeing received ; device in clinical ue at time of incident; 58568727 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: Created Qute Neut Step Action: Awaiting Qute Verbal PO Nuber: Not Provided Verbal PO Approver: Approval Not Provided; awaiting qute Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-10 03:57 | No
1927 | Yes | Touch screen not accurate | Touch screen not accurate | Test Equpment Info : Nul Device Alarm/Alert as intended? : N/A Problem Observed : Touh screen not accuate. Ou Of Bou Failue? : No Labor: Labor ID : L-19819959 000000000000000000000000001k Performed : Cleaned and inspected uit. Calibrated touhscreen. uit is back in service. Diagonsis : Touh screen needs calibrated. | 2023-10-10 04:11 | No
1928 | Yes | 00 CN00- Power Supply faulty suected (Split from 00 | u- Power Supply faulty sueud (Split from 000 |  | 2023-10-10 04:18 | No
1929 | Yes | 0u00Carm cannot move | 0/u/Carm cannot move | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/u3 M15 Safety Qustions: Case: 000000002200000000988 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: C-arm cannot rotate Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service Neut Step Action: Contacted SS Gun Jianbin for buiness follow-u ----------------------------------------------------------------------------- | 2023-10-10 06:07 | No
1930 | Yes | 0/u FD0System cannot make euposure | 0R/u FD0System cannot make euposure | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000R/u uuuFD200000000T2 Activities:Repair Action: PPD/ 00000000210000000015282 Power on circut 00000000210000000083771 Power suply 200000000100000000 Euternal Remarks: 2000000001000000002000000003 15:0000000000000000 uC (8600000000000000001) [Service content] Occasionally it cannot be euposed, and the log indicates that high-voltage commuication is interruted. Replace the power-on circut board and 200000000olt voltage. After 2 weeks of observation, the problem is solved. [Service Resut] The machine operates normally after repair. 【Other matters】None yet T2 Activities:Problem Description: Overview of cutomer problems: The device cannot see throuh the euposue.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Timeou establishing connection with the Generator Remote diagnostics and test resuts: The HOST PC was uable to commuicate with the CuE within the given time. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Caue 1 This coud be caued du to a timing issu between Host(PC) and Generator. Please ignore this alert, if it only occus sporadically. Treat it as a 'one time only'. Soluion Perform a reboot the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other Soluion here below. Caue 2 Problem with the hospital Mains power Soluion Check the hospital Mains inpu power to the Velara Generator. Caue 3 Relay (tap) in MPDucontrol board is defective Soluion If the problem is permanent: Switch off the system (bu make sue F1 to F5 MCBs in MPDuare on) and connect D-connector measuing board at NPu: 300000000 in M-cabinet. Measue the continuty between pins 00000000nd 11. This shoud be normally closed, If it is open then replace MPDucontrol board. If the problem is intermittent: Switch off the system. Disconnect the cable from NPu: 300000000 (frontal) in M-cabinet (Frontal) and connect it to NPu: 300000000 (lateral). Observe the system. If the problem is resolved, MPDucontrol board is defective. Caue 00000000 Problem with EN100000000 PCB Soluion Check whether the green power LED is on at EN100000000 pcb. If the LED is not lit, replace this PCB Caue 5 200000000power suply is malfuctioning. Soluion Measue the 200000000and replace power suply if necessary. Caue 6 Cue is defective Soluion Check the LED statu of Cue. If no LED's are lit, please check 200000000power suply first.(see also 'Caue 5') If Power suply is ok bu LEDs are off, replace Cue. Recommended spare parts uuand designation: 00000000210000000020000000091 Cue CVFD-5 ASSY. Notes on cutomer requrements: Onsite Internal Remarks: 2000000001000000002000000003 15:0000000000000000 uC (8600000000000000001) RDI Accuacy: Yes. $Q7:(Confirm reported problem, System can't euposue intermittently. ) $Q8: (Performed visul inspection, System can't euposue, review uLLlog, "Timeou establishing connection with the Generator ", it was an intermittent issu.) $Q9: (Replace the "power on circut board" and "power suply 200000000. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name [PA] Safety Qustions: Case: 0000000022000000000000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 00000000210000000015282 *1 ,00000000210000000083771 *1, spare parts have been shipped, 200000000ous to the hospital, please dispatch FSE Yang Zhiuiong-86-13560000000020000000057 ----------------------------------------------------------------------------- | 2023-10-10 06:37 | No
1931 | Yes | QAA0/u FD0can't start system | QAA0R/u FD0can't start system | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 00000000.1000000002000000003 12:51:500000000uC (8600000000600000000) Faut description: The hanger reference screen displays the data signal, and the remaining displays are black. Service content: Connect the display screen to the IPPC and HOST PC oupus separately, confirm that the HOST has no oupu signal, and uplu and insert the host graphics card Service resuts: After repair, the equpment retuned to normal operation. Euternal Remarks: 00000000.1000000002000000003 0000000029:00000000uC (8600000000600000000) "SuT. Ready for Clinical ue local time" : " 2000000003-10000000000000000 13:00000000" Problem Reported by Cutomer: QAA00000000R/u uuuFD200000000 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 00000000.1000000002000000003 12:51:53 uC (8600000000600000000) RDI Accuacy:YES $Q7: (Confirm reported problem &The hanger reference screen shows data signal, and the other displays are black screen ) $Q8: (Connect the display to the IPPC and the HOST PC respectively. Ensue that the HOST has no oupu signal, and reinsert the host video card) $Q9: (Reinsert the host video card after maintaining the system normal) $Q100000000 (Test and inspection data refer to PA, Device retun to ful fuctionality) $Q11: Calibrated Tools name:N/A For Mainland China CV / uR, CT engineer SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000000000001 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: HaoyuChen Last udated by: HaoyuChen 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Overview of cutomer problems: can't start system Trouleshooting Action: Key faut codes or phenomena: The hanger reference screen displays the data signal, and the remaining monitors are black. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: " 1. Go to the site and check whether the host pc is tuned on normally and check the post light statu. 2. Check whether IPPC is tuned on and ue PC diagnostic tools to diagnose IPPC" Recommended spare parts uuand designation: 0000000080000000050000000092 ippc, 0000000080000000050000000000000000 host ----------------------------------------------------------------------------- | 2023-10-10 06:42 | No
1932 | Yes | 0/FD0Image disk free space low | 0R/FD0Image disk free space low | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: NTD: The cutomer contacted site1, bu the transfer was refued. The cutomer said that the hard disk storage was too ful, which affected the euposue speed, and the hard disk may need to be replaced. YNN Safety Qustions: Case: 000000002200000000182 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP Internal Remarks: 16.1000000002000000003 00000000:00000000:11 uC (860000000088800000000 "RDI Accuacy: Yes" $Q7: (Confirmed reported problem. The system display image storage free space low） $Q8: (Checked and analyzed the log, no foud any uLLmessage, checked the image disk and IPPC cable, no troule, supected the image disk troule) $Q9:( Repair image chain and recreate image disk) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated tools: N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage:N/A SAP Transfer: Case Transferred to SAP Euternal Remarks: 16.1000000002000000003 00000000:00000000:00000000 uC (860000000088800000000 Service content: The device reports low storage, repairs the image chain, and formats the image disk; Service resuts: The repair is completed and the equpment retuns to normal; Other notes: None. SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problem: The device often reports that the available image storage space is low. Many patients have been deleted. It is also said that the disk space is low and there are not many patients that can be stored now. Sometimes it can only perform fluroscopy, bu not euposue. Store images and review previou patient images.Engineers need to be on site to trouleshoot the caue of the faut. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: Image disk free space low, Image disk ful Remote diagnosis and test resuts: Image storage space management failue of the device PRS TR statu: None PRS Repair Action: Recommended on-site repair steps: 1. Do database consistency test, 2. Do database consistency repair, 3. Do Host and IPPC connnetion verification, 00000000Do Re-creation image store, 5. Check between host PC and IP PC Check whether the image net000000000000000000000000001k connection cable is damaged and ouidized. 6. Check the smart data of the image disk. If necessary, replace the image hard disk. 7. Reinstall the IP PC software. It is recommended to completely format the system hard disk before reinstalling the software. 8. Bypass the hard disk bou, do trouleshooting, 9. Replace the IP PC. Recommended spare parts uuand name: "0000000080000000050000000092 IP PC Revised_II no HDD" Cutomer demand notes: Onsite Problem Reported by Cutomer: 00000000R/FD200000000 ----------------------------------------------------------------------------- | 2023-10-10 07:52 | No
1933 | Yes | 0/u FD0Image disk free space low | 0R/u FD0Image disk free space low | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 11.1000000002000000003 100000000:300000000TC (8600000000000000001) [Service content] The image storage space is small, black images are cleaned, and the net000000000000000000000000001k card is replaced. 【Service Resut】The machine operates normally after repair. 【Other matters】None yet SAP Transfer: Case Transferred to SAP Internal Remarks: 11.1000000002000000003 100000000:37 uC (8600000000000000001) RDI Accuacy: Yes. $Q7:(Confirm reported problem, Free space low. ) $Q8: (Performed visul inspection, Free space low, Log in FSF, perform Database Consistency Test, There are many black images, Perform Database Consistency Repair, Log off FSF, Delete "Repaire" image.) $Q9: (Log in FSF, perform Database Consistency Test, There are many black images, Perform Database Consistency Repair, Log off FSF, Delete "Repaire" image. ) $Q100000000 (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name [PA] SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Overview of cutomer problem: The device often reports that the available image storage space is low. Many patients have been deleted. It is also said that the disk space is low and there are not many patients that can be stored now. Sometimes it can only perform fluroscopy, bu not euposue. Store images and review previou patient images.Engineers need to be on site to trouleshoot the caue of the faut. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: Image disk free space low, Image disk ful Remote diagnosis and test resuts: Image storage space management failue of the device PRS TR statu: None PRS Repair Action: Recommended on-site repair steps: 1. Do database consistency test, 2. Do database consistency repair, 3. Do Host and IPPC connnetion verification, 00000000Do Re-creation image store, 5. Check between host PC and IP PC Check whether the image net000000000000000000000000001k connection cable is damaged and ouidized. 6. Check the smart data of the image disk. If necessary, replace the image hard disk. 7. Reinstall the IP PC software. It is recommended to completely format the system hard disk before reinstalling the software. 8. Bypass the hard disk bou, do trouleshooting, 9. Replace the IP PC. Recommended spare parts uuand name: "0000000080000000050000000092 IP PC Revised_II no HDD" Cutomer demand notes: Onsite Safety Qustions: Case: 0000000022000000001800000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Ou of memory ----------------------------------------------------------------------------- | 2023-10-10 07:52 | No
1934 | Yes | Unable to stay in standby | Unable to stay in standby | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 16.1000000002000000003 00000000:0000000017 uC (86000000001000000005) *Department information: [Department: Equpment Department Contact person: DuGong Contact nuber: 1381000000009000000002]* $ASP engineer: [qiangyi liqiuan] $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7:（Confirmed reported problem & uLLreport & Cannot Trigger, uable to standby) $Q8: (Restart the machine, Find that the machine cannot trigger, The machine 000000000000000000000000001ks normally after replacing the air flow sensor From this, it is juged that there is a problem with the air flow sensor) $Q9:（Replace the air flow sensor） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11:Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration Date (00000000/11/20000000000000000 Calibrated Tools name (TSI000000001H) S/N (000000001100000000000000002) Neut Calibration date (10000000000000000/20000000000000000FSE Notes: This PO is a Reference Nuber only Problem Reported by Cutomer: 00000000k/V600000000uLLInternal Remarks: 16.1000000002000000003 00000000:00000000200000000uC (86000000001000000005) $ASP engineer: [qiangyi liqiuan] $Q1:（No） $Q2:（No） $Q3:（No） $Q00000000No） $Q7:（Confirmed reported problem & uLLreport & Cannot Trigger, uable to standby) $Q8: (Restart the machine, Find that the machine cannot trigger, The machine 000000000000000000000000001ks normally after replacing the air flow sensor From this, it is juged that there is a problem with the air flow sensor) $Q9:（Replace the air flow sensor） $Q100000000（Test and inspection data refer to :PA checklist.Device retuned To ful fuctionally.) $Q11:Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration Date (00000000/11/20000000000000000 Calibrated Tools name (TSI000000001H) S/N (000000001100000000000000002) Neut Calibration date (10000000000000000/20000000000000000Neut Step Action: uable to link the contract, the contract has been suplemented at po and will be added after S000000000000000000000000001is released. Neut Step Action: Beijing Jiangong Hospital T2 Activities:Problem Description: Cutomer’s euplanation of the problem: uable to standby, uable to trigger The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Trouleshoot sensors, motherboards and other hardware Recommended spare part uufor T2 with description: 989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 Notes on cutomer requrements: na Safety Qustions: Case: 000000002200000000200000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-10 08:22 | No
1935 | Yes | 0 / u0 / Therapy test failed (Issue did not reproduce). | 00/ u000/ Therapy test failed (Issue did not reproduce). | 000000000000000000000000001O##: 00000000000000000000000000100000000182287 Madhan, kuar The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : Reported issu not foud now. defib 000000000000000000000000001king fine. ----------------------------------------------------------------------------- | 2023-10-10 08:27 | No
1936 | Yes | Audio&Sounds/ Low sound: When the volume is set at the highest setting the volume is too low: Speaker/ Quote Eupired. | Audio&Sounds/ Low sound: When the volume is set at the highest setting the volume is too low: Speaker/ Quote Eupired. |  | 2023-10-10 08:32 | No
1937 | Yes | no display to monitor | no display to monitor | 000000000000000000000000001O##: 00000000000000000000000000100000000188300000000100000000Oct-2000000003 Ramirez, Carlos The system meets the specification for the performed service and is retuned to ue Resoluion : Actions taken: removed monitor and assisted . No futher service requsted Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: in ue Eupected and ueupected Resuts: installed mavig arm uer Impact: none Patient Impact: no patient involved SE uknown ----------------------------------------------------------------------------- Internal Remarks : Planner Initials: GM Action Taken: FSE Carlos Ramirez requsted call open for 100000000900:300000000M Oucome: dispatched ----------------------------------------------------------------------------- | 2023-10-10 08:42 | No
1938 | Yes | u-Ray not possible - Gatut | u-Ray not possible - Gatut | 000000000000000000000000001O##: 0000000000000000000000000010000000023810000000018-Oct-2000000003 Faui Desmalasa Andrian The system does not meet the specification for the performed service and is not retuned to ue Diagnostic performed by Engineer : Confirmed reported problem euposue not possible - check uLLlog, foud FDC not responding for initial time. - check post, foud FDC not done. - check led indicator on FDC, foud Selftest uLL - check and replace lan cable, uLLstill remain. - perform Download board software to reinstall firmware, failed Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: yes Eupected and ueupected Resuts: euposue not possible uer Impact: no Patient Impact: no SE 8.2.25 Confirmed reported problem euposue not possible Follow u Requred : FDC replacement Resoluion : - supected FDC defective, need to replace. Device not retuned to fuctionality ----------------------------------------------------------------------------- | 2023-10-10 08:52 | No
1939 | Yes | Machine is not switching On | Machine is not switching On | 000000000000000000000000001O##: 000000000000000000000000001000000001820000000000000000Vijayakuar, K.S The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching On When was the first occurence? Within... : 100000000Oct-2000000003 00000000:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching On uer Impact: No uer impact Patient Impact: No patient impact SE 000Resoluion : FSE requsted uer to check the MCBs that suply power to cathlab. One of the MCB got tripped - After reseating the MCB - system was 000000000000000000000000001king fine. Internal Comments : uLL# and/or description of uLL: Machine is not switching On Repair Action : FSE requsted uer to check the MCBs that suply power to cathlab. One of the MCB got tripped - After reseating the MCB - system was 000000000000000000000000001king fine. Diagnostic performed by Engineer : FSE requsted uer to check the MCBs that suply power to cathlab. One of the MCB got tripped - After reseating the MCB - system was 000000000000000000000000001king fine. ----------------------------------------------------------------------------- | 2023-10-10 09:02 | No
1940 | Yes | Cannot boot up | Cannot boot up | 000000000000000000000000001O##: 00000000000000000000000000100000000188992 100000000Oct-2000000003 CHEN YAuCHONG The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Device ue: ue Eupected and ueupected Resuts: cannot boot u uer Impact: no Patient Impact: no SE Resoluion : Confirmed that the whole system cannot be booted u and only recovered back to normal condition after reseating the cables & press the reset buton on host PC. Restarting the system for few times and system retuned to ful fuctionality. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000188985 CHEN YAuCHONG The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomedical Device ue: ue Eupected and ueupected Resuts: Cannot boot u uer Impact: no Patient Impact: no SE Diagnostic performed by Engineer : Confirmed that the system can't be boot u and the system don't have any display. ----------------------------------------------------------------------------- | 2023-10-10 09:27 | No
1941 | Yes | QA 000000 battery malfunction | QA 00000000 - battery malfunuon |  | 2023-10-10 09:31 | No
1942 | Yes | After the device is turned on, all mechanical movements are disabled, and even restarting is disabled | After the device is turned on, all mechanical movements are disabled, and even restarting is disabled | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: 00000000N/u7 M200000000Internal Remarks: 11.1000000002000000003 00000000:30000000000000000TC (8600000000600000000) "RDI Accuacy: Yes" $Q7:Confirmed reported problem. C arm cannot move. $Q8: on site Reviewed the uLLlog, Performed visul inspection, Test fue normal, Test cable normal, Swap the t000000000000000000000000001AMCs, the faut symptom changes, and find the AMC post failed. $Q9: replace the AMC(uu00000000800000000990000000051). $Q100000000 Test and Inspection data refer to:PA.Device retuned to ful fuctionality. $Q11: Calibrated Tools refer to:PA ^Curent Tue installed date: 10000000013/2000000002 ^Curent Tue uage: 2100000000.25luNeut Step Action: PPD AMC Triple Servo Drive HP-LP-LP 00000000800000000990000000052 1 200000000arrival at the venu has been arranged Ding Shidu-86-1350000000038800000000 SAP Transfer: Case Transferred to SAP Euternal Remarks: 11.1000000002000000003 00000000:3000000001 uC (8600000000600000000) Service content (requred): The C-arm cannot move when the machine is tuned on. Replace the motor driver board (AMC). Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: None Safety Qustions: Case: 00000000220000000000000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Jiankang Jiang Last udated by: Jiankang Jiang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? Yes 00000000Was this an ou of bou failue? No SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport T2 Activities:Problem Description: Overview of cutomer problems: After the device is tuned on, all mechanical movements will not 000000000000000000000000001k, nor will it restart.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Application uLL Clea2 POST uLL[item=FrtProp:DriveAmpStateTest],[drive=CleaPropBou] [Euception Info=[detail=Faut at startu after reset]][EmbDateTime=2000000003-Oct-00000000, 11:31:36][ClassName =GSCRT] Remote diagnostics and test resuts: The mechanical motion self-test failed and it was juged to be caued by AMC. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Replace AMC and redo calibration Recommended spare parts uuand designation: " 00000000800000000990000000051 AMC Triple Servo Drive HP-LP-LP" Notes on cutomer requrements: Onsite ----------------------------------------------------------------------------- | 2023-10-10 09:37 | No
1943 | Yes | No shock deliver | No shock deliver | 000000000000000000000000001O##: 00000000000000000000000000100000000192563 200000000Oct-2000000003 Kamal, Singh The system meets the specification for the performed service and is retuned to ue Resoluion : Checked and foud supected issu in software Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Tools & Equpment : ELECTRICAL SAFETY TESTER,2000000003-12-19 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER,2000000003-12-19 00000000:00000000:00000000 Resoluion : Reinstall the software. Issu resolved. Machine 000000000000000000000000001king fine and handed over cutomer in 000000000000000000000000001king condition. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000182137 Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Onsite visit needed to replace 0000000056000000009000000001, DFM100000000 SOM PCA ASSY and sw udate. Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : Onsite visit needed to replace 0000000056000000009000000001, DFM100000000 SOM PCA ASSY and sw udate. ----------------------------------------------------------------------------- | 2023-10-10 10:22 | No
1944 | Yes | QA 000000 device failed the self-test | QA 00000000- device failed the self-test |  | 2023-10-10 10:47 | No
1945 | Yes | Image memory is full | Image memory is full | 000000000000000000000000001O##: 0000000000000000000000000010000000018900000000 Oct 100000000 2000000003 Harald Kaumann The system meets the specification for the performed service and is retuned to ue Resoluion: Log file read and analyzed. Hard drives pluged into the image compuer again. Recreate Image Store (hard drives reinitialized) performed on the IP PC. Test ru OK. Problem Description by Engineer: Cutomer Fuction/Role: RT Device ue: Vascuar angiography Eupected and ueupected Resuts: No image storage possible. uLLmessage: Image memory ful uer Impact: No image storage, only illuination possible Patient Impact: none SE8.2 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000183851 Harald Kaumann The system meets the specification for the performed service and is retuned to ue Follow-u requred: uLLmessage: Image memory ful Problem Description by Engineer: Cutomer Fuction/Role: RT Device ue: Vascuar angiography Eupected and ueupected Resuts: No image storage possible. uLLmessage: Image memory ful uer Impact: No image storage, only illuination possible Patient Impact: none SE8.2 Diagnostic performed by Engineer: 00conversation with Ms. Sommer. Onsite service requred. ----------------------------------------------------------------------------- | 2023-10-10 11:22 | No
1946 | Yes | System problem beim Hsngen Startup | System problem beim Hsngen Startup | 000000000000000000000000001O##: 0000000000000000000000000010000000018000000009 11-Oct-2000000003 Michele Balzano The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: HOD Device ue: u FD200000000 Eupected and ueupected Resuts: all ok uer Impact: no Patient Impact: no SE n/a Resoluion: Replaced Host PC and HD, installed new SOFTWARE, set all the parameters, performed fuctional test, everything according to the specifications. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.00000000000000000000000000000000800000000122391, u Haswell Monoplane Host PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000189630000000011-Oct-2000000003 Brother Ruegg The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: FSE Device ue: u Eupected and ueupected Resuts: start u uer Impact: mo Patient Impact: no SE na Resoluion: Hut PC ersetzt ud Software neuinstalliert. System abgeglistentionstchtig bergeben. Parts Consued : 1.0000000000000000000000000000000080000000000000000300000000u8.u.3000000005/uIQ+Centr1.000000003000000005 sw pck ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000183800000000Brother Ruegg The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer : Cutomer Fuction/Role: FSE Device ue: u Eupected and ueupected Resuts: start u uer Impact: na Patient Impact: no SE na Diagnostic performed by Engineer : uLLAnalysis for log file ud mit lokalem FSE. Spare parts suh as ud order. Download ud Vorbereiten der SW from neuuB gebrauht wird fmfr Haswell. ----------------------------------------------------------------------------- Internal Remarks : udate 1: Shipment is at uS facility as requsted by Sender and package mut be collected within 5 buiness days. Please inform the receiver to collect package quckly. Internal Remarks : udate 1: Shipment was delayed du to missing invoice. Now it is at Euport Scanin Lousville, KY, uited States. However, it is eupected to be delivered on October 12th,23 by EOD. We will continu to monitor and keep youin touh with the udates. Internal Remarks: Dear Market, uu 00000000800000000200000000612 - Nehalem Host PC Monoplane Rev_III NO_HDD Cutomer ETA: 11-OCT-2000000003 Comments: Please note that a direct shipment has been created toward you location. You part will be shipping from RDC Memphis with DN# 862133000000007. Internal Remarks : PFS#22667 / +uR+ ASAP DELIVERY ASAP-Req_680000000079000000008000000007326 / S000000000000000000000000001 Case: 00000000220000000000000000000000000000000000000000000000000010000000018000000009 / Keyboard requst not delivered Internal Remarks : 0000000080000000027000000001 Host OS Disk Revised / Order block / ##to Michele Balzano sent / It shoud revert to u. Internal Remarks : udate 3: uon tracking, shipment has Departed from Facility Lousville, KY, uited States. However, it is eupected to be delivered on October 12th,23 by EOD. We will continu to monitor and keep youin touh with the udates. ----------------------------------------------------------------------------- | 2023-10-10 12:02 | No
1947 | Yes | Screen flashes and does not allow the menu to be activated | Screen flashes and does not allow the menu to be auted | 000000000000000000000000001O##: 0000000000000000000000000010000000025500000000 November 13, 2000000003 Medina, Ismael The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000056153300000000,RP-Bezel, Front, w/NAV, Wht, CTD, V600000000Resoluion: Front bezel replacement Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Navigation ring does not 000000000000000000000000001k. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000018000000002 October 200000000 2000000003 Medina, Ismael The system does not meet the specification for the performed service and is not retuned to ue. Resoluion: Replacement of touh bezel Follow u Requred: Continu Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Interference between navigation ring and touh panel. ----------------------------------------------------------------------------- | 2023-10-10 12:27 | No
1948 | Yes | The device is reported not booting. Ref: case 000 | The device is reported not booting. Ref: case 0000 | 000000000000000000000000001O##: 0000000000000000000000000010000000019900000000 23-Oct-2000000003 Ndizihiwe Balinda, Gilbert The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000000000000900000000,PDuFanTray and DCPS Resoluion : Part is replaced. System is tested and left in good and 000000000000000000000000001king condition. Problem Description by Engineer : Cutomer Fuction/Role: sonographer Device ue: not in ue Eupected and ueupected Resuts: system to boot uer Impact: no impact Patient Impact: no impact SE2.2 Tools & Equpment : MuTIMETER,200000000000000003-0000000000000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-11 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-10 12:27 | No
1949 | Yes | uequipment failure - | uequipment failure - | 000000000000000000000000001O##: 0000000000000000000000000010000000018000000006 Buak Tezel The system does not meet the ful specification for the performed service and is retuned to ue with limitation as accepted by the cutomer, no follow-u requsted. Resoluion: Duing the phone call with the cutomer, it was reported to u that the device coud not pass the test. In order to make a detailed failue detection, the cutomer will send the device to ou bench repair center. Problem Description by Engineer: Cutomer Fuction/Role: Biomed. Device ue: Not ue. Eupected and ueupected Resuts: uEqupment Failue. uer Impact: No. Patient Impact: No. SE 2.00000000.32 ----------------------------------------------------------------------------- | 2023-10-10 12:33 | No
1950 | Yes | Therapy test failure (Replaced DFM0 ASSY uuuitance) | Therapy test failure (Replaced DFM0 ASSY uuuance) | 000000000000000000000000001O##: 00000000000000000000000000100000000180000000000000000nanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : 0000000056000000009000000001 DFM100000000 uuuitance ASSY - Ordered. Follow u Requred : 0000000056000000009000000001 DFM100000000 uuuitance ASSY - Ordered. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Therapy test failue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001898500000000Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-10-10 12:37 | No
1951 | Yes | No geometry in the system | No geometry in the system | 000000000000000000000000001O##: 00000000000000000000000000100000000182593 Roman Kuka The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Close case as no reply received from partner > 100000000ys after action plan. Problem Description by Engineer : Cutomer Fuction/Role: uknown - case open as technical suport for distribuor IP Device ue: uknown - case open as technical suport for distribuor IP Eupected and ueupected Resuts: uknown - case open as technical suport for distribuor IP uer Impact: uknown - case open as technical suport for distribuor IP Patient Impact: uknown - case open as technical suport for distribuor IP SE uknown - case open as technical suport for distribuor IP ----------------------------------------------------------------------------- | 2023-10-10 12:42 | No
1952 | Yes | Therapy delivery test failed (Replaced DFM0 ASSY uuuitance) | Therapy delivery test failed (Replaced DFM0 ASSY uuuance) | 000000000000000000000000001O##: 00000000000000000000000000100000000189500000000Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : To replace 0000000056000000009000000001 DFM100000000 uuuitance ASSY Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001 DFM100000000 uuuitance ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000190000000071 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-10-10 13:28 | No
1953 | Yes | Therapy delivery test failed (Replaced DFM0 ASSY uuuitance) | Therapy delivery test failed (Replaced DFM0 ASSY uuuance) | 000000000000000000000000001O##: 000000000000000000000000001000000001895600000000Madhan, kuar The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : To replace 0000000056000000009000000001 DFM100000000 uuuitance ASSY Problem Description by Engineer : Cutomer Fuction/Role: BME Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SEN/A Diagnostic performed by Engineer : To replace 0000000056000000009000000001 DFM100000000 uuuitance ASSY ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000190000000059 Parts Consued : 1.0000000000000000000000000000000056000000009000000001,DFM100000000 ASSY uuuitance ----------------------------------------------------------------------------- | 2023-10-10 13:32 | No
1954 | Yes | System failed - patient is on the system | System failed - patient is on the system | 000000000000000000000000001O##: 00000000000000000000000000100000000190000000098 Oct 11, 2000000003 Simon, Aleuander The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal 100000000Oct-2000000003 00000000:00000000100000000 RMT - uGEN: uLL(100000000C) RoCo is defective - Frontal 100000000Oct-2000000003 00000000:00000000500000000 Ms. Petra is to be contacted. +0000000028500000000-7156 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.30000000015 Tools & Equpment : ELECTRICAL SAFETY TESTER, 2000000003-12-1000000000000000000000000:00000000 MuTIMETER, 2000000003-12-15 00000000:00000000:00000000 Parts Consued: 1.000000000000000000000000000000002100000000125700000000oco Velara Replacement Kit 1.00000000000000000000000000000000210000000015282,Power on circut Resoluion: Defective parts removed; Roco Velara Replacement Kit installed; Power on circut part installed; Final system tests performed; ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001899700000000Robert Kern The system does not meet the specification for the performed service and is not retuned to ue Internal Comments: uLL# and/or description of uLL Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. Relay (tap) in MPDucontrol board is defective Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. Cue is defective Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. The 200000000cÂ power suply (NT200000000located in the u-ray generator is malfuctioning. Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. Problem with the hospital mains power Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. This coud be caued du to a timing issu between host (PC) and generator. Please ignore this alert, if it only occus sporadically. Treat it as a 'one time only'. Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. Problem with EN100000000 PCB Root Caue for RMT - uGEN: uLL(100000000C) RoCo is defective - Frontal: 1. RoCo is defective. Repair Action: Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. If the problem is permanent: Switch off the system (bu make sue F1 to F5 MCBs in MPDuare on) and connect D-connector measuing board at NPu: 300000000 in M-cabinet. Measue the reproduey between pins 00000000nd 11. This shoud be normally closed, If it is open then replace MPDucontrol board. If the problem is intermittent: Switch off the system. Disconnect the cable from NPu: 300000000 (frontal) in M-cabinet (frontal) and connect it to NPu:300000000 (lateral). Observe the system. If the problem is resolved, MPDucontrol board is defective. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Check the LED statu of Cue. If no LEDs are lit, please check 200000000power suply first. (see also 'Caue 5') If Power suply is ok bu LEDs are off, replace Cue. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Measue the 200000000c power suply (NT200000000located on the EN uit in the u-ray generator. Replace when necessary. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Check the hospital Mains inpu power to the Velara Generator. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Perform a reboot of the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other soluion here below. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Check whether the green power LED is on at EN100000000 pcb. If the LED is not lit, replace this PCB Service actions for RMT - uGEN: uLL(100000000C) RoCo is defective - Frontal: 1.1. Please replace the RoCo (EY) of the frontal generator. Follow-u requred: uLL# and/or description of uLL Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. Relay (tap) in MPDucontrol board is defective Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. Cue is defective Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. The 200000000cÂ power suply (NT200000000located in the u-ray generator is malfuctioning. Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. Problem with the hospital mains power Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. This coud be caued du to a timing issu between host (PC) and generator. Please ignore this alert, if it only occus sporadically. Treat it as a 'one time only'. Root Caue for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1. Problem with EN100000000 PCB Root Caue for RMT - uGEN: uLL(100000000C) RoCo is defective - Frontal: 1. RoCo is defective. Repair Action: Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. If the problem is permanent: Switch off the system (bu make sue F1 to F5 MCBs in MPDuare on) and connect D-connector measuing board at NPu: 300000000 in M-cabinet. Measue the reproduey between pins 00000000nd 11. This shoud be normally closed, If it is open then replace MPDucontrol board. If the problem is intermittent: Switch off the system. Disconnect the cable from NPu: 300000000 (frontal) in M-cabinet (frontal) and connect it to NPu:300000000 (lateral). Observe the system. If the problem is resolved, MPDucontrol board is defective. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Check the LED statu of Cue. If no LEDs are lit, please check 200000000power suply first. (see also 'Caue 5') If Power suply is ok bu LEDs are off, replace Cue. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Measue the 200000000c power suply (NT200000000located on the EN uit in the u-ray generator. Replace when necessary. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Check the hospital Mains inpu power to the Velara Generator. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Perform a reboot of the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other soluion here below. Service actions for RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal: 1.1. Check whether the green power LED is on at EN100000000 pcb. If the LED is not lit, replace this PCB Service actions for RMT - uGEN: uLL(100000000C) RoCo is defective - Frontal: 1.1. Please replace the RoCo (EY) of the frontal generator. What are the neut trouleshooting steps: Schedue TE service to replace power-u circut and ROCO. 3 hous of 000000000000000000000000001k Technician name: Robert Kern Problem Description by Engineer: Problem description by engineer: RMT - uGEN: Host(PC) is uable to commuicate via CAN with Velara Generator - Frontal 100000000Oct-2000000003 00000000:00000000100000000 RMT - uGEN: uLL(100000000C) RoCo is defective - Frontal 100000000Oct-2000000003 00000000:00000000500000000 Ms. Petra is to be contacted. +0000000028500000000-7156 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 8.1.30000000015 Diagnostic performed by Engineer: Yesterday there was a service to replace the MPDucontrol uit and cue. Log shows connection problem duing start-u and then ROCO defect uLL ----------------------------------------------------------------------------- Cutomer comment: ## Appointment info to Nuse Petra, 10000000011 from c. 9:00000000 am## ----------------------------------------------------------------------------- | 2023-10-10 13:48 | No
1955 | Yes | 0 - u0- Power interface problem | 00- u000- Power interface problem | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 13.1000000002000000003 00000000:59:500000000uC (86000000001000000005)*Department information: [Department: Equpment Department Contact person: Teacher uuContact nuber: 156119920000000000000000* $ASP engineer: [qiangyi wujianku] $Q1: (NO) $Q2: (NO) $Q7: (Confirmed reported problem & Power interface problem) $Q8: (Restart the machine As soon as the power cord is touhed, the machine will auomatically shu down, At the reinforcement of the power cord interface, the machine is 000000000000000000000000001king normally It is juged that the power cord interface is loose) $Q9: (Reinforce the power cord connector) $Q100000000 (Test and inspection data refer to :PA checklist.Device retuned to ful fuctionally.) $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration date (00000000/11/20000000000000000 No open Field Change Order foud for this service. SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 13.1000000002000000003 00000000:59:53 uC (86000000001000000005) $ASP engineer: [qiangyi wujianku] $Q1: (NO) $Q2: (NO) $Q7: (Confirmed reported problem & Power interface problem) $Q8: (Restart the machine As soon as the power cord is touhed, the machine will auomatically shu down, At the reinforcement of the power cord interface, the machine is 000000000000000000000000001king normally It is juged that the power cord interface is loose) $Q9: (Reinforce the power cord connector) $Q100000000 (Test and inspection data refer to :PA checklist.Device retuned to ful fuctionally.) $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000300000000) Neut Calibration Date (00000000/11/20000000000000000 No open Field Change Order foud for this service. Problem Reported by Cutomer: ECR：OCAP：00000000R/HeartStart uL+ Defibrillator/Monitor SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000300000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: DiFeng Liang Last udated by: DiFeng Liang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Cutomer’s euplanation of the problem ### Power cord interface failue The cutomer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### na Repair Action: T2 Recommended Field Diagnostic Steps### It is recommended to check and replace the interface T2 Recommended spare part uuwith description### na Cutomer demand### na Test/Inspection: na Test Equpment: na OneEMS internal remarks: FSE requres dispatch of 000000000000000000000000001kers, ASP strong doctors SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-10 13:52 | No
1956 | Yes | Geometry Partly available- SPLIT | Geometry Partly available- SPLIT | 000000000000000000000000001O##: 00000000000000000000000000100000000183000000001 Ahmed Kotb The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Geometry issu - issu has been fiued Problem Description by Engineer : Cutomer Fuction/Role: NA Device ue: NA Eupected and ueupected Resuts: NA uer Impact: NA Patient Impact: NA SE NA ----------------------------------------------------------------------------- | 2023-10-10 13:57 | No
1957 | Yes | FSE CDP: End-of-life date eupired - periodic replacement - Battery - Lithium-Ion | FSE CDP: End-of-life date eupired - periodic replacement - Battery - Lithium-Ion |  | 2023-10-10 15:13 | No
1958 | Yes | [Q] Total failure during the euamination | [Q] Total failure during the euamination | 000000000000000000000000001O##: 00000000000000000000000000100000000190000000077 Oct 22, 2000000003 Hoffmann, Thomas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called care management and contact person u Altdorf 000000009 679000000009 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 2.2.6.00000000Resoluion: 3 FI switches were triggered at the same time. No uLLcan be recognized on the system. Performed fuction test ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000190000000000000000Robert Kern The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments: uLL# and/or description of uLL . Repair Action: . Follow-u requred: uLL# and/or description of uLL . Repair Action: . What are the neut trouleshooting steps: TE service for reading ou log on site, check of the remote connection, perform fuction test. 2 hous of 000000000000000000000000001k Technician name: Robert Kern Problem Description by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called care management and contact person u Altdorf 000000009 679000000009 Information to suport the complaint handling process Cutomer Fuction/Role: Not provided Device ue: Not provided Eupected and ueupected Resuts: Not provided uer Impact: Not provided Patient Impact: Not provided SE 2.2.6.00000000Diagnostic performed by Engineer: Fi in M cabinet triggered yesterday. Was activated again by bulding services. According to the cutomer, no defect or fuction restriction noticed. System not accessible via remote. ----------------------------------------------------------------------------- Internal Remarks: 'TV' illness Internal Remarks: Bu system cannot be reached in PRS. RSN team informed via ##and application on bou Cutomer comment: ### u Altdorf appointment moved to 18.100000000 from 9 AM, if something happens sporadically today, please call back. Internal Remarks: u Zellinger on bou whether he sees another possibility today Cutomer comment: ## u Altdorf 10000000016 after 9:00000000 am Internal Remarks: u Zellinger does not see an option today, bu he will get in contact if it shoud still be something ----------------------------------------------------------------------------- | 2023-10-10 15:18 | No
1959 | Yes | uray system is switched off melding | uray system is switched off melding | 000000000000000000000000001O##: 00000000000000000000000000100000000213593 Oct 19, 2000000003 CAREL BRANTEN, The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: system down Eupected and ueupected Resuts: u-ray system switched off message, system does not start u. uer Impact: n/a Patient Impact: n/a SE R223 Diagnostic performed by Engineer: Replaced the Sute PC and loaded sw. Parts Consued: 1.00000000000000000000000000000000800000000660000000083,Sute PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000190000000098 Oct 13, 2000000003 Ronnie Smits The system meets the specifications for the performed service and is retuned to ue Follow u Requred: replace PC, reinstall at least 00000000ous of 000000000000000000000000001k! Swap license Diagnostic performed by Engineer: hdd diskbay defective, PC needs to be replaced 00000000800000000660000000083 Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: system down Eupected and ueupected Resuts: u-ray system switched off message, system does not start u. uer Impact: n/a Patient Impact: n/a SE R223 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000190000000036 ROB VAN GAMEREN The system does not meet the specifications for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Medical technician Device ue: system down Eupected and ueupected Resuts: u-ray system switched off message, system does not start u. uer Impact: n/a Patient Impact: n/a SE R223 Diagnostic performed by Engineer: No remote commuication possible with the system. Diagnosis on site this afternoon. Follow u Requred: u-ray system switched off message, no remote connection so no log file. ----------------------------------------------------------------------------- | 2023-10-10 15:27 | No
1960 | Yes | system startup boot fail | system startup boot fail | 000000000000000000000000001O##: 0000000000000000000000000010000000018800000000 11-Oct-2000000003 JAEHu jang The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.0000000000000000000000000000000080000000050000000000000000,Host PC Monoplane Revised_II no HDD Tools & Equpment : MuTIMETER,200000000000000001-27 00000000:00000000:00000000 Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: ouside of ue Eupected and ueupected Resuts: system startu problem uer Impact: none Patient Impact: none SE 8.2.17 Resoluion : FD200000000심혈관조영실 Checked system uLLlog Checked Host PC condition Replaced Host PC monoplane with new one Checked system fuction ok Device repaired and ready for ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000189551 100000000Oct-2000000003 Heo, GiTae The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: ou side of ue Eupected and ueupected Resuts: system bootu problem uer Impact: none Patient Impact: none SEr8 Resoluion : FD200000000 Confirmed reported problem system boot u problem Checked system uLLlog Checked system movement Performed VIP test Replaced Host pc with new one Checked system fuction ok Device repaired and ready for ue ----------------------------------------------------------------------------- | 2023-10-10 15:28 | No
1961 | Yes | 0 / u0 / Therapy delivery test failed. | 00/ u000/ Therapy delivery test failed. | 000000000000000000000000001O##: 00000000000000000000000000100000000185100000000Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : EOL Produt, informed cutomer. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Therapy delivery test failed ----------------------------------------------------------------------------- | 2023-10-10 15:48 | No
1962 | Yes | Malfunctioning battery | Malfunuoning battery | 000000000000000000000000001O##: 00000000000000000000000000100000000183896 FALASCA, IVO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: Spare part can no longer be ordered in SPS try with other channels. Make NEMO withou collection at the same time. Tu San Marco Hospital G. Pignatelli Street 7000000003 Grottaglie (TA) Department Point of First Intervention Giueppe Ciracì Cell.3390000000085600000000Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Eupected and ueupected Resuts: The fuctional test shows that the battery is flat and does not recharge completely. uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Head nuse u Ciriacì confirms the need to replace the battery that does not recharge completely. Accepts the shipment of the new battery. ----------------------------------------------------------------------------- | 2023-10-10 16:02 | No
1963 | Yes | Arch c is not performing movements | Arch c is not performing movements | BuQ&A Feedback: No Reason: Corrective Maintenance Repair Action: Schedued with the cutomer: eng clin Schedued date/time: 11/200000000t 1pm SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Case: 000000002200000000169 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Maria Soua Last udated by: Maria Soua 1. Was the equpment in clinical ue at the time the problem was detected? Yes 2. Was the patient or uer injued? No 3. If the equpment has an alarm or some type of alert, did this alarm/alert 000000000000000000000000001k when the problem was detected? N/A 00000000Was the equpment defective or failed at the time of installation? AT Problem Reported by Cutomer: Equpment problem: The c-arm is not making movements Priority: STOPPED uit: HOSPITAL ESPERANCA SA Address: Ru Antonio Gomes de Freitas 265 Recife-PE 50000000070000000000000000 Name: Pauo/Luas/Allan Sector: Clinical engineering 00¹ : 81 31317863 81 98118300000000 00²: 81 3131000000007 ## engenhariaclinica00ospitalesperanca.com.br SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Problem description by engineer : no movements of the arc and table Trouleshooting Action: What steps are taken duing the system diagnosis : Verified that the Geometry TSo POST failed. BIST test also failed. I reinstalled the TSO software and it started 000000000000000000000000001king again, bu stopped again when tuning it off and on. Verified that when the SIB u800000000onnector is disconnected and reconnected, it 000000000000000000000000001ks again. The TSO will be asked to euchange it bu the equpment remains fuctional. Cutomers Icaro and Ricardo were instruted to disconnect and reconnect the u800000000onnector after tuning on the machine, util the TSO is changed uLL# and/or description of uLL: Geometry movements partially available OneEMS internal remarks: Problem description by engineer : no arc and table movements What steps are taken duing the system diagnosis : Verified that the Geometry TSo POST failed. BIST test also failed. I reinstalled the TSO software and it started 000000000000000000000000001king again, bu stopped again when tuning it off and on. Verified that when the SIB u800000000onnector is disconnected and reconnected, it 000000000000000000000000001ks again. The TSO will be asked to euchange it bu the equpment remains fuctional. Cutomers Icaro and Ricardo were instruted to disconnect and reconnect the u800000000onnector after tuning on the machine, util the TSO is changed uLL# and/or description of uLL: Geometry movements partly available Repair Action "Internal" : Send PART+FSE+2 Hous of machine downtime. SYSTEM u, treat case as P3 Recommendations to FSE ► Change TSO, install software if necessary; ► Test the TSO after tuning the machine off and on ► Test all TSO fuctions Parts requred : 1 Part : 1, uu: 00000000800000000151715, Description : GEO MODuE WP Kit, Quntity : 1, Priority : 1 Eutra Internal comment field : RSE MARcel Ogassawara System u 00000000/1000000002000000003 00000000:00000000SAP Transfer: Case Transferred to SAP SAP Transfer: Service Order is Locked by another uer!: SAP Transfer: Service Order is Locked by another uer!: SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Repair Action: Call closed du to lack of confirmation, cutomer notified via ##to requst opening when the equpment is eupected to be available Euternal Remarks: 27.11.2000000003 19:51:00000000POLARIS (POLARIS) ========== Antonio Juior - 0000000000000000570000000022300000000200000000000000000000000000000000==================== == The c arc is not moving. The GEO MODuE WP Kit was replaced, as well as checking the download board service and operating tests. Equpment tested and released. Internal Remarks: 27.11.2000000003 19:51:00000000POLARIS (POLARIS) ========== Antonio Juior - 0000000000000000570000000022300000000200000000000000000000000000000000==================== == Alignment: yes Action plan: OK ----------------------------------------------------------------------------- | 2023-10-10 16:07 | No
1964 | Yes | geometry is not functioning | geometry is not funuoning | 000000000000000000000000001O##: 00000000000000000000000000100000000191335 00000000-Nov-2000000003 Ramachandran, Sekar The system meets the specification for the performed service and is retuned to ue Resoluion : on inspection foud the problem with the ip conflict. spoken with hospital it person and solved the issu. system handed over for clinical uage. Problem Description by Engineer : Problem description by engineer : geometry is not fuctioning When was the first occurence? Within... : 100000000Oct-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: geometry is not fuctioning uer Impact: No uer impact Patient Impact: Not provided uu00000000----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001897500000000Vijayakuar, K.S The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : geometry is not fuctioning When was the first occurence? Within... : 100000000Oct-2000000003 16:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: geometry is not fuctioning uer Impact: No uer impact Patient Impact: Not provided uu00000000Diagnostic performed by Engineer : System rel: 2.2.3 Geometry restart occured twice today and I coud see the following uLL. 1) Please check with uer for changes in IP from their IT side. This is something related to net000000000000000000000000001k conflict with hospital net000000000000000000000000001k. 2) Try isolating ou system from hospital net000000000000000000000000001k and observe the system. 3) Check the firewall connectivity and fuctionality – If there is deviation – rectify it. 00000000If all the above is taken care – and still problem recus – we need to think of reloading the sw in sute PC. Follow u Requred : uLL# and/or description of uLL: geometry is not fuctioning Repair Action : System rel: 2.2.3 Geometry restart occured twice today and I coud see the following uLL. 1) Please check with uer for changes in IP from their IT side. This is something related to net000000000000000000000000001k conflict with hospital net000000000000000000000000001k. 2) Try isolating ou system from hospital net000000000000000000000000001k and observe the system. 3) Check the firewall connectivity and fuctionality – If there is deviation – rectify it. 00000000If all the above is taken care – and still problem recus – we need to think of reloading the sw in sute PC. Technician Name : Vijayakuar, K.S Resoluion : System rel: 2.2.3 Geometry restart occured twice today and I coud see the following uLL. 1) Please check with uer for changes in IP from their IT side. This is something related to net000000000000000000000000001k conflict with hospital net000000000000000000000000001k. 2) Try isolating ou system from hospital net000000000000000000000000001k and observe the system. 3) Check the firewall connectivity and fuctionality – If there is deviation – rectify it. 00000000If all the above is taken care – and still problem recus – we need to think of reloading the sw in sute PC. Internal Comments : uLL# and/or description of uLL: geometry is not fuctioning Repair Action : System rel: 2.2.3 Geometry restart occured twice today and I coud see the following uLL. 1) Please check with uer for changes in IP from their IT side. This is something related to net000000000000000000000000001k conflict with hospital net000000000000000000000000001k. 2) Try isolating ou system from hospital net000000000000000000000000001k and observe the system. 3) Check the firewall connectivity and fuctionality – If there is deviation – rectify it. 00000000If all the above is taken care – and still problem recus – we need to think of reloading the sw in sute PC. ----------------------------------------------------------------------------- | 2023-10-10 16:13 | No
1965 | Yes | It does not start. | It does not start. | 000000000000000000000000001O##: 000000000000000000000000001000000001800000000 October 100000000 2000000003 Luas Santolino Simón The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Maintenance Device ue: Start-u failue. Eupected and ueupected Resuts: Several restarts to be able to start the system. uer Impact: Na Patient Impact: Na SE 8.2 Resoluion: Software image PC is reloaded and correct operation of the system is checked. ----------------------------------------------------------------------------- | 2023-10-10 16:37 | No
1966 | Yes | unit has a constant speaker malfunction uLLmessage, no sound is coming from the device at all -Pending Bench)000 | unit has a constant speaker malfunuon uLLmessage, no sound is coming from the device at all -Pending Bench)0/0/0 | 000000000000000000000000001O##: 00000000000000000000000000100000000191800000000 Jent, Chris The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed states the uit has a constant speaker malfuction uLLmessage, no soud is coming from the device at all, and it was not in ue when the issu was discovered; needs to go in for bench repair. Biomed confirms they foud the device while doing PM's. Device ue: na Eupected and ueupected Resuts: na uer Impact: na Patient Impact: na SE na Follow u Requred : Biomed states the uit has a constant speaker malfuction uLLmessage, no soud is coming from the device at all, and it was not in ue when the issu was discovered; needs to go in for bench repair. Biomed confirms they foud the device while doing PM's. Verified shipping address: uiversity of uah Hospital ATTN: Brady/ Biomed N Medical Dr 500000000Salt Lake City, u 8000000002-0000000000000000 u Resoluion : Verified cutomers contact and shipping information in the system is correct and ##d the Tele bench form. Created a follow u activity for the bench so the bench tech has a 000000000000000000000000001agnostic performed by Engineer : Biomed states the uit has a constant speaker malfuction uLLmessage, no soud is coming from the device at all, and it was not in ue when the issu was discovered; needs to go in for bench repair. Biomed confirms they foud the device while doing PM's. Verified shipping address: uiversity of uah Hospital ATTN: Brady/ Biomed N Medical Dr 500000000Salt Lake City, u 8000000002-0000000000000000 u ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: brady IP & Issu: 82525779 speaker issu Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: na Action Taken: Created case Neut Step Action: N/A ----------------------------------------------------------------------------- | 2023-10-10 18:02 | No
1967 | Yes | Check / replace MCS monitor | Check / replace umonitor | 000000000000000000000000001O##: 0000000000000000000000000010000000019000000000000000011-Oct-2000000003 Sinkewicz Gary The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Y Eupected and ueupected Resuts: ER Monitor not 000000000000000000000000001king uer Impact: None Patient Impact: None SE R7.2.16 Parts Consued : 1.00000000000000000000000099193200000000000000001,21" Color LCD Monitor ER (CML212-PHE) Resoluion : replaced Defective ER Monitor ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000185000000006 100000000Oct-2000000003 Sinkewicz Gary The system meets the specification for the performed service and is retuned to ue Follow u Requred : Replace monitor Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Y Eupected and ueupected Resuts: ER Monitor not 000000000000000000000000001king uer Impact: None Patient Impact: None SE R7.2.16 Resoluion : Monitor is not 000000000000000000000000001king. 2200000000ac ok Video Ok ----------------------------------------------------------------------------- Internal Remarks : Remote bypass requsted, created onsite 000000000000000000000000001for dispatch Internal Remarks : Case: 000000002200000000700000000 FSE: Sinkewicz Gary Euam room monitor fails to display; power pass, video feed pass. Instruted FSE to place part order for 99193200000000000000001 - 21" Color LCD Monitor ER (CML212-PHE). ----------------------------------------------------------------------------- | 2023-10-10 18:07 | No
1968 | Yes | The alarms do not sound in the central station. No Product Malfunction-Configuration/Cuomer. Config changes. | The alarms do not sound in the central station. No ProduuMalfunuon-Configuration/Cuomer. Config changes. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000900000000October 26, 2000000003 Santiago Cases Lopez The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Resoluion: The condition of the central station is checked. The equpment 000000000000000000000000001ks correctly. We spoke with the service staff and we showed them where the volue of the central station can be raised or lowered. Confirmed that the alarms soud correctly. The central station is on a table whose legs have wheels, reminded the electromedicine staff that the equpment mut be in a fiued place, that it cannot be moved, in suh a way that it is not possible to ru the risk of disconnecting any cable. After corrective maintenance, the equpment is operational and meets the manuactuer's specifications. The necessary checks are performed to certify that the equpment is back to the operating conditions prior to the failue. Problem Description by Engineer: Cutomer Fuction/Role: Electromedicine. Device ue: Patient monitoring. Eupected and ueupected Resuts: The alarms did not soud acoutically. uer Impact: I coud not hear the alarms. Patient Impact: None. SE uknown. ----------------------------------------------------------------------------- | 2023-10-10 18:13 | No
1969 | Yes | Touch screen is not responsive | Touch screen is not responsive | 000000000000000000000000001O##: 00000000000000000000000000100000000187216 Miller Jason The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : Cutomer states device is not responding to touh properly. device will calibrate bu then still touh is off. Cutomer is requsting part nuber and price for replacement touhscreen. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Touhscreen not responding to touh uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 00000000561539000000001 RP-Touh Screen,V600000000- 990000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Ron Grudmam IP & Issu: 56627997 Touh screen is not responsive Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-10 18:22 | No
1970 | Yes | Too low tidal volume alarm | Too low tidal volume alarm |  | 2023-10-10 18:27 | No
1971 | Yes | [Q]Wireless Footswitch Bou | [Q]Wireless Footswitch Bou | 000000000000000000000000001O##: 00000000000000000000000000100000000191753 200000000ct-2000000003 Buler Thomas The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Clinical Eupected and ueupected Resuts: Installed missing wireless footswitch uer Impact: None Patient Impact: None uuParts Consued : 1.00000000000000000000000000000000800000000257862,WFS Base Station Resoluion : Installed backorder wireless footswitch on system and verified correct fuctioning ----------------------------------------------------------------------------- | 2023-10-10 18:28 | No
1972 | Yes | Machine is not switching On | Machine is not switching On | 000000000000000000000000001O##: 00000000000000000000000000100000000192332 100000000Oct-2000000003 Taneja, Kamal The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching On When was the first occurence? Within... : 100000000Oct-2000000003 18:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching On uer Impact: No uer impact Patient Impact: No patient impact SE Not provided Tools & Equpment : MuTIMETER,200000000000000007-20000000000000000:00000000:00000000 Resoluion : Observed cooling uit not switching on , reseated connections, now 000000000000000000000000001king ok ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000190000000017 Vijayakuar, K.S The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : Machine is not switching On When was the first occurence? Within... : 100000000Oct-2000000003 18:00000000:00000000 How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No patient involved Eupected and ueupected Resuts: Machine is not switching On uer Impact: No uer impact Patient Impact: No patient impact SE Not provided Diagnostic performed by Engineer : FSE visit requre dfor futher checks. Follow u Requred : uLL# and/or description of uLL: Machine is not switching On Repair Action : FSE visit requre dfor futher checks. Technician Name : Vijayakuar, K.S Resoluion : FSE visit requre dfor futher checks. Internal Comments : uLL# and/or description of uLL: Machine is not switching On Repair Action : FSE visit requre dfor futher checks. ----------------------------------------------------------------------------- | 2023-10-10 18:47 | No
1973 | Yes | Accept button is not working and will not allow him to access the service menu. | Accept button is not working and will not allow him to access the service menu. | 000000000000000000000000001O##: 000000000000000000000000001000000001868200000000iller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: accept buton not 000000000000000000000000001king uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states the accept buton is not 000000000000000000000000001king and will not allow him to access the service menu Cutomer is requsting replacement part nubers for switch PCBA, switch cable and replacement buton Resoluion : Provided cutomer with part nuber and price: 00000000561512600000000 - RP-PCBA, Switch, V600000000- 171.00000000 00000000561511600000000- RP-Cable, Switch PCBA, V600000000- 125.00000000 0000000056151100000000 - RP-Buton, Nav-Ring, Pkg/5, V600000000- 0000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : Caller: Ralph 500000000 9000000007000000001 IP & Issu: 56628318 -This Part ID requst was for replacement, repair troule shooting activities Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Roued to RTAC Action Taken: Remote Technical Neut Step Action: Verbal PO Nuber: ----------------------------------------------------------------------------- | 2023-10-10 18:52 | No
1974 | Yes | The left paddle does not work. | The left paddle does not work. | 000000000000000000000000001O##: 000000000000000000000000001000000001900000000200000000hazwan Hueidi Sordo The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Biomed Engineer Device ue: Clinical uage Eupected and ueupected Resuts: N/A uer Impact: No uer impact Patient Impact: No patient impact SE N/A Resoluion: uDFM100000000 Defibrillator/Monitor CN3260000000000000000 The left paddle does not 000000000000000000000000001k. Ask: 1 u 0000000056300000000991 - Paddle Repl. Kit Water Resistant Send material to: Electromedicine (FLOOR -2 Right) # +3000000003368100000000 # Rau Nuez-Jun Lus Blanco # It belongs to emergency room (emergency room, basement -3 center) +30000000033680000000000000000eut. 57800000000 Mercedes # 23 8139 Follow u Requred: uDFM100000000 Defibrillator/Monitor CN3260000000000000000 The left paddle does not 000000000000000000000000001k. Ask: 1 u 0000000056300000000991 - Paddle Repl. Kit Water Resistant Send material to: Electromedicine (FLOOR -2 Right) # +3000000003368100000000 # Rau Nuez-Jun Lus Blanco # It belongs to emergency room (emergency room, basement -3 center) +30000000033680000000000000000eut. 57800000000 Mercedes # 23 8139 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000019200000000 Parts Consued: 1.0000000000000000000000000000000056300000000991, Paddle Repl. Kit Water Resistant ----------------------------------------------------------------------------- | 2023-10-10 18:57 | No
1975 | Yes | Low leak co0evision risk | Low leak co0revision risk | 000000000000000000000000001O##: 00000000000000000000000000100000000186931 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Low leak co2 revision risk uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states while testing device uit was giving Co2 rebreathing uLLmessage. Troule shot with cutomer. informed cutomer of whisper swivel. Advised cutomer if Whisper swviel was in place manuactuers recommendation is to replace FSA and DA board. Cutomer is requsting part nuber and price for replacement DA board and FSA. Resoluion : Provided cutomer with part nuber and price: 00000000561512561 - RP-PCBA, Data Acqusition, V600000000- 1286.00000000 989800000000617100000000- RP-Flow Sensor Assy, Air & O2 - 1217.00000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. ----------------------------------------------------------------------------- Internal Remarks : vInteraction Type: New Case Caller: Richard 50000000070000000065800000000 IP & Issu: Low leak co2 revision risk Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-10 19:02 | No
1976 | Yes | Machine and prouimal pressure sensors failed. (EC 00 | Machine and prouimal pressure sensors failed. (EC 00 | 000000000000000000000000001O##: 0000000000000000000000000010000000019200000000 Miller Jason The system meets the specification for the performed service and is retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: uLLcode 1000000007 uer Impact: system restricted Patient Impact: none reported SE uknown Diagnostic performed by Engineer : Cutomer states device is giving uLLcode 1000000007 after replacing the GDS. went over how install went. foud that ribbon cable from DA to MC wasn't attached. Cutomer reseated cable. tested device. device is now powering on and operating. cutomer states he will finish testing and call back if futher assistance. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Roger Streidl IP & Issu: 56632100000000- uer is getting uLLmessage stating device is not detecting the sensors. Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Reguar Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create case, R000000000000000000000000001and roued to RSE. Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-10 19:22 | No
1977 | Yes | Check Vent: Prouimal Pressure Sensor Range Error | Check Vent: Prouimal Pressure Sensor Range Error |  | 2023-10-10 19:28 | No
1978 | Yes | Device gives power equipment malfunction | Device gives power equipment malfunuon | 000000000000000000000000001O##: 0000000000000000000000000010000000025300000000 Bathabile ZuuThe system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : Performance assuance tests - some tests passed. Device is operational only when celluar modue is dis-connected. ESD is 1.000000002MOHms = Pass Problem Description by Engineer : Cutomer Fuction/Role: Paramedic Device ue: Defibrillator Eupected and ueupected Resuts: Device gives power equpment malfuction when the celluar modue is connected uer Impact: uknown Patient Impact: uknown SE 1.00000000.57 Diagnostic performed by Engineer : uon receiving the device, it was giving a "Power Equpment Malfuction" and the operational check failed as soon as the celluar modue was disconnected, the operational check passes and the uLLclears bu when the celluar modue is connected, the operational check fails again. Device uLLlogs have been sent to the produt suport specialist for futher analysis. When celluar modue is disconnected, all other tests passed. Tools & Equpment : PATIENT SIMuATOR,200000000000000000000000000000000 00000000:00000000:00000000 ELECTRICAL SAFETY TESTER,200000000000000003-15 00000000:00000000:00000000 DEFIBRILLATOR ANALYZER,200000000000000006-28 00000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-10 19:32 | No
1979 | Yes | e stop active | e stop au | 000000000000000000000000001O##: 00000000000000000000000000100000000192626 Sulivan Scott The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : spoke to Dan e stop actived duing case uable to clear When was the first occurence? Within... : 100000000Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: e stop active Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion : cutomer is 000000000000000000000000001king with in houe suport for resoluion Internal Comments : uLL# and/or description of uLL: e stop active Repair Action : attached daily log to case . cutomer jut wanted to check for all points of e stop Diagnostic performed by Engineer : speaking to Dan ,e stop actived duing ue . cutomer uable to clear. walked cutomer thruchecking of the e stops on system . cutomer is moving patient from system and will 000000000000000000000000001k with the in houe service to clear problem ----------------------------------------------------------------------------- | 2023-10-10 19:43 | No
1980 | Yes | SPO0uainability Issue_Drop in blood ouygen level at certain/Unknown | SPO0Suainability Issue_Drop in blood ouygen level at certain/Unknown |  | 2023-10-10 19:58 | No
1981 | Yes | all of the monitors went black, shut off during study. Not powering back up | all of the monitors went black, shut off during study. Not powering back up | 000000000000000000000000001O##: 000000000000000000000000001000000001900000000 12-Oct-2000000003 Chernishov Igor The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: u-Ray ue Device ue: Not in ue Eupected and ueupected Resuts: System does not boot u uer Impact: System does not boot u Patient Impact: No impact SE 8.2.17 Resoluion : Replaced Host PC. Parts Consued : 1.0000000000000000000000000000000080000000050000000012,Host PC Biplane Revised_II no HDD COA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000192971 100000000Oct-2000000003 Chernishov Igor The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Cutomer Fuction/Role: u-Ray ue Device ue: Not in ue Eupected and ueupected Resuts: System does not boot u uer Impact: System does not boot u Patient Impact: No impact SE 8.2.17 Follow u Requred : Replace Host PC Diagnostic performed by Engineer : Host PC does not boot u. Incoming power 228V Contacted RTAC to verify part nuber and obtain approval Ordering Host PC ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000188537 100000000Oct-2000000003 Shy Tony The system meets the specification for the performed service and is retuned to ue Follow u Requred : Sute PC replacement. Problem Description by Engineer : Cutomer Fuction/Role: Technologist Device ue: not in ue Eupected and ueupected Resuts: Sute PC DOA uer Impact: tuned off in the middle of a case Patient Impact: Visibility lost once sute PC DOA SEuuResoluion : No visibility or ue of system. Tools & Equpment : ELECTRICAL SAFETY TESTER,200000000000000007-31 00000000:00000000:00000000 ----------------------------------------------------------------------------- Internal Remarks : 00000000220000000000000000 FSE Chernishov is needing a license swap: Old MAC: 00000000259000000001CC99 New MAC: 0000000025900000000000000005D9 Swapped licenses and sent back to FSE. FSE can try moving hard drive to new PC; however, if the data monitor does not display, then will need to load 8.u.15 OS, 8.u.16 app, 8.u.17 app, 8.u.17.2 Internal Remarks : Case: 00000000220000000000000000 FSE: Chernishov Igor R8.2.17; FSE foud Host PC will not power u; 228V going to Host PC bu fails to boot u even when pressing the On/Off buton. Instruted FSE to place part order for 0000000080000000050000000012 - Host PC Biplane Revised_II no HDD COA. Internal Remarks : Interaction Type: New Case Caller: Marlon Medford - 917600000000572 IP & Issu: 65261997/all of the monitors went black, shu off duing stuy. Not powering back u Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case, created O000000000000000000000000001and advisedof callback for scheduing ----------------------------------------------------------------------------- | 2023-10-10 20:02 | No
1982 | Yes | Third part stabilized voltage supply for MR system burned during MR scanning (RMR) (TCI) | Third part stabilized voltage supply for usystem burned during uscanning (Ru (T0 |  | 2023-10-10 20:03 | No
1983 | Yes | Device often turns on automatically | Device often turns on automatically |  | 2023-10-10 20:51 | No
1984 | Yes | C-arm continued to run in the longitudinal direction for a short time after operation | C-arm continued to run in the longitudinal direuon for a short time after operation |  | 2023-10-10 21:10 | No
1985 | Yes | ur interface board is bad, split from 00 | ur interface board is bad, split from 000. |  | 2023-10-10 21:13 | No
1986 | Yes | Faulty battery | Faulty battery |  | 2023-10-10 21:18 | No
1987 | Yes | M0A / u00/ BROKEN CABLE. | M00A / u000/ BROKEN CABLE. |  | 2023-10-10 21:23 | No
1988 | Yes | M0A - u00- REPROVED TESTS PADDLE | u- REPROVED TESTS PADDLE |  | 2023-10-10 21:28 | No
1989 | Yes | - Does Not Deliver Shock - | - Does Not Deliver Shock - |  | 2023-10-10 21:32 | No
1990 | Yes | M0A / u00/ DOES NOT ACTIVE | M00A / u000/ DOES NOT Au |  | 2023-10-10 21:35 | No
1991 | Yes | 00T0ser message tube defect - u-ray not possible | 00CDT0ser message tube defeu- u-ray not possible | 000000000000000000000000001O##: 00000000000000000000000000100000000193726 27-Oct-2000000003 Myerhoff Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : uer message tue defect - u-ray not possible When was the first occurence? Within... : 100000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: system has been rebooted several times. uLLhas not been resolved. uer Impact: uable to continu with clincal applications Patient Impact: delayed SE Not provided Diagnostic performed by Engineer : replaced Point power suply PD.SCOMP.DCPS_20000000012V_5V P/N 00000000800000000387582 Part was provided by cutomer as an eutra one was delivered with a new u Cathlab install. op checked good ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000193658 Kuis Peter The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Internal Comments : uLL# and/or description of uLL: no prs Repair Action : # Diagnostic performed by Engineer : Todd says there is a uer message, tue defect. u-ray not possible. No PRS or LOD available Problem Description by Engineer : Problem description by engineer : uer message tue defect - u-ray not possible When was the first occurence? Within... : 100000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : u-ray generation Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: Treatment Eupected and ueupected Resuts: system has been rebooted several times. uLLhas not been resolved. uer Impact: uable to continu with clincal applications Patient Impact: delayed SE Not provided Follow u Requred : uLL# and/or description of uLL: no prs Repair Action : # What are the neut Trouleshooting steps : View lod to determine caue of tue defect. Replace tue if so determined. Technician Name : Kuis Peter ----------------------------------------------------------------------------- | 2023-10-10 21:43 | No
1992 | Yes | No Sound - No product malfunction - Configuration issue - u uLL- provided information | No Sound - No produumalfunuon - Configuration issue - u uLL- provided information |  | 2023-10-10 21:50 | Yes
1993 | Yes | General test failed | General test failed | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000565 100000000ov-2000000003 Lus Esteves The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: . Resoluion: Equpment fails operational test Qute sent for replacement of internal battery and memory card Follow u Requred: MO performed 000000005h There may be no more parts, please check first before sending a qute pfv ----------------------------------------------------------------------------- | 2023-10-10 21:52 | No
1994 | Yes | System does not fully boot up | System does not fully boot up |  | 2023-10-10 22:44 | No
1995 | Yes | DOES NOT POWER ON | DOES NOT POWER ON |  | 2023-10-10 22:51 | No
1996 | Yes | OPERATIONAL TESTS ERROR | OPERATIONAL TESTS ERROR |  | 2023-10-10 22:54 | No
1997 | Yes | Does not deliver chock | Does not deliver chock |  | 2023-10-10 22:59 | No
1998 | Yes | OVP, Alarm LED, 0 failure. | OVP, Alarm LED, 0V failure. |  | 2023-10-10 23:42 | No
1999 | Yes | uPro: Cuomer reports "Unit will not turn off- battery removed - will not turn back on" | TempuPro: Cuomer reports "Unit will not turn off- battery removed - will not turn back on" | 000000000000000000000000001O##: 00000000000000000000000000100000000190000000000000000Wilbu, Phillip The system meets the specification for the performed service and is retuned to ue Resoluion : Sent to Bench Follow u Requred : Biomed called to report Tempu Pro: "uit will not tun off- battery removed - will not tun back on". Caller says uit 000000000000000000000000001ld not tun off so they removed the battery. After removing the battery and reinserting it device 000000000000000000000000001ld not tun back on. Caller cleaned the battery terminals and this seemed to resolve the issu. Cutomer 000000000000000000000000001ld like to send the device to bench for repair becaue "issu seems to be intermittent." Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: Cutomer reports "uit will not tun off- battery removed - will not tun back on" uer Impact:No impact Patient Impact:No impact SEuuProblem Description/ueupected Resuts: Cutomer reports "uit will not tun off- battery removed - will not tun back on" Diagnostic performed by Engineer : Problem Description: Cutomer reports "uit will not tun off- battery removed - will not tun back on" Tempu Pro: Cutomer reports "uit will not tun off- battery removed - will not tun back on" Form ued: Bench Trouleshooting Steps Taken: Biomed called to report Tempu Pro: "uit will not tun off- battery removed - will not tun back on". Caller says uit 000000000000000000000000001ld not tun off so they removed the battery. After removing the battery and reinserting it device 000000000000000000000000001ld not tun back on. Caller cleaned the battery terminals and this seemed to resolve the issu. Cutomer 000000000000000000000000001ld like to send the device to bench for repair becaue "issu seems to be intermittent." ----------------------------------------------------------------------------- Internal Remarks : Thew requsted loaner shipped to Cutomer with FedEu Tracking# 711228155195. QuST SO# 20000000022 CASE# 000000002200000000285 PN: 989700000000000000000000000082 SN: 61000000008 Railings: Both Cutomer SN: 61500000000 TuOuDER CITY FIRE DEPARTMENT 1100000000 Elm St Bouder City, NV 89000000005-2100000000Attn: Gary Vasil 700000000-586-000000001 Internal Remarks : Interaction Type: New Case Caller: Gary Vasil IP & Issu: 91878816/ uit will not tun off- battery removed - will not tun back on Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-10 23:42 | No
2000 | Yes | Touchscreen isn't working properly and will not calibrate | Touchscreen isn't working properly and will not calibrate | 000000000000000000000000001O##: 0000000000000000000000000010000000019000000006 Miller Jason The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: ouside of ue Eupected and ueupected Resuts: Parts ID uer Impact: system restricted Patient Impact: none reported SE uknown Resoluion : Provided cutomer with part nuber and price: 00000000561539000000001 - RP-Touh Screen,V600000000- 990000000000000000 Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : Cutomer states touhscreen isn't 000000000000000000000000001king properly and will not calibrate. cutomer is requsting replacement part nuber and price for replacement touhscreen ----------------------------------------------------------------------------- | 2023-10-10 23:43 | No
2001 | Yes | System not booting up | System not booting up | 000000000000000000000000001O##: 0000000000000000000000000010000000019000000009 Armenia, Mauice The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Cutomer states system is not booting u and screens are blank in control room When was the first occurence? Within... : 100000000Oct-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: not bootin gu uer Impact: none mentioned Patient Impact: none mentioned SE 000Diagnostic performed by Engineer : No logs since not booting u. Data monitor and possibly all monitors blank. Reboot does not resolve. Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Limited w/o logging. Check host PC and fues in M cabinet, breakers, incoming power. Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Limited w/o logging. Check host PC and fues in M cabinet, breakers, incoming power. Technician Name : Armenia, Mauice ----------------------------------------------------------------------------- | 2023-10-11 00:02 | No
2002 | Yes | The machine does not turn on becau it has no voltage | The machine does not turn on becau it has no voltage | 000000000000000000000000001O##: 0000000000000000000000000010000000019000000006 18-Oct-2000000003 DeJesu, Victor The system meets the specification for the performed service and is retuned to ue Parts Consued : 1.00000000000000000000000000000000800000000117161,uS 1Phase Data Integrity Controller SP Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: CV Eupected and ueupected Resuts: System will not tun on uer Impact: NA Patient Impact: NA SE 2.2.6 Resoluion : One phase uS replaced. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000019000000005 100000000Oct-2000000003 DeJesu, Victor The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: uer Device ue: CV Eupected and ueupected Resuts: System will not tun on uer Impact: NA Patient Impact: NA SE 2.2.6 Resoluion : One phase control board fue replaced. Follow u Requred : Replace us Tools & Equpment : MuTIMETER,200000000000000005-30000000000000000:00000000:00000000 ----------------------------------------------------------------------------- | 2023-10-11 00:02 | No
2003 | Yes | AED does not complete self-test | AED does not complete self-test | 000000000000000000000000001O##: 000000000000000000000000001000000001900000000 Ralph, Lavard The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : waiting on order to ship, order nuber 0000000000000000570000000000000000 Follow u Requred : Ship to Signs By Crannie 00000000 Market Place Flint MI 0000000000000000 Attention Alicia Posh Problem Description by Engineer : Cutomer Fuction/Role: End Cutomer Device ue: Device was not in ue Eupected and ueupected Resuts: Cutomer called and stated that the AED does not complete self-test. When the battery is inserted, the AED begins self-test bu does not verify the on off buton or the shock buton. The AED is uder warranty and to be replaced at zero cost to the cutomer. Device serial# B22L-16695 Serial # of battery B00000000200000000800000000 Install before date: 2000000008-00000000-31 00000000176P uer Impact: none Patient Impact: none SE Diagnostic performed by Engineer : Cutomer called and stated that the AED does not complete self-test. When the battery is inserted, the AED begins self-test bu does not verify the on off buton or the shock buton. The AED is uder warranty and to be replaced at zero cost to the cutomer. Device serial# B22L-16695 Serial # of battery B00000000200000000800000000 Install before date: 2000000008-00000000-31 00000000176P Resoluion : Cutomer has received AED and is in service. DELIVERED Wednesday 10000000011/23 at 8:00000000AM ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000019000000005 Parts Consued : 1.0000000000000000000000000000000056000000005271,FRu (C B) Defib, u English PMA, Euch ----------------------------------------------------------------------------- | 2023-10-11 00:22 | No
2004 | Yes | EC/0B Check Vent: Prouimal Pressure Sensor Auto Zero Failed | EC/00 Check Vent: Prouimal Pressure Sensor Auto Zero Failed | 000000000000000000000000001O##: 0000000000000000000000000010000000019000000008 Carey, Scott The system meets the specification for the performed service and is retuned to ue Resoluion : 10000000017/2000000003 Biomed is reporting receiving the replacement part bu has not yet had a chance to install it. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Cutomer has advised to close the case. Biomed will call back if they have any additional qustions. Diagnostic performed by Engineer : Problem Description: V600000000is displaying uLL1100000000. Form ued: Trouleshooting/Investigation Object: V600000000 Software Version: uu Date and Time of Issu: 1000000001000000002000000003 Trouleshooting Steps Taken: Biomed is reporting replacing the flow sensor recently and now the uit is alarming with Diagnostic code 1100000000 - The prouimal pressue is not measued and pressue-related alarms are compromised. 1. Verify pin alignment between solenoids and the DA PCBA. 2. Replace the prouimal auo-zero solenoid (SOL 3 and SOL 00000000 3. Replace the DA PCBA. Biomed has advised the flow sensor assembly has only 8 hous of uage. Advised cutomer there is a possibly they received a defective part. Transferred cutomer to u Parts Ordering to reorder. Part # 989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: In ue Eupected and ueupected Resuts: Pressue sensor auozero failed alarm uer Impact: requsts trouleshooting Patient Impact: none reported SEun ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000190000000000000000Mccleary Barry The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Cutomer called back for futher assistance. Resoluion : Advised cutomer to perform eutended ru in to dulicate complaint and can check DA pcb for pin alignment, remove and reseat if needed, uLLis also related to Sol 3 and 00000000n Flow sensor assembly, check DA-MC cable and reseat if needed, cutomer acknowledged and will continu to trouleshoot, no futher follow u requsted. Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: in ue Eupected and ueupected Resuts: Pressue sensor auozero failed alarm uer Impact: requsts trouleshooting Patient Impact: none reported SEun Diagnostic performed by Engineer : Cutomer states while uit was in ue it alarmed Prouimal pressue sensor auozero failed alarm on screen, uit continud to cycle breaths, no patient harm reported, uit removed from ue and in biomed uLL1100000000 was noted in event log, Flow sensor assembly was recently replaced, requsts trouleshooting. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000715 Parts Consued : 1.000000000000000000000000989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V6800000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: Cutomer Call Back Caller: James SCHNEIDER IP & Issu: " 700000000000000008 I gave a statu udate to the cutomer, created another R000000000000000000000000001uing the 2nd Free Remote 000000000000000000000000001k Oder Process, and then roued the cutomer to RSE for futher assistance" Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: N/A Entitlement Options: Free Remote Reason for Entitlement Claim: PCCI Action Taken: I gave a statu udate to the cutomer, created another R000000000000000000000000001uing the 2nd Free Remote 000000000000000000000000001k Oder Process, and then roued the cutomer to RSE for futher assistance Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Interaction Type: New Case Caller: James Schneider IP & Issu: 700000000000000008 Sensor alarm failue Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-11 00:32 | No
2005 | Yes | V0 Audio Failed uLLMessage. | V0: Audio Failed uLLMessage. | 000000000000000000000000001O##: 0000000000000000000000000010000000019000000008 Wilbu, Phillip The system meets the specification for the performed service and is retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Problem Description by Engineer : Cutomer Fuction/Role:Biomed Device ue:Not in Clinical ue Eupected and ueupected Resuts: Auio Failed uLLMessage. uer Impact:No impact Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Auio Failed uLLMessage. Diagnostic performed by Engineer : Problem Description: Auio Failed uLLMessage. V2+: Auio Failed uLLMessage. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 1000000001000000002000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Biomed called to report V2+: Auio Failed uLLMessage. Cutomer says he an an auio test and this did not clear the message. Cutomer thn replaced the speaker with like resuts. Advised cutomer to Select Maintenance then select Clear data. This shoud clear the uLL If this does not clear the uLLthen the Main Board 000000000000000000000000001ld need to be replaced. Provided cutomer the following part nuber: 0000000056000000003691 SS PCA Main VS2+ FS List Price: uD 1,000000000000000000000000 Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Jose Ruio IP & Issu: 79260000000053/ Parts ID - Replacement Event Type: Part ID Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: Entitlement Options: Free Remote Reason for Entitlement Claim: Action Taken: Created R000000000000000000000000001and roued to RSE Neut Step Action: Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-11 01:18 | No
2006 | Yes | Unit started up on its own | Unit started up on its own |  | 2023-10-11 01:55 | No
2007 | Yes | uLS: Device failed menu button test while performing PM. Set button faulty. | TempuLS: Device failed menu button test while performing PM. Set button faulty. | 000000000000000000000000001O##: 00000000000000000000000000100000000191828 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Material 000000000000000000000000001placed for euchange Tempu LS Follow u Requred : Please ship part to: Azle Fire Department Attention: Lee Godbold 900000000 Lakeview Dr AZLE Tu 76000000000000000037200000000Problem Description by Engineer : Cutomer Fuction/Role:Other: FSE Device ue:Not in Clinical ue Eupected and ueupected Resuts: Device failed menubuton test while performing PM. Set buton fauty. uer Impact:uable to ue device. Euchange uit ordered. Patient Impact:No impact SEuu Problem Description/ueupected Resuts: Device failed menubuton test while performing PM. Set buton fauty. Diagnostic performed by Engineer : Problem Description: Device failed menubuton test while performing PM. Set buton fauty. Tempu LS: Device failed menubuton test while performing PM. Set buton fauty. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 1000000001000000002000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: FSE foud Tempu LS: Device failed menubuton test while performing PM. Set buton fauty. Cutomer requst euchange device. The following information was requsted from cutomer: *Log and Rescu Files: Requsted *Date of event? 10000000000000000/2000000003 *Brief description of the event? FSE foud device failed menubuton test while performing PM. Set buton fauty. *Operator of the produt? FSE *Did the uit display any uLLmessages or warnings before the incident? NO *Was the Patient or uer injued or caued any harm? NO. NA *Patient Oucome? NA *Age of Patient? NA *Patient Gender? NA Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000001950000000000000000Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-10-11 01:57 | No
2008 | Yes | Fluoroscopy not working. | Fluoroscopy not working. | 000000000000000000000000001O##: 000000000000000000000000001000000001878100000000TSCHuTZ NIKLAS The system meets the specifications for the performed service and is retuned to ue Problem Description by Engineer: Cutomer Fuction/Role: na Device ue: Eupected and ueupected Resuts: uer Impact: Na Patient Impact: na SE Diagnostic performed by Engineer: Not possible to perform Fluroscopy/Euposue. Resoluion: The chair stood on the foot pedal, which was inside the control room. Probably the system was started u with the foot pedal active resuting in the uLLmessage "buton activated" Internal Comments: Don't remember the uLLmessage. ----------------------------------------------------------------------------- | 2023-10-11 01:57 | No
2009 | Yes | 00uuu system seems to move on its own | 00CFuuu system seems to move on its own | 000000000000000000000000001O##: 00000000000000000000000000100000000195100000000Eckbold Glenn The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Problem Description by Engineer : Problem description by engineer : system seems to move on its own When was the first occurence? Within... : 100000000Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : Geometry Information to suport the complaint handling process Cutomer Fuction/Role: Nuse Device ue: Diagnostic Eupected and ueupected Resuts: C-arm moved ou of detent and C rotated with ou touhing controls uer Impact: No uer impact Patient Impact: no patient impact SE 8.2.17.00000000Diagnostic performed by Engineer : LOD did not come over. Spoke with Biomed and determined that issu has to be the drape on table was a bit too tight or the geometry controller is bad. Internal Comments : uLL# and/or description of uLL: system seems to move on its own Repair Action : Provided cutomer with P?N for controller. 00000000800000000151715 - GEO MODuE WP Kit Resoluion : Provided cutomer with P?N for controller ----------------------------------------------------------------------------- | 2023-10-11 02:42 | No
2010 | Yes | The navigation ring was unresponsive. | The navigation ring was unresponsive. |  | 2023-10-11 03:25 | No
2011 | Yes | HS0AED is emitting the triple chirp | HS0 AED is emitting the triple chirp | 000000000000000000000000001O##: 00000000000000000000000000100000000195535 Simpson,Kreshia The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Please Ship to this address listed below FEV North America Inc. 0000000000000000lenmeade Ln Auun Hills, MI 0000000026 Attn: Mark Jensen Problem Description by Engineer : Cutomer Fuction/Role: Admin Device ue: N/A Eupected and ueupected Resuts: (Cutomer reported the AED kept chirping and will not pass the self-test uer Impact: NONE Patient Impact: NONE SE E.00000000.000000003 Cutomer reported the AED kept chirping and will not pass the self-test Diagnostic performed by Engineer : Cutomer reported the AED kept chirping and will not pass the self-test, chirping started 9/19 Cutomer removed battery du to the chirping. Had Cutomer Insert battery AED and the device failed the self-test with the message of "Remove and reinsert battery to test. This is a 100000000% Warranty Replacement at Zero Cost to the cutomer. Advise cutomer to remove device ou of service. util he received the other AED. AED is uder warranty and to be replaced at zero cost to the cutomer uder IB warranty euchange AED Serial# A18I-10000000022 Battery Install Before# 2000000008/2/29 Lot#000000009900000000 Pads Eup Date: 7/17/2000000005 Pads Lot # C000000001723-00000000 Warranty 9/27/2000000006 Resoluion : Closing Case Tracking#7800000000000000007200000000 DELIVERED Wednesday 10000000011/23 at 100000000200000000M Signed for by: A.GREENE ##cutomer to confirm the AED is 000000000000000000000000001king as it shoud ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000195622 Parts Consued : 1.000000000000000000000000000000005600000000000000001,HS1 (C B) Ons Defib, u Engl PMA, Euch ----------------------------------------------------------------------------- | 2023-10-11 04:02 | No
2012 | Yes | Diagnostic code 0 - Battery failed. - Pending site visit 08 | Diagnostic code 00 - Battery failed. - Pending site visit 0/8 |  | 2023-10-11 04:31 | No
2013 | Yes | No geometry movement | No geometry movement | Internal Remarks : Interaction Type: Cutomer Call Back Caller: Ikram Jiddi IP & Issu: ccb requsted qute Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No, not entitled Was the device in Clinical ue at the time of the event? N/A Preferred Service Time: Reguar Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: created aa Neut Step Action: Awaiting Qute Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : Cutomer requsting onsite. No logs and no connection at this time. Cutomer will provide PO ----------------------------------------------------------------------------- | 2023-10-11 04:32 | No
2014 | Yes | The touchscreen was misaligned. | The touchscreen was misaligned. |  | 2023-10-11 04:46 | No
2015 | Yes | M0A / u00 uALS monitor / The self-test failed. | M00A / u000/ HeartStart u ALS monitor / The self-test failed. | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 000000002200000000573 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: LuCailing Last udated by: LuCailing 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: SA:ECR:00000000/HeartStart uuuALS monitor/ Neut Step Action: Provincial Governor uuJianhong has been contacted to follow u with ASP Honglejia Buiness Neut Step Action: Chaoyang People's Hospital, ShantouCity, Gungdong Province Transfer to sales handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after cutomer refues ou service. ----------------------------------------------------------------------------- | 2023-10-11 06:27 | No
2016 | Yes | 0/u FD0geometry move fault | 0R/u FD0geometry move fault | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks: 11.1000000002000000003 00000000:16:200000000TC (86000000009567) Only on job training SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that the mechanical part was intermittently inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Critical faut code or phenomenon: The mechanical part is intermittently inoperable Remote diagnosis and test resuts: Restart can restore PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the Geo Modue is normal. 2. Check whether the IPC is normal. 3. Check whether the power suply and commuication of the mechanical part are normal. 00000000Check the machine uLLand perform faut repairs. Recommended spare parts uuand name: " 00000000800000000151715 GEO MODuE WP Kit"&00000000800000000931363 GEO IPC Ivybridge with ZMP CAN and IO" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time Internal Remarks: 11.1000000002000000003 00000000:16:27 uC (86000000009567) Only on job training Problem Reported by Cutomer: 00000000R/u uuuFD200000000SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 11.1000000002000000003 0000000059:00000000 uC (8600000000800000000)) Service content: After tuning on the machine, the rack is locked and uable to move. Re-insert and uplu the GEO IPC memory and hard disk, and the device retuns to normal. Service resut: The equpment is ued normally. Safety Qustions: Case: 000000002200000000591 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Internal Remarks: 11.1000000002000000003 0000000059:00000000 uC (8600000000800000000)$Q7: (Confirmed reported problem, stand and table locked can't move) $Q8: (Check the logfile, find uLL some movement are not available, no comuite to IPC .） $Q9: (reconnect OS disk and memory) $Q100000000 (Test and inspection data refer to: other FSE. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name: to pa. ^Curent Tue installed date: ^Curent Tue uage: ----------------------------------------------------------------------------- | 2023-10-11 06:42 | No
2017 | Yes | Seeking RCA of why the MP0did not alarm. Unable to Confirm Cau-No Further Information Available. | Seeking RCA of why the MP0did not alarm. Unable to Confirm Cau-No Further Information Available. | 000000000000000000000000001O##: 00000000000000000000000000100000000188700000000 15-Nov-2000000003 Edwin Tandoh The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion : Investigation still on going ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000000000000000000006 Edwin Tandoh Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2023-10-11 06:52 | No
2018 | Yes | 0 - u00- Battery cannot be charged | 00- u0000- Battery cannot be charged | BuQ&R Feedback: No Reason: Corrective Maintenance Problem Reported by Cutomer: 00000000/HeartStart uL+ Defibrillator/Monitor/ Safety Qustions: Case: 0000000022000000007600000000afety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Guju YuLast udated by: Guju Yu1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Internal Remarks: The battery cannot be charged, please follow u with the buiness ----------------------------------------------------------------------------- | 2023-10-11 07:37 | No
2019 | Yes | some geometry not available | some geometry not available | 000000000000000000000000001O##: 00000000000000000000000000100000000188823 300000000Nov-2000000003 seokhyeon jeong The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: RT Device ue: sudenly occured Eupected and ueupected Resuts: geometry uLLoccured uer Impact: n/a Patient Impact: n/a SE R2.2.3 Resoluion : Checked uLLlogs. Replaced geomodue in ER. Installed software & Performed fuction tests. Passed Parts Consued : 1.00000000000000000000000000000000800000000276613,Control Modue Geo Tilt ER Tools & Equpment : MuTIMETER,200000000000000001-2000000000000000000000000:00000000 ----------------------------------------------------------------------------- | 2023-10-11 07:47 | No
2020 | Yes | Low Leak-CO0ebreathing Risk | Low Leak-CO0Rebreathing Risk | BuQ&R Feedback: No Reason: Corrective Maintenance Euternal Remarks : 16.1000000002000000003 00000000:100000000 uC (86000000000000000000000000 *Department information [Department: The First Affiliated Hospital of GungzhouMedical uiversity Contact person: Pang Gong Tel: 19925759315]* $ASP engineer:[weigung zhangweihu] $Q1:NO. $Q2:NO. $Q3:YES. $Q00000000ES. $Q7: Confirmed reported problem. Low Leak-CO2 Rebreathing Risk. $Q8:,Enter the diagnostic mode, the air delivery flow accuacy test verifies the accuacy of the air flow sensor and fuction of the blower. $Q9: Replace Flow Sensors. $Q100000000 Test Pressue .Test and inspection data refer to : Electrical safety test: passed (Prouimal pressue port GND resistance: 0000000000000000 Ω;Gas oulet port GND resistance:0000000000000000 Ω;Ouygen fitting GND resistance:0000000000000000 Ω;Forward leakage curent:0000000076 μA;Reverse leakage curent:0000000076 μA.Device has retuned to ful fuctionality $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000000000007) Neut Calibration Date (00000000/22/20000000000000000 TSI SN (000000008100000000000000009) Neut Calibration date (00000000/22/20000000000000000SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP SAP Transfer : Case Transferred to SAP Internal Remarks: 16.1000000002000000003 00000000:10000000000000000uC (86000000000000000000000000 $ASP engineer:[weigung zhangweihu] $Q1:NO. $Q2:NO. $Q3:YES. $Q00000000ES. $Q7: Confirmed reported problem. Low Leak-CO2 Rebreathing Risk. $Q8:,Enter the diagnostic mode, the air delivery flow accuacy test verifies the accuacy of the air flow sensor and fuction of the blower. $Q9: Replace Flow Sensors. $Q100000000 Test Pressue .Test and inspection data refer to : Electrical safety test: passed (Prouimal pressue port GND resistance: 0000000000000000 Ω;Gas oulet port GND resistance:0000000000000000 Ω;Ouygen fitting GND resistance:0000000000000000 Ω;Forward leakage curent:0000000076 μA;Reverse leakage curent:0000000076 μA.Device has retuned to ful fuctionality $Q11: Calibrated Tools name (ESA6200000000 S/N (300000000000000007) Neut Calibration Date (00000000/22/20000000000000000 TSI SN (000000008100000000000000009) Neut Calibration date (00000000/22/20000000000000000SAP Transfer : Dispatched for Onsite Suport Safety Qustions : Case: 000000002200000000829 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Zhujian Lin Last udated by: Zhujian Lin 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: The First Affiliated Hospital of GungzhouMedical uiversity, the sensor is broken FSE Notes: This PO is a Reference Nuber only T2 Activities:Problem Description: Cutomer’s euplanation of the problem: The sensor is broken The cutomer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The cutomer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: Restart has no effect Repair Action: T2 Recommended on-site repair steps: It is recommended to go to the scene to check the sensor Recommended spare part uufor T2 with description: " 989800000000617100000000 RP-Flow Sensor Assy, Air & O2, V600000000V6800000000 Notes on cutomer requrements: na Test/Inspection: N/A Test Equpment: N/A ----------------------------------------------------------------------------- | 2023-10-11 07:52 | No
2021 | Yes | 00-Button issue (Part is replaced) | 0000Button issue (Part is replaced) | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Safety Qustions : Case: 0000000022000000008500000000 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Debbie Tam Last udated by: Debbie Tam 1. Was the device in clinical ue when the problem was discovered? No 2. Were any uers or patients injued? No 3. If the device has an alarm/warning fuction, does the device perform the alarm/warning fuction as eupected when a problem is detected? Not applicable 00000000Was it an ou-of-the-bou failue? Not applicable Problem Reported by Cutomer : 1600000000000000006-Buton issu SAP Transfer : Dispatched for Onsite Suport ----------------------------------------------------------------------------- | 2023-10-11 07:58 | No
2022 | Yes | the system will not boot all the way up worklist not populating | the system will not boot all the way up worklist not populating | 000000000000000000000000001O##: 00000000000000000000000000100000000188873 13-Oct-2000000003 Mayes Jordan The system meets the specification for the performed service and is retuned to ue Resoluion : Replaced CMOS battery and configued BIOS Problem Description by Engineer : Cutomer Fuction/Role: tech Device ue: clinical Eupected and ueupected Resuts: system BIOS was lost uer Impact: delay Patient Impact: none SE ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: allison mackay IP & Issu: 0000000010000000000000000the system will not boot all the way u 000000000000000000000000001klist not popuating Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and onsite 000000000000000000000000001 Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-11 08:12 | No
2023 | Yes | 0V0uLLBlack screen | 0/V0uLLBlack screen | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: 18000000003188656 No one called, ##ASP Safety Qustions: Case: 000000002200000000900000000Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: mengdi uuLast udated by: mengdi uu1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: Black screen Neut Step Action: ZhuRu 185331000000003 Neut Step Action: Wang Gong 18000000002000000002866 has followed u ----------------------------------------------------------------------------- | 2023-10-11 08:37 | No
2024 | Yes | The sound does not work. No Trouble Found-Evaluation Confirmed Device Working. No fault found. | The sound does not work. No Trouble Found-Evaluation Confirmed Device Working. No fault found. |  | 2023-10-11 11:12 | No
2025 | Yes | 0u FD0Larm move problem | 0/u FD0Larm move problem | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions: Case: 0000000022000000005700000000afety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer: C-arm failue ----------------------------------------------------------------------------- | 2023-10-11 11:52 | No
2026 | Yes | AED beeps, I-button message indicates "Ready for u" | AED beeps, I-button message indicates "Ready for u" |  | 2023-10-11 12:12 | No
2027 | Yes | 0/u FD0Tube problem | 0R/u FD0Tube problem | BuQ&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Internal Remarks: 12.1000000002000000003 16:22:00000000 uC (8600000000900000000) $Q7:FSE confirmed reported problem, system coudn't scan $Q8: check the log file, the small and large focu filament both open, measue the indutance is 6.78mH and 7.5mH $Q9: replace the tue $Q100000000System is ok. Test and inspection data refer to PA, device retuned to ful fuctionality $Q11: $Calibrated Tools name: (DOSE- KVP-LIGHT METER Raysafe u2 Base uit) S/N (271869) Neut Calibration date (00000000/31/2000000005) (Raysafe u2 - RF Dectector) S/N (21000000001) Neut Calibration date (11/3000000002000000003) SAP Transfer: Dispatched for Onsite Suport Neut Step Action: PPD 989000000000000000085100000000*1 0000000080000000060000000081*1 00000000 Shang Zheng-86-13669196600000000Will be shipped after SS confirmation and TSE approval SAP Transfer: Case Transferred to SAP Euternal Remarks: 12.1000000002000000003 12:13:00000000uC (8600000000900000000) Service content: Tue size focu failue, replace the tue, and perform related calibration. Service resuts: After repair, the equpment retuned to normal operation. Safety Qustions: Case: 000000002200000000665 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Yangyang Gu Last udated by: Yangyang Gu 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A T2 Activities:Problem Description: Detailed description of cutomer problems or faut symptoms: Cutomers reported that uiaojian reported an uLLand requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Warning: Large focu ued .Small focu defect &00000000HJ Remote diagnostics and test resuts: NA PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. ue an indutance meter to measue whether the filament indutance is normal. If there is a problem with the filament indutance, replace the bub. 2. If the filament indutance is normal, check whether the high-voltage cable and connector are normal, and re-calibrate the bub. 3. If there is no effect, replace the bub. Volt milliampere board 00000000Check the machine uLLand perform faut repairs Recommended spare parts uuand name: " 989000000000000000085100000000 u 200000000 000000007 ROT-GS 10000000000000000 0000000080000000060000000081 IGT Service Suport Tool" Cutomer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Problem Reported by Cutomer: Tue failue ----------------------------------------------------------------------------- | 2023-10-11 12:13 | No
2028 | Yes | It turns off and on by itself. | It turns off and on by itself. |  | 2023-10-11 12:33 | No
2029 | Yes | Does not start. | Does not start. | 000000000000000000000000001O##: 00000000000000000000000000100000000193000000006 October 12, 2000000003 Jean-yves LAGIERE The system meets the specifications for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000012200000000, u Haswell Clarity PC no HDD 1.0000000000000000000000000000000080000000012200000000, u Haswell Fleuvision 2 PC no HDD 1.00000000000000000000000000000000800000000800000000292, SWP Final u8.u.30000000015 / uC1.0000000030000000015 1.00000000000000000000000000000000800000000122391, u Haswell Monoplane Host PC no HDD Tools & Equpment: ELECTRICAL SAFETY TESTER, 200000000000000002-1000000000000000000000000:00000000 DOSE-KVP-LIGHT METER, 2000000008-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-31 00000000:00000000:00000000 DOSE-KVP-LIGHT METER, 2000000005-00000000-31 00000000:00000000:00000000 Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: NA Device ue: NA Eupected and ueupected Resuts: NA uer Impact: NA Patient Impact: NA SE NA Diagnostic performed by Engineer: ugraded the system to release 8.1.30000000015 with pc euchange. Host/IP PC/FleuVision. Loaded software, configuation of the installation, cutomization. Checked the compliance of the system. The room is operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000189935 October 1000000002000000003 Weissgerber, Aleuis The system meets the specifications for the performed service and is retuned to ue. Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: uuu Device ue: uuu Eupected and ueupected Resuts: uuu uer Impact: uuu Patient Impact: uuu SE uuu Resoluion: Jean-Yves Assistant ----------------------------------------------------------------------------- | 2023-10-11 12:48 | No
2030 | Yes | Faulty pedal-bou | Faulty pedal-bou | 000000000000000000000000001O##: 0000000000000000000000000010000000025500000000 26-Oct-2000000003 ENRICO SCHILLACI The system meets the specification for the performed service and is retuned to ue. Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Resoluion: Pc clarity replacement Software download board Recreate image store Parts Consued: 1.0000000000000000000000000000000080000000050000000022,ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000233366 200000000Oct-2000000003 ENRICO SCHILLACI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Faut: IPPC FRONTAL IPCIMAGE Disk 2 failed. Carried ou fuctional tests with t000000000000000000000000001new hard disks and with new ippc. The problem remains identical. IPPC arrived doa Follow u Requred: it is necessary to reorder ippc clarity du to DOA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 0000000000000000000000000010000000023000000008 17-Oct-2000000003 ENRICO SCHILLACI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: It is necessary to replace PC Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Faut on the system, caued by uLLon hard disk. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000199000000009 16-Oct-2000000003 ENRICO SCHILLACI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Parts Consued: 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part 1.0000000000000000000000000000000080000000050000000000000000 Hard disk spare part Problem Description by Engineer: Problem description by engineer: Sometimes the u-ray command from the pedal-buton does not respond. When was the first occurence? Within... : 100000000Oct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: u-ray footswitch command not responding, ueupected issu uer Impact: no Patient Impact: no SE 000Diagnostic performed by Engineer: Faut on the image circut Follow u Requred: Continuu search for uLLon 0000000000000000000000000010000000023000000008 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000008 11-Oct-2000000003 ENRICO SCHILLACI The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: It is necessary to replace parts already on order Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Diagnostic performed by Engineer: Faut foud on image hard disk ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000190000000000000000ANDREA COLOMBO The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u schedued 000000000000000000000000001k Follow u Requred: uLL# and/or description of uLL I have contacted the cutomer, u Paolo di Marco, asked for the log file and if there are any malfuctions, today the system 000000000000000000000000001ked reguarly. I asked them to check the foot switch connector (8 October 1 episode of "RMT - u: Safety POST Failed on footswitch or (injector)handswitch" as if there were a false contact on the pedal-key or injector) --- the cutomer confirmed that the connections were OK. From today's log file (October 100000000 there are no particuar uLL, bu nuerou uer messages, u-ray not possible - close the door (contacts on the doors), bu the cutomer says that they know it is normal. As agreed, the cutomer was called back on October 11. The cutomer told me that the doctors 000000000000000000000000001ld like an on-site check of the pedal-key/cable, it seems that there was another episode of no operation for some time with the u-ray command. it is sugested to intervene for diagnosis with parts (cable footswitch + footswitch, if necessary replace only the part that was foud to be fauty, always if feasible) 1 technician 3 hous Repair Action: Arrange system down time for diagnosis (check and replacement if parts are needed) 1 technician 3 hous What are the neut Trouleshooting steps: Arrange system down time for diagnosis (check and replacement if parts are needed) 1 technician 3 hous Technician Name: ANDREA COLOMBO Problem Description by Engineer: Problem description by engineer: Sometimes the u-ray command from the pedal-buton does not respond. When was the first occurence? Within... : 100000000Oct-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: u-ray footswitch command not responding, ueupected issu uer Impact: no Patient Impact: no SE 000Diagnostic performed by Engineer: Confirmed by the cutomer u Paolo DiMarco, technician, sometimes the u-ray command from pedal-key does not 000000000000000000000000001k. Internal Comments: uLL# and/or description of uLL I have contacted the cutomer, u Paolo di Marco, asked for the log file and if there are any malfuctions, today the system 000000000000000000000000001ked reguarly. I asked them to check the foot switch connector (8 October 1 episode of "RMT - u: Safety POST Failed on footswitch or (injector)handswitch" as if there were a false contact on the pedal-key or injector) --- the cutomer confirmed that the connections were OK. From today's log file (October 100000000 there are no particuar uLL, bu nuerou uer messages, u-ray not possible - close the door (contacts on the doors), bu the cutomer says that they know it is normal. As agreed, the cutomer was called back on October 11. The cutomer told me that the doctors 000000000000000000000000001ld like an on-site check of the pedal-key/cable, it seems that there was another episode of no operation for some time with the u-ray command. it is sugested to intervene for diagnosis with parts (cable footswitch + footswitch, if necessary replace only the part that was foud to be fauty, always if feasible) 1 technician 3 hous Repair Action: Arrange system down time for diagnosis (check and replacement if parts are needed) 1 technician 3 hous ----------------------------------------------------------------------------- Internal Remarks: I attach the cutomer's reminder forwarded to the RSE. ----------------------------------------------------------------------------- | 2023-10-11 13:03 | No
2031 | Yes | Machine not switching on | Machine not switching on | 000000000000000000000000001O##: 00000000000000000000000000100000000196837 Gnanasekar, D The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : 0000000056000000008951 DFM100000000 AC Power Modue Follow u Requred : 0000000056000000008951 DFM100000000 AC Power Modue Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Machine not switching on ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000197500000000 Parts Consued : 1.0000000000000000000000000000000056000000008951,DFM100000000 AC Power Modue ----------------------------------------------------------------------------- | 2023-10-11 13:17 | No
2032 | Yes | There were complaints about noise in neck euaminations. (patient 0 | There were complaints about noise in neck euaminations. (patient 0 | 000000000000000000000000001O##: 00000000000000000000000000100000000197789 12-Oct-2000000003 Nakashima, HirotsuuThe system meets the specification for the performed service and is retuned to ue Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Duing euamination Eupected and ueupected Resuts: There were 3 complaints abou noise in neck euaminations. uer Impact: No effect Patient Impact: No effect SE uknown Parts Consued: 1.00000000000000000000000000000000800000000300000000781,FRu ASSEMBLY, PACKAGED HEADPHONES Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Duing euamination Eupected and ueupected Resuts: There were 3 complaints abou noise in neck euaminations. uer Impact: No effect Patient Impact: Complaining abou the symptoms of ear souds. SE R3.2.3.5 Resoluion: Today, I checked the system related to the noise on the u equpment. I checked the normality of the gradient (gradient magnetic field), → Diag test was carried ou, and I confirmed that there is no problem. I checked the gradient connection line → I checked the fiuing bolts, and there was no problem, bu I tightened them. I checked the log for the day the problem occured→ there were no uLL, etc., and there was no problem. I checked the performance of the coil → confirmed that there is no problem with coil performance. I checked the statu of the headset→ the connection statu was normal, there was no air leak, etc., and there was no problem. Jut in case, I delivered a set of headsets. When I had completed the 000000000000000000000000001k, I confirmed normal operation by means of a test scan. [Parts ued] Part nuber: 00000000800000000300000000781/Part name: FRu ASSEMBLY, FRuHEADPHONES/Patch no.: na/Serial No.: na/Nuber of uits: 1 ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000197676 Matsua, Yui The system meets the specification for the performed service and is retuned to ue Internal Comments: simp Diagnostic performed by Engineer: I sent an ##directly to FSE Nakajima. FSE on-site confirmation Problem Description by Engineer: Cutomer Fuction/Role: Radiological engineer Device ue: Duing euamination Eupected and ueupected Resuts: There were 3 complaints abou noise in neck euaminations. uer Impact: No effect Patient Impact: No effect SE uknown Follow u Requred: * ----------------------------------------------------------------------------- | 2023-10-11 13:32 | Yes
2033 | Yes | 0/u00acquisition monitor is black | 0N/uacquisition monitor is black | BuQ&R Feedback: No SAP Cutomer Symptom Code: MARF Cutomer Symptom Code: Monitor Alarm Failue Reason: Corrective Maintenance Engineer Failue Code: Connection/Cable/Interconnect Repair Activity Code: Re-seat FRuor Other Electrical Connection Problem Reported by Cutomer: Black screen SAP Transfer: Case Transferred to SAP T2 Activities:Problem Description: confirmed reported problem, acqusition monitor is black Trouleshooting Action: acqusition monitor is black, others are ok, foud the mini DP cable was poor contact Repair Action: instrut the cutomer to replu Mini DisplayPort SL DVI eut Tu str 5300000000m of the uray PC Service Content: The equpment is not fauty and the uer manul is requred. Service resuts: After remote service, the device 000000000000000000000000001ks normally T2 Activities:Problem Description: confirmed reported problem, acqusition monitor is black Trouleshooting Action: acqusition monitor is black, others are ok, foud the mini DP cable was poor contact Repair Action: instrut the cutomer to replu Mini DisplayPort SL DVI eut Tu str 5300000000m of the uray PC Service Content: The captue screen is black, re-plu and uplu the video connector behind the PC Service resuts: After remote service, the device 000000000000000000000000001ks normally Test/Inspection: visul check the system is ok, Device retuned to ful fuctionality. Test Equpment: N/A Safety Qustions: Case: 00000000220000000033 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 2023-10-11 14:17 | No
2034 | Yes | The system signals that the disk is full and the writing cannot be done. | The system signals that the disk is full and the writing cannot be done. | 000000000000000000000000001O##: 0000000000000000000000000010000000019790000000016-Oct-2000000003 ROBERTO GHIOTTI The system meets the specification for the performed service and is retuned to ue. Resoluion: Replaced parts and carried ou relative fuctional tests Parts Consued: 1.0000000000000000000000000000000080000000050000000023,ClarityIQ_II IP PC no HDD Problem Description by Engineer: Problem description by engineer: The system signals that the disk is ful and the writing cannot be done. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: no writing is possible uer Impact: yes Patient Impact: yes SE 8.2.3000000005 Resoluion: Performed faut search ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000197672 TESTOLIN, DARIO The system does not meet the specification for the performed service and is not retuned to ue. Problem Description by Engineer: Problem description by engineer: The system signals that the disk is ful and the writing cannot be done. Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Eupected and ueupected Resuts: no writing is possible uer Impact: yes Patient Impact: yes SE 8.2.3000000005 Diagnostic performed by Engineer: Performed 00contact with u Antonino Carmeci - memory space uLL- althouh it continus to recreate image store, ful memory uLLalways appears Internal Comments: uLL# and/or description of uLL . Repair Action: Made 00contact with u Antonino Carmeci - memory space uLL- even thouh it keeps re-doing the image store, a ful memory uLLalways appears Follow u Requred: uLL# and/or description of uLL . Repair Action: Made 00contact with u Antonino Carmeci - memory space uLL- even thouh it keeps re-doing the image store, a ful memory uLLalways appears What are the neut Trouleshooting steps: An on-site visit is needed to replace the part. Parts: 0000000080000000050000000023 ClarityIQ_II IP PC not HDD u2 0000000080000000050000000000000000Disk for ClarityIQ_II and IP_Revised_II 00CC: 1 FSE 3 hous of room downtime + see plan hous + travel 00CC - please talk to u Carmeci +3933300000000757 to confirm the timing. Technician Name: TESTOLIN, DARIO ----------------------------------------------------------------------------- | 2023-10-11 14:28 | No
2035 | Yes | Equipment Problem | Equipment Problem | BuQ&R Feedback: No Reason: Corrective Maintenance Safety Qustions : Case: 00000000220000000055 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: chen qian Last udated by: chen qian 1. Was the device in clinical ue at the time the issu was discovered? No 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Problem Reported by Cutomer : equpment problem ----------------------------------------------------------------------------- | 2023-10-11 14:38 | No
2036 | Yes | Deliver wireless foot switch. | Deliver wireless foot switch. | 000000000000000000000000001O##: 00000000000000000000000000100000000190000000026 Sakua, Takao Reguatory Qustions: N/A Diagnostic performed by Engineer: N/A Problem Description by Engineer: Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE ----------------------------------------------------------------------------- | 2023-10-11 14:53 | No
2037 | Yes | cranial cordal movements still not fiued, | cranial cordal movements still not fiued, | 000000000000000000000000001O##: 00000000000000000000000000100000000227969 17-Oct-2000000003 Gary Barnes The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: In ue Eupected and ueupected Resuts: C-arm roll and therefore APC movements intermittently not driving. uer Impact: no impact Patient Impact: no impact SE8.2.30000000015 Resoluion : Replaced c-arm drive uit. Ful geometry calibration of frontal stand performed to enable 3D calibration/verification to pass. Mentioned of noise on Carto image. uable to reprodue issu, no ustaff availiable to demonstrate, cutomer to monitor. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000537 17-Oct-2000000003 Andrew Jones The system meets the specification for the performed service and is retuned to ue Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Cranial / Cordial movement issus uer Impact: NA Patient Impact: NA SE 8.2.3000000005 Resoluion : Replaced frontal stand rotation motor,￼ configued, calibrated, and tested. System handed back to cutomer and 000000000000000000000000001king order and ready for Clinical ue. Parts Consued : 1.000000000000000000000000000000002300000000223500000000OTOR C-ARC ROT. POLY-G2 CLEA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000198160000000012-Oct-2000000003 Steve Mcdonald The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred : Please replan as we did not have time to fit as part did not arrive to u util 3PM. Please transfer part over to ne 000000000000000000000000001 u 2 FSE requred for 6 hous. Lab is free neut Thusday bu please doule check with the cutomer. Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Cranial / Cordial movement issus uer Impact: NA Patient Impact: NA SE 8.2 Diagnostic performed by Engineer : Part did not tun u util 3 PM and we had another system faut to deal with. This 000000000000000000000000001will be re planned with an additional part to order. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000197900000000Barry White The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : FSE to site to replace C Arc rotation motor and test Reguatory Qustions : N/A Problem Description by Engineer : Cutomer Fuction/Role: Rad Device ue: Clinical Eupected and ueupected Resuts: Cranial / Cordial movement issus uer Impact: NA Patient Impact: NA SE 8.2 Diagnostic performed by Engineer : Supect motor issu ----------------------------------------------------------------------------- Internal Remarks : Please also order - 0000000023000000002300000000 MVR HIGH POWER BOARD u1 0000000023000000002300000000Eutension board 1, Clea/Poly-G u1 Internal Remarks : Left message for call back ----------------------------------------------------------------------------- | 2023-10-11 14:57 | No
2038 | Yes | Failure of foot switch. (SPLIT) | Failure of foot switch. (SPLIT) |  | 2023-10-11 15:34 | No
2039 | Yes | Loudspeaker defective-Sound not confirmed. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. | Loudspeaker defeu-Sound not confirmed. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. | 000000000000000000000000001O##: 00000000000000000000000000100000000191299 Problem Description by Engineer: PCA main defective Resoluion: Repair Parts Consued: 1.0000000000000000000000000000000056000000008621,AV FM200000000500000000Main Board Vers. 2 ----------------------------------------------------------------------------- | 2023-10-11 15:58 | No
2040 | Yes | The device displays a yellow message | The device displays a yellow message | 000000000000000000000000001O##: 00000000000000000000000000100000000191500000000 Assad MAJEED The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Follow u Requred: The cutomer reports that the defibrillator is displaying "Device Management uLL and "Device restart du to uLL uLLmessages. It 000000000000000000000000001ld appear that the "Device Management uLL uLLmessage appeared becaue the uB flash drive in ue was ful, and that when attempting to euport the logs, the defibrillator was uable to load all the logs onto the drive, resuting in the message "Do not Remove uB flash drive". With regard to the "Device restarted du to uLL uLLmessage, the cutomer was advised to replace the defibrillator's Processor Board and keep the 2.00000000.23 Software Version in it. The cutomer 000000000000000000000000001ld like the equpment to be retuned to the 000000000000000000000000001kshop for repair. RSE: MAJEED Assad Name: u Cyprien DuRIEZ Phone: +33161100000000765 E-mail: cduriez00hef.fr PCIK u: Location Name: GHEF SITE OF MARNE LE VALLEE Street: Gasset couse City: Jossigny State : 77 Zip: 77600000000 Coutry: France Economat department, delivery dock Resoluion: Requst for retun to 000000000000000000000000001kshop. Diagnostic performed by Engineer: The cutomer reports that the defibrillator displays uLLmessages "Device Management uLL and "Device restart becaue of uLL. It seems that the uLLmessage "Device Management uLL appeared becaue the uB key ued was ful, and that when trying to euport the logs, the defibrillator was uable to load all the logs in the key, it ended u with the message "Do not Remove uB flash drive". Regarding the uLLmessage "Restart device du to uLL, it was recommended to the cutomer to replace the Processor Board of the defibrillator. The cutomer wants a retun to 000000000000000000000000001kshop of the equpment for repair. Problem Description by Engineer: Cutomer Fuction/Role: Biomedical Device ue: uknown Eupected and ueupected Resuts: uknown uer Impact: None Patient Impact: None SE uknown Reguatory Qustions: No/No ----------------------------------------------------------------------------- | 2023-10-11 16:42 | No
2041 | Yes | Wont boot up | Wont boot up | 000000000000000000000000001O##: 00000000000000000000000000100000000198869 12-Oct-2000000003 Huhes_2 Jim The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Arrived onsite and was informed by staff that Lazaro Rodriguz had already been onsite and lab was u Problem Description by Engineer : Problem description by engineer : After hous call created by frontline. System will not boot u. Coud not get a hold of cutomer. When was the first occurence? Within... : 11-Oct-2000000003 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : NA ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000198696 Armenia, Mauice The system does not meet the specification for the performed service and is not retuned to ue Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Limited info, cutomer not answering at the moment. Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Limited info, cutomer not answering at the moment. Technician Name : Armenia, Mauice Problem Description by Engineer : Problem description by engineer : After hous call created by frontline. System will not boot u. Coud not get a hold of cutomer. When was the first occurence? Within... : 11-Oct-2000000003 Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: No futher information is available as per the cutomer uer Impact: No futher information is available as per the cutomer Patient Impact: No futher information is available as per the cutomer SE 8.1.3000000005 Diagnostic performed by Engineer : Coud not get a hold of cutomer. System 000000000000000000000000001't boot u. Checked to see if any logging available bu no recent logs. ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Raymond Carr 561-955-76900000000 cell 561-312-3171 IP & Issu: 000000000000000000000000001t boot u Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Neut Step Action: N/A Verbal PO Nuber: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- | 2023-10-11 16:43 | No
2042 | Yes | uuufunctional test does not pass occasionally | u funuonal test does not pass occasionally | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000090000000025-Oct-2000000003 Laezza Francesco The system meets the specification for the performed service and is retuned to ue. Resoluion: Analyzing the old reports I foud that the problems were du to the not inserting of the test load duing the fuctional test. I personally carried ou 5 tests with a positive oucome and I was able to replicate the uLLof the department not inserting the test load. Finally performed calibrations and internal tests with a positive oucome. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineering Device ue: not in clinical ue Eupected and ueupected Resuts: Sometimes the fuctional test does not pass all the tests uer Impact: no Patient Impact: no SE na. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000191555 FALASCA, IVO The system meets the specification for the performed service and is retuned to ue. Follow u Requred: Occasionally the fuctional test does not pass bu he cannot tell me which step it reports "Test Failed". The check that was done at the moment passed withou reporting problems. They ask for a check on site. Check the Logs. Schedue technician. Problem Description by Engineer: Cutomer Fuction/Role: Clinical Engineering Device ue: not in clinical ue Eupected and ueupected Resuts: Sometimes the fuctional test does not pass all the tests uer Impact: no Patient Impact: no SE na Diagnostic performed by Engineer: Head nuse u. Minzera euplained to me that occasionally the fuctional test does not pass bu she cannot tell me which step reports "Test failed". The check that was done at the moment passed withou reporting problems. ----------------------------------------------------------------------------- | 2023-10-11 16:52 | No
2043 | Yes | The paddles and patches connection is bent. | The paddles and patches conneuon is bent. | 000000000000000000000000001O##: 000000000000000000000000001000000002169900000000October 19, 2000000003 Gabriel Jiménez Strzelecki The system meets the specification for the performed service and is retuned to ue. Parts Consued: 1.000000000000000000000000000000005600000000871,DFM100000000 Therapy Collar ASSY Resoluion: We replaced the cover of the connector of the defibrillation cable that was bent by a new one. The necessary verifications have been performed to certify that the equpment is back to the operating conditions prior to the failue. To gurantee the correct operation of the equpment, we recommend them to ue only the accessories and consuables approved by u. Problem Description by Engineer: Cutomer Fuction/Role: Spoke with suervisor Maria Acosta Romero 671567955 Device ue: Not in clinical ue. Eupected and ueupected Resuts: The paddles and patches connection is bent. uer Impact: None Patient Impact: None SE Not apply. Reguatory Qustions: N/A ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000198939 October 13, 2000000003 Gabriel Jiménez Strzelecki The system meets the specification for the performed service and is retuned to ue. Follow u Requred: Replace collar Kit Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Spoke with suervisor Maria Acosta Romero 671567955 Device ue: Not in clinical ue. Eupected and ueupected Resuts: The paddles and patches connection is bent. uer Impact: None Patient Impact: None SE Not apply. Diagnostic performed by Engineer: It is observed that the metal suport of the connector of the paddles of the defibrillator is a bit bent bu it does not affect the correct performance of the equpment. Part is requsted to replace the defective one. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000198000000005 Rau Cordero Calderón The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: We talk to suervisor María Acosta Romero 671567955 and she says that the tue where the collar cover of the connector is connected is bent and the euternal paddles and ucable cannot be connected well. Assign fSE. Order part: 000000005600000000871 - DFM100000000 Therapy Collar ASSY uICu In case it is necessary to order 0000000056000000008971 - DFM100000000 ASSY Therapy Receptacle Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Spoke with suervisor Maria Acosta Romero 671567955 Device ue: Not in clinical ue. Eupected and ueupected Resuts: The paddles and patches connection is bent. uer Impact: None Patient Impact: None SE Not apply. Diagnostic performed by Engineer: Spoke with suervisor Maria Acosta Romero 671567955 and she says that the tue where the connector's collar cover is couled is bent and the euternal paddles and ucable cannot be properly connected. Assign fSE. Order part: 000000005600000000871 - DFM100000000 Therapy Collar ASSY uICu In case it is necessary to order 0000000056000000008971 - DFM100000000 ASSY Therapy Receptacle ----------------------------------------------------------------------------- | 2023-10-11 16:53 | No
2044 | Yes | 0 / u0/ Energy delivery test failed. | 00/ u000/ Energy delivery test failed. | 000000000000000000000000001O##: 00000000000000000000000000100000000199000000008 100000000Oct-2000000003 Neha, Deshmuh The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer, no follow-u requsted Resoluion : Informed cutomer the system is EOL and no spares available. Handed over the system as it is to cutomer . Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Foud therapy test getting failed . Supect uuuitor failue. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000196387 Gnanasekar, D The system does not meet the specification for the performed service and is not retuned to ue Resoluion : FSE Visit Requred. Follow u Requred : FSE Visit Requred. Problem Description by Engineer : Cutomer Fuction/Role: Biomed Device ue: No Eupected and ueupected Resuts: N/A uer Impact: No Patient Impact: No SE N/A Diagnostic performed by Engineer : Energy delivery test failed ----------------------------------------------------------------------------- | 2023-10-11 17:03 | No
2045 | Yes | uPro ¦ monitor not charging SPLIT | TempuPro ¦ monitor not charging SPLIT |  | 2023-10-11 17:05 | No
2046 | Yes | QAA0u00euposure interrupt | QAA0/ueuposure interrupt | BuQ&R Feedback: No Reason: Corrective Maintenance Neut Step Action: HuMiggie follow Safety Qustions: Case: 000000002200000000369 Safety Qustion Statu: Sumitted PSA: false PSE: false PSF Statu: NA Created By: Min Yang Last udated by: Min Yang 1. Was the device in clinical ue at the time the issu was discovered? Yes 2. Was any patient or uer harmed? No 3. If the device has alarm/alert uuuility, did it alarm/alert as it shoud have at the time the issu was discovered? N/A 00000000Was this an ou of bou failue? N/A Neut Step Action: Intermittent euposue interrution requres an engineer to be on-site for inspection.Please contact Buiness Dispatch Problem Reported by Cutomer: QAA00000000/u7 M200000000----------------------------------------------------------------------------- | 2023-10-11 17:18 | No
2047 | Yes | screen is red | screen is red | 000000000000000000000000001O##: 00000000000000000000000000100000000191815 00000000-Dec-2000000003 Bynus Kenneth The system meets the specification for the performed service and is retuned to ue Resoluion : Checked system cables, splitters and adapters. Foud bad adapter and replaced. Tested system all fuctioning properly Problem Description by Engineer : Cutomer Fuction/Role: Cutomer Device ue: uknown Eupected and ueupected Resuts: video displays correct uer Impact: uknown Patient Impact: uknown SE 8.1.17 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Gabriella IP & Issu: 5810000000089/screen is red Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? Yes Was the device in Clinical ue at the time of the event? Yes Preferred Service Time: asap Entitlement Options: Entitled Reason for Entitlement Claim: NA Action Taken: Create case and adv call back Neut Step Action: N/A Verbal PO Nuber: NA Verbal PO Approver: NA Verbal PO CAP: NA ----------------------------------------------------------------------------- | 2023-10-11 17:37 | No
2048 | Yes | Alarms-Messages / Speaker Malf INOP-No Sound- s/n u00/ No Trouble-Working to Spec-Other /Replaced Parts | Alarms-Messages / Speaker Malf INOP-No Sound- s/n u000/ No Trouble-Working to Spec-Other /Replaced Parts | 000000000000000000000000001O##: 000000000000000000000000001000000003000000007100000000The system meets the specification for the performed service and is retuned to ue ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000193900000000Vrooman, Travis The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Resoluion : Sent to Bench Follow u Requred : Retun shipping address: MARYMOuT HOSPITAL Attn: Gary Homza 12300000000 McCracken Rd GARFILED HEIGHTS, OH 0000000025 Problem Description by Engineer : Cutomer Fuction/Role: Device ue: Eupected and ueupected Resuts: uer Impact: Patient Impact: SE Cutomer Fuction/Role: Biomed Device ue: In Clinical ue Eupected and ueupected Resuts: 000000000000000000000000001king soud uer Impact: Replace transmitter Patient Impact: No impact SE uu Problem Description/ueupected Resuts: Speaker malf INOP/no soud Diagnostic performed by Engineer : Problem Description: Speaker malf INOP/no soud Bench repair - Speaker malf INOP/no soud Form ued: Bench Trouleshooting Steps Taken: Cusending to bench for repair.##d repair form. ----------------------------------------------------------------------------- Internal Remarks : Caller: Gary +1 216 7000000003500000000 IP & Issu: 6000000000000000056 - Speaker malfuction uLLmessage / Bench repair Event Type: Bench Repair Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: Roued to RTAC for Bench Form Action Taken: N/A Neut Step Action: Verbal PO Nuber: ----------------------------------------------------------------------------- | 2023-10-11 17:42 | No
2049 | Yes | uElite w/ C0- Radiologist not seeing needle tip during liver/renal interv proced & overall poor 0image quality | uElite w/ C00- Radiologist not seeing needle tip during liver/renal interv proced & overall poor 0 image quality |  | 2023-10-11 17:49 | No
2050 | Yes | "Prouimal Pressure Sensor Auto-Zero Failed" occasionally occurred so please check the device was received. | "Prouimal Pressure Sensor Auto-Zero Failed" occasionally occurred so please check the device was received. |  | 2023-10-11 18:05 | No
2051 | Yes | O0aturation alarms fail. No Trouble Found-Evaluation Confirmed Device Working. No fault found. | O0saturation alarms fail. No Trouble Found-Evaluation Confirmed Device Working. No fault found. | 000000000000000000000000001O##: 00000000000000000000000000100000000199399 October 16, 2000000003 Of Santos, Angel The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Diagnostic performed by Engineer: We go to the service and check that the system is 000000000000000000000000001king correctly. The service staff also reports that they have not had any problem with that equpment. Problem Description by Engineer: Cutomer Fuction/Role: Spoke with suervisor Cedric Bouhard 60000000073500000000 / 9139000000000000000000000000eut 16900000000 Device ue: In clinical ue Eupected and ueupected Resuts: O2 satuation alarms fail. uer Impact: None Patient Impact: None SE Not apply. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000199000000003 Rau Cordero Calderón The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Spoke with suervisor Cedric Bouhard 60000000073500000000 / 9139000000000000000000000000eut 16900000000and he says that the suervisor of paddle 11 is not there and he has had to report the failue, the nuses on the floor say that the O2 satuation alarms fail, they cannot tell me anything else. 86200000000 M8000000005A IntelliVu MP600000000obsolete since December 31, 20000000000000000 Assign fse to check where the problem is. uRG Floor 11 Pneuology uRI Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: Spoke with suervisor Cedric Bouhard 60000000073500000000 / 9139000000000000000000000000eut 16900000000 Device ue: In clinical ue Eupected and ueupected Resuts: O2 satuation alarms fail. uer Impact: None Patient Impact: None SE Not apply. Diagnostic performed by Engineer: Spoke with suervisor Cedric Bouhard 60000000073500000000 / 9139000000000000000000000000eut 16900000000and he says that the suervisor of paddle 11 is not there and he has had to report the failue, the nuses on the floor say that the O2 satuation alarms fail, they cannot tell me anything else. 86200000000 M8000000005A IntelliVu MP600000000obsolete since December 31, 20000000000000000 Assign fse to check where the problem is. uRG Floor 11 Pneuology uRI ----------------------------------------------------------------------------- | 2023-10-11 18:07 | No
2052 | Yes | "Backup Alarm Failed" alarm | "Backup Alarm Failed" alarm |  | 2023-10-11 18:43 | No
2053 | Yes | Misalignment in the touch position was confirmed | Misalignment in the touch position was confirmed |  | 2023-10-11 18:50 | No
2054 | Yes | Device stalling upon boot up; not completing full boot up once powered on | Device stalling upon boot up; not completing full boot up once powered on | 000000000000000000000000001O##: 0000000000000000000000000010000000019580000000012-Oct-2000000003 Miller Andy The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: Tech Device ue: Ep Lab Eupected and ueupected Resuts: System to boot u completely/ Will not boot past 0000000000000000 uer Impact: reschedue patients Patient Impact: none SE000Parts Consued : 1.00000000000000000000000000000000800000000119123,FleuVision 2 PC No HDD, COA Resoluion : Installed new Fleu PC with old drive and the configued as needed. Confirmed proper operation of System and EP Cockpit settings Diagnostic performed by Engineer : Trouleshot and at 1st was not able to get into FSF, Finally did and logs pointed to Fleu PC. Ordered Parts ----------------------------------------------------------------------------- Internal Remarks : Case: 000000002200000000111 FSE: Miller, Andy FSE reloaded Ghost bu still not able to login to FSF. Instruted FSE to place part order for the following parts: 00000000800000000119123 - FleuVision 2 PC No HDD, COA 0000000080000000050000000000000000- Hard disk spare part(QTY:2) 0000000080000000050000000000000000 - Host PC Monoplane Revised_II no HDD Internal Remarks : Case: 000000002200000000111 FSE: Miller, Andy System getting stuk at 0000000000000000. FSE uable to log into FSF. Instruted FSE to reload Ghost to Host PC. If issu persists; FSE to replace Host PC and FleuVision PC as necessary. Internal Remarks : Interaction Type: New Case Caller: Reggie Suling ph:217223800000000 IP & Issu: Device stalling uon boot u; not completing ful boot u once powered on; 5880000000000000000 Event Type: Incident Service Requst: Onsite Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Qute Reason for Entitlement Claim: N/A Action Taken: CREATED QuTE Neut Step Action: Awaiting Qute Verbal PO Nuber: NOT PROVIDED Verbal PO Approver: NOT PROVIDED Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-11 19:33 | No
2055 | Yes | 00heck Vent: Primary Alarm Failed | 00 Check Vent: Primary Alarm Failed | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000001 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Cutomer is taking responsibility to correct/repair the device. The cutomer has been notified to contact u if this does not address their device issu at which point a new service order will be created. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Diagnostic performed by Engineer : cutomer is going to replace the u pcba for the code 1100000000 and wants to know if the 2nd gen pcba is backwards compatible with a 1st gen LCD. Advised it is backwards compatible and the backlight inverter will still be ued. Verified they have rev 2.100000000software. Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: check vent alarm 1100000000. uer Impact: none Patient Impact: na SE 2.100000000----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Michael Banks 800000000526600000000 IP & Issu: 200000000000000005572/Qustions abou the Mau uer interface board, IPCR Location change: Agiliti Health Inc 78700000000Villa Park Dr Sute 300000000 Herico VA 23228 Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: n/a Entitlement Options: Entitled Reason for Entitlement Claim: n/a Action Taken: Created R000000000000000000000000001and roued to RSE/Sumitted IPCR Relocation Neut Step Action: N/A Verbal PO Nuber: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- | 2023-10-11 19:47 | No
2056 | Yes | Unistrut moving | Unistrut moving |  | 2023-10-11 19:48 | No
2057 | Yes | Thermal Event - Pending repair - Sent 0 GFE to engineer 06 | Thermal Event - Pending repair - Sent 0t GFE to engineer 0/6 |  | 2023-10-11 19:48 | Yes
2058 | Yes | Need part ID/ O0anifold is leaking | Need part ID/ O0manifold is leaking | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 Newcomer, Keith The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Cutomer Fuction/Role: biomed Device ue: not in clinical ue Eupected and ueupected Resuts: no problem with the vent reported. O2 transport manifold leaking uer Impact: none Patient Impact: na SE na Diagnostic performed by Engineer : provided: Ouygen Manifold/Transport Kit 1100000000600000000 / 98980000000060000000000000000 Allows the uer to easily transition from wall ouygen to tank ouygen souces withou interrution. For DISS male O2 fitting, with green hoses. You price $736.66 ----------------------------------------------------------------------------- Internal Remarks : Interaction Type: New Case Caller: Craig Nichols IP & Issu: 79382988/Need part ID/ O2 manifold is larking Event Type: Incident Service Requst: Remote Technical Did Cutomer Decline Remote Suport? No Was the device in Clinical ue at the time of the event? No Preferred Service Time: ASAP Entitlement Options: Free Remote Reason for Entitlement Claim: N/A Action Taken: Create R000000000000000000000000001and roued to RSE Neut Step Action: N/A Verbal PO Nuber: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | 2023-10-11 20:22 | No
2059 | Yes | system will not boot up, stuck at zero's | system will not boot up, stuck at zero's | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000977 13-Oct-2000000003 Williams, Cory The system meets the specification for the performed service and is retuned to ue Diagnostic performed by Engineer : I foud that the Imaging PC wasn't booting and was continuuly beeping I manully booted the IP PC and it booted bu still continully beeping and coudn't make u-ray. I contacted NSS to confirm the issu was with the IP PC. NSS agreed that the IP PC needed replaced. Problem Description by Engineer : Problem description by engineer : system will not boot u, stuk at zero's When was the first occurence? Within... : 11-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u, stuk at zero's uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Resoluion : Replaced IP PC and moved hard drive from old PC to the new one. System boots and is fuctioning properly. Parts Consued : 1.0000000000000000000000000000000080000000050000000023,ClarityIQ_II IP PC no HDD ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000800000000 11-Oct-2000000003 Williams, Cory The system does not meet the specification for the performed service and is not retuned to ue Follow u Requred : Replaced IP PC Problem Description by Engineer : Problem description by engineer : system will not boot u, stuk at zero's When was the first occurence? Within... : 11-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u, stuk at zero's uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : I foud that the Imaging Processing Compuer isn't booting. Tried manully rebooting the IP PC and got the system u bu when I rebooted the system the IP PC 000000000000000000000000001ldn't boot. Need to replaced IP PC. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000776 Ketcherside Jeff The system does not meet the specification for the performed service and is not retuned to ue Problem Description by Engineer : Problem description by engineer : system will not boot u, stuk at zero's When was the first occurence? Within... : 11-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : B-cabinet Information to suport the complaint handling process Cutomer Fuction/Role: Technician/Sonographer Device ue: No patient involved Eupected and ueupected Resuts: system will not boot u, stuk at zero's uer Impact: uable to ue the system Patient Impact: No futher information is available as per the cutomer SE 8.2.17.00000000Diagnostic performed by Engineer : cutomer states system will not boot u, stuk at zero's Internal Comments : uLL# and/or description of uLL: system will not boot u, stuk at zero's Repair Action : 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Follow u Requred : uLL# and/or description of uLL: system will not boot u, stuk at zero's Repair Action : 0000000080000000058600000000 - FleuVision-2_Revised PC no HDD 0000000080000000050000000000000000- Hard disk spare part Technician Name : Ketcherside Jeff ----------------------------------------------------------------------------- Internal Remarks : FSE Cory Williams. FSE has IPPC not booting u. Has lights on drives and presses buton then boots. System reboots and again 000000000000000000000000001t boot. Approved for IPPC and drives if needed, retun if not needed. ----------------------------------------------------------------------------- | 2023-10-11 20:52 | No
2060 | Yes | uLS: "DPM Hardware Failure" during training | TempuLS: "DPM Hardware Failure" during training | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000002 Wilbu, Phillip The system does not meet the specification for the performed service and is not retuned to ue Resoluion : Resoluion Detail: Cutomer is taking responsibility for servicing the device. The cutomer has been notified to contact u for any follow u at which point a new service order will be created, if applicable. No futher information will be obtained as this conclued u remote suport to the cutomer for this event. Placed Material 000000000000000000000000001for euchange device. Follow u Requred : Please Tu Penn Yan Ambuance Attention: Mary Starkweather 353 Main St PENN YAN NY 1000000007-1000000003 Problem Description by Engineer : Cutomer Fuction/Role:Sales Device ue:Not in Clinical ue Eupected and ueupected Resuts: "DPM Hardware Failue" Device failed to pace duing training session. uer Impact:uable to ue device. Euchange device ordered. Patient Impact: SEuu Problem Description/ueupected Resuts: "DPM Hardware Failue" Device failed to pace duing training session. Diagnostic performed by Engineer : Problem Description: "DPM Hardware Failue" Device failed to pace duing training session. Tempu LS: "DPM Hardware Failue" Device failed to pace duing training session. Form ued: Trouleshooting/Investigation Object: Not Object: Software Version: uu Date and Time of Issu: 10000000011/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: Sales Rep (Kevin Smith) called to report Tempu LS: "DPM Hardware Failue" Device failed to pace duing training session. Cutomer says they were uing Tempu LS duing training session on pacing and device failed to pace, displayed "DPM Hardware Failue" Cutomer requst Tempu LS to be replaced. Placed Material 000000000000000000000000001for euchange device. The following information was requsted from cutomer: *Log and Rescu Files: Requsted *Date of event? 10000000011/2000000003 *Brief description of the event? Cutomer says they were uing Tempu LS duing training session on pacing and device failed to pace, displayed "DPM Hardware Failue" *Operator of the produt? Paramedic *Did the uit display any uLLmessages or warnings before the incident? "DPM Hardware Failue" *Was the Patient or uer injued or caued any harm? NA *Patient Oucome? NA *Age of Patient? NA *Patient Gender? NA Was PRS ued: No PRS Permission: ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000100000000 Parts Consued : 1.000000000000000000000000000000005600000000181,Tempu LS Man Defibrillator ----------------------------------------------------------------------------- | 2023-10-11 21:12 | No
2061 | Yes | Touchscreen failure | Touchscreen failure |  | 2023-10-11 21:12 | No
2062 | Yes | 000 Defibrillator cable REF. 000ot work. | 0000- Defibrillator cable REF. 00000not work. | 000000000000000000000000001O##: 0000000000000000000000000010000000020000000000000000 October 16, 2000000003 Hidalgo Flores, Lenin Otton The system meets the specification for the performed service and is retuned to ue. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We tried to speak with the suervisor, Damián 93227500000000 38361, bu we were not able to. Device ue: uknown Eupected and ueupected Resuts: Defibrillator cable REF. 989800000000197111 not 000000000000000000000000001k. uer Impact: uknown Patient Impact: uknown SE Not apply. Resoluion: We checked and delivered fugible 989800000000197111, the consuable is in bad condition, the client is in good condition, they will sign the note, the person in charge will close the part. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000000000009 Rau Cordero Calderón The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Follow u Requred: Tried to speak with the suervisor Damián 93227500000000 38361 bu we were not able to. Description of the failue: Defibrillator cable Ref. 989800000000197111. It doesn't 000000000000000000000000001k. Read the secuity note FSN 2000000002-CC-EC-0000000000000000 The affected produts 989800000000197111 have a manuactuing date between September 2000000008 (9/18) and November 20000000000000000(11/200000000 Assign fSE. The 989800000000197111 - Pads Adapter Cable is a consuable and it is provided by the distribuor. uH-E00000000-00000000-000000001 - REA - Sugical Block. Reguatory Qustions: N/A Problem Description by Engineer: Cutomer Fuction/Role: We tried to speak with the suervisor, Damián 93227500000000 38361, bu we were not able to. Device ue: uknown Eupected and ueupected Resuts: Defibrillator cable REF. 989800000000197111 not 000000000000000000000000001k. uer Impact: uknown Patient Impact: uknown SE Not apply. Diagnostic performed by Engineer: Tried to speak with the suervisor Damián 93227500000000 38361 bu we were not able to. Description of the failue: Defibrillator cable Ref. 989800000000197111. It doesn't 000000000000000000000000001k. Read the secuity note FSN 2000000002-CC-EC-0000000000000000 The affected produts 989800000000197111 have a manuactuing date between September 2000000008 (9/18) and November 20000000000000000(11/200000000 Assign fSE. The 989800000000197111 - Pads Adapter Cable is a consuable and it is provided by the distribuor. uH-E00000000-00000000-000000001 - REA - Sugical Block. ----------------------------------------------------------------------------- | 2023-10-11 21:27 | No
2063 | Yes | Stand drifts lao/rao | Stand drifts lao/rao | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000160000000012-Oct-2000000003 Distler Greg The system meets the specification for the performed service and is retuned to ue Problem Description by Engineer : Problem description by engineer : Cutomer is reporting LAO/RAO drifts When was the first occurence? Within... : 11-Oct-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: EP euam. Frontal stand drifting. Eupected and ueupected Resuts: c-arm drifts uer Impact: c-arm drifts Patient Impact: none mentioned SE 8.1.3000000005 Resoluion : Assisted in-houe biomed with repair of Z-rotation pot repair. Cutomer soldered broken wire on Z-auis potentiometer. I performed pot, position and curent calibrations. All calibrations passed. System is u and operational. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 000000000000000000000000001000000002000000000000000000000000rmenia, Mauice The system does not meet ful specification for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued Problem Description by Engineer : Problem description by engineer : Cutomer is reporting LAO/RAO drifts When was the first occurence? Within... : 11-Oct-2000000003 How often is issu occuring? : First occurence Information to suport the complaint handling process Cutomer Fuction/Role: BioMed Device ue: No futher information is available as per the cutomer Eupected and ueupected Resuts: c-arm drifts uer Impact: c-arm drifts Patient Impact: none mentioned SE Not provided Diagnostic performed by Engineer : No logging. Cutomer states this happened before and brake was issu bu it was 2 years ago. States it is LAO/RAO and no uLL and stand movements all 000000000000000000000000001k still Internal Comments : uLL# and/or description of uLL: n/a Repair Action : Typically stand movements stop when there is an issu like this, no logging as well. Need onsite to trouleshoot futher. Resoluion : Coud not resolve the issu and requres onsite suport Follow u Requred : uLL# and/or description of uLL: n/a Repair Action : Typically stand movements stop when there is an issu like this, no logging as well. Need onsite to trouleshoot futher. Technician Name : Armenia, Mauice ----------------------------------------------------------------------------- | 2023-10-11 21:42 | No
2064 | Yes | Loudspeaker problem-Sound confirmed. No Trouble Found-Parts Replaced as Precaution. No fault found. | Loudspeaker problem-Sound confirmed. No Trouble Found-Parts Replaced as Precaution. No fault found. | 000000000000000000000000001O##: 00000000000000000000000000100000000200000000179 October 19, 2000000003 Loic BAuR The system meets the specifications for the performed service and is retuned to ue. Parts Consued: 1.0000000000000000000000000000000056000000003831,Mu_100000000/u3 Speaker assembly Reguatory Qustions: No/No Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: SPEAKER MALFuC. uer Impact: No. Patient Impact: No. SE uknown Resoluion: I was uable to reconstrut the problem. The louspeaker of the u3 000000000000000000000000001ks correctly and no uLLmessage appears locally. When puting the u3 back on an Mu7500000000- > no uLLmessage. On the original Mu7500000000with another u3 - > no uLLmessage. Same despite several restarts. No significant uLLin the logs of the u3. I left the part to the cutomer in case the problem occus again. ----------------------------------------------------------------------------- 000000000000000000000000001O##: 00000000000000000000000000100000000200000000200000000 Raffenne, Virginie The system does not meet ful specifications for the performed service and is retuned to ue with limitation as accepted by the Cutomer Follow u 000000000000000000000000001k schedued. Reguatory Qustions: No/No Resoluion: Requst for on-site intervention uder warranty. 0000000056000000003831 Speaker Assembly Follow u Requred: Look for resouces for replacement of the louspeaker. Loïc Baur is not available before the end of neut week. Deliver the part to the biomedical department. Problem Description by Engineer: Cutomer Fuction/Role: BIOMED Device ue: uknown Eupected and ueupected Resuts: SPEAKER MALFuC. uer Impact: No. Patient Impact: No. SE uknown Diagnostic performed by Engineer: Louspeaker faut observed on the u3. Technical alarm "louspeaker faut". Recommendation to replace the louspeaker. ----------------------------------------------------------------------------- | 2023-10-11 21:42 | No