## Holdout_Results
model | smote | recall | precision | f1 | pr_auc | threshold
LightGBM | no-smote | 0.9 | 0.081081 | 0.14876 | 0.57775 | 0.0058124
Ensemble-4 | smoteenn | 0.9 | 0.080357 | 0.14754 | 0.61001 | 0.035603
LightGBM | smoteenn | 0.9 | 0.082192 | 0.15063 | 0.4882 | 0.006837
Ensemble-4 | no-smote | 0.9 | 0.082949 | 0.1519 | 0.63491 | 0.040435
DistilBERT | weighted-loss | 0.9 | 0.082949 | 0.1519 | 0.60642 | 0.0037772
XGBoost | no-smote | 0.85 | 0.098266 | 0.17617 | 0.61186 | 0.019184
LR | no-smote | 0.75 | 0.13043 | 0.22222 | 0.19325 | 0.5725
LR | smoteenn | 0.75 | 0.11628 | 0.20134 | 0.18951 | 0.59789
XGBoost | smoteenn | 0.7 | 0.4375 | 0.53846 | 0.5093 | 0.0046239

## Dev_OOF_Results
model | smote | recall | precision | f1 | pr_auc | threshold
LightGBM | smoteenn (dev) | 1 | 0.1 | 0.18182 | 0.88737 | 0.006837
XGBoost | smoteenn (dev) | 1 | 0.63483 | 0.77663 | 0.92035 | 0.0046239
LightGBM | no-smote (OOF) | 0.93805 | 0.1 | 0.18073 | 0.43888 | 0.0058124
DistilBERT | weighted-loss (val) | 0.91304 | 0.1 | 0.18026 | 0.63051 | 0.0037772
XGBoost | no-smote (OOF) | 0.81416 | 0.1 | 0.17812 | 0.43728 | 0.019184
LR | smoteenn (dev) | 0.67257 | 0.1 | 0.17411 | 0.11828 | 0.59789
LR | no-smote (OOF) | 0.56637 | 0.1 | 0.16999 | 0.10091 | 0.5725

## NestedCV_noSMOTE
fold | threshold | recall | f1
1 | 0.0021451 | 1 | 0.1388
2 | 0.0013585 | 1 | 0.12885
3 | 0.0018459 | 1 | 0.14791
4 | 0.0020238 | 0.91304 | 0.12963
5 | 0.0019172 | 1 | 0.14239

## NestedCV_SMOTEENN
fold | threshold | recall | f1
1 | 7.2675e-06 | 1 | 0.063037
2 | 6.4923e-07 | 1 | 0.065714
3 | 1.0213e-06 | 1 | 0.065714
4 | 4.3337e-06 | 1 | 0.065808
5 | 1.0961e-06 | 1 | 0.063037

## Holdout_Predictions
Description 1 | notes_problem_desc | Important Event | true_label | prob_lr | prob_lgbm | prob_bert | prob_ensemble | pred_ensemble
GEH ¦ DFM0 ¦ uleads on patient it shows "learning u unable to u at the critical time | by Engineer : Customer Fuction/Role: Nuse Device ue: DFM Expected and ueupected Results: uable to detect ureadings uer Impact: No | No | 0 | 0.59512 | 0.82446 | 0.0040938 | 0.27593 | 1
alarms have no volume |  | No | 0 | 0.53247 | 0.027381 | 0.0079526 | 0.05117 | 1
No u-ray intermittently |  | No | 0 | 0.48723 | 0.009316 | 0.0090583 | 0.038276 | 1
u/Geometry move problem | : Detailed description of customer problems or symptoms: Customer feedback that there is a problem with the mechanical part, and engineers need to be on site for repairs Is the device itself fauty: Yes Trouleshooting Action: Critical Faut Code or Symptom: POSTItem: Force Sensor Done/Failed &Application uLL ENMV_at_startu_HIGH Remote diagnosis and test resuts: restarting has no effect PRS TR Statu: All Pass Repair Action: Sugested field repair steps: 1. Check whether the force sensor is damaged, and replace it if there is any problem. 2. Check whether the Geo Modue is normal. 3. Check the uLLreport of the machine and perform trouleshooting Sugested spare part uuand name: AD7 FORCE SENSOR & GEO MODuE WP Kit Remarks for customer needs: contact the customer in time to confirm the Onsite time Safety Qustions: Ca | No | 0 | 0.43394 | 0.002915 | 0.0026047 | 0.030195 | 0
00 CN00- Defective speaker |  | No | 0 | 0.53153 | 0.096288 | 0.0044273 | 0.062985 | 1
u reference failed, device shut down | by Engineer : Customer Fuction/Role: biomed Device ue: in ue Expected and ueupected Results:Vent inop while on patient uer Impact: not uable needs repair | No | 0 | 0.44851 | 0.00010422 | 0.0024543 | 0.030283 | 0
Can not boot up | : Overview of customer problems: can't start system Trouleshooting Action: Key faut codes or phenomena: After powering on again after tripping, it cannot be tuned on and stays at : PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: " 1. Enter FS mode and check post trouleshooting 2. Reinstall fluvisison pc software and observe" Recommended spare parts uuand designation: "ppd/ FleuVision-2_Revised PC no HDD/18h to hospital/site1" SAP Transfer: Dispatched for Onsite Suport SAP Transfer: uLLduing processing of BAPI methods: SAP Transfer: Case Transferred to SAP SAP Transfer: Maintain duation to calcuate : SAP Transfer: Maintain duation to calcuate : SAP Transfer: Case Transferred to SAP Neut Step Action: R | No | 0 | 0.33718 | 0.0036498 | 0.0024747 | 0.025818 | 0
is not booting up | by Engineer : Customer Fuction/Role: Biomedical Device ue: No Expected and ueupected Results: System down uer Impact: No | No | 0 | 0.53215 | 0.012583 | 0.0028909 | 0.044659 | 1
Not powering on when the battery is inserted | by Engineer : Customer Fuction/Role: End Customer Device ue: Device was not in ue Customer called and stated that the replacement AED is not powering on when the battery is inserted. The AED does not make a soud when the customer inserts the battery. I spoke With Kelly and he informed that the AED A23B- has never been placed into service and can be replaced with a new AED. Battery SN B Battery install before date29- Expected and ueupected Results: uer Impact: none | No | 0 | 0.43681 | 0.00082333 | 0.0029221 | 0.031665 | 0
button activated during startup | by Engineer: Customer Fuction/Role: n/a Device ue: n/a Expected and ueupected Results: n/a uer Impact: n/a | No | 0 | 0.49718 | 0.0039492 | 0.0028142 | 0.036265 | 0
Live Monitor Display is Blank | by Engineer : Problem description by engineer : Live Monitor Display is Blank When was the first occurence? Within... : 200000000Sep-2000000003 13:: How often is issu occuring? : First occurence Malfuction area : Monitors/Displays Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: Live Monitor Display is Blank uer Impact: No uer impact | No | 0 | 0.39148 | 0.001878 | 0.0025789 | 0.028514 | 0
New AED has no sound or power/DEFOA | by Engineer : Customer Fuction/Role: end uer Device ue: ouside of clinical uage Expected and ueupected Results: Customer puchased new AED and when battery is inserted there is no soud and no power. AED and battery are uder IB warranty and to be replaced at zero cost to the customer. Device serial# A23G- Serial # of battery: B Install before date: uer Impact: none | No | 0 | 0.47031 | 0.00076355 | 0.0025106 | 0.031386 | 0
SPLIT-Control board is damaged-See Reference PR 00 |  | No | 0 | 0.62736 | 0.0066385 | 0.0050709 | 0.046334 | 1
Failing self-test. | by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device u | No | 0 | 0.39914 | 0.00039855 | 0.0027369 | 0.026236 | 0
Footswitch problem | : Detailed description of customer problems or faut symptoms: Cutomers reported that the machine intermittently not emit radiation when the euposue foot switch was pressed, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine intermittently does not emit rays when the euposue foot switch is pressed. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether there is a problem with the foot brake and try to adjut the switch. 2. Check the machine for uLL and perform faut repairs. Recommended spare part uuand name: CV 3p Customer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Ca | No | 0 | 0.36826 | 0.00027619 | 0.003013 | 0.024373 | 0
Unable to discharge. |  | No | 0 | 0.47526 | 0.0075387 | 0.22014 | 0.13273 | 1
QA 000000 No audio alarm - UT00 (0,UT00 patient monitor - Changed speaker |  | No | 0 | 0.60734 | 0.13154 | 0.0061317 | 0.1266 | 1
00uuu Foot Pedal | by Engineer : Problem description by engineer : wireless X-ray foot pedals are not fuctional / lose connectivity Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: Not provided uer Impact: No futher information is available as per the customer | No | 0 | 0.48312 | 0.0023024 | 0.0026327 | 0.033532 | 0
Tube defect | : Overview of customer problems: Tue failue Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: n/a Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Caue 1 Small focu is defect of Frontal X-ray Tue or kV/mA modue is malfuctioning. Soluion 1 Check the filament indutance uing LCR Meter at Frontal u-Ray Tue side. When the filament is open: Youwill measue abou 7.5 mH. 2. If filament indutances are ok, check the HT cables and its contact pins. If the gap at HT pins are not ok, ue pin gage tool to achieve the requred gap. 3. If HT cables are OK bu problem persists, replace kV/mA modue. Soluion 2 HT Cable fauty If filament indutances are ok, check the HT cables and its contact pins. If the gap at HT pins are not ok, ue pin gaue tool to achieve the requred gap. Caue 2 Doserate limit cannot be kept, Q below Qmin. To keep to the dose rate limit, EDL have to go below the Qmin valu. NOTE: This can be caued as collateral of a malfuctioning X-ray tue. (e.g Gridswitch uLL). Then look at that CAT pattern. Soluion 1 Perform Tue adaptation and EDL calibration. Soluion 2 Check EPu data for correct valus. Soluion 3 Check for any filter in X-ray beam. Soluion If problem still euists measue Dose oupu of the X-ray tue. Soluion 5 If to less dose oupu try new X-ray tue. Recommended spare parts uuand designation: uu u rot-GS 10000000000000000nternal Remarks: .1000000002000000003 11:12: uC ( "RDI Accuacy:Yes" $Q7:(Confirmed reported problem, the system fluroscopy and euposue failed, and uLLno X-ray,tue defected" display on data monitor) $Q8:(test the small and large focu filament with LCR meter, they are 7.5mH, the small and large focu filament are open ) $Q9:(replaced the tue and performed the system adjutments) $Q (Test and inspection data refer to PA, performing the cardiac scanning, perfectedimages display on monitors, Device retuned to ful fuctionality) $Q11:Calibrated Tools refer toPA For Mainland China CV / uR, CT engineers only.Please fill in: ^Curent Tue installed date: ^Curent Tue uage: scan second_ Neut Step Action: SS gun jianbin 18600000000299918 Call to inqure abou repairs SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP | No | 0 | 0.24403 | 0.00076087 | 0.0024655 | 0.017355 | 0
DFM0 not getting power- Duplicate of PR 00 |  | No | 0 | 0.6078 | 0.11862 | 0.0046495 | 0.11057 | 1
NZA0leuvision black frame in the screen as with Dos Windows | by Engineer: Customer Fuction/Role: Device ue: ue Expected and ueupected Results: uer Impact: system can stop any sec, black screen on fleuvision System does not start u, it hangs at | No | 0 | 0.46295 | 0.00038365 | 0.002799 | 0.032212 | 0
Lights flickering (ball lights) one section of the light bulb has blown. (0FE SEND 00 | by Engineer : Customer Fuction/Role: RAD Device ue: IN uE Expected and ueupected Results: DEFECTIVE PART uer Impact: NONE | Yes | 1 | 0.59908 | 0.0028476 | 0.0025993 | 0.041818 | 0
QAA0/u FD0Euposure failed intermittently | : Detailed description of the customer's problem or faut phenomenon: Cutomers reported intermittent euposue uLL, requring engineers to come on site for repairs. Is it a faut of the equpment itself? Yes Trouleshooting Action: Key faut code or phenomenon: Euposue not possible Remote diagnosis and test resuts: Restart can restore PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC is normal 2. Check whether the MPD control board is normal 3. Check whether the high-voltage commuication is normal whether the FDC commuication is normal 5. Check the machine uLLand perform faut repairs Customer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Ca | No | 0 | 0.38742 | 0.0072028 | 0.0025439 | 0.028612 | 1
tuns on loads and then freezes | by Engineer : Customer Fuction/Role: Nuse Device ue: Cath Expected and ueupected Results: System not booting uer Impact: system down | No | 0 | 0.50578 | 0.00035282 | 0.0024458 | 0.036363 | 0
Battery faulty | : Customer’s euplanation of the problem ### Battery uLLreported when booting The customer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### na Repair Action: T2 Recommended Field Diagnostic Steps### It is recommended to re-plu the battery and measue the battery voltage. T2 Recommended spare part uuwith description### 989800000000626900000000Assy Battery Li-Ion,11ah,V Customer demand### na Test/Inspection: na Test Equpment: na Euternal Remarks: .2000000003 :57:200000000TC (86000000000000000085) *Department information: [Department: Respiratory Department Contact Person: Teacher Gu Tel: ] * $ASP Engineer: [Chongqing WuBEI Jia ZHOuuLu] $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q(Not applicable) $Q7: (confirmed reported problem, battery failue) $Q8: (Download the "DRPT" report, display uLLcode 11200000000juge it is battery failue) $Q9: (replace battery) $Q (Refer to PA checklist for test and inspection data. The equpment is restored to ful fuction.) $Q11: Not applicable Internal Remarks: .2000000003 :57:26 uC (86000000000000000085) $ASP Engineer: [Chongqing WuBEI Jia ZHOuuLu] $Q1: (No) $Q2: (No) $Q3: (Not applicable) $Q(Not applicable) $Q7: (confirmed reported problem, battery failue) $Q8: (Download the "DRPT" report, display uLLcode 11200000000juge it is battery failue) $Q9: (replace battery) $Q (Refer to PA checklist for test and inspection data. The equpment is restored to ful fuction.) $Q11: Not applicable SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP | No | 0 | 0.36496 | 5.9174e-05 | 0.0024227 | 0.023875 | 0
Screen is broken. | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Confirmed uer Impact: uable to ue | No | 0 | 0.36087 | 0.00014316 | 0.002403 | 0.025179 | 0
The technon ung the detector slipped on his hands and fell on the floor. It doesn't work. | by Engineer : Problem description by engineer : Tecnologo manejando el detector se le resbalo de las manos y cayo al piso. NO fuciona. When was the first occurence? Within... : 23-Oct-2000000003 : Information to suport the complaint handling process Customer Fuction/Role: uer Device ue: No futher information is available as per the customer Expected and ueupected Results: No futher information is available as per the customer uer Impact: Not provided | No | 0 | 0.39969 | 0.0013536 | 0.0041345 | 0.029955 | 0
u/ Fork electricity issue (suited battery failure). |  | No | 0 | 0.50374 | 0.00035997 | 0.0040538 | 0.037273 | 0
Screen froze |  | No | 0 | 0.57047 | 0.032338 | 0.0067699 | 0.050167 | 1
00T0fter power up, nothing comes up on screen, connection to host is lost | by Engineer : Problem description by engineer : after power u, nothing comes u on screen, connection to host is lost When was the first occurence? Within... : 26-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : M-cabinet Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: No patient involved Expected and ueupected Results: after power u, nothing comes u on screen, connection to host is lost uer Impact: uable to ue the system | No | 0 | 0.50012 | 0.0025027 | 0.0024245 | 0.037459 | 0
AED emits single beep, “To test, remove the battery” voice message | by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Expected and ueupected Results: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. | No | 0 | 0.48171 | 0.00027676 | 0.0026378 | 0.032677 | 0
QA#000000uL cardiac conductance line of the device is faulty |  | No | 0 | 0.62707 | 0.30957 | 0.0095496 | 0.13797 | 1
IC resp rec'd ¦ uCVu - machine randomly freeze and shuts down and restarts | by Engineer : Customer Fuction/Role: Sonographer Device ue: Diagnostic Expected and ueupected Results: The system is freezing u and has to forced to reboot. uer Impact: delay in performing euams | No | 0 | 0.48011 | 0.0023288 | 0.002877 | 0.034317 | 0
M0A - S000 - UPS failed | by Engineer : Customer Fuction/Role: Biomed Kelvin Shankle Device ue: In ue, ued for uS protection for net000000000000000000000000001k switch Expected and ueupected Results: uS not power on, main board failue uer Impact: Switch had to be pluged into backu uS util this uS was replaced | No | 0 | 0.48167 | 0.00069109 | 0.0025138 | 0.041014 | 0
DOA HS0 device does not pass first Battery INsertion Test at cuomer, fault teut is: pads not uble |  | No | 0 | 0.67345 | 0.00092234 | 0.0036928 | 0.045102 | 0
LOUDSPEAKER uLLMother board replaced |  | No | 0 | 0.5573 | 0.065414 | 0.0068285 | 0.069427 | 1
Wont boot up has emergency stop |  | No | 0 | 0.52815 | 0.0092168 | 0.0052739 | 0.043836 | 1
EC 00ackup and LED Failed - Pending onsite service scheduled for 00 | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service | No | 0 | 0.4732 | 0.04709 | 0.0027137 | 0.053228 | 1
Wless foot switch does not work | by Engineer: Customer Fuction/Role: MTRA Device ue: yes Expected and ueupected Results: Wifi foot switch defective uer Impact: no | No | 0 | 0.53305 | 0.00012144 | 0.0025212 | 0.037168 | 0
EM: Ac power won't turn on (Ashhibes Municipal Hospital) | by Engineer: Customer Fuction/Role: Fuua Denshi Technical Service Device ue: Not for clinical ue Expected and ueupected Results: Does not start on AC power uer Impact: None | No | 0 | 0.48422 | 0.00013464 | 0.0032877 | 0.032859 | 0
PSAPSE_Legal_Death_Telemety on Wrd 0 GFE0Sent 000-Escalated to Paul 00 |  | Yes | 1 | 0.64566 | 0.9975 | 0.99117 | 0.95713 | 1
post flow sensor cal error | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: flow calibration uLL uer Impact: none | No | 0 | 0.46523 | 0.00056844 | 0.0022839 | 0.034218 | 0
Machine is not booting | by Engineer : Problem description by engineer : Machine is not booting When was the first occurence? Within... : 19-Sep-2000000003 :: How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: Machine is not booting uer Impact: No uer impact | No | 0 | 0.42498 | 0.0009553 | 0.0026137 | 0.029559 | 0
CO0ebreathing risk alarm. | by Engineer: Customer Fuction/Role: BIOMED Device ue: uknown Expected and ueupected Results: uLLCODE LOW LEAK RISK OF CO2 REBREATHING RISK uer Impact: No | No | 0 | 0.46873 | 0.00044438 | 0.0025784 | 0.030822 | 0
uPro ¦ Door on the side is broken off (capnometer door) | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Confirmed uer Impact: uable to ue | No | 0 | 0.32312 | 0.0096464 | 0.0028213 | 0.026897 | 1
Battery not charging |  | No | 0 | 0.51616 | 0.0033932 | 0.0040541 | 0.035641 | 0
Host PC bootup failure SPLIT_ | by Engineer: Customer Fuction/Role: biomedDevice ue: noEupected and ueupected Results: noneuer Impact: none | No | 0 | 0.47146 | 0.00011844 | 0.0032226 | 0.034854 | 0
Warranty replacement Pads | by Engineer : Customer Fuction/Role: End uer Device ue: Device was not in ue Expected and ueupected Results: New AED Pads not being recognized by the AED. Customer inserts the pads into the AED and the aed states to insert pads cartridge. With the pads inserted and the Blu I buton pressed the AED states to insert pads cartridge. Device serial# A00000000b- Serial # of battery Battery lot#: Install before date: Pads lot# c122922- Eup date: uer Impact: none | No | 0 | 0.48396 | 0.0034071 | 0.0026954 | 0.033037 | 0
Not supposing. detector sky plate heavy shock in log. Customer had dropped the detector. | by Engineer : Problem description by engineer : Not euposing, same issu as Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Expected and ueupected Results: Not provided uer Impact: Not provided | No | 0 | 0.41052 | 0.0023513 | 0.0056485 | 0.041881 | 0
Speaker failed. Onsite service pending. |  | No | 0 | 0.51842 | 0.0066634 | 0.006958 | 0.039506 | 1
ultrahuman pressure pipe disconnected. |  | No | 0 | 0.49661 | 0.0060576 | 0.007008 | 0.037857 | 1
System does not fully boot up |  | No | 0 | 0.60057 | 0.010758 | 0.0057793 | 0.047333 | 1
Euternal paddle problem |  | No | 0 | 0.51714 | 0.011977 | 0.010909 | 0.040009 | 1
check mark/enter button is not functional | by Engineer : Customer Fuction/Role: Biomed Device ue: Not in Clinical ue Expected and ueupected Results: V600000000has a fuctional check mark enter buton. uer Impact: V600000000has been removed from service | No | 0 | 0.42001 | 0.00010182 | 0.002351 | 0.030874 | 0
will not stay in standby - Final 00 | by Engineer : Customer Fuction/Role: biomed Device ue: in ue Expected and ueupected Results: will not stay in standby shu down duing case uer Impact: system restricted | No | 0 | 0.40544 | 6.0436e-05 | 0.0025073 | 0.038376 | 0
Direct to sale: Black screen | by Engineer : Customer Fuction/Role: tech Device ue: no Expected and ueupected Results: no uer Impact: no | No | 0 | 0.39307 | 0.0002517 | 0.0027191 | 0.025808 | 0
uPro - Not charging from side or bottom terminals | by Engineer : Customer Fuction/Role: Medical Engineer Device ue: not in ue Expected and ueupected Results: When fitted into the cradle for charging the terminals have failed in the cradle du to poor design and have created a hotspot uer Impact: n/a | No | 0 | 0.46605 | 0.00069921 | 0.0026257 | 0.031614 | 0
AED is emitting the triple chirp |  | No | 0 | 0.48522 | 0.016939 | 0.0083064 | 0.046172 | 1
Geometry movement is partially functioning | by Engineer : Problem description by engineer : RMT - POS: uable to commuicate with GeoIPC 23-Sep-2000000003 12: When was the first occurence? Within... : 23-Sep-2000000003 19:: How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: Not provided uer Impact: NO uer impact | No | 0 | 0.38271 | 0.018787 | 0.0031975 | 0.031131 | 1
The power does not come on. | by Engineer: Customer Fuction/Role: u Kagoshima of Fuda Electronics Device ue: Not ued for clinical ue. Expected and ueupected Results: Loaner uer Impact: None | No | 0 | 0.49756 | 0.00019633 | 0.0025942 | 0.032708 | 0
The system did not start. | by Engineer: Customer Fuction/Role: Radiological engineer Device ue: Not ued Expected and ueupected Results: The system does not start. uer Impact: None | No | 0 | 0.45946 | 0.00061502 | 0.0026237 | 0.030709 | 0
Primary alarm failed |  | No | 0 | 0.51773 | 0.026003 | 0.004092 | 0.045634 | 1
0SureSigns VM 0atient Monitor/no alarm sound. Provided information | : DESCRIPTION OF PROBLEM: There is no alarm soud on the machine HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? na Trouleshooting Action: NOTE uLL# AND / OR DESCRIPTION OF uLL na WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Reconnect the leads and sensors Repair Action: WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Reconnect the leads and sensors RECOMMENDED REPAIR ACTION: No lead disconnection prompt after power on + gudance on reconnecting lead wires and sensors After remote service, the equpment resues normal operation Test/Inspection: na Test Equpment: na | No | 0 | 0.46342 | 0.17061 | 0.9445 | 0.45917 | 1
The equipment does not complete start-up. | by Engineer: Problem description by engineer: The equpment does not finish the start-u sequnce, so it is not possible to access remotely to check uLL; We need to go and check it. Information to suport the complaint handling process Customer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Expected and ueupected Results: The equpment does not complete the start-u sequnce. uer Impact: The equpment does not start. | No | 0 | 0.43682 | 0.00027648 | 0.0025667 | 0.028768 | 0
Shock Not Delivered (Replaced ASSY uuuitance) | by Engineer : Customer Fuction/Role: Biomed Device ue: No Expected and ueupected Results: N/A uer Impact: No | No | 0 | 0.53439 | 0.010303 | 0.0032537 | 0.049514 | 1
uLS Manual ¦ 'Hardware Malfunction' uLLdisplayed on screen during daily check | by Engineer : Customer Fuction/Role:Other: EMS Site Manager Device ue:Not in Clinical ue Expected and ueupected Results: "Hardware Malfuction uLL displayed on screen duing daily check. uer Impact:uable to ue device, euchange ordered | No | 0 | 0.3414 | 0.00097356 | 0.0028362 | 0.025225 | 0
CV u uLLFLUOSTR-IMAGEDS | by Engineer: Customer Fuction/Role: Angio Device ue: Angio Expected and ueupected Results: uknown uer Impact: uknown | No | 0 | 0.50157 | 0.0052117 | 0.0027412 | 0.035111 | 0
0 / u0/ Therapy delivery test failed. | by Engineer : Customer Fuction/Role: Biomed Device ue: No Expected and ueupected Results: N/A uer Impact: No | No | 0 | 0.5248 | 0.00017447 | 0.0026515 | 0.034892 | 0
black screen | : Detailed description of customer problems or faut symptoms: Customer feedback that the large screen has no display Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: No display on the large screen Remote diagnosis and test resuts: The large-screen compuer does not tun on normally, the power pointer is on, and the hard disk indicator light has no display. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Try to connect directly to the uuuhard drive 2. Check whether the software is normal and try to reinstall it 3. Check the machine uLLand perform faut repairs Recommended spare part uuand name: PC Notes on customer requrements: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Safety Qustions: Ca | No | 0 | 0.45021 | 0.00018975 | 0.0026869 | 0.029632 | 0
System not euposing | by Engineer : Problem description by engineer : Need a new footswitch When was the first occurence? Within... : 26-Oct-2000000003 How often is issu occuring? : Intermittent Malfuction area : uer interface Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: Diagnostic Expected and ueupected Results: Footswitch to be good/Footswitch is not - intermittent uer Impact: footswitch is not releasing urays | No | 0 | 0.46713 | 0.00033279 | 0.0027607 | 0.031634 | 0
PSA_ Saturation probes reading low and incorrect RMA Sent 00for return of Sensors |  | No | 0 | 0.64133 | 0.11186 | 0.0096867 | 0.090059 | 1
During PM service, the unit cannot be powered on | by Engineer : Customer Fuction/Role: ASP Device ue: not in ue Expected and ueupected Results: does not power on uer Impact: none | No | 0 | 0.45233 | 0.00018766 | 0.0023767 | 0.031379 | 0
Primary alarm failed. | by Engineer : Customer Fuction/Role: biomed Device ue: ouside of ue Expected and ueupected Results: Primary alarm failed. uer Impact: system restricted | No | 0 | 0.45283 | 0.00053887 | 0.0024247 | 0.029522 | 0
QA 000000 "temperature is too high" alarm shows |  | No | 0 | 0.57819 | 0.068656 | 0.0053276 | 0.064958 | 1
Wireless foot switch losing connectivity | by Engineer : Customer Fuction/Role: none Device ue: none Expected and ueupected Results:none uer Impact:none | No | 0 | 0.5534 | 0.00029221 | 0.0028277 | 0.036797 | 0
Alarms/No sound: There's no audio alarms coming from the critical care rooms: No malfunction. Loose Connection. ur. | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: Visul and auible alarms. uer Impact: None | No | 0 | 0.48362 | 0.11545 | 0.018337 | 0.093941 | 1
hot case B:0/u00no u-ray | : Overview of customer problems: No X-ray Trouleshooting Action: Key faut codes or phenomena: The patient is on the stage and no radiation can be emitted. The customer did not provide detailed uLLinformation and said that he had contacted the engineer. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: " 1. Go to the scene and check the event log to find ou whether it is caued by the en_u signal or other reasons. 2. The customer eupressed dissatisfaction with the equpment, saying that there were many repair requsts, it was ustable, and it needed detailed inspection." Euternal Remarks: .2000000003 :38 uC (860000000063600000000Service content (requred): Cannot be euposed; replace the uSC modue in the high voltage generator. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: none. Internal Remarks: .2000000003 : ( $Q7: (Confirmed reported problem, no X-ray, system prompt "X-ray not possible. Tue problem. Contact technical suport." ) $Q8: (Check system log, foud message "Coud not access X-ray generator (ubufered) via uRayGeneration service"; uSC has problem, need to be replaced. ) $Q9: (Replace uSC Certeray(uu).) $Q (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11:Calibrated Tools: Refer to PA For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage: N/A T2 Activities:Repair Action: ppd/ uSC Certeray/200000000to hospital/site1 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 1. Shipped LiuChao-86-13979783177 SAP Transfer: Case Transferred to SAP Euternal Remarks: .2000000003 :13: (860000000063600000000SuT. Ready for Clinical ue local time" : " 2000000003- 17:" | No | 0 | 0.36785 | 0.00038726 | 0.0031267 | 0.024689 | 0
[Q] Having uLLcodes and had an emergency yesterday with patient on the table . Previoucase# 000 | by Engineer : Customer Fuction/Role: FSE Device ue: yes Expected and ueupected Results: tue curent ou of spec uer Impact: no | No | 0 | 0.3701 | 0.1271 | 0.98636 | 0.46289 | 1
Failing prouimal pressure sensor auto zero, GFE00EC0*** | by Engineer : Customer Fuction/Role: biomed Device ue: in ue Expected and ueupected Results:Prouimal pressue sensor auozero failed alarm uer Impact: removed from service | No | 0 | 0.48137 | 0.00066533 | 0.0024972 | 0.047151 | 0
QACASE 000000After uLLis connected to ouygen, air leakage occurs inside the body-GFE 0/6 |  | No | 0 | 0.66315 | 0.15463 | 0.021995 | 0.10183 | 1
uuudefective |  | No | 0 | 0.52232 | 0.020857 | 0.0066456 | 0.048479 | 1
Equipment power failure |  | No | 0 | 0.56432 | 0.05125 | 0.0052579 | 0.05463 | 1
0/u FD0/image space low | : Detailed description of customer problems or faut symptoms: Cutomers reported that image storage space is low and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Free disk space is too low. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Euecue Database Consistency Test and Repair, delete the Repaired patient, and see if the space is restored. 2. If there is no effect, euecue Recreate Image Store. 3. If there is no effect, try Bypass hard disk, and then reinstall the IPPC software. there is no effect, ue PC. Diagnostic Tool performs diagnostic tests and replaces corresponding spare parts 5. Check machine uLLreports and perform faut repairs Recommended spare part uuand name: " IP PC Revised_II no HDD & disk spare part " Customer demand notes: Contact cutomers promptly to confirm Onsite time Internal Remarks: 31.1000000002000000003 1000000005:25 uC () $Q1: (Y) $Q2: (N) $Q3: (NA) $Q(NA) $Q5: (NA) $Q6:（Image cannot be saved ) $Q7: (Confirmed reported problem, Image cannot be saved) $Q8: (image disk damaged, cannot repair) $Q9: (replaced the spare part, tun normal) $Q (Test and Inspection data refer to PA. Device retuned to ful fuctionality.) | No | 0 | 0.38466 | 0.0020104 | 0.0026786 | 0.027342 | 0
Low Tidal volume | by Engineer : Customer Fuction/Role: biomed Device ue: in ue Expected and ueupected Results:leak and alarms uer Impact: not uable needs repair | No | 0 | 0.42876 | 0.0023803 | 0.0025685 | 0.030167 | 0
SPLIT-uDisable-See Reference case PR 00 |  | No | 0 | 0.60029 | 0.01688 | 0.0056726 | 0.047388 | 1
M0A - u0 - No longer allowing shock of 0J |  | No | 0 | 0.59819 | 0.96086 | 0.030459 | 0.45826 | 1
device beeping with good battery | by Engineer : Customer Fuction/Role: Admin Device ue: no Expected and ueupected Results: device beeping with good battery uer Impact: none | No | 0 | 0.50914 | 0.0014275 | 0.0025938 | 0.034505 | 0
Audio failure. Can Not Duplicate-No Trouble Found. No fault found. | by Engineer: Customer Fuction/Role: Biomed Engineer Device ue: Clinical uage Expected and ueupected Results: N/A uer Impact: No uer impact | No | 0 | 0.51617 | 0.0043769 | 0.0028747 | 0.037876 | 0
Prouimal sensor auto zero failure | by Engineer : Customer Fuction/Role: biomed Device ue: ouside of ue Expected and ueupected Results: Prouimal sensor auo zero failue uer Impact: system restricted | No | 0 | 0.47267 | 0.00043653 | 0.0024677 | 0.035188 | 0
EC 00 Backup alarm failed | by Engineer : Customer Fuction/Role: biomed Device ue: uit set u Expected and ueupected Results:uit alarming uer Impact: not uable needs repair | No | 0 | 0.41597 | 0.00058078 | 0.0024317 | 0.02749 | 0
HS0ED is emitting the triple chirp (Removing the AED from service) | by Engineer : Customer Fuction/Role: Admin Device ue: N/A Expected and ueupected Results: Customer reported that they changed the battery and pads when they did the self-test AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE | No | 0 | 0.51484 | 0.002286 | 0.0023515 | 0.035408 | 0
Failed Testing .. | by Engineer: Customer Fuction/Role: clinical ing Device ue: uuuplate test Expected and ueupected Results: fauty plate houing uer Impact: no impact | No | 0 | 0.38995 | 9.3987e-05 | 0.0025055 | 0.030894 | 0
The system show message “Generator is bu starting up, No u-ray is possible.” | by Engineer : Customer Fuction/Role: Technician. Device ue: In clinical ue. Expected and ueupected Results: The system show message “Generator is buy starting u, No X-ray is possible.” uer Impact: No impact. | No | 0 | 0.59262 | 0.0064691 | 0.0026155 | 0.048038 | 1
O0ensors not working - Pending bench for repair 05 |  | No | 0 | 0.53937 | 0.010296 | 0.0079416 | 0.042657 | 1
C arc no longer moves into the cranial angulated positions | by Engineer: Problem description by engineer : The reported uLLis confirmed Information to suport the complaint handling process Customer Fuction/Role: Not provided Device ue: Not provided Expected and ueupected Results: Not provided uer Impact: Not provided | No | 0 | 0.18574 | 0.0017974 | 0.011404 | 0.021055 | 0
0/u FD0Unable to start the machine |  | No | 0 | 0.48028 | 0.0050857 | 0.0061416 | 0.034816 | 0
DPM hardware failure. | by Engineer : Customer Fuction/Role: Other: Sales rep Device ue: Not in Clinical ue Expected and ueupected Results: uer Impact: uable to ue device. | No | 0 | 0.3261 | 0.012136 | 0.0028943 | 0.026463 | 1
Questioning the management of supraventricular tachycardias by the SMART ANALYSE software |  | Yes | 1 | 0.63019 | 0.19048 | 0.96885 | 0.49248 | 1
AED Warranty Replacement | by Engineer : Customer Fuction/Role: End Customer Device ue: Device was not in ue Expected and ueupected Results: Customer called and stated that the AED is not properly. The AED states instrutions for uing the AED pads on a patient when inserted. The customer informed that these are brand when inserted. The pads are uder warranty and have been replaced in case . When the customer received the pads replacement, we spoke and foud that the AED still consistent with immediately stating instrutions for placing pads on the patient. The new set of pads have not been opened and the AED is not operating properly. This is an indication that the AED is the issu. The AED is uder warranty and to be replaced at zero cost to the customer Device serial# A19a-10000000038 uer Impact: none | No | 0 | 0.34643 | 3.3122e-05 | 0.0028989 | 0.02403 | 0
Alarm problem | by Engineer: Customer Fuction/Role: Biomed Device ue: uknown Expected and ueupected Results: uknown uer Impact: uknown | No | 0 | 0.50383 | 0.0010753 | 0.002784 | 0.033278 | 0
Death:Alarms/No Sound: No sound on the monitor in the nurses station:No Malfunction/ur uLLReplaced Euternal Speaker | : Not getting auio alarm tone Not getting auio alarm tone Form ued: Trouleshooting/Investigation Object: N2OBSSuV00000000C Not Object: The other suveillance stations in other floors Software Version: C.. Date and Time of Issu: 1000000006/2000000003 uer Actions: Where is the Deviation: Trouleshooting Steps Taken: - Customer says the patient in room 200000000was in alarm condition aroud 9PM last night bu there was no soud at the central station and the patient eupired - Customer says the central station showed the visul alarm bu there was no soud - Customer does not know what the alarm was - Patient was on an Mu - Customer did not have the Mu00000000in qustion - I had customer simuate an red and yellow alarm on another Mu00000000and there was no soud - The alarm volue is set to 100000000on the suveillance station - The PC is at the nusing station desk - I had customer check the euternal speaker and there is no speaker present - Customer got a spare speaker and connected it to the PC and can now hear the soud Was PRS ued: No PRS Permission: Problem Description by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: Hear alarm tone uer Impact: Not getting auio alarm tone | Yes | 1 | 0.47892 | 0.24051 | 0.14677 | 0.30314 | 1
u / Operational Test Fail with therapy disable uLL | by Engineer : Customer Fuction/Role: Bme Device ue: no Expected and ueupected Results:na uer Impact: no | No | 0 | 0.46173 | 6.8093e-05 | 0.0026259 | 0.030436 | 0
battery does not charge | by Engineer : Customer Fuction/Role: Biomed Device ue: Ouside of ue Expected and ueupected Results: Pass uer Impact: N/A | No | 0 | 0.35561 | 4.3341e-05 | 0.0024159 | 0.027713 | 0
V0 - High tidal volume |  | No | 0 | 0.5856 | 0.18542 | 0.0061365 | 0.10392 | 1
Failing self test |  | No | 0 | 0.48426 | 0.0028701 | 0.0052816 | 0.033519 | 0
Accept button is not working and will not allow him to access the service menu. | by Engineer : Customer Fuction/Role: biomed Device ue: ouside of ue Expected and ueupected Results: accept buton not uer Impact: system restricted | No | 0 | 0.54113 | 0.0035575 | 0.0026674 | 0.042944 | 0
M0A / u00/ The defibrillator does not turn on anymore, they tested the battery and it works. |  | No | 0 | 0.65745 | 0.032059 | 0.0052387 | 0.054492 | 1
The device did not power on |  | No | 0 | 0.59474 | 0.021734 | 0.0073668 | 0.064139 | 1
u - unable to get tracing |  | Yes | 1 | 0.5765 | 0.9416 | 0.023978 | 0.46838 | 1
M0A / u0 / Cardioverter doesn't recognize the blades. | : DESCRIção DO PROBLEMA: Cardioverter doesn't recognize the blades WHEN WAS THE FIRST OCCuRENCE? HA... Week HOW FREQuNT IS THE PROBLEM? First Occurence WHERE IN THE PROCESS DOES THE PROBLEM OCCu? Pre-Acqustion/Scan WHAT RECENT EVENTS HAVE OCCuRED? Duing Normal ue CuTOMER HAS CREATED A SYSTEM-SPECIFIC REPORT FOR THE ISSu? No BACKu PHONE IF DIFFERENT FROM ONEEMS: No Trouleshooting Action: NOTE # uLLAND/OR DESCRIPTION OF THE PROBLEM: Cardioverter doesn't recognize the blades WHAT ARE THE STEPS FOLLOWED DuING DIAGNOSIS? IN WHAT ARE THE NEuT TROuLESHOOTING STEPS FOR THE ISSu? IN Estimated downtime: Repair Action: IF NECESSARY PART(S) - 12 NC RECOMMENDED : N/A Recommended Repair Action(s) IN RECOMMENDED ACTIONS FOR THE ESF PRIOR TO THE VISIT: N/A SAP Transfer: Dispatched for Onsite Suport Safety Qustions: Ca | No | 0 | 0.31279 | 0.00015196 | 0.002792 | 0.025202 | 0
Blower motor stalled message | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: blower stalled uer Impact: none | No | 0 | 0.46311 | 0.0021839 | 0.0023923 | 0.031734 | 0
System is not powering on | by Engineer : Customer Fuction/Role: GE service Device ue: clinical Expected and ueupected Results: will not power up uer Impact: same | No | 0 | 0.47193 | 0.0001829 | 0.0027895 | 0.03163 | 0
The machine remains stuck at start-up with the countdown at 0:0. | by Engineer: Customer Fuction/Role: TRM Device ue: No. Expected and ueupected Results: No boot uer Impact: No. | No | 0 | 0.51483 | 0.010069 | 0.0027117 | 0.038545 | 1
SPLIT - Unit does not get past the self-test (Ref 00 |  | No | 0 | 0.60034 | 0.00056831 | 0.0054642 | 0.042399 | 0
the host computer crashed is not coming up | by Engineer : Customer Fuction/Role: FSE Device ue: no Expected and ueupected Results: host pc will not power u blowing F11 fue uer Impact: no | No | 0 | 0.4811 | 0.00016104 | 0.0027094 | 0.035737 | 0
system is not working, not getting video. May be hard drive issue. | by Engineer : Customer Fuction/Role: GE FSE Device ue: u uuuFD Expected and ueupected Results: System will not boot u uer Impact: none | No | 0 | 0.25305 | 0.0031431 | 0.0026627 | 0.023192 | 0
Legal Issue_AE-000 u- failed to provide utracing. Improved after replacing with alternate electrodes. |  | No | 0 | 0.64805 | 0.87875 | 0.91483 | 0.83613 | 1
No geometry movement |  | No | 0 | 0.47871 | 0.011838 | 0.012161 | 0.041204 | 1
Speaker malfunction-Sound not confirmed. Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. | by Engineer : Customer Fuction/Role: Biomedical Engineer Device ue: Ouside of ue Expected and ueupected Results: Speaker malfuction occuring. uer Impact: Low | No | 0 | 0.51734 | 0.00041903 | 0.0026739 | 0.042156 | 0
PSA_PSE_ISCV Contains a bug which delaying archiving of images to PACS - Cuomer has reported a patient has passed away |  | Yes | 1 | 0.69827 | 0.56747 | 0.99123 | 0.72485 | 1
Machine is automatically shutting down | by Engineer : Problem description by engineer : Machine is auomatically shuting down When was the first occurence? Within... : 16:: How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: Machine is auomatically shuting down uer Impact: No uer impact | No | 0 | 0.24686 | 7.5597e-05 | 0.0028617 | 0.016638 | 0
Touchscreen is nonresponsive. |  | No | 0 | 0.57443 | 0.68744 | 0.013455 | 0.24052 | 1
System cannot be started | by Engineer: Customer Fuction/Role: Heidelberg uiversity Hospital Device ue: u M12 Expected and ueupected Results: uer Impact: | No | 0 | 0.38908 | 0.00052854 | 0.0025094 | 0.025818 | 0
Wireless footswitch defective | by Engineer : Customer Fuction/Role: uknown Device ue: uknown Expected and ueupected Results: follow u for FCO uer Impact: no | No | 0 | 0.46971 | 8.072e-05 | 0.0026501 | 0.031675 | 0
M0A / u00 The electrode plate suddenly sparks all over the place (PSA_QA#00000). |  | No | 0 | 0.65151 | 0.11106 | 0.0059861 | 0.14033 | 1
0/u FD0Host PC cannot power on | : Overview of customer problem: Hos tPC device cannot be tuned on. Detailed description of customer problem or faut phenomenon: Hos tPC device cannot be tuned on. Is it a faut of the device itself? Yes Trouleshooting Action: Key faut codes or phenomena: The Hos tPC device cannot be powered on. Remote diagnosis and test resuts: The Hos tPC device cannot be powered on. Check that the power suply is normal. It is supected that the hostpc is fauty. PRS TR statu: All pass Repair Action: Recommended on-site repair steps: 1. Check the HostPC power suply, 2. Reformat the HostPC OS disk and reinstall the system, 3. Replace the host PC. Recommended spare parts uuand name: " Host PC Monoplane Revised_II no HDD" Customer demand Remarks: Onsite Problem Reported by Customer: The system cannot be started Safety Qustions: Ca | No | 0 | 0.36653 | 0.00070841 | 0.0023755 | 0.024132 | 0
00 CN00- Red u and Beep - Ran OPs Check to Resolve |  | No | 0 | 0.48343 | 0.002497 | 0.0054196 | 0.033152 | 0
Failing self test |  | No | 0 | 0.48704 | 0.0023918 | 0.0052816 | 0.034021 | 0
The installed udoes not pass the self-test |  | No | 0 | 0.50841 | 0.0067317 | 0.0027247 | 0.035086 | 1
LS - 00 | by Engineer : Customer Fuction/Role: in-houe imaging tech Device ue: first boot of the day Expected and ueupected Results: system did not boot u uer Impact: inconvenience | No | 0 | 0.4458 | 0.0004208 | 0.0024218 | 0.029047 | 0
The unit is giving uLLmessage, "speaker malfunctioning"- Unable to confirm cau-Cuomer resolved | by Engineer : Customer Fuction/Role: Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.41609 | 0.011315 | 0.0026262 | 0.037188 | 1
No uray image displayed on fleuvision. | by Engineer : Customer Fuction/Role: Technologist Device ue: uknown Expected and ueupected Results: Live image not available on fleuvision monitor. uer Impact: Delay | No | 0 | 0.44323 | 0.0010977 | 0.0028257 | 0.029955 | 0
00he device is not going through the test. | by Engineer: Customer Fuction/Role: Biomed. Device ue: Not ue. Expected and ueupected Results: The device does not pass the test. uer Impact: No. | No | 0 | 0.51636 | 0.00041216 | 0.0029287 | 0.034345 | 0
Sky Plate dropped ( calibration unsuccessful replacement needed) | by Engineer : Customer Fuction/Role: an Device u | No | 0 | 0.54989 | 0.0046614 | 0.0029511 | 0.039364 | 0
Device is switching between ac power to internal power simultaneouy | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: uit is switching from AC to battery uer Impact: none | No | 0 | 0.51509 | 0.00027674 | 0.0024377 | 0.033363 | 0
uLLcodes 00000, 0 |  | No | 0 | 0.56549 | 0.050494 | 0.006071 | 0.054079 | 1
Equipment alarm | : Customer’s euplanation of the problem: Cutomers reported that the device kept alarming The customer observed the uLLmessage displayed on the monitor: NA Trouleshooting Action: The customer observed the uLLmessage displayed on the monitor: NA T2 Remote Diagnostics and Testing: Replacement of electrode bou and battery bu ineffective Repair Action: T2 Recommended on-site repair steps: It is recommended to check the host Recommended spare part uufor T2 with description: # Notes on customer requrements: NA Test/Inspection: NA Test Equpment: NA | No | 0 | 0.37423 | 0.00014496 | 0.0025866 | 0.027419 | 0
Unable to charge |  | No | 0 | 0.49944 | 0.0029729 | 0.0060407 | 0.049029 | 0
Fluoro storage was not possible | by Engineer : Problem description by engineer : remote alert of CV u: uLLFLuSTR-IMAGEDS The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog: The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog:The following alarm has been generated by asset CV_722000000002-17100000000Alarm name=CV u: uLLFLuSTR-IMAGEDS, Description=uLLode=uLLFLuSTR-IMAGEDSK Abstract=EventLog: Flu storage not poss. Image disk problem Image disk problem When was the first occurence? Within... : 31-Oct-2000000003 How often is issu occuring? : Mutiple times per month Malfuction area : Image processing Information to suport the complaint handling process Customer Fuction/Role: M2M/Radar Device ue: No futher information is available as per the customer Expected and ueupected Results: No futher information is available as per the customer uer Impact: Flu storage not poss. | No | 0 | 0.21919 | 0.00086803 | 0.0025049 | 0.014978 | 0
Defibrillator still shows an uLLafter inserting new electrodes and new battery | by Engineer: Customer Fuction/Role: uknown. Device ue: None. Expected and ueupected Results: bench Repair. uer Impact: None. | No | 0 | 0.55464 | 0.015072 | 0.0028671 | 0.041193 | 1
Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 Other HW-Speaker-HW / Replaced Parts |  | No | 0 | 0.62884 | 0.0035598 | 0.0035948 | 0.042993 | 0
Display is getting blurred | by Engineer : Customer Fuction/Role: Bme Device ue: yes Expected and ueupected Results: na uer Impact: no | No | 0 | 0.5457 | 0.0038711 | 0.0030272 | 0.050433 | 0
Battery defect mu be replaced - GFE 0/8 |  | No | 0 | 0.55685 | 0.044018 | 0.0037137 | 0.049566 | 1
The wireless footswitch is still not operational - Medimover Storage, Burrfields Road, P00 | by Engineer : Customer Fuction/Role: Rad Device ue: Expected and ueupected Results: The wireless footswitch is still not operational - Medimover Storage, Burfields Road, P 5NN uer Impact: | No | 0 | 0.552 | 0.00387 | 0.0024483 | 0.038263 | 0
Wireless footswitch does not work, they want to connect another one that they have, question how this works | by Engineer: Customer Fuction/Role: biomed Device ue: y Expected and ueupected Results: footswitch defect uer Impact: low | No | 0 | 0.63755 | 0.0008999 | 0.0026868 | 0.044193 | 0
The three-beep sounded - | by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device u | No | 0 | 0.40768 | 0.0014569 | 0.0028139 | 0.027764 | 0
Getting alarm , AUu alarm, supply failed | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: V600000000not display uLLcodes uer Impact: V600000000has been removed from service | No | 0 | 0.27391 | 0.00094379 | 0.0025486 | 0.031707 | 0
Won't charge battery although it will run on battery power | : 't charge battery althouh it will ru on battery power 't charge battery althouh it will ru on battery power Form ued: Bench Trouleshooting Steps Taken: Received Teams message from NSS - Travis Haines asking me if I suport the Tempu Pro monitor, informing me that 3rd party F | No | 0 | 0.38968 | 0.00093441 | 0.0025308 | 0.027305 | 0
u - It beeps all the time. | by Engineer: Customer Fuction/Role: Biomed Engineer Device ue: Clinical uage Expected and ueupected Results: N/A uer Impact: No uer impact | No | 0 | 0.48464 | 0.0048735 | 0.0025825 | 0.034779 | 0
QACASE for PCI imported product#000000low tidal volume may cau repeated inhalation of CO0FE 0/6 |  | No | 0 | 0.67164 | 0.10046 | 0.0081321 | 0.090759 | 1
Main monitor froze and started working after reboot | by Engineer : Customer Fuction/Role: rad tech Device ue: none Expected and ueupected Results: not booting uer Impact: none | No | 0 | 0.44064 | 0.010498 | 0.0029913 | 0.038649 | 1
[Q] Wireless footswitch not operational | by Engineer : Customer Fuction/Role: Device ue: Expected and ueupected Results: replace footswitch du to lost connection uer Impact: | No | 0 | 0.4142 | 0.0024285 | 0.0027511 | 0.030628 | 0
Speaker Issue/Malfunction-No Speaker sound for alarm-Device-other-HW/Replacement sent | by Engineer : Customer Fuction/Role: Biomed/ clinical engineering Device ue: in ue Expected and ueupected Results: speaker malfuction displayed on the device uer Impact: no impact, device is sent ou for euchange | No | 0 | 0.5471 | 0.00047792 | 0.0024755 | 0.036354 | 0
AED will not go through self-test | by Engineer : Customer Fuction/Role: Admin Device ue: N/A Expected and ueupected Results: Customer reported that AED will not go throuh self-test uer Impact: NONE | No | 0 | 0.36267 | 0.00046301 | 0.0029281 | 0.025179 | 0
Unit alarming for "Prouimal Line Disconnect". | by Engineer : Customer Fuction/Role: RT Device ue: Set u - not on patient Expected and ueupected Results: uit alarming for "Prouimal Line Disconnect". uer Impact: none | No | 0 | 0.50057 | 0.00011223 | 0.0028987 | 0.035005 | 0
Tidal Volume Issue - in progress / onsite unscheduled | by Engineer : Customer Fuction/Role: biomed Device ue: on patient Expected and ueupected Results: not reaching set volues uer Impact: swapped vent | No | 0 | 0.46679 | 0.0021707 | 0.0030784 | 0.045035 | 0
Device alarmed due to no Ouygen supply | by Engineer : Customer Fuction/Role: biomedical engineer Device ue: ou of ue Expected and ueupected Results: V600000000Displaying 'No Ouygen Suply' uer Impact: nil | No | 0 | 0.47887 | 0.0015938 | 0.0023983 | 0.034309 | 0
00N0RM will not u-ray, generator cabinet will not come on. | by Engineer : Problem description by engineer : Frontal Generator is not booting u properly. When was the first occurence? Within... : 200000000ct-2000000003 How often is issu occuring? : First occurence Malfuction area : X-ray generation Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: Treatment Expected and ueupected Results: No futher information is available as per the customer uer Impact: No futher information is available as per the customer | No | 0 | 0.14563 | 0.0017235 | 0.0026188 | 0.013514 | 0
Screen is frozen. | by Engineer : Problem description by engineer : System not booting When was the first occurence? Within... : 11-Oct-2000000003 15:10000000056 How often is issu occuring? : First occurence Malfuction area : Others Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: No patient involved Expected and ueupected Results: System not booting uer Impact: No futher information is available as per the customer | No | 0 | 0.23521 | 0.00046299 | 0.0026854 | 0.016211 | 0
Power cord fails EST |  | No | 0 | 0.53047 | 0.024262 | 0.0045386 | 0.043298 | 1
SA:0/u FD00System cannot boot up | : Overview of customer problem: The device couts down to : and cannot enter the system normally. I restarted it three times, bu it still doesn't .Engineers are requred to be present. Is the faut caued by the equpment itself? Yes Trouleshooting Action: Remote diagnosis and test resuts: The device couts down to : and cannot enter the system normally. I restarted it three times, bu it still doesn't .Requres engineer presence PRS TR Statu: No PRS Repair Action: Recommended on-site repair steps: Check whether IPPC or IPC is not tuned on, check whether the PC's power suply or PC's BIOS battery is discharged Customer demand notes: Onsite Problem Reported by Customer: SA:/u uuuFD100000000100000000Large screen display problem Euternal Remarks: .11.2000000003 : uC (8600000000000000001)"SuT. Ready for Clinical ue local time" : " 2000000003-11- 17: [Service Content] Large-screen PC software crashes, reinstall the system and software and retun to normal. 【Service Resut】The machine operates normally after repair. 【Other matters】None yet Safety Qustions: Ca | No | 0 | 0.42122 | 0.0051731 | 0.0024815 | 0.029093 | 0
The instrument cannot be turned on. |  | No | 0 | 0.59112 | 0.50784 | 0.0084866 | 0.19706 | 1
SPLIT - 00- Paddle cable issue |  | No | 0 | 0.50633 | 0.00077304 | 0.0050394 | 0.034841 | 0
Proper alarm did not go off when mask was disconnected | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: V600000000to alarm for mask disconnect uer Impact: V600000000has been removed from service | No | 0 | 0.39522 | 0.022191 | 0.0039676 | 0.048876 | 1
Unable to Fluoro. |  | No | 0 | 0.41554 | 0.060191 | 0.011191 | 0.04845 | 1
Touchscreen not responding | by Engineer : Customer Fuction/Role: Technician Device ue: Duing PM Expected and ueupected Results: system touh screen not coud not perform PM uer Impact: Cannot monitor | No | 0 | 0.5001 | 0.0020186 | 0.0024564 | 0.034246 | 0
u- The defibrillator plate corrosion was found in regular testing, |  | No | 0 | 0.61286 | 0.35187 | 0.11756 | 0.1939 | 1
Device Fails Self-Test | by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Expected and ueupected Results: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. | No | 0 | 0.38105 | 0.00079014 | 0.0024531 | 0.025096 | 0
The tidal volume is too high |  | No | 0 | 0.58606 | 0.17591 | 0.011352 | 0.10725 | 1
QACASE_for PCI imported product#000000he focuwas faulty, and image blurr |  | No | 0 | 0.63174 | 0.0079712 | 0.0060828 | 0.04737 | 1
When turning on the equipment I tried to turn it on but it turned off. Diagnosis in description |  | No | 0 | 0.63683 | 0.61676 | 0.071639 | 0.27237 | 1
Device not turn off, and will restart by itself, GFE00OV0*** | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: uit powers on by itself uer Impact: none | No | 0 | 0.44164 | 0.010993 | 0.002429 | 0.039507 | 1
QA000000 - Device cannot turn on when rescuing the patient |  | Yes | 1 | 0.5847 | 0.71801 | 0.9896 | 0.75824 | 1
Capno Not Working | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Confiimed uer Impact: uable to ue | No | 0 | 0.41019 | 0.0021732 | 0.0026244 | 0.029385 | 0
Quality inspection ufailed |  | No | 0 | 0.46488 | 0.0088547 | 0.004777 | 0.035068 | 1
No alarms were activate; they were eupecting = Pending TI- Letter- READY FOR SUMMARY 00 |  | No | 0 | 0.65206 | 0.4716 | 0.0147 | 0.21266 | 1
The system stopped during a photon. | by Engineer: The system stopped duing a 3-bee-shot. u/Ms. Sada hiro uhin: 7722 ・Duing the operation of a patient weighing 1600000000g between 9: and yesterday (November 6). Some message was displayed. He like u to check the log. Customer Fuction/Role: Radiology Department Device ue: Duing euamination Expected and ueupected Results: The system stopped duing a 3-beg. u/Ms. Sada hiro uhin: 7722 ・Duing the operation of a patient weighing 1600000000g between 9: and yesterday (November 6). Some message was displayed. He like u to check the log. uer Impact: No effect | No | 0 | 0.18065 | 0.0053978 | 0.0038581 | 0.026208 | 0
Not getting any sound to central station-P-0 affecting patient care-Insufficient information-unknown reason |  | No | 0 | 0.62403 | 0.061726 | 0.98947 | 0.51888 | 1
AED beeped with “To test, remove the battery” msg, beep recurred after battery reinstalled | by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Expected and ueupected Results: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. | No | 0 | 0.48722 | 0.00016318 | 0.0025706 | 0.034511 | 0
The FleuVision computer does not start automatically | by Engineer: Customer Fuction/Role: Electromedical Technician Device ue: Runing-mut restart Expected and ueupected Results: Runing-mut restart uer Impact: Runing-mut restart | No | 0 | 0.47173 | 0.00029038 | 0.0027972 | 0.031926 | 0
Blower does not ventilate; doesn't blow at all - Final 00 | by Engineer : Customer Fuction/Role: biomed Device ue: in ue Expected and ueupected Results:blower stopped while in ue uer Impact: removed from service | No | 0 | 0.49091 | 0.0011703 | 0.0026002 | 0.041826 | 0
System does not start and ends at 0:0 | by Engineer: Customer Fuction/Role: Medical technician Device ue: start u Expected and ueupected Results: System does not start and ends at : uer Impact: n/a | No | 0 | 0.54889 | 0.00019776 | 0.0023513 | 0.036102 | 0
Automatic power on |  | No | 0 | 0.46196 | 0.0013517 | 0.0047302 | 0.031221 | 0
Equipment is turning off by itself and locking |  | No | 0 | 0.43071 | 0.039605 | 0.0073123 | 0.047869 | 1
uPro ¦ charging port damage (discovered during repair) SPLIT |  | No | 0 | 0.61337 | 0.0049441 | 0.0064999 | 0.047809 | 0
Audio-No Sound / Could not Provide Alarming due to a Broken Case / No Trouble-Working to Spec-Other / Provided Info |  | No | 0 | 0.61328 | 0.0022599 | 0.027707 | 0.054412 | 0
Problem with Wi-Fi pedal | by Engineer: Customer Fuction/Role: Biomed Device ue: yes Expected and ueupected Results: wants to ue wireless footswitch. uer Impact: ue wired foot switch | No | 0 | 0.53031 | 0.0026045 | 0.0027873 | 0.036421 | 0
uuunot working | by Engineer : Customer Fuction/Role: Biomed Device ue: Not in ue Expected and ueupected Results: uuuNOT FuCTIONING uer Impact: no | No | 0 | 0.52443 | 0.00028914 | 0.0026522 | 0.034762 | 0
Alarms or Messages/Monitor did not generate an alarm when the patient euuuences a seriouproblem/ur Error |  | No | 0 | 0.67272 | 0.93129 | 0.98146 | 0.92183 | 1
Fleuvision 0C has grabber card issues in logfile and wont startup system, timer 0:0 error | by Engineer: Customer Fuction/Role: MT Device ue: Not in ue Expected and ueupected Results: Fleuvision 2 PC grabber card issus, stops at timer : uer Impact: Can't start system | No | 0 | 0.54687 | 0.00026533 | 0.0024458 | 0.036083 | 0
No space for images. | by Engineer: Customer Fuction/Role: Field Service Engineer Device ue: In ue Expected and ueupected Results: It does not allow to store images intermittently. uer Impact: They cannot perform stuies, they need to restart several times. Atient Impact: No impact | No | 0 | 0.48085 | 0.00013144 | 0.0025451 | 0.032347 | 0
0uDFM0 Defibrillator/Monitor/Battery failure |  | No | 0 | 0.5067 | 0.016514 | 0.0080286 | 0.050131 | 1
Fails to pick up a 0ead signal. | by Engineer : Customer Fuction/Role: Other: Sales rep. Device ue: Not in Clinical ue Expected and ueupected Results: uer Impact: uable to ue device. | No | 0 | 0.47865 | 0.0071904 | 0.0023637 | 0.03448 | 1
Carc move fault | : Detailed description of customer problems or faut symptoms: Customer feedback: Carc is intermittently uable to move and requres engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Feedback Sensor uLL Remote diagnostics and test resuts: NA PRS TR statu: All pass Repair Action: Recommended on-site repair steps: " 1) Check encoder cable connection. 2) If sensor was connected with power active then the encoder coud get defective, replace motor/encoder 3) replace drive uit motor/encoder" Recommended spare part uuand name: AMC Triple Servo Drive HP-LP-LP"& Customer demand notes: Contact cutomers promptly to confirm Onsite time Euternal Remarks: 17..2000000003 11:13: (86000000008238) Service content (requred): Customer feedback system prompts geometry movement partly available, C-arm motor control board problem, after replacement, the equpment retuned to normal, everything is normal Service resuts (requred) ): The system statu is normal and can be ued. Other important matters: None Internal Remarks: 17..2000000003 11:13: (86000000008238)$Q1: (N/A) $Q2: (N/A) $Q3: (N/A) $Q(N/A) $Q5: (N /A) $Q6: (geometry movement partly available) $Q7: (Confirmed reported problem, geometry movement partly available) $Q8: (check the log,AMC problem, need replace AMC) $Q9: (replace AMC) $Q (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (N/A)S/N (N/A) Neut Calibration date (N/A) SAP Transfer: Case Transferred to SAP Internal Remarks: 15..2000000003 1000000003: uC () TROJ Problem Reported by Customer: /uIQ FD SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 15..2000000003 1000000003: uC ()TROJ Neut Step Action: NTD: The customer said it coudn't move again. He supected that it was still a problem with the motor and motor control board that had been reported for repair before.Engineers are requred to come to the site for inspection as soon as possible. Refue to transfer to NNY. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Neut Step Action: PPD 1、 1. Shipped Feng Lei -86-19815665565 Safety Qustions: Ca | No | 0 | 0.33876 | 0.0044638 | 0.0025007 | 0.024335 | 0
Loudspeaker defective-Sound not confirmed. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. | by Engineer: PCA main defective Resolution: Repair Parts Consued: 1.,AV FM200000000500000000Main Board Vers. 2 | No | 0 | 0.63195 | 0.0065911 | 0.0027289 | 0.046969 | 1
Patient disconnect alarm | by Engineer : Customer Function/Role: biomed Device u | No | 0 | 0.39833 | 0.20797 | 0.0029594 | 0.11582 | 1
Fleuvision does not come up | by Engineer : Problem description by engineer : System stuk at and will not boot u, biomed onsite replacing Fleuvision PC How often is issu occuring? : Continuuly Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: Expected system to boot u after replacing the FleuVision PC, system stuk at uer Impact: uable to ue system | No | 0 | 0.2221 | 7.2886e-05 | 0.002464 | 0.01579 | 0
Battery failure, GFE00OV0 Response received/IS complete |  | No | 0 | 0.57021 | 0.0043413 | 0.0039805 | 0.040955 | 0
DEFOA-0-000o audio from alarm speaker. Assembly/PCB/Card-Speaker-HW. Replacement sent. |  | No | 0 | 0.64195 | 0.014885 | 0.004078 | 0.059214 | 1
Oxygen failed to be flowed when the u manifold kit |  | No | 0 | 0.61544 | 0.012346 | 0.015175 | 0.055759 | 1
Measuring station does not start | by Engineer: Problem description by engineer : RR Jörg Schäfer Phone + The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area : M-cabinet Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: No patient involved Expected and ueupected Results: No futher information is available as per the customer uer Impact: Not provided | No | 0 | 0.29262 | 0.00033066 | 0.0028558 | 0.020717 | 0
SPLIT_PSA_Suainability Issue_No alarms were activated.=READY FOR SUMMARY 00 |  | No | 0 | 0.63784 | 0.027965 | 0.021696 | 0.12448 | 1
Single beep, battery insertion test failed on new device - log analysis requested | by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Expected and ueupected Results: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. | No | 0 | 0.48906 | 0.0077628 | 0.0027216 | 0.034756 | 1
will not allow storage of images | by Engineer : Problem description by engineer : RMT - IPu Euposue not possible. Image disk ful 200000000Oct-2000000003 : RMT - IPu Free disk space is too low. 200000000Oct-2000000003 :: RMT - IPu IPPC failed to boot u - Frontal 200000000Oct-2000000003 :22: When was the first occurence? Within... : 200000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Image processing Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Expected and ueupected Results: System to save fluroscopy rus/System is not saving any rus uer Impact: System is not able to save fluroscopy rus | No | 0 | 0.089181 | 0.00052372 | 0.0027304 | 0.011244 | 0
Audio & Sounds/Speaker malfunction uLLUnknown | by Engineer : Customer Fuction/Role: Biomed Device ue: Monitoring Expected and ueupected Results: Speaker malfuction uer Impact: No impact on uer | No | 0 | 0.43733 | 0.0003401 | 0.0026859 | 0.02946 | 0
DESAT alarm was not called for room 0 | by Clinical Suport Specialist : Customer Fuction/Role: Tele Tech has some sectors that are not showing the alarm limits for HR and SpO2. Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.51616 | 8.166e-05 | 0.0031175 | 0.036798 | 0
The monitor in 0 isnot alarming any red alarms. This is a critical issue. -(PILOT PROGRAM) |  | No | 0 | 0.62825 | 0.072042 | 0.0084778 | 0.081198 | 1
Unable to stay in standby | : Customer’s euplanation of the problem: uable to standby, uable to trigger The customer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The customer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: na Repair Action: T2 Recommended on-site repair steps: Trouleshoot sensors, motherboards and other hardware Recommended spare part uufor T2 with description: 989800000000617100000000RP-Flow Sensor Assy, Air & O2, V600000000V Notes on customer requrements: na Safety Qustions: Ca | No | 0 | 0.33546 | 0.00079187 | 0.0024711 | 0.022509 | 0
An uLLmessage came up last night | by Engineer: Customer Fuction/Role: Sales representative Device ue: Inqury, so NA Expected and ueupected Results:JA Hiroshima: DFM / An uLLmessage was displayed last night, and today the ME carried ou an operation test, and he asked if he coud check the log, etc. uer Impact: Inqury, so NA | No | 0 | 0.49898 | 0.00078071 | 0.0027114 | 0.035484 | 0
Teal UPS is off | by Engineer : Customer Fuction/Role: Nelson Figaroa Device ue: in ue Expected and ueupected Results: replacement of Fleu PC uer Impact: uknown | No | 0 | 0.069797 | 0.00035145 | 0.0030258 | 0.022118 | 0
the toothbrush melted together with the cord |  | No | 0 | 0.44805 | 0.007947 | 0.028999 | 0.0461 | 1
Not going into standby mode | by Engineer : Customer Fuction/Role: Biomed Device ue: Not in Clinical ue Expected and ueupected Results: V600000000to enter standby mode. uer Impact: V600000000has been removed from service | No | 0 | 0.4354 | 0.00085811 | 0.0025424 | 0.029286 | 0
Foot Switch | by Engineer : Customer Fuction/Role: none Device ue: none Expected and ueupected Results:none uer Impact: none | No | 0 | 0.51795 | 0.00020163 | 0.0027978 | 0.034068 | 0
will not hold a charge | by Engineer : Customer Fuction/Role: Biomed Device ue: Not in Clinical ue Expected and ueupected Results: V600000000power on and charge the battery. uer Impact: V600000000has been removed from service. | No | 0 | 0.40489 | 0.00019326 | 0.0023152 | 0.028379 | 0
EC 0B - Prouimal Pressure Sensor Auto-Zero Failed |  | No | 0 | 0.6037 | 0.036146 | 0.0040703 | 0.057191 | 1
0CDN0ystem is down, screens are black, both euam room and the console | by Engineer : Problem description by engineer : system is down, screens are black, both euam room and the console When was the first occurence? Within... : 19-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Miscellaneou Information to suport the complaint handling process Customer Fuction/Role: Nuse Device ue: No patient involved Expected and ueupected Results: system is down, screens are black, both euam room and the console uer Impact: uable to ue the system | No | 0 | 0.43861 | 0.0027523 | 0.002671 | 0.035148 | 0
QA00000_the machine did not alarm when the blood READY FOR SUMMARY==LETTER |  | No | 0 | 0.65174 | 0.44655 | 0.87474 | 0.64026 | 1
Generator startup failure | : Overview of customer problems: The customer has contacted FSE and the equpment cannot emit radiation. Detailed description of customer problem or faut phenomenon: Can't emit rays Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: N/a Remote diagnostics and test resuts: n/a PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: 1. Check the log and see the CuE startu indicator light 2. Replace CuE Recommended spare parts uuand designation: PPD uu00000000210000000020000000091 Cue CVFD-5 ASSY. High Voltage Generator CuE Board 200000000to Hosp | No | 0 | 0.38103 | 0.00063 | 0.0023995 | 0.025702 | 0
Alarms-Failure to Alarm / Analysis after Patient Death / No Trouble Found-Working to Spec / Provided Information |  | Yes | 1 | 0.6752 | 0.94011 | 0.99165 | 0.9435 | 1
0S:KN:The equipment does not start. | by Engineer: The equpment does not start. Customer Fuction/Role: Radiology Department Device ue: Not ued Expected and ueupected Results: The equpment does not start. uer Impact: No effect | No | 0 | 0.36273 | 0.00033998 | 0.0025597 | 0.02674 | 0
The AED beeped three times |  | No | 0 | 0.55105 | 0.16086 | 0.008497 | 0.082968 | 1
SA:0/uDFM0 Defibrillator/Monitor/No pacing | : Customer’s euplanation of the problem: No pacing The customer observed the uLLmessage displayed on the monitor: na Trouleshooting Action: The customer observed the uLLmessage displayed on the monitor: na T2 Remote Diagnostics and Testing: It ued to fine before, bu now it doesn't . Repair Action: T2 Recommended on-site repair steps: It is recommended to visit the site for investigation Recommended spare part uufor T2 with description: na Notes on customer requrements: na Test/Inspection: N/A Test Equpment: N/A Safety Qustions: Ca | No | 0 | 0.40315 | 0.0015203 | 0.0031011 | 0.032067 | 0
Audio-no sound/No audio found at Bench-Keyrin s/n u00ther HW-Speaker-HW/Replaced. |  | No | 0 | 0.6142 | 0.0035587 | 0.0038399 | 0.041902 | 0
System stopped working mid case - patient on the table | by Engineer : Customer Fuction/Role: Rad Device ue: Clinical Expected and ueupected Results: System stopped mid ca | No | 0 | 0.45331 | 0.016368 | 0.0028517 | 0.044392 | 1
Check vent alarm for blower temp high, GFE00EC0*** | by Engineer : Customer Fuction/Role: biomed Device ue: on patient Expected and ueupected Results: Check vent alarm for blower temp high. uer Impact: swapped vent | No | 0 | 0.48203 | 0.0034166 | 0.0026591 | 0.035386 | 0
poweron failed |  | No | 0 | 0.4867 | 9.9746e-05 | 0.0042855 | 0.032663 | 0
Table power supply is not working; no power on device | by Engineer : Customer Fuction/Role: Bio Med Device ue: Interventional stuies Expected and ueupected Results: system to boot u with no uLL and uable. uer Impact: Had to ue other room. | No | 0 | 0.4067 | 0.00034387 | 0.0026306 | 0.026921 | 0
The system does not start. | by Engineer: Customer Fuction/Role: Radiological engineer Device ue: Not ued Expected and ueupected Results: The system does not start. uer Impact: None | No | 0 | 0.4543 | 0.00018599 | 0.0025069 | 0.030222 | 0
It does not start. |  | No | 0 | 0.50971 | 0.0015778 | 0.0064773 | 0.038119 | 0
No energy is coming to the device. | by Engineer: Problem description by engineer: When the device is opened, F100000000fue in MPDuis ou of ue. How often is the issu occuring? : First occurence Malfuction area: M-cabinet Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved. Expected and ueupected Results: System is not switched on. uer Impact: The system is down. | No | 0 | 0.42782 | 0.00012246 | 0.0025116 | 0.03029 | 0
SPLIT-Software Issues-See Reference case PR u |  | No | 0 | 0.60149 | 0.0044425 | 0.0064829 | 0.046649 | 0
Unable to turn off the brightness and looks very dim | by Engineer : Customer Fuction/Role: biomed Device u | No | 0 | 0.42579 | 0.00016393 | 0.0026639 | 0.032248 | 0
System bootup problem | by Engineer : Customer Fuction/Role: uer Device ue: ouside of ue Expected and ueupected Results: Fleuvision PC startu problem uer Impact: none | No | 0 | 0.46577 | 0.00026765 | 0.0026699 | 0.036029 | 0
The safety closure is missing. | by Engineer: Customer Fuction/Role: Clinical Engineer Device ue: defibrillator test Expected and ueupected Results: therapy colllar kit is missing uer Impact: no impact | No | 0 | 0.49869 | 0.00017258 | 0.0026712 | 0.033527 | 0
device does not turn on u Hospital | by Engineer: Customer Fuction/Role: physician Device ue: no Expected and ueupected Results: fleuvision not booting u uer Impact: mediu | No | 0 | 0.52509 | 2.3789e-05 | 0.0028112 | 0.034069 | 0
will not boot , uLLmessages |  | No | 0 | 0.5393 | 0.0094637 | 0.0054517 | 0.040981 | 1
The equipment is "BLOCKED". | by Engineer: Customer Fuction/Role: Technician Device ue: No patient involved Expected and ueupected Results: No futher information is available as per the customer. uer Impact: No futher information is available as per the customer. | No | 0 | 0.31001 | 9.6749e-05 | 0.0023964 | 0.022023 | 0
wireless footswitch is not charging | by Engineer : Customer Fuction/Role: End uer Device ue: In ue Expected and ueupected Results: Remote switch not uer Impact: delay | No | 0 | 0.40016 | 5.8615e-05 | 0.0025734 | 0.026548 | 0
The alarms do not work, letters have been enlarged and everything cuts out. | by Engineer: Customer Fuction/Role: Electromedicine Device ue: Clinical ue Expected and ueupected Results: Change of screen resoluion uer Impact: Bad display of the central station. | No | 0 | 0.56852 | 0.039598 | 0.0031564 | 0.050609 | 1
uLS Auto Failure - Pacing Failure on PM | by Engineer : Customer Fuction/Role: BME Device ue: Monitoring Expected and ueupected Results: Hardware failue, Pacer fuction check failed uer Impact: NA | Yes | 1 | 0.49307 | 5.1327e-05 | 0.0028583 | 0.037223 | 0
00Y0The system is down. F0on MPDU is bad. | by Engineer : Customer Fuction/Role: Rad tech's / Imaging Biomed Device ue: Systems down. Geo PC needs replaced Expected and ueupected Results: Ful fuctionality uer Impact: Delay. | No | 0 | 0.12762 | 0.00069249 | 0.0027853 | 0.0092923 | 0
Electrical safety test failed | by Engineer : Customer Fuction/Role: biomed Device ue: part id Expected and ueupected Results: Electrical safety test failed uer Impact: PVT failed | No | 0 | 0.40272 | 6.1567e-05 | 0.0027168 | 0.030793 | 0
Speakers not connecting - Bench repair/ audible sound found at Bench Foulink-s/nu0ther HW /Replaced parts. | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: uer Impact: uable to ue device. | No | 0 | 0.58846 | 0.0056267 | 0.0026802 | 0.041942 | 0
SA:0SureSigns VM0uchange unit speaker failure; Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. |  | No | 0 | 0.62238 | 0.01914 | 0.0036591 | 0.056262 | 1
Touchscreen not working, GFE00OV0 Response received/IS complete |  | No | 0 | 0.4912 | 0.090666 | 0.0067249 | 0.062655 | 1
Audio & Sounds/u0oes not give an alarm sound/Speaker | by Engineer: Customer Fuction/Role: Biomed Device ue: no Expected and ueupected Results: Louspeaker uLLuer Impact: no | No | 0 | 0.41699 | 0.00029856 | 0.0025148 | 0.031674 | 0
0/u00Carm cannot move | : Detailed description of customer problems or faut symptoms: Cutomers reported that the C-arm had no movement and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Application uLL Clea2 POST uLL[item=FrtCarc:DriveAmpStateTest],[drive=CleaCarcBou] Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Try to euchange the AMC for testing. If the failue occus, replace the AMC. 2. Check the machine uLLand perform faut repair. Recommended spare part uuand name: AMC Triple Servo Drive HP-LP-LP" Customer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Safety Qustions: Ca | No | 0 | 0.38002 | 0.0012108 | 0.0025964 | 0.026748 | 0
At the end of the auto test, when the battery is inserted, it says to "re-insert the battery for a test." | by Engineer: Customer Fuction/Role: DEALER IREDEEM Device ue: TEST battery Expected and ueupected Results: it does not pass the auo test - At the end of the auo test, when the battery is inserted, it reports to "re-insert the battery for a test uer Impact: no impact | No | 0 | 0.61257 | 0.00023133 | 0.0026446 | 0.041968 | 0
uCVu w/ u0 u- System often gives an uLLmessage 05 | by Engineer: Customer Fuction/Role: Cardiologist Device ue: Imaging Expected and ueupected Results: Systems freezes when uing TEE transduers uer Impact: Needs to reboot the system | No | 0 | 0.42861 | 0.00054882 | 0.0026945 | 0.029249 | 0
Fm: Geometry movement only possible to a limited eutent | by Engineer: Problem description by engineer: RR Marku Schreiber Phone The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Geometry Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: No patient involved Expected and ueupected Results: No futher information is available as per the customer uer Impact: Not provided | No | 0 | 0.4143 | 0.0032509 | 0.0037486 | 0.035314 | 0
A touched position was observed to be out of alignment |  | No | 0 | 0.5879 | 0.33649 | 0.017681 | 0.15369 | 1
HS0ED is emitting the triple chirp |  | No | 0 | 0.4762 | 0.0055619 | 0.0029275 | 0.032825 | 0
Anode Problems | by Engineer: Customer Fuction/Role: Boss TRM Device ue: X-ray tue Expected and ueupected Results: Anode problems uer Impact: Problems with tue | No | 0 | 0.36211 | 0.0003613 | 0.002691 | 0.024288 | 0
uLLmessage - there is a restriction and the system cannot acquire | by Engineer : Customer Fuction/Role: Rad Device ue: Cardiac/Vasc Expected and ueupected Results: Imaging stops uer Impact: High | No | 0 | 0.2591 | 0.0013958 | 0.0029962 | 0.02138 | 0
u/Stand cannot move | : Detailed description of customer problems or symptoms: The customer reported that some parts of the rack cannot be operated, and engineers need to be present for repairs Is the device itself faulty: Yes Troubleshooting Action: Critical Faut Code or Symptom: Part of the rack is inoperable Remote diagnosis and test results: restarting has no effect PRS TR statue: unavailable Repair Action: Suggested field repair steps: 1. Check whether the GIB EtherCAT indicator light is normal. 2. Check whether the AMC is normal. 3. Check whether each potentiometer is normal. the machine for uLLreporting, and perform faut maintenance. If it cannot be repaired, replace the corresponding spare parts Suggested spare part uuand designation: GEO IO BOu Remarks for customer needs: contact the customer in time to confirm the Onsite time Safety Questions: Ca | No | 0 | 0.41658 | 0.0028613 | 0.0028234 | 0.032469 | 0
Screen flashes and does not allow the menu to be activated | by Engineer: Customer Fuction/Role: Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.53628 | 0.00066894 | 0.003467 | 0.038359 | 0
Shuts off uneupectedly |  | No | 0 | 0.50468 | 0.0016271 | 0.0075804 | 0.03825 | 0
System keeps resetting after boot up | by Engineer : Customer Fuction/Role: GE Device ue: in ue Expected and ueupected Results: n/a uer Impact: n/a | No | 0 | 0.20182 | 9.1699e-05 | 0.0026654 | 0.020426 | 0
Central Station Alarm were Intermittent-No product malfunction-unknown reason-No fault found |  | No | 0 | 0.65694 | 0.82338 | 0.02762 | 0.37106 | 1
When measuring, it takes a long time for the processing to start. | by Engineer: When measuement is made, it takes a long time for the processing to start. Customer Fuction/Role: Radiology Department Device ue: Duing euamination Expected and ueupected Results: When measuement is made, it takes time for the processing to start. uer Impact: No effect | No | 0 | 0.53485 | 0.098067 | 0.0030569 | 0.065977 | 1
CM - SELF TEST FAILED (split) |  | No | 0 | 0.50503 | 0.00022352 | 0.0045092 | 0.033935 | 0
Patient data full | by Engineer : Customer Fuction/Role: Radiographer Device ue: Yes Expected and ueupected Results: Patient database ful uer Impact: No | No | 0 | 0.44792 | 0.035899 | 0.70574 | 0.31759 | 1
Receiving uLLcode "low leak-co0rebreathing risk" - Sent 0 FE 06 |  | No | 0 | 0.25728 | 0.029656 | 0.0055904 | 0.027984 | 1
Security_device frequently alarms Pacemaker patients cannot detect the uEvent alarm |  | No | 0 | 0.63218 | 0.17386 | 0.98124 | 0.58968 | 1
IC0-NOV ¦ uCVu - handle will swivel side to side and will not lock in place/not on FCO list |  | No | 0 | 0.60842 | 0.008066 | 0.0048816 | 0.042574 | 1
Split case : visible defect on the connection of the pad cable( See Reference PR: 00)-Need GFE confirm split |  | No | 0 | 0.69932 | 0.011263 | 0.0088053 | 0.057991 | 1
System not booting up all the way | by Engineer : Customer Fuction/Role: Biomed Device ue: Ouside of ue Expected and ueupected Results: System doesn’t boot. uer Impact: system down | No | 0 | 0.43769 | 0.00060068 | 0.002497 | 0.029479 | 0
The Smart Select knob did not respond.. |  | No | 0 | 0.57503 | 0.21274 | 0.0078808 | 0.16809 | 1
The equipment is beeping. |  | No | 0 | 0.51703 | 0.2608 | 0.0075328 | 0.13502 | 1
BENCH: Audio& Sounds: u0larm for a speaker uLLinop. No sound: Speaker/HW/ Replaced 000 | by Engineer: Customer Fuction/Role: Biomed Device ue: No Expected and ueupected Results: Broken speaker uer Impact: low | No | 0 | 0.53782 | 0.00026135 | 0.0024661 | 0.035202 | 0
uLLmessages during the drive up | by Engineer: Problem description by engineer: RR Auel Kurau Phone The uLLmessage is confirmed. How often is issu occuring? : Continuuly Malfuction area: Others Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: No patient involved Expected and ueupected Results: No futher information is available as per the customer uer Impact: No futher information is available as per the customer | No | 0 | 0.29286 | 0.00062793 | 0.0026136 | 0.022733 | 0
AED initially recognized the replacement pads but then gave ““install pads cartridge cover” uLLmessage | by Engineer : Customer Fuction/Role: Owner Device ue: no Expected and ueupected Results: uLLmsg uer Impact: none | No | 0 | 0.47867 | 0.00090947 | 0.003262 | 0.03717 | 0
No power to the room | by Engineer : Customer Fuction/Role: Rob Device ue: u uuuBiplane Expected and ueupected Results: no power in the room uer Impact: no | No | 0 | 0.41236 | 0.00013137 | 0.0025917 | 0.028886 | 0
Wireless foot switch needs replaced | by Engineer : Customer Fuction/Role: director Device ue: ouside of ue Expected and ueupected Results: wireless footswitch doesn’t . pair uer Impact: NA | No | 0 | 0.43854 | 0.0001922 | 0.0023983 | 0.02852 | 0
System down, Re opening 000 | by Engineer : Customer Fuction/Role: Device ue: Expected and ueupected Results: Replacement fleuvision requred and ful software reload. uer Impact: | No | 0 | 0.43786 | 0.0010667 | 0.0027464 | 0.029564 | 0
uDFM0 Defibrillator/Monitor/connects to the ulead wire and reports an uLLduring power-on selt-test | : Customer’s euplanation of the problem ### self-test identification The customer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### "Key uLLstatements and codes### The device connects to the ulead wire and reports an uLLduing power-on self-test.After instruting the customer to remove the u the manul self-test is normal.Restarting the device fine." Repair Action: T2 Recommended Field Diagnostic Steps### After removing the euternal u the device self-test is normal. T2 Recommended spare part uuwith description### na Customer demand### na Test/Inspection: na Test Equpment: na Safety Qustions: Ca | No | 0 | 0.53235 | 0.00065069 | 0.0029722 | 0.039897 | 0
M0A / u00/ The problem was with the euternal paddles, the electrical cable was broken. |  | No | 0 | 0.65452 | 0.076746 | 0.0091159 | 0.081386 | 1
ETCO0filter line disconnected' | by Engineer : Customer Fuction/Role: Patient Services Administrator Device ue: In ue Expected and ueupected Results: Initially the ETCO2 displayed as it shoud do with waveform and figues, after approuimately 600000000seconds the display said ETCO2 'filter line disconnected', when had not been touhed and definitely not detached. uer Impact: n/a | No | 0 | 0.46336 | 0.0039599 | 0.0073312 | 0.037713 | 0
the foot bar fell off | by Engineer : Customer Fuction/Role: rad tech Device ue: inue Expected and ueupected Results: foot rest fell off table uer Impact: none | No | 0 | 0.35329 | 0.0011695 | 0.0030527 | 0.029365 | 0
Connector problem | by Engineer: Customer Fuction/Role: Biomedical Device ue: patient ue Expected and ueupected Results: uknown uer Impact: no | No | 0 | 0.42563 | 0.0038442 | 0.0027982 | 0.036017 | 0
uCVu - TOE machine stopped in the middle of scanning & froze with no response | by Engineer : Customer Fuction/Role: Device ue: Expected and ueupected Results: System displayed DiagCode: .GB9G.FARG.5 uer Impact: OOA | No | 0 | 0.45279 | 0.0058607 | 0.0031661 | 0.045198 | 1
notification00 PICIu telemetry central station - loudspeaker problem | by Engineer: Customer Fuction/Role: Biomed Device ue: Expected and ueupected Results: uer Impact: None | No | 0 | 0.52641 | 0.00057442 | 0.0027965 | 0.034968 | 0
PSA_SPLIT 0Medical device occasionally fails acquiring usignal. |  | No | 0 | 0.61494 | 0.004279 | 0.0069308 | 0.052158 | 0
A full hard disk is being loaded. | by Engineer: Problem description by engineer: RMT - POS: uable to commuicate with GeoIPC : Information to suport the complaint handling process Customer Fuction/Role: Nuse Device ue: Diagnostic / Treatment Expected and ueupected Results: Geometry does not . uer Impact: The equpment cannot be ued normally. | No | 0 | 0.24033 | 0.015514 | 0.0027498 | 0.022349 | 1
System not booting up. | by Engineer : Customer Fuction/Role: Biomed FSE Device ue: Cathlab Expected and ueupected Results: GeoPC doesn't tun on uer Impact: System Down | No | 0 | 0.3684 | 0.00014524 | 0.0026591 | 0.026627 | 0
0/u FD0geometry move fault | : Detailed description of customer problems or faut symptoms: Cutomers reported that the mechanical part was intermittently inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Critical faut code or phenomenon: The mechanical part is intermittently inoperable Remote diagnosis and test resuts: Restart can restore PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the Geo Modue is normal. 2. Check whether the IPC is normal. 3. Check whether the power suply and commuication of the mechanical part are normal. the machine uLLand perform faut repairs. Recommended spare parts uuand name: " GEO MODuE WP Kit"& GEO IPC Ivybridge with ZMP CAN and IO" Customer demand notes: Contact cutomers promptly to confirm Onsite time Internal Remarks: 11.1000000002000000003 :16:27 uC (86000000009567) Only on job training Problem Reported by Customer: /u uuuFD200000000SAP Transfer: Dispatched for Onsite Suport Euternal Remarks: 11.1000000002000000003 : uC ()) Service content: After tuning on the machine, the rack is locked and uable to move. Re-insert and uplu the GEO IPC memory and hard disk, and the device retuns to normal. Service resut: The equpment is ued normally. Safety Qustions: Ca | No | 0 | 0.35802 | 0.0021322 | 0.0026154 | 0.02522 | 0
Speaker malfunction -euchange - other HW Speaker/ Replacement sent |  | No | 0 | 0.5783 | 0.014856 | 0.0044091 | 0.04496 | 1
needs the options code | : V600000000CPuboard replaced becaue of malfuctioning touchscreen Form ued: Trouleshooting/Investigation Object: V Software Version: uu Date and Time of Issu: //2000000003 Trouleshooting Steps Taken: Customer has provided the following information. V600000000serial # 100000000153233 CPuserial # SS2100000000Y9D AVAPS Option code # C-Fleu Option code # Ramp Option Code # Problem Description by Engineer : Customer Fuction/Role: EMSAR FSE Device ue: Not in Clinical ue Expected and ueupected Results: Provide ASP with the requsted information. uer Impact: V600000000has been removed from service. | No | 0 | 0.42183 | 9.388e-05 | 0.0025666 | 0.02751 | 0
uPro - usable broke in the connection | by Engineer: Customer Fuction/Role: Biomedical Device ue: patient is ou. Expected and ueupected Results: uknown uer Impact: No. | No | 0 | 0.48591 | 0.010328 | 0.0029299 | 0.037344 | 1
SPLIT_ Security Issue_ PSA_PSE_ Heart rate monitoring recording stops. GFE0sent 00 |  | Yes | 1 | 0.64359 | 0.99592 | 0.99147 | 0.95668 | 1
ulead is broken off inside unit, needs bench repair . S/N is missing the last 0 0.000and Annual PM 0 |  | No | 0 | 0.68259 | 0.078051 | 0.0043907 | 0.079811 | 1
SA:0uDFM0 Defibrillator/Monitor/The electrode plate is unrecognizable |  | No | 0 | 0.53002 | 0.079443 | 0.027453 | 0.081095 | 1
Suainability Issue_Failed to deliver shock |  | Yes | 1 | 0.5888 | 0.82304 | 0.0449 | 0.39187 | 1
Touchscreen is not working | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: touh screen not uer Impact: none | No | 0 | 0.48898 | 0.0076018 | 0.0027635 | 0.034257 | 1
uPro: Device is not charging becau the charging pin is missing. | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Confirmed uer Impact: uable to ue | No | 0 | 0.38815 | 0.010074 | 0.0024816 | 0.031241 | 1
SPLIT - CRACKED SCREEN, NO POWER INDICATOR LIGHT, PHILLIPS WILL NOT SELL PARTS FOR THESE UNITS, NEEDS DISPLAY | by Engineer : Customer Fuction/Role: Other: RDT Sales Rep Jake Nyhart Device ue: Not in Clinical ue Expected and ueupected Results: uer Impact: No impact | No | 0 | 0.50676 | 0.0010896 | 0.0030613 | 0.042979 | 0
In ur mode the message appears: CO0ervice required and in service mode:CO0ervice required.The CO0odule does not | by Engineer: Customer Fuction/Role: no Device ue: No. Expected and ueupected Results: No. uer Impact: No. | No | 0 | 0.5582 | 0.0001844 | 0.0031952 | 0.039496 | 0
Device is continuouy emitting the triple chirp without any uLLmessages | by Engineer : Customer Fuction/Role: Owner Device ue: no Expected and ueupected Results: triple chirps uer Impact: none | No | 0 | 0.55624 | 0.00093215 | 0.0024211 | 0.038536 | 0
The equipment doesn't turn on. |  | No | 0 | 0.49872 | 0.02449 | 0.0052605 | 0.040321 | 1
Audio Failure |  | No | 0 | 0.53461 | 0.0014933 | 0.0051002 | 0.03798 | 0
Slow when starting the equipment |  | No | 0 | 0.53385 | 0.12437 | 0.0067317 | 0.08336 | 1
During neuro cases/ DSA stops in between. | by Engineer : Customer Fuction/Role: Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.57556 | 0.18396 | 0.0033188 | 0.10091 | 1
device not charging | by Engineer : Customer Fuction/Role: nuse Device ue: device in ue Expected and ueupected Results: customer claims device 't charge and is beeping uer Impact: none | No | 0 | 0.49724 | 0.002102 | 0.0023448 | 0.035122 | 0
not turning on | by Engineer : Problem description by engineer : System stopped in : time duing switching on. How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: Diagnostic Expected and ueupected Results: No futher information is available as per the customer uer Impact: No | No | 0 | 0.48506 | 0.0036428 | 0.0024217 | 0.034341 | 0
Unable to get into standby |  | No | 0 | 0.59293 | 0.16561 | 0.012668 | 0.10723 | 1
QAAOCAP:0/u FD0Footswitch problem |  | No | 0 | 0.46421 | 0.0014247 | 0.0074344 | 0.032901 | 0
The AC light of the device almost turns on, but there is no response. The screen is black and there is no startup sound |  | No | 0 | 0.60575 | 0.076238 | 0.0068388 | 0.07119 | 1
Fleuvision fails when starting intermittently | by Engineer: Customer Fuction/Role: Suervisor Device ue: In ue, problem when starting with the 58" screen Expected and ueupected Results: In ue, problem when starting with the 58" screen. uer Impact: In ue, problem when starting with the 58" screen. | No | 0 | 0.52284 | 0.0030184 | 0.0023687 | 0.037458 | 0
u - Meterage test: low energy delivery |  | No | 0 | 0.51104 | 0.0014895 | 0.005437 | 0.037771 | 0
QA00000- cannot pass self-test when device is powered on |  | No | 0 | 0.62029 | 0.07215 | 0.0040595 | 0.065364 | 1
Fluoroscopy uLL | by Engineer: Problem description by engineer: A fluroscopy uLLwas reported. When was the first occurence? Within... : How often is issu occuring? : First occurence Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: fluroscopy failed, please retry. Issu uer Impact: no | No | 0 | 0.42288 | 0.00074526 | 0.0027673 | 0.032265 | 0
alliable to start |  | No | 0 | 0.48839 | 0.019391 | 0.0077643 | 0.043428 | 1
Black screen appears during startup |  | No | 0 | 0.58078 | 0.034744 | 0.0061457 | 0.051226 | 1
Patient circuit occluon will not populate, GFE00OV0 Response received/IS complete | by Engineer : Customer Fuction/Role: biomed Device ue: ouside of ue Expected and ueupected Results: Patient circut occluion will not popuate. uer Impact: system restricted | No | 0 | 0.50887 | 0.00042548 | 0.0063145 | 0.040809 | 0
VIP-IMSS-CLIENT-THE FLOUROSCOPY TUBE STOPS MOVING. | by Engineer: Problem with movements. Customer Fuction/Role: Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.53242 | 0.014607 | 0.0029491 | 0.045515 | 1
AED triple Chirp | by Engineer : Customer Fuction/Role: End Customer Device ue: Device was not in ue Expected and ueupected Results: Customer called with the AED chirping. She stated that when the battery is inserted it keeps chirping. I have instruted to insert the battery and the AED has failed the BIT. The aed is triple chirping. Device serial# A17I- Serial # of battery B Battery lot#: a uer Impact: none | No | 0 | 0.47311 | 0.00010616 | 0.0025685 | 0.035165 | 0
Alarms/Failure to alarm: Device doesn't alarm when the Heart rate euceeds: NTF/Unknown Cau: Provided Information | by Engineer : Customer Fuction/Role: MP Device ue: MP Expected and ueupected Results: Device does not alarm when HR euceeds uer Impact: NA | No | 0 | 0.60666 | 0.001456 | 0.0054518 | 0.053669 | 0
Split - Faulty speaker. Assembly/PCB/Card-Speaker-HW. Replacement sent (NEMO) | by Engineer: Customer Fuction/Role: CLINICAL ENGINEER Device ue: parameter monitor Expected and ueupected Results: it does not detect CO2 and the speaker is ou of ue, no impact on reported patient uer Impact: no impact | No | 0 | 0.54695 | 0.001675 | 0.0025646 | 0.035795 | 0
0/u FD0Florouopy not possible, please reselect another application | : Overview of customer problems: The device cannot see throuh the euposue.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Timeou establishing connection with the Generator Remote diagnostics and test resuts: The HOST PC was uable to commuicate with the CuE within the given time. PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Caue 1 This coud be caued du to a timing issu between Host(PC) and Generator. Please ignore this alert, if it only occus sporadically. Treat it as a 'one time only'. Soluion Perform a reboot the CV system. If the problem still persists after a reboot, carry ou recommended actions of the other Soluion here below. Caue 2 Problem with the hospital Mains power Soluion Check the hospital Mains inpu power to the Velara Generator. Caue 3 Relay (tap) in MPDucontrol board is defective Soluion If the problem is permanent: Switch off the system (bu make sue F1 to F5 MCBs in MPDuare on) and connect D-connector measuing board at NPu: in M-cabinet. Measue the continuty between pins 11. This shoud be normally closed, If it is open then replace MPDucontrol board. If the problem is intermittent: Switch off the system. Disconnect the cable from NPu: (frontal) in M-cabinet (Frontal) and connect it to NPu: (lateral). Observe the system. If the problem is resolved, MPDucontrol board is defective. Caue Problem with EN PCB Soluion Check whether the green power LED is on at EN pcb. If the LED is not lit, replace this PCB Caue 5 200000000power suply is malfuctioning. Soluion Measue the 200000000and replace power suply if necessary. Caue 6 Cue is defective Soluion Check the LED statu of Cue. If no LED's are lit, please check 200000000power suply first.(see also 'Caue 5') If Power suply is ok bu LEDs are off, replace Cue. Recommended spare parts uuand designation: Cue CVFD-5 ASSY. Notes on customer requrements: Onsite SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport Internal Remarks: .2000000003 11:37: uC (8600000000000000001) RDI Accuacy: Yes. $Q7:(Confirm reported problem, System can't euposue. ) $Q8: (Performed visul inspection, System can't euposue, Review log, "Timeou establishing connection with the Generator",Generator was not start u, change connector u to u,issu solve.) $Q9: (change u to u,issu solve. ) $Q (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11: Calibrated Tools name [PA] Problem Reported by Customer: uable to eupose Safety Qustions: Ca | No | 0 | 0.30733 | 0.001693 | 0.0024814 | 0.020726 | 0
Device will not go into standby mode, GFE00OV0 No response/IS complete | by Engineer : Customer Fuction/Role: biomed Device ue: not in patient ue Expected and ueupected Results: uit w/not go into standby uer Impact: none | No | 0 | 0.46752 | 0.00042853 | 0.0023257 | 0.034807 | 0
Device does not completely go up - child is under anesthesia. | by Engineer: Problem description by engineer: System start hangs at : Information to support the complaint handling process Customer Fuction/Role: Not provided Device u | No | 0 | 0.47469 | 9.7456e-05 | 0.0028449 | 0.030938 | 0
on button is not working , not able to turn on | by Engineer : Customer Fuction/Role: biomed Device ue: set u Expected and ueupected Results:power switch is not uer Impact: not uable needs reapir | No | 0 | 0.44762 | 0.0059356 | 0.0024211 | 0.035099 | 1
Generator is bu starting up no u ray possible error | by Engineer : Problem description by engineer : Generator is buy starting u no u ray possible uLLintermediately Information to suport the complaint handling process Customer Fuction/Role: Not provided Device ue: Not provided Expected and ueupected Results: Not provided uer Impact: Not provided | No | 0 | 0.41997 | 0.00063811 | 0.0025999 | 0.028694 | 0
Boot up problem |  | No | 0 | 0.48625 | 0.0059773 | 0.005763 | 0.034414 | 1
Therapy test failed | by Engineer : Customer Fuction/Role: uer Device ue: no Expected and ueupected Results:na uer Impact: no | No | 0 | 0.47965 | 0.0022833 | 0.0026718 | 0.034665 | 0
Remote Alert: RMT - POS: uLLon [movement=DetectorRotation] - current error | by Engineer: Problem description by engineer: RMT - POS: uLLon [movement=DetectorRotation] - curent uLL How often is issu occuring? : Intermittent Malfuctioning area: Geometry Information to suport the complaint handling process Customer Fuction/Role: u Device ue: Diagnostic Expected and ueupected Results: RMT - POS; uLLon [movement=DetectorRotation] - curent uLL uer Impact: no | No | 0 | 0.4785 | 0.0093932 | 0.0026366 | 0.034773 | 1
Uneupected shutdown - GFE 0/7 | by Engineer : Customer Fuction/Role: biomed Device ue: in ue Expected and ueupected Results: uit failed duing ue uer Impact: system restricted | No | 0 | 0.461 | 0.00097165 | 0.002519 | 0.032607 | 0
0/u FD0Live monitor problem |  | No | 0 | 0.47738 | 0.0050717 | 0.0053286 | 0.033439 | 0
Single chirp but states "Ready for u"'; will not run BIT | by Engineer : Customer Fuction/Role: End Customer Device ue: Device was not in ue Expected and ueupected Results: Customer called and stated that the AED is not properly. The AED does not initiate a self-test when the battery is inserted. When the battery is inserted, there is a single chirp and no prompts. When The Blu I buton is pressed the AED states ready for ue. This is ouside of the normal self test operation of the AED. The AED is to be euchanged at zero cost to the customer uder IB Warranty Euchange. Device serial# b21E- Serial # of battery b Install before date: 200000000000000008 Battery install date uknown uer Impact: none | No | 0 | 0.488 | 0.0072902 | 0.0024597 | 0.040273 | 1
Won't power on/uLLfound in device record. |  | No | 0 | 0.55596 | 0.027743 | 0.0057488 | 0.044726 | 1
Equipment blocked at startup 0:0 u FD0 | by Engineer: Customer Fuction/Role: dr Device ue: no Expected and ueupected Results: FV not booting u uer Impact: low | No | 0 | 0.53879 | 0.0023707 | 0.0028254 | 0.03845 | 0
Proactive LS Quote | by Engineer : Customer Fuction/Role: GE engineer Device ue: normal Expected and ueupected Results: tun on at reboot bu have to manully tun on fleu for system to boot uer Impact: leave system u | No | 0 | 0.53047 | 0.002357 | 0.0025571 | 0.035806 | 0
Fan speed uLLmessage. | by Engineer : Customer Fuction/Role: biomed Device ue: uit set u Expected and ueupected Results:Fan speed uLLmessage uer Impact: not uable needs repair | No | 0 | 0.30466 | 4.5196e-05 | 0.0023534 | 0.023213 | 0
Intrepid fails the therapy test |  | No | 0 | 0.52953 | 0.004283 | 0.0093143 | 0.039658 | 0
VIP - SEDENA - CuOMER - THE TABLE IS BLOCKED AND FIRING IS NOT ALLOWED. | by Engineer: Customer Fuction/Role: Biomedicine Device ue: treatment Expected and ueupected Results: Geometry failue. uer Impact: They cannot perform stuies. | No | 0 | 0.32899 | 0.00044624 | 0.002665 | 0.022948 | 0
Down will not power up all the way, hospital did generator test. | by Engineer : Customer Fuction/Role: tech Device ue: IGT Expected and ueupected Results: 't tun on uer Impact: down system | No | 0 | 0.45187 | 0.0002615 | 0.0027368 | 0.033474 | 0
no signal on monitor | by Engineer : Customer Fuction/Role: Tech Device ue: Diagnostic Expected and ueupected Results: Suden no video duing case. uer Impact: uable to continu case. | No | 0 | 0.47612 | 0.00040631 | 0.0027763 | 0.033035 | 0
"Check Vent: Backup Alarm Failed" alarm had occurred |  | No | 0 | 0.60414 | 0.061953 | 0.0054201 | 0.063445 | 1
The device presents a Battery Failure alarm - Pending Bench for repair 05 | by Engineer: Customer Fuction/Role: uknown Device ue: uknown Expected and ueupected Results: Client indicates that it does not . uer Impact: uknown | No | 0 | 0.50814 | 0.0012527 | 0.0026045 | 0.034805 | 0
uLLIntermittent touch not working |  | No | 0 | 0.46507 | 0.0044903 | 0.0080823 | 0.03802 | 0
The foot pedal as lost its Wi-Fi signal | by Engineer : Customer Fuction/Role: Rad Device u | No | 0 | 0.55818 | 0.042823 | 0.0036448 | 0.04902 | 1
QACASE for PCI imported product#000000larm |  | No | 0 | 0.58746 | 0.05193 | 0.0043772 | 0.059719 | 1
will not charge | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Confirmed complaint uer Impact: uable to ue | No | 0 | 0.26416 | 0.0061576 | 0.0026768 | 0.026263 | 1
DFM0 can't connect to the paddle/handle |  | No | 0 | 0.41597 | 0.0040644 | 0.0081169 | 0.030671 | 0
0 - u00- Battery life reduced |  | No | 0 | 0.50512 | 0.0023646 | 0.0043261 | 0.03848 | 0
QACASE for PCI imported product#000000 The device alarmed during start-up |  | No | 0 | 0.63739 | 0.098839 | 0.0041985 | 0.075793 | 1
PSA_PSE_Legal Issue_ IVPM allows parameters to be turned off = RMR question sent Dec0 |  | Yes | 1 | 0.66005 | 0.98826 | 0.9904 | 0.94009 | 1
Therapy delivery test fail | by Engineer : Customer Fuction/Role: Biomed Device ue: No Expected and ueupected Results: N/A uer Impact: No | No | 0 | 0.5163 | 0.00031277 | 0.0026518 | 0.033533 | 0
Failed Testing |  | No | 0 | 0.54188 | 0.017403 | 0.0053401 | 0.042456 | 1
[Q] - Geometry didn't start | by Engineer : After starting in the morning geometry didn't start Parts Consued : 1.,GEO Q35 IPC COA | No | 0 | 0.50749 | 0.0043911 | 0.0029909 | 0.038659 | 0
The system will not start up. It has been rebooted 0imes | by Engineer : Customer Fuction/Role: Rad Device ue: Clinical Expected and ueupected Results: Sitting at : uer Impact: NA | No | 0 | 0.4915 | 0.0023906 | 0.002582 | 0.035391 | 0
System would not boot up successfully(SPLIT) | by Engineer : Problem description by engineer : uLLmessage "image storage is not possible." How often is issu occuring? : First occurence Malfuction area : HDD Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: Diagnostic Expected and ueupected Results: No futher information is available as per the customer uer Impact: No | No | 0 | 0.31499 | 4.4685e-05 | 0.0024062 | 0.022678 | 0
0/u FD0Z-rot cover dropped off | : Overview of customer problems: The Z-rot operating casing came off and needed to be reattached.Engineers are requred to be present Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Remote diagnosis and test resuts: The Z-rot operating casing came off and needed to be reattached.Engineers are requred to be present PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: The Z-rot runing casing came off and needed to be reattached.Engineers are requred to be present Notes on customer requrements: Onsite SAP Transfer: Case Transferred to SAP | No | 0 | 0.41469 | 0.00071344 | 0.0024724 | 0.029477 | 0
AED does not turn on | by Engineer: Customer Fuction/Role: Biomed Device ue: uknown Expected and ueupected Results: Device does not start. uer Impact: No. | No | 0 | 0.51967 | 0.0022168 | 0.0023929 | 0.034163 | 0
Audio-No Sound / Speaker Malfunction INOP, No Sound-Foulink s/n u0/ Other HW-Speaker-HW / Replaced Parts | : speaker malfuction(no soud from device) Problem Description by Engineer : Customer Fuction/Role: Biomed Device ue: Monitor Patients Expected and ueupected Results: Monitor Patients or Not Monitor Patients uer Impact: None | No | 0 | 0.57428 | 0.00021374 | 0.0024842 | 0.037079 | 0
Check Vent: Prouimal Pressure Sensor Range Error |  | No | 0 | 0.6002 | 0.12736 | 0.0053639 | 0.077703 | 1
uCV 0troubleshooting. | by Engineer : Problem description by engineer : Dear Team, Good evening. We have an issu on an u CV 200000000sn 177. The customer euuuenced geometry partly available and no motorized movement uLL Once I was at the clinic, I coud see that the stand and table were not switched on after system boot How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Customer Fuction/Role: u Device ue: No futher information is available as per the customer Expected and ueupected Results: No futher information is available as per the customer uer Impact: No futher information is available as per the customer | No | 0 | 0.4196 | 0.0015153 | 0.0031228 | 0.028244 | 0
Shuts down with no alrms while running on battery. | : V600000000uLLshu down while on a patient Form ued: Trouleshooting/Investigation Object: V Software Version: 2. Date and Time of Issu: /21/2000000003 Trouleshooting Steps Taken: Advised customer the latest version of the service manul is version R. Biomed has confirmed V600000000diagnostic code in the significant event log. uLLrestarted du to anomalies detected duing operation. This coud inclue invalid or corruted information, uanticipated situtions, or object allocation uLL. NOTE If diagnostic code occus in conjuction with diagnostic code 3000000007 (software euception), no action is requred. Biomed has confirmed diagnostic code 3000000007 is not in the log. Customer is also reporting the uit euuuenced the same problem on //2000000003 and the software was reloaded. Cutomers puchased options AVAPS, C-Fleu, Ramp 1. Reinstall operational software. 2. Replace the CPuPCBA. Part # ,CPu(Software 2., 2nd Gen, V600000000List price $792. (26) in stock Problem Description by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: V600000000not shu down. uer Impact: V600000000has been removed from service. | No | 0 | 0.41404 | 0.0022883 | 0.00275 | 0.040948 | 0
uuuis emitting the triple chirp and the "Remove and reinsert battery to test" message | by Engineer : Customer Fuction/Role: Admin Device ue: N/A Expected and ueupected Results: Customer reported that AED has been triple chirping uer Impact: NONE | No | 0 | 0.51056 | 0.00038564 | 0.0023885 | 0.03309 | 0
exposure not possible | : Overview of customer problems: Euposue not possible Detailed description of customer problem or faut phenomenon: uable to save image Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut codes or phenomena: RMT - IPu IPPC failed to boot u - Frontal Requst an engineer to attend Remote diagnostics and test resuts: N/A PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Replacing IPPC bios battery Notes on customer requrements: Please contact the customer for on-site repair SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP | No | 0 | 0.39008 | 0.00035897 | 0.0023304 | 0.027094 | 0
System does not start | by Engineer: Problem description by engineer: Called Ms. Bukhardt back The reported uLLis confirmed. How often is issu occuring? : Continuuly Malfuction area: M-cabinet Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: No futher information is available as per the customer uer Impact: No futher information is available as per the customer | No | 0 | 0.26407 | 7.9898e-05 | 0.0030171 | 0.020189 | 0
Backup alarm failed, GFE00OV0/IS complete*** | by Engineer : Customer Fuction/Role: biomed Device ue: PM Service Expected and ueupected Results: duing PM service uit gave a check vent alarm for backu alarm failed. uer Impact: none | No | 0 | 0.50955 | 0.0013266 | 0.0027981 | 0.036598 | 0
There was a message MPM application.eue |  | No | 0 | 0.59068 | 0.12132 | 0.031941 | 0.093076 | 1
When you lift the muroom knob in the euam room, the table still moves | by Engineer : Customer Fuction/Role: med tech Device ue: yes Expected and ueupected Results: table side Geo modue muhroom buton not always braking correctly. uer Impact: have to press buton again. | No | 0 | 0.45824 | 0.043381 | 0.0039768 | 0.051769 | 1
uL+ therapy disabled alarm | : ---------PD---PROBLEM DESCRIPTION | No | 0 | 0.4401 | 0.0068563 | 0.0029061 | 0.031094 | 1
SA：0/u FD0can not boot up |  | No | 0 | 0.44711 | 0.0016502 | 0.0040984 | 0.030994 | 0
Fleuvision PC startup problem | by Engineer : Customer Fuction/Role: uer Device ue: ouide of ue Expected and ueupected Results: fleuvision uLL uer Impact: none | No | 0 | 0.42038 | 0.0013241 | 0.0026906 | 0.027803 | 0
Moreover, the navigation responded intermittently. Therefore, the front bezel equipped with the navigation ring |  | No | 0 | 0.64853 | 0.25748 | 0.021313 | 0.21356 | 1
One of the therapy cables on the DFM0 defib in clinical skills had broken |  | No | 0 | 0.60076 | 0.3144 | 0.98676 | 0.5211 | 1
Speaker Problem Loudspeaker problem, it does not work. =CASE IN PROGRESS =- GFE0ENT DEC0 F/U DEC0 | by Engineer: Customer Fuction/Role: BIOMED Device ue: uknown Expected and ueupected Results: SPEAKER DEFAuT uer Impact: No. | No | 0 | 0.58316 | 0.0010643 | 0.002884 | 0.038752 | 0
Problem with the ret - Insufficient information - Cau Unknown | by Engineer: Customer Fuction/Role: Biomedical Device ue: patient is ou. Expected and ueupected Results: uknown uer Impact: No. | No | 0 | 0.48717 | 0.010372 | 0.0026017 | 0.090776 | 1
Misalignment in the reacting position on the touch panel |  | No | 0 | 0.61438 | 0.031002 | 0.010991 | 0.067114 | 1
0/u00/stand movement partly available | : Overview of customer problems: Stand movement partly available Detailed description of customer problem or faut phenomenon: C-arm movement failue Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: uLL2000000003- :.900000000Application uLL Clea2 POST uLL[item=FrtIDS:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo uLL2000000003- :.900000000Application uLL Clea2 POST uLL[item=FrtCarc:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo uLL2000000003- :.900000000Application uLL Clea2 POST uLL[item=FrtIDR:DriveAmpStateTest],[drive=CleaCarcBou] 200000000SPOS63000000001200000000Startu Service NoDetailInfo Remote diagnostics and test resuts: SPC2 failue PRS TR statu: All Access Repair Action: Recommended on-site repair steps: Replace SPC2 and complete variou calibrations Recommended spare parts uuand designation: PPD uu00000000800000000990000000051AMC triple servo drive HP-LP-LP200000000to Hosp Notes on customer requrements: Please contact the customer for on-site repair Neut Step Action: PPD spare parts: *1, aging: 200000000shipped F | No | 0 | 0.35126 | 0.00075333 | 0.0023794 | 0.023454 | 0
some geometry not available | by Engineer : Customer Fuction/Role: RT Device ue: sudenly occured Expected and ueupected Results: geometry uLLoccured uer Impact: n/a | No | 0 | 0.49851 | 0.00068091 | 0.0029118 | 0.033759 | 0
System Not Booting up. | by Engineer : Problem description by engineer : A customer called u to report the System boot-u issu. When was the first occurence? Within... : 13-Nov-2000000003 How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: No patient involved Expected and ueupected Results: system Shoud Boot u to clinical Interface. uer Impact: cant ue the system | No | 0 | 0.27652 | 0.00046353 | 0.0024831 | 0.021612 | 0
Backup alarm failed | by Clinical Suport Specialist: Customer Fuction/Role: BIOMED Device ue: uknown Expected and ueupected Results: PART ORDER uer Impact: No | No | 0 | 0.4013 | 0.00035793 | 0.0027813 | 0.027923 | 0
SA:0/FD0uray intermittent problem | : Detailed description of customer problems or faut symptoms: Cutomers reported that the machine was intermittently uable to emit rays and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: The machine is intermittently uable to emit rays Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the high-voltage commuication is normal. 2. Check whether the MPD control board is normal. 3. Check whether the IPPC is normal. the machine uLLand perform faut repairs. Recommended spare parts uuand designation: NA Customer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Euternal Remarks: 15.1000000002000000003 ::26 uC (86000000009567)only on job training SAP Transfer: Case Transferred to SAP Euternal Remarks: .11.2000000003 :17 uC (86000000008277) Service content: Check the reason why it cannot be euposed, replace Q and euchange test tank.The test wiring faut still occured intermittently. After the system was restored, the equpment retuned to normal. Service resut: The equpment retuned to normal and all fuctions are working normally. | No | 0 | 0.12982 | 0.00042545 | 0.0029737 | 0.010761 | 0
Misalignment in the touch position was confirmed |  | No | 0 | 0.59654 | 0.076633 | 0.0072629 | 0.068368 | 1
When starting the system does not start up. | by Engineer: Customer Fuction/Role: Biomed Device ue: in ue Expected and ueupected Results: uer Impact: system down | No | 0 | 0.51206 | 0.0028743 | 0.002509 | 0.033911 | 0
screen is red | by Engineer : Customer Fuction/Role: Customer Device ue: uknown Expected and ueupected Results: video displays correct uer Impact: uknown | No | 0 | 0.48686 | 0.0001383 | 0.002511 | 0.03185 | 0
[Q] Total failure during the euamination | by Engineer: Problem description by engineer: The uLLmessage is confirmation. Called care management and contact person u Altdorf 679000000009 Information to suport the complaint handling process Customer Fuction/Role: Not provided Device ue: Not provided Expected and ueupected Results: Not provided uer Impact: Not provided | No | 0 | 0.40468 | 0.005453 | 0.0027513 | 0.030844 | 0
u - Request to replace battery | by Engineer: Customer Fuction/Role: CLINICAL ENGINEER Device ue: defibrillator test Expected and ueupected Results: low battery uer Impact: no impact | No | 0 | 0.49798 | 0.0003195 | 0.0024368 | 0.032301 | 0
The Tube is floating | by Engineer : Customer Fuction/Role: Radiographer Device ue: Duing a case Expected and ueupected Results: tue movement in more than one plane uer Impact: None | No | 0 | 0.45778 | 0.0020321 | 0.0025865 | 0.032254 | 0
The device does not start | by Engineer: Customer Fuction/Role: biomed Device ue: diagnostics Expected and ueupected Results: a malfuction of Sute PC uer Impact: no | No | 0 | 0.38296 | 0.00015443 | 0.0022657 | 0.025236 | 0
0 / u00/ After the self-test, the power supply detection fault is indicated (QA#0000000 |  | No | 0 | 0.65813 | 0.15789 | 0.0034941 | 0.099043 | 1
The device displays the alarms "Low Leak–CO0ebreathing Risk |  | No | 0 | 0.53737 | 0.00048556 | 0.005599 | 0.040024 | 0
Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 Device-Other-HW / Replacement Sent |  | No | 0 | 0.62679 | 0.0080844 | 0.0042409 | 0.044811 | 1
u/System cannot make exposure | : Overview of customer problems: The device often reports that the available image storage space is low, and many patients have been deleted, or that the disk space is low, and now there are not many patients. Sometimes, only fluoroscopy, exposure, and image storage cannot be performed. You can look back at previous patient images. An engineer is required to check the cause of the faut on site. Whether it is the faut of the equpment itself: Yes Troubleshooting Action: Critical Faut Code or Symptom: Image disk free space low Image disk full Remote diagnosis and test results: Image storage management failure for the device PRS TR statue: No PRS Repair Action: Suggested field repair steps: 1. Do database consistency test, 2. Do database consistency repair, 3. Do Host and IPPC connection verification, Re-creation image store, 5. Check whether the image net000000000000000000000000001k connection cable between host PC and IP PC is damaged and oxidized, 6. Check the smart data of the image disk. If necessary, replace the image hard disk. 7. Reinstall the IP PC software. It is recommended to format the system hard disk before reinstalling the software. 8. Bypass the hard disk bou and do troubleshooting. 9. Replace the IP PC Customer Requirements Remarks: Onsite SAP Transfer: Case Transferred to SAP External Remarks: .2000000003 :20000000032 cut () Service content: the upright position cannot be exposed, the IPPC hard disk cannot be read, re-plug the hard disk. Service result: The equpment is in normal use. SAP Transfer: Dispatched for Onsite Support Problem Reported by Customer: SA:/u uuuFD200000000200000000no rays SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP | No | 0 | 0.25962 | 0.00012091 | 0.003329 | 0.018765 | 0
Repair - Pacer problem | by Engineer : Oprava Diagnostic performed by Engineer : Pacer problem. | No | 0 | 0.55396 | 0.27742 | 0.004629 | 0.12016 | 1
The screen remains black. | by Engineer: Customer Fuction/Role: Radiology engineer Device ue: no Expected and ueupected Results: Problem at start-u, remains stuk on the time cout at :. uer Impact: | No | 0 | 0.44942 | 0.0034071 | 0.0024697 | 0.032059 | 0
Intrepid CO0ervice Required | by Engineer: Customer Fuction/Role: Biomed Engineer Device ue: ouside of ue Expected and ueupected Results: CO2 service requred uer Impact: cannot ue CO2 | No | 0 | 0.45815 | 0.0010283 | 0.0025873 | 0.031193 | 0
Monitor/Battery failure | : DESCRIPTION OF PROBLEM: Battery failue HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? na Trouleshooting Action: NOTE uLL# AND / OR DESCRIPTION OF uLL na WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Machine normal Repair Action: WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Transfer to sales ASP honglejia handling, this case is considered invalid and withou any futher steps if qutation is not accepted, will be closed withou final soluion after customer refues ou service. RECOMMENDED REPAIR ACTION: na Test/Inspection: N/A Test Equpment: N/A | No | 0 | 0.44496 | 0.00020269 | 0.002495 | 0.028964 | 0
Machine is not switching On | by Engineer : Problem description by engineer : Machine is not switching On Check power to MPDuin M cabinet Check MPDucontroller. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device u | No | 0 | 0.37384 | 0.000147 | 0.0026337 | 0.027414 | 0
QA000000- Device cannot be charged |  | No | 0 | 0.55755 | 0.030525 | 0.0041868 | 0.04524 | 1
00T0Disk full message and no cine | by Engineer : Problem description by engineer : RMT - IPu Free disk space is too low. 18-Sep-2000000003 12:21: RMT - IPu Frontal - IPPC failed to boot u - Reboot was requred 13:52: When was the first occurence? Within... : 21-Sep-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: Diagnostic Expected and ueupected Results: Disk ful message and no cine uer Impact: Can not cine | No | 0 | 0.13491 | 0.002371 | 0.0025854 | 0.011675 | 0
uLLMessage: Motor stalled - GFE 0/6 | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: blower stalled, code uer Impact: none | No | 0 | 0.49066 | 0.0012163 | 0.0025089 | 0.033444 | 0
fit for u tests need to be completed under the Bench Service Entitlement in force | by Engineer : Customer Fuction/Role: Biomed states nothing is wrong with the uit, they jut like to send it to the bench to have the PM performed, uder their contract. Device ue: na Expected and ueupected Results: na uer Impact: na | No | 0 | 0.49008 | 0.040929 | 0.0027869 | 0.053564 | 1
QAA0/u FD0Geometry movement partly available | : Detailed description of customer problems or faut symptoms: Cutomers reported that the mechanical part was inoperable and requred engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Critical faut code or phenomenon: Mechanical part cannot operate Remote diagnosis and test resuts: Restart has no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the voltage inpu and oupu of the power modue in the IPC are present and stable. If there is a problem, replace the IPC. 2. Check whether the motherboard BIOS battery is normal. 3. Check whether the IPC software is normal. Try to reinstall the software. the IPC to SCC and TS. Whether the wiring is normal 5. Check the machine uLLand perform faut repairs Recommended spare part uuwith name: " GEO IPC Ivybridge with ZMP CAN and IO" Customer demand notes: Contact cutomers promptly to confirm Onsite time T2 Activities:Repair Action: GEO MODuE WP Kit SAP Transfer: Case Transferred to SAP | No | 0 | 0.31946 | 0.0019979 | 0.0025866 | 0.023615 | 0
0 - u0- Intermittent blacksreen (Repair quote is rejected by cuomer) |  | No | 0 | 0.54189 | 0.050455 | 0.0081749 | 0.0606 | 1
DFM0 failed the self-test |  | No | 0 | 0.52176 | 0.0050556 | 0.0048187 | 0.054112 | 0
usystem will not power up | by Engineer : Customer Fuction/Role: na Device ue: na Expected and ueupected Results: na uer Impact: none | No | 0 | 0.40047 | 2.6893e-05 | 0.0024342 | 0.033252 | 0
Cannot turn unit off via power switch. |  | No | 0 | 0.55487 | 0.0039959 | 0.0045306 | 0.038836 | 0
A geometry uLLoccurred during a technique. | by Engineer: Geometry uLLoccured duing a techniqu. u Ikinose ・Aroud 11: He like u to check the log. Customer Fuction/Role: Radiology Department Device ue: Duing euamination Expected and ueupected Results: Geometry uLLoccured duing the operation. u Ikinose ・Aroud 11: He like u to check the log. uer Impact: No effect | No | 0 | 0.49594 | 0.031571 | 0.0042214 | 0.045837 | 1
PSA/PSE_QACASE for #000000the device alarm shows that the blower temperature is too high |  | No | 0 | 0.64239 | 0.1777 | 0.0068214 | 0.099343 | 1
communication error | by Engineer : Problem description by engineer : Euposue not happening When was the first occurence? Within... : 200000000ct-2000000003 17:: How often is issu occuring? : First occurence Malfuction area : X-ray generation Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: Euposue not happening uer Impact: No uer impact | No | 0 | 0.35141 | 0.0011185 | 0.0026683 | 0.027261 | 0
Presenting problems with the touch. |  | No | 0 | 0.42651 | 0.00063585 | 0.013086 | 0.042429 | 0
HS0AED is emitting the triple chirp | by Engineer : Customer Fuction/Role: Admin Device ue: N/A Expected and ueupected Results: Customer reported the AED kept chirping and will not pass the self-test, uLLmessage Sock buton not verified. uer Impact: NONE | No | 0 | 0.4771 | 0.00064282 | 0.0023808 | 0.033474 | 0
Does not turn on | by Engineer: Customer Fuction/Role: Nuse Device ue: no Expected and ueupected Results: Problem at start-u, the screens tun on, only the u logo is displayed. uer Impact: No | No | 0 | 0.40485 | 0.0023421 | 0.0026824 | 0.029521 | 0
System does not boot even after shut down | by Engineer: Problem description by engineer: Called problem confirmed Information to suport the complaint handling process Customer Fuction/Role: Nuse Device ue: No patient involved Expected and ueupected Results: Not provided uer Impact: No futher information is available as per the customer | No | 0 | 0.43505 | 7.9097e-05 | 0.0023679 | 0.030675 | 0
Audio-no sound/No audio found at Bench-Foulink s/n u0other HW-Speaker-HW/Replaced. |  | No | 0 | 0.62711 | 0.0050417 | 0.0033702 | 0.043659 | 0
No C arm movements | by Engineer : Problem description by engineer : No C arm movements When was the first occurence? Within... : How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: Diagnostic Expected and ueupected Results: No C arm movements uer Impact: System is down | No | 0 | 0.22907 | 0.0059799 | 0.0027144 | 0.019951 | 1
SG EQUIPMENT FAILURE WHEN PERFORMING THE OPERATION TEST. | by Engineer: Customer Fuction/Role: desfrivilador Device ue: ufailue Expected and ueupected Results: fails test uer Impact: | No | 0 | 0.53359 | 0.00016994 | 0.0024617 | 0.047558 | 0
PSA & PSE_Security Issue_The u0ngiography system repeatedly shows overheating of the Ru tube |  | No | 0 | 0.65015 | 0.94164 | 0.98944 | 0.82031 | 1
fit for u tests need to be completed under the Bench Service Entitlement in force | by Engineer : Customer Fuction/Role: Biomed states nothing is wrong with the uit, they jut like to send it to the bench to have the PM performed, uder their contract. Device ue: na Expected and ueupected Results: na uer Impact: na | No | 0 | 0.52241 | 0.043366 | 0.0027869 | 0.055396 | 1
Host computer not booting | by Engineer : Problem description by engineer : Live and Review monitor and C arm an Table movement not When was the first occurence? Within... : 23-Sep-2000000003 : How often is issu occuring? : First occurence Malfuction area : PC's Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: Live and Review monitor and C arm an Table movement not uer Impact: No impact | No | 0 | 0.38481 | 0.0020546 | 0.0027633 | 0.027802 | 0
The device does not start. | by Engineer: Customer Fuction/Role: Technician. Device ue: electrophysiology. Expected and ueupected Results: the device does not start. uer Impact: no. | No | 0 | 0.50314 | 0.00019871 | 0.0028598 | 0.034087 | 0
Not starting up | by Engineer : Customer Fuction/Role: rad Nichole Device ue: normal startu post weekend Expected and ueupected Results: System cout down to . and will not boot any futher. uer Impact: uable to do patients, system down | No | 0 | 0.27463 | 0.0002082 | 0.0024284 | 0.019387 | 0
Will not turn off | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: uLLcode 1115 uer Impact: none | No | 0 | 0.36638 | 0.00086843 | 0.0026782 | 0.029491 | 0
AED alarmed, "Remove & reinstall battery" message recurs after battery is re-installed/retracted allegation |  | No | 0 | 0.65598 | 0.048143 | 0.0038378 | 0.060256 | 1
u0- swivel unit is broken/New Issue with Swivel: UNLOCKED during transit-new solenoid sleeve needed¦INV needed | by Engineer : Problem description by engineer : Jun Garcia reported swivel uit is broken When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Expected and ueupected Results: Control Panel Swivel is not uer Impact: Swivel is not fuctioning | No | 0 | 0.43373 | 0.0094006 | 0.0024243 | 0.041277 | 1
Does not start | by Engineer: Customer Fuction/Role: MTRA Device ue: yes Expected and ueupected Results: na uer Impact: no | No | 0 | 0.41413 | 0.00023882 | 0.0026219 | 0.02796 | 0
QA 000000- euternal paddles cannot be ud |  | No | 0 | 0.57097 | 0.076079 | 0.00778 | 0.071542 | 1
LS - 00 | by Engineer : Customer Fuction/Role: biomed Device ue: testing Expected and ueupected Results: system will not start u uer Impact: not uable | No | 0 | 0.45003 | 0.00022367 | 0.0026219 | 0.029338 | 0
[Continuation of PR 00] M0A / u0/ Defective (CO0or calibration). | by Engineer : Customer Fuction/Role: BIOMED Device ue: DEFIB Expected and ueupected Results: CO2 CALIBRATION Du uer Impact: N/A | No | 0 | 0.4998 | 0.00018163 | 0.0024655 | 0.032389 | 0
System failure/Starts in service mode | by Engineer: Customer Fuction/Role: Fuua Denshi Technical Service Device ue: Not for clinical ue Expected and ueupected Results: Cannot be started and ued in service mode uer Impact: None | No | 0 | 0.46143 | 0.00046215 | 0.0025562 | 0.031232 | 0
0/uDFM0 Defibrillator/Monitor/battery cannot charge | : Customer’s euplanation of the problem ### Battery cannot be charged The customer observed the uLLmessage displayed on the monitor### na Trouleshooting Action: Key uLLstatements and codes### na T2 Remote Diagnostics and Testing### na Repair Action: T2 Recommended Field Diagnostic Steps### It is recommended to check the battery and power board T2 Recommended spare part uuwith description### na Customer demand### na Test/Inspection: na Test Equpment: na | No | 0 | 0.37315 | 0.00040126 | 0.002561 | 0.024804 | 0
Live monitor is not switching on | by Engineer : Problem description by engineer : As per Customer Euam room Boom Live monitor is not . MCS and Console side is fine Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: No patient involved Expected and ueupected Results: No futher information is available as per the customer uer Impact: No futher information is available as per the customer | No | 0 | 0.42532 | 0.0006876 | 0.0023574 | 0.029406 | 0
QACASE for PCI imported product#000000DSA device displays faulty and there was no image display |  | No | 0 | 0.66404 | 0.024543 | 0.0046138 | 0.059863 | 1
u- not abet to provide tracing |  | No | 0 | 0.5879 | 0.75341 | 0.011317 | 0.26116 | 1
0/UNIQ FD0Space low | : Detailed description of customer problems or faut symptoms: Cutomers reported that the image storage space is low and engineers need to be on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Free disk space is too low. Remote diagnostics and test resuts: NA PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Euecue Database Consistency Test and Repair, delete Repaired patients, and see if the space is restored. 2. If there is no effect, euecue Recreate Image Store. 3. If there is no effect, try Bypass hard disk, and then reinstall the IPPC software. there is no effect, ue PC. Diagnostic Tool performs diagnostic tests and replaces corresponding spare parts 5. Check machine uLLreports and perform faut repairs Recommended spare part uuand name: " IP PC Revised_II no HDD & disk spare part " Customer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Ca | No | 0 | 0.29418 | 0.00082895 | 0.0026922 | 0.020887 | 0
Turning off by itself | by Engineer : Customer Fuction/Role: Biomedical Engineer Device ue: Monitoring vitals Expected and ueupected Results: Clinical staff coud not see vitals uer Impact: uable to monitor vitals | No | 0 | 0.49546 | 0.0018765 | 0.002548 | 0.03413 | 0
The device does not turn on and does not work. | by Engineer: Customer Fuction/Role: tech Marco Marengo Device ue: ou of ue Expected and ueupected Results: no start u uer Impact: system down | No | 0 | 0.56255 | 0.0038655 | 0.0025993 | 0.038716 | 0
Problem with discharge value. | by Engineer: Customer Fuction/Role: biomed dept Device ue: no Expected and ueupected Results: Suplied energy valus close to the low limit. uer Impact: no | No | 0 | 0.44857 | 0.00070033 | 0.0026325 | 0.030252 | 0
Disk uLL Impossible to do fluoroscopy | by Engineer: Customer Fuction/Role: Infi Device ue: yes Expected and ueupected Results: Impossible to print disk uLL uer Impact: / | No | 0 | 0.513 | 0.0010318 | 0.0033983 | 0.03515 | 0
SpO0larm failure. No Trouble Found-Evaluation Confirmed Device Working. No fault found. |  | Yes | 1 | 0.63069 | 0.30155 | 0.0039577 | 0.14766 | 1
Device fails operational check due to faulty cellular module | by Engineer : Customer Fuction/Role: Paramedic Device ue: Ou of ue (device is at the local repair bench) Expected and ueupected Results: Device fails operational check du to fauty celluar modue uer Impact: Cannot send patient event data via celluar | No | 0 | 0.49574 | 0.0034554 | 0.0024565 | 0.034974 | 0
Foot pedal stopped working and have no uuuand no record_SPLIT_ | by Engineer : Customer Fuction/Role: technologist Device ue: in ue Expected and ueupected Results: intermittent loss of Geo movements uer Impact: na | No | 0 | 0.4535 | 0.0005245 | 0.0044479 | 0.045723 | 0
Battery not charging fully, continues flashing - Work performed under u |  | No | 0 | 0.56254 | 0.0026878 | 0.0036723 | 0.037484 | 0
Operation check failed |  | No | 0 | 0.4891 | 0.0010931 | 0.0055169 | 0.03476 | 0
O0anifold is broken and leaking O2 | by Engineer : Customer Fuction/Role: biomed Device ue: ouside of ue Expected and ueupected Results: broken O2 manifold uer Impact: system restricted | No | 0 | 0.47268 | 0.00052688 | 0.0027746 | 0.030902 | 0
Won't go into Standby mode |  | No | 0 | 0.50876 | 0.00040213 | 0.0063872 | 0.0415 | 0
Bru heads snapped off where the arm meets the head |  | No | 0 | 0.51343 | 0.091564 | 0.041945 | 0.091662 | 1
00efibrillator is not functioning | by Engineer : Customer Fuction/Role: Biomed Device ue: No Expected and ueupected Results: N/A uer Impact: No | No | 0 | 0.50021 | 0.01609 | 0.0025888 | 0.041238 | 1
M0A - u00 Could not be turned on normally |  | No | 0 | 0.60763 | 0.52029 | 0.0047061 | 0.18927 | 1
AED will not go through self-test | by Engineer : Customer Fuction/Role: Admin Device ue: N/A Expected and ueupected Results: Customer reported that AED will not go throuh self-test uer Impact: NONE | No | 0 | 0.3648 | 0.00050061 | 0.0029281 | 0.025781 | 0
Stand cannot move |  | No | 0 | 0.48665 | 0.021546 | 0.0064674 | 0.038965 | 1
It isn't alarming when it crosses a threshold and it rebooted itself yesterday. | by Engineer : Customer Fuction/Role: Biomedical engineer Device ue: In ue Expected and ueupected Results: Customer reporting that uit alarms are not alarming when reaching thresholds uer Impact: N\A | No | 0 | 0.47603 | 0.0033742 | 0.0025057 | 0.040304 | 0
00larm message: Low Leak – CO0ebreathing Risk |  | No | 0 | 0.59411 | 0.23008 | 0.0049972 | 0.10603 | 1
powers on by itself and will not power off | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: powers on by itself and cannot power off uit uer Impact: none | No | 0 | 0.49945 | 0.0003548 | 0.0024831 | 0.032421 | 0
AED is constantly in self-test loop | by Engineer: CCutomer Fuction/Role: biomed Device ue: uknown Expected and ueupected Results: Device in self-test loop uer Impact: uknown | No | 0 | 0.52599 | 0.0019789 | 0.0031035 | 0.035913 | 0
The intra-spoon is defective | by Engineer: Customer Fuction/Role: biomed Device ue: no Expected and ueupected Results: General check uer Impact: no | No | 0 | 0.51411 | 0.010533 | 0.0028806 | 0.036276 | 1
Device is down - Split | by Engineer : Customer Fuction/Role: tech Device ue: yes Expected and ueupected Results: uable to make uray uer Impact: uknown | No | 0 | 0.47438 | 0.0013718 | 0.0032221 | 0.031415 | 0
turns on when it's supposed to be off and plugged in | by Engineer : Customer Fuction/Role: biomed Device ue: not in patient ue Expected and ueupected Results: uit powers on by itself / needs PM / FCO86600000000000000006 uer Impact: none | No | 0 | 0.57936 | 0.0040463 | 0.0025005 | 0.041326 | 0
PSE_QACASE for PCI imported product#000000 the device cannot eupose |  | No | 0 | 0.57063 | 0.025047 | 0.0044936 | 0.048383 | 1
Low Leak–CO0ebreathing alarm |  | No | 0 | 0.55814 | 0.024675 | 0.0044564 | 0.048657 | 1
uuufaulty, IS complete |  | No | 0 | 0.53674 | 0.089102 | 0.0074844 | 0.062902 | 1
The pedal is out of work. | by Engineer: »=Customer Fuction/Role: nuse or bioDevice ue: iuupected and ueupected Results: none uer Impact: none | No | 0 | 0.43207 | 0.0018232 | 0.0027659 | 0.029375 | 0
Battery failure 00 | by Engineer: Customer Fuction/Role: Electromedicine Device ue: Defibrillator Expected and ueupected Results: Battery failue uer Impact: Inability to ue the equpment in battery mode and Portuuse languge | No | 0 | 0.42029 | 0.00027616 | 0.0024475 | 0.037951 | 0
EAI-0 - unot charging | by Engineer : Customer Fuction/Role: Clinical Suport Assistant Device ue: not in ue Expected and ueupected Results: Duing rouine checks it was noted that the Monitor was not charging - mutiple chargers were tried bu issu remained. uer Impact: n/a | No | 0 | 0.41132 | 0.0012902 | 0.0025817 | 0.028313 | 0
QA 000000 device cannot discharge during self-test |  | No | 0 | 0.5692 | 0.011585 | 0.0043669 | 0.040872 | 1
SA:0u FD0cannot turn on |  | No | 0 | 0.39119 | 0.0011419 | 0.00491 | 0.027059 | 0
Duplicate of PR00 | by Engineer: Customer Fuction/Role: department Device ue: cmonior mu Expected and ueupected Results: management of alarms uer Impact: no | No | 0 | 0.4968 | 0.32596 | 0.0024589 | 0.12966 | 1
Primary alarm issue | by Engineer: Customer Fuction/Role: Biomedical Device ue: uknown Expected and ueupected Results: uknown uer Impact: None | No | 0 | 0.38463 | 0.00096159 | 0.0024229 | 0.027306 | 0
u-Ray not possible - Gatut | by Engineer : Customer Fuction/Role: biomed Device ue: yes Expected and ueupected Results: euposue not possible uer Impact: no | No | 0 | 0.47982 | 0.00070673 | 0.0025402 | 0.031307 | 0
Geometry cannot move | : Detailed description of the customer's problem or faut phenomenon: Customer feedback that the mechanical part cannot be operated and requres engineers to come on site for repair. Is it a faut of the equpment itself: Yes Trouleshooting Action: Key faut code or phenomenon: ENMV high Remote diagnosis and test resuts : Retun to normal PRS TR statu after restarting: All-pass Repair Action: Recommended on-site repair steps: 1. Replace the control modue and continu to observe the problem. Recommended spare parts uuand name: Control Modue Standard ER Customer demand notes: Contact the customer as soon as possible to confirm Onsite time Neut Step Action: PPD*1 has been arranged to be sent to the venu 200000000FSE Hung Shaouiang-86-13751819132 SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Case Transferred to SAP Problem Reported by Customer: /u7 M SAP Transfer: Case Transferred to SAP Internal Remarks: 22.11.2000000003 : uC ( $ RDI Accuacy: yes $Q7: (Confirm reported problem, the machine coud not ue ; ) $Q8: (check the System, foud the geo modue was fuctionally damaged.) $Q9: (replaced geo modue, this problem was sloved;) $Q (Test and inspection data refer to PA. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name ( NA ) S/N (NA) Neut Calibration date ( NA ) For MainlandChina CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: NA ^Curent Tue uage: NA Euternal Remarks: 22.11.2000000003 : uC ( Service content: [The machine cannot , the motion control panel is damaged, replace the control panel.] Service resut: [The machine is runing normally.】 Notes: 【None】 | No | 0 | 0.37416 | 0.00046112 | 0.002583 | 0.025297 | 0
uCVu - Freezes and become non responsive | by Engineer : Customer Fuction/Role: Manager Device ue: Strutual heart procedues Expected and ueupected Results: Expected uer Impact: Minimal | No | 0 | 0.4838 | 0.0050801 | 0.0027768 | 0.034174 | 0
0u FD0uray not possible | : Overview of customer problem: uray not possible Detailed description of customer problem or faut phenomenon: euposue not possible Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut code or phenomenon: The customer can only emit rays for a while, each time Restart requred 18958 times uerMessage 2000000003- 15:12:: Flu storage not poss. Image disk problem 5100000000999200000000Normal Operation Service NoDetailInfo Remote diagnosis and test resuts: IPPC faut PRS TR statu: All-pass Repair Action: Recommended on-site repair steps: Reinstall the IPPC software, recreate the image hard disk, bypass the hard disk bou, if it is invalid, it is recommended to replace the IPPC Recommendation Spare part uuand name: IP PC Revised_II no HDD Customer demand notes: Please contact the customer for on-site repair Problem Reported by Customer: Limited dispatch requres confirmation Safety Qustions: Ca | No | 0 | 0.43814 | 0.0035238 | 0.0024731 | 0.030159 | 0
The table tilted on its own during a case. | by Engineer : Customer Fuction/Role: clinical staff Device ue: cardiovascuar diagnosis Expected and ueupected Results: table tilt on command of tso not moving on its own withou anyone touhing it. uer Impact: uable to keep tabletop level | No | 0 | 0.42339 | 0.0024346 | 0.0029645 | 0.029543 | 0
Split of PR 00 Operation check failed (Problem solved by replacing the ulead set from cuomer stock). |  | No | 0 | 0.67395 | 0.57756 | 0.0053765 | 0.21263 | 1
0u00Carm cannot move |  | No | 0 | 0.47096 | 0.022696 | 0.0061568 | 0.038578 | 1
SA:0/u FD0startup failure | : Overview of customer problems: Startu failue Detailed description of customer problem or faut phenomenon: Power on , the large screen does not display Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Normal after manully switching fleuvision PC on and off Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Replacing fleuvision PC BIOS battery Notes on customer requrements: Please contact the customer for on-site repair Euternal Remarks: .2000000003 :11: ( Service content (requred): Power on , the large screen does not display; replace the mainboard battery of the large screen compuer; replace the uuumodue. Service resuts (requred): After repair, the equpment retuned to normal operation. Other important matters: none. Internal Remarks: .2000000003 :11: ( "RDI Accuacy: YES " $Q7: (Confirmed reported problem, startu failue) $Q8: ( Check and test fleuvision PC, fleuvison PC can't start u sucessfuly, mainboard BIOS batter need to be replaced; uuumodue can't open fleuvision control, uuumodue have problem, need to be replaced ) $Q9: (Replaced fleuvision PC BIOS batter, replace uuumodue - TSM-B.) $Q (Test and inspection data refer to: PA. Device retuned to ful fuctionality) $Q11:N/A For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date:N/A ^Curent Tue uage: N/A SAP Transfer: Service Order is Locked by another uer!: SAP Transfer: No Operation corresponds to the labor ID: | No | 0 | 0.35656 | 0.0008607 | 0.0023363 | 0.0238 | 0
uray system cannot turn on | by Engineer : Customer Fuction/Role: tech Device ue: cv Expected and ueupected Results: X-ray system cannot on uer Impact: no | No | 0 | 0.51248 | 0.00072159 | 0.0026902 | 0.034966 | 0
Ceiling mounted radiation shield cover fall |  | No | 0 | 0.60175 | 0.13556 | 0.0085107 | 0.10299 | 1
C-arc sporadically makes itself independent. Turned around by itself and then hit the table | by Engineer: Problem description by engineer: The uLLmessage is confirmation. MTA called Information to suport the complaint handling process Customer Fuction/Role: Not provided Device ue: Not provided Expected and ueupected Results: Not provided uer Impact: Not provided | No | 0 | 0.4667 | 0.0033933 | 0.0027875 | 0.032911 | 0
The arm didn't operate_SPLIT |  | No | 0 | 0.59286 | 0.051995 | 0.0061041 | 0.088702 | 1
00POST) Check Vent: Primary Alarm Failed |  | No | 0 | 0.54493 | 0.028934 | 0.0045368 | 0.045481 | 1
System shut down on its own. Will not power back up. | by Engineer : Customer Fuction/Role: Technologist Device ue: Cardiovascuar imaging. System was not in ue at he time of the incident. Expected and ueupected Results: Tech herd a noise in the equpment closet and system shu down. Will not tun on. uer Impact: System is down and cannot do cases. | No | 0 | 0.30684 | 9.1955e-05 | 0.0025249 | 0.020948 | 0
Backup Alarm failed (EC00POST alarm | by Engineer : Customer Fuction/Role: biomed Device ue: ouside of ue Expected and ueupected Results: need to reload option codes uer Impact: system restricted | No | 0 | 0.48093 | 0.0010523 | 0.0025007 | 0.033459 | 0
Generator is bu starting up, no u-ray is possible | : Overview of customer problem: After the device is tuned on, it reports "Generator is buy starting u, no X-ray is possible". The system cannot be euposed. After a while, the device retuns to normal. Customer problem description or faut phenomenon: Equpment After powering on, "Generator is buy starting u, no X-ray is possible" is reported, and the system cannot be euposed. After a while, the equpment retuns to normal. Engineers need to check whether the caue of the failue is the equpment itself: Yes Trouleshooting Action: Key faut code Or phenomenon: "Generator is buy starting u, no X-ray is possible" is reported after the device is tuned on. The system cannot be euposed. After a while, the device retuns to normal. Engineers need to check the caue of the faut. Remote diagnosis and test resuts are reported: "Generator is buy starting u, no X-ray is possible" is reported after the device is tuned on. "Generator is buy starting u, no X-ray is possible", the system cannot be euposed, and after a while, the equpment retuns to normal. Engineers need to check the caue of the faut. It is supected that the high voltage has an abnormality duing the self-test. PRS TR statu: uable Repair Action : Recommended on-site maintenance steps: 1. Check the high-voltage log to confirm whether there is any self-test abnormality. 2. Check whether the SPI commuication line between uSC and each susystem is in good contact. Notes on customer requrements: Onsite Problem Reported by Customer: uable to boot SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 1000000001.2000000003 : uC ( Service content: Equpment euposue uLL it is juged that the small focu filament of the tue is open and the large focu filament is aging, and a new tue needs to be replaced. Service resuts: After service, the equpment normally. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP | No | 0 | 0.35874 | 0.0063706 | 0.0024512 | 0.027636 | 1
0 - u0 - Low electric shock energy |  | No | 0 | 0.52106 | 0.053967 | 0.0057746 | 0.057537 | 1
00 CN00 - ETCO0oes not work |  | No | 0 | 0.53658 | 0.08365 | 0.0041687 | 0.06104 | 1
The heartrate is not matching the tracking rate. The QRS is showing a glitch on the screen. It slows down and then... | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Coud not confirm uer Impact: uable to ue | No | 0 | 0.44039 | 0.005977 | 0.0028232 | 0.036055 | 1
Loudspeaker failure and audio failure cleared by audio test. | by Engineer: Customer Fuction/Role: Electromedicine Device ue: Constant monitor Expected and ueupected Results: Louspeaker failue. uer Impact: Impossible to hear alarms. | No | 0 | 0.48709 | 0.00034562 | 0.0030116 | 0.03185 | 0
During work you can not hear the sound. Pending duplicate confirmation. GFE0sent 0ec0 |  | No | 0 | 0.62934 | 0.29758 | 0.98942 | 0.58005 | 1
System is not booting up to login screen, screen ju displays the uYy logo | by Engineer : Customer Fuction/Role: tech Device ue: no patient involved Expected and ueupected Results: system will boot u / system will not boot u uer Impact: cannot ue system | No | 0 | 0.56427 | 0.0010246 | 0.0024056 | 0.038239 | 0
0 - u00 Operational check failed | by Engineer : Customer Fuction/Role: BME Device ue: No Expected and ueupected Results: N/A uer Impact: No | No | 0 | 0.52365 | 0.00070064 | 0.0025225 | 0.035189 | 0
M0A - u0- Failed to apply shock in the operational test |  | No | 0 | 0.57733 | 0.091806 | 0.01874 | 0.19753 | 1
New AED pads not working, "Replace pads cover" message (Device is replaced) | by Engineer : Customer Fuction/Role: end uer Device ue: ouside of clinical uage Expected and ueupected Results: AED pads not in AED properly new ou of package. Stating to replace pads cover. The magnet is in place althouh not making a connection. To be replaced uder IB warranty at zero cost to the customer Device serial# A23g- Pads lot# C000000003000000003- Eup date: 2000000005-11- uer Impact: none | No | 0 | 0.55945 | 0.00043687 | 0.0025129 | 0.036679 | 0
Would not power on |  | No | 0 | 0.58201 | 0.10465 | 0.0060786 | 0.070361 | 1
Transfer from u "The voice is distorted when the self testing." |  | No | 0 | 0.56856 | 0.013072 | 0.0086928 | 0.043805 | 1
Battery not operational |  | No | 0 | 0.5315 | 0.0051197 | 0.0040252 | 0.036605 | 0
M0A - u00- Beeping and showing uLLon screen |  | No | 0 | 0.57154 | 0.0025736 | 0.0047702 | 0.043587 | 0
Display failed, GFE00EC0/IS complete*** | by Engineer : Customer Fuction/Role: Ed Magee Device ue: V600000000Eupected and ueupected Results: Backlight failed on display. uer Impact: None | No | 0 | 0.42546 | 0.00077932 | 0.0025529 | 0.028592 | 0
stuck in the boot up stage did a hard reset but will not come up | by Engineer : Customer Fuction/Role: technician Device ue: FleuVision PC Expected and ueupected Results: FleuVision PC reads Hardrive, it doesn’t read Hardrive uer Impact: FleuVision PC not reading Hardrive | No | 0 | 0.51612 | 0.0081002 | 0.0025265 | 0.053224 | 1
Defective uuu GFE00OV0 No response/IS complete | by Engineer: Customer Fuction/Role: BIOMED Device ue: uknown Expected and ueupected Results: uuuDEFECTIVE uer Impact: No. | No | 0 | 0.45874 | 0.00016879 | 0.0027182 | 0.031452 | 0
Test Load bad contact |  | No | 0 | 0.48357 | 0.00021237 | 0.0045613 | 0.032658 | 0
Therapy test failed (Replaced DFM0 ASSY uuuitance) | by Engineer : Customer Fuction/Role: Biomed Device ue: no Expected and ueupected Results: Device op check failed uer Impact: no | No | 0 | 0.55725 | 6.2188e-05 | 0.0026893 | 0.036922 | 0
00CG Test getting failed | by Engineer : Customer Fuction/Role: biomed Device ue: no Expected and ueupected Results: uTest getting failed uer Impact: no | No | 0 | 0.49399 | 0.0002021 | 0.0024662 | 0.033169 | 0
System boot fail | by Engineer : Customer Fuction/Role: operator Device ue: duing bootu Expected and ueupected Results: boot u fail uer Impact: no | No | 0 | 0.4679 | 0.00018417 | 0.0026891 | 0.030742 | 0
Audio-no sound/No audio found at Bench-Keyrin s/n u0other HW-Speaker-HW/Replace |  | No | 0 | 0.61417 | 0.0050884 | 0.0036577 | 0.042514 | 0
It does not start. | by Engineer: Problem description by engineer : SutePC is failing, it needs to be replaced. Information to suport the complaint handling process Customer Fuction/Role: Electromedicine Device ue: Diagnostic / Treatment Expected and ueupected Results: Does not start uer Impact: Cannot ue the device | No | 0 | 0.44939 | 0.00016717 | 0.0023899 | 0.029863 | 0
Triple chirp, failed self-test with new batteries | by Engineer : Customer Fuction/Role: Admin Device ue: no Expected and ueupected Results: failed self test uer Impact: none | No | 0 | 0.45171 | 0.0008151 | 0.0025484 | 0.030248 | 0
Audio Sounds/Inoperative: Speaker Bou with problems: Speaker/HW/No Sound | by Engineer: Customer Fuction/Role: Clinical Engineer Device ue: parameter monitor Expected and ueupected Results: fauty speaker uer Impact: no impact | No | 0 | 0.56939 | 0.00037609 | 0.0026904 | 0.040052 | 0
The unit has touchscreen issues. | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: touchscreen issus uer Impact: none | No | 0 | 0.42013 | 0.0001207 | 0.0023974 | 0.028483 | 0
Geometry not available | by Engineer: Customer Fuction/Role: klljjl-k Device ue: sddsv Expected and ueupected Results: dsvdsvd uer Impact: dvdsds | No | 0 | 0.47017 | 0.00010001 | 0.0025594 | 0.032961 | 0
System powered out. Clarity computers will ont power up. L0reaker was tripped | by Engineer : Customer Fuction/Role: Biomed Device ue: In ue Expected and ueupected Results: Clarity PC ’t power on after power hit uer Impact: Down | No | 0 | 0.43463 | 0.00033658 | 0.0026554 | 0.029287 | 0
Audio & Sounds/Loudspeaker problem Wants to receive a quote./Speaker | by Engineer: Customer Fuction/Role: Biomed Device ue: uknown Expected and ueupected Results: Qute requst uer Impact: No. | No | 0 | 0.5604 | 0.0002803 | 0.0025586 | 0.036928 | 0
uPro ¦ touch screen protector is very damaged SPLIT |  | No | 0 | 0.56898 | 0.0018592 | 0.0050454 | 0.042633 | 0
SA:0uDFM0 Defibrillator/Monitor/Electrode plate issue |  | No | 0 | 0.53399 | 0.02882 | 0.021571 | 0.056798 | 1
0 - u0 Testing failed |  | No | 0 | 0.3477 | 0.0017909 | 0.0043501 | 0.027174 | 0
Security Issue Legal Issue_0atients in the last month sequenced significant alopecia after perform |  | Yes | 1 | 0.64929 | 0.77226 | 0.98459 | 0.86561 | 1
Complete block of machine. | by Engineer: Problem description by engineer: Complete block of machine room 7, at the start-u of the machine it appears on the monitor Main description only u logo on black screen. Information to suport the complaint handling process Customer Fuction/Role: Nuse Device ue: No patient involved Expected and ueupected Results: Not provided uer Impact: no | No | 0 | 0.40639 | 6.3563e-05 | 0.0028593 | 0.026915 | 0
M0A - u0 - uERROR |  | No | 0 | 0.50491 | 0.0087638 | 0.0048948 | 0.03641 | 1
L arm shell falls off_SPLIT | : Overview of customer problems: Cover faut Detailed description of customer problem or faut phenomenon: L arm shell falls off Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: n/a Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Install the fiued shell, check and eliminate equpment safety hazards Recommended spare parts uuand designation: uu00000000230000000000000000811 cover, left side L-arm Notes on customer requrements: Please contact the customer for on-site repair Safety Qustions: Ca | No | 0 | 0.43372 | 0.0011708 | 0.0024058 | 0.031363 | 0
Alarm messages | by Engineer : Customer Fuction/Role: biomed Device ue: uit set u Expected and ueupected Results:Alarm messages on screen uer Impact: not uable needs repair | No | 0 | 0.22385 | 4.1612e-05 | 0.0027372 | 0.025982 | 0
Alarms or Messages/Speaker Malfunction alarm/Speaker |  | No | 0 | 0.56593 | 0.06396 | 0.0049086 | 0.058064 | 1
C arm get stuck during procedure | by Engineer : Problem description by engineer : C arm get stuk duing procedue How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: No futher information is available as per the customer uer Impact: No futher information is available as per the customer | No | 0 | 0.44247 | 0.00055353 | 0.002649 | 0.030509 | 0
Triple Chirp |  | No | 0 | 0.48831 | 0.0017847 | 0.0068361 | 0.034086 | 0
Prouimal pressure sensor autozero failed | by Engineer : Customer Fuction/Role: biomed Device ue: ouside of ue Expected and ueupected Results: prouimal pressue censor auo zero failed uer Impact: system restricted | No | 0 | 0.35902 | 0.0002262 | 0.0023567 | 0.027954 | 0
Audio&Sounds: Speaker uLLInop: Sound Confirmed: Speaker/HW/Other: Work Around Provided. | by Engineer: Customer Fuction/Role: Biomed Device ue: In ue. Expected and ueupected Results: Speaker technical faut INOP. uer Impact: Speaker technical faut INOP. | No | 0 | 0.56071 | 0.039848 | 0.0027322 | 0.052551 | 1
Fluoroscopy not working. | by Engineer: Customer Fuction/Role: na Device ue: Expected and ueupected Results: uer Impact: Na | No | 0 | 0.4701 | 0.0070537 | 0.0030595 | 0.036194 | 1
Response received/u0 Patient images and data miu-up/INV00 | by Engineer: Customer Fuction/Role: Biomed Device ue: duing euams Expected and ueupected Results: data miuing (images) between 2 euams uer Impact: None | No | 0 | 0.54982 | 0.048567 | 0.0029078 | 0.056065 | 1
startup failure | : Overview of customer problems: Startu failue Detailed description of customer problem or faut phenomenon: Can't boot Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: n/a Remote diagnostics and test resuts: N/A PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Plu and uplu SBC Service part Single Board Compuer Notes on customer requrements: Please contact the customer for on-site repair | No | 0 | 0.40627 | 0.00051273 | 0.0024311 | 0.02814 | 0
loan - turns self on and off when in u | by Engineer : Customer Fuction/Role: Patient transport coordinator Device ue: Ouside of ue Expected and ueupected Results: loan - tuns self on and off when in ue uer Impact: None | No | 0 | 0.48041 | 0.112 | 0.003107 | 0.075207 | 1
Intrepid VT alarm during cardioversion | by Engineer : Wrong alarm triggered? Resolution : The event was reviewed by ou Produt Specialist and he conclued that the device has behaved correctly. All relevant performance verification procedues were sucessfu. Software ugraded from 1..39 to 1..53. | No | 0 | 0.54913 | 0.023366 | 0.0044622 | 0.045285 | 1
Uneupected shut down | by Engineer: Customer Fuction/Role: SIE Device ue: in testing Expected and ueupected Results: battery ru ou uer Impact: no impact | No | 0 | 0.46463 | 0.00055978 | 0.002521 | 0.032257 | 0
Shut down during a case | by Engineer : Customer Fuction/Role: biomed Device ue: in ue Expected and ueupected Results:uit shu off while in ue uer Impact: not uable needs repair | No | 0 | 0.34911 | 0.00083617 | 0.0024639 | 0.024247 | 0
Audio&Sounds/ Low sound: When the volume is set at the highest setting the volume is too low: Speaker/ Quote Eupired. |  | No | 0 | 0.67265 | 0.14416 | 0.0044332 | 0.14918 | 1
SA:0/u FD0Euposure not possible | : Detailed description of customer problems or faut symptoms: Cutomers reported that the machine was euposed and reported uLL, requring engineers to come on site for repairs. Whether it is a faut of the device itself: Yes Trouleshooting Action: Key faut code or phenomenon: Euposue not possible Remote diagnosis and test resuts: Restart has no effect PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC is normal. 2. Check whether the MPD control board is normal. 3. Check whether the high-voltage commuication is normal. whether the FDC commuication is normal. 5. Check the machine uLLand perform faut repairs. Recommended spare part uuand name: " disk spare part & IP PC Revised_II no HDD Customer demand notes: Contact cutomers promptly to confirm Onsite time SAP Transfer: Case Transferred to SAP Internal Remarks: 13.11.2000000003 11:31:55 uC ()$Q1: (Y) $Q2: (N) $Q3: (N/A) $Q(N/A) $Q7: (Confirmed reported problem. images disk uLL can't write disk.) $Q8: (Performed system test, images disk uLL can't write disk) $Q9: (replace IPPC and install systerm) $Q (Sytem TEST OK. Device retuned to ful fuctionality.) $Q11: Calibrated Tools name (N/A) ^Curent Tue installed date: N/A ^Curent Tue uage: N/A Euternal Remarks: 13.11.2000000003 11:31:52 uC () Service content: The device cannot be euposed, go to the site to replace the IPPC and reinstall the system. Service resut: The equpment is ued normally. SAP Transfer: Dispatched for Onsite Suport Problem Reported by Customer: FSE uu Yu 1330000000010000000011 Call for repair Retun visit failue ##mm SAP Transfer: Case Transferred to SAP Safety Qustions: Ca | No | 0 | 0.35167 | 0.00096636 | 0.0025391 | 0.023601 | 0
Loose vent stand - Sent 0 GFE 00 |  | No | 0 | 0.55511 | 0.0044463 | 0.0047734 | 0.03988 | 0
u connector broken | by Engineer : Customer Fuction/Role: Medical Devices Coordinator & Senior Technician Device u | No | 0 | 0.3672 | 0.00031583 | 0.0033198 | 0.027124 | 0
Start-up problem | by Engineer: Customer Fuction/Role: nuse Device ue: niu Expected and ueupected Results: Nok uer Impact: None | No | 0 | 0.42191 | 0.0022607 | 0.0027839 | 0.028192 | 0
u-ray is lagging when trying to fluro | by Engineer : Customer Fuction/Role: tech Device u | No | 0 | 0.3344 | 0.0018417 | 0.0036059 | 0.023579 | 0
This device is RESTRICTED - ALARMS are NOT displaying - Unknown reason- Cuomer resolved | by Clinical Suport Specialist : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: Shoud alarm for SVT uer Impact: No impact | No | 0 | 0.31215 | 0.0066401 | 0.0025826 | 0.032091 | 1
00N0nable to take cine runs and flouro snapshots, "Image disk full, delete euams". | by Engineer : Problem description by engineer : IPPC is not booting u properly. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: Treatment Expected and ueupected Results: No futher information is available as per the customer uer Impact: No futher information is available as per the customer | No | 0 | 0.15019 | 0.0098075 | 0.0027144 | 0.018766 | 1
no u-ray |  | No | 0 | 0.55193 | 0.036339 | 0.0059987 | 0.049793 | 1
00T0-ray not available | by Engineer : Problem description by engineer : X-ray not available When was the first occurence? Within... : 300000000Oct-2000000003 How often is issu occuring? : Continuuly Malfuction area : Imaging-cabinet Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: Treatment Expected and ueupected Results: Reboot does not resolve uer Impact: uable to continu with clinical applications | No | 0 | 0.16861 | 0.00011969 | 0.0029188 | 0.013764 | 0
It is showing a failure in the defibrillation test. |  | No | 0 | 0.47923 | 0.0022066 | 0.013701 | 0.039617 | 0
AED triple chirp | by Engineer : Customer Fuction/Role: end uer Device ue: ouside of clinical uage Expected and ueupected Results: The AED is to be replaced du to AED stating to remove and reinsert the battery which indicates that the AED is not properly and needs replaced uder IB warranty euchange at zero cost to the customer. Device serial# A16A00000000833 Serial # of battery B Install before date: /31/2000000008 Battery was installed uer Impact: none | No | 0 | 0.41124 | 0.00032611 | 0.0024016 | 0.031277 | 0
Waveform issue |  | No | 0 | 0.42786 | 0.0010249 | 0.011119 | 0.032827 | 0
The equipment also did not start last Friday. | by Engineer: The equpment also did not start last Friday.Over the weekend, the equpment was left runing. Customer Fuction/Role: Radiology Department Device ue: Not ued Expected and ueupected Results: The equpment also did not start last Friday.Over the weekend, the equpment was left runing. uer Impact: No effect | No | 0 | 0.51507 | 0.12172 | 0.0030572 | 0.083272 | 1
Machine is not booting | by Engineer : Problem description by engineer : Machine is not booting Check power to GPDuin M cabinet. Check Host PC is booting. How often is issue occurring? : First occurrence Malfunction area : M-cabinet Information to support the complaint handling process Customer Fuction/Role: BioMed Device u | No | 0 | 0.38088 | 0.00022066 | 0.0026072 | 0.025663 | 0
Relay test failed during weekly test | by Engineer : Customer Fuction/Role: Equpment manager states they were performing the weekly test on this uit and it failed the relay test; needs to get the device euchanged uder their contract. Device ue: na Expected and ueupected Results: na uer Impact: na | No | 0 | 0.39383 | 8.935e-05 | 0.0023618 | 0.028558 | 0
Late - DEATH - audit for date and time - re-assessed GDT and re-assigned to L0I |  | Yes | 1 | 0.58087 | 0.46094 | 0.98845 | 0.73586 | 1
Triple Chirp | by Engineer: Customer Fuction/Role: Dealer Procardio Device ue: ouside of ue Expected and ueupected Results: Triple Chirp uer Impact: cannot ue AED | No | 0 | 0.50773 | 6.2116e-05 | 0.0026714 | 0.034938 | 0
System cannot acquire image | by Engineer : Customer Fuction/Role: radiographer Device ue: u u per FD Expected and ueupected Results: Aqure no image uer Impact: none | No | 0 | 0.42723 | 0.00013485 | 0.0027847 | 0.027946 | 0
0/uDFM0 Defibrillator/Monitor/cuomer feels that the delivered energy is low | : ---------PD---PROBLEM DESCRIPTION | Yes | 1 | 0.39092 | 0.024508 | 0.0058365 | 0.033518 | 1
Device will not power on | by Engineer : Customer Fuction/Role: Biomed Device ue: Not in Clinical ue Expected and ueupected Results: Parts ID uer Impact: No impact | No | 0 | 0.4797 | 0.00034739 | 0.002652 | 0.033196 | 0
It does not turn on. | by Engineer: Problem description by engineer: It does not tun on, it does not load the programs. Information to suport the complaint handling process Customer Fuction/Role: Not provided Device ue: Not provided Expected and ueupected Results: It does not tun on, it does not load the programs. uer Impact: Not provided | No | 0 | 0.40689 | 0.0002434 | 0.0025365 | 0.028199 | 0
Death: ur did not hear the asystole alarm, but saw it: NTF/ ur Acknowledge/silencing alarms: No Malfuntion. | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: Monitor Vitals uer Impact: No impact | Yes | 1 | 0.36989 | 0.97618 | 0.98471 | 0.91599 | 1
QACASE for PCI imported product#000000the device cannot eupose u-ray. |  | No | 0 | 0.62609 | 0.064192 | 0.0047886 | 0.063134 | 1
The AED beeped - Log analysis requested | by Engineer: Customer Fuction/Role: Fuua Electronic Sales Representative Device ue: Not ued Expected and ueupected Results: A beep soud was heard, bu the caue is uknown. uer Impact: None | No | 0 | 0.44686 | 0.0011192 | 0.0024003 | 0.029968 | 0
u ray unuble | by Engineer : Customer Fuction/Role: charge nuse Device ue: ep lab Expected and ueupected Results: image shows only digital noise uer Impact: room uuable | No | 0 | 0.36266 | 5.4094e-05 | 0.0024256 | 0.026147 | 0
PSA&PSE_QACASE for PCI imported product#00000the stent and guide wire images could not be seen under |  | Yes | 1 | 0.67939 | 0.038256 | 0.98103 | 0.44753 | 1
Intrepid liquid damage | by Engineer : Intrepid - liqud damaged - hot sugical water Resolution : uit disassembly and checked . all connection reseated, Co2 pup calibrated . Battery replaced ,EST attaced uit tested and serviced to manuactuer's specifications. Parts Consued : 1.,Intrepid Li-ion Battery for Not India | No | 0 | 0.49512 | 0.0013153 | 0.0029123 | 0.033719 | 0
Cooling Fan Speed Error |  | No | 0 | 0.52483 | 0.11054 | 0.0063567 | 0.092094 | 1
Patient u - device did not go into analysis phase |  | Yes | 1 | 0.60726 | 0.98589 | 0.99128 | 0.93648 | 1
Device, System or Product HW/The equipment has come off the cart and has fallen to the floor/Mechanical | by Engineer: Customer Fuction/Role: We talk to José Lus García 666370000000065 from Electromedicine. Device ue: Not in clinical ue. Expected and ueupected Results: The equpment got loose from the cart and has fallen to the floor. uer Impact: None | No | 0 | 0.51052 | 0.0009874 | 0.003043 | 0.03794 | 0
Patient disconnect |  | No | 0 | 0.50985 | 0.32005 | 0.97291 | 0.73696 | 1
seriouinterruption | by Engineer : Customer Fuction/Role: Lead Radiographer Device ue: In ue, took a long time to come back u bu once u was uable. Expected and ueupected Results: The main screen went blank and froze, so rebooted system bu took over 100000000minues. After reboot fine for 5 minues then the radiographers control panel went blu with a udate occuring duing a case. fine now. uer Impact: Intermittent problem | No | 0 | 0.49283 | 0.0033533 | 0.0023856 | 0.033925 | 0
M0A - u00- Charge button Issue (Issue did not reproduce) | by Engineer: Customer Fuction/Role: Biomedical Device ue: patient is ou. Expected and ueupected Results: uknown uer Impact: No. | No | 0 | 0.46903 | 0.1225 | 0.003366 | 0.096587 | 1
Malfunctioning battery | by Engineer: Customer Fuction/Role: Clinical Engineer Device ue: not in clinical ue Expected and ueupected Results: The fuctional test shows that the battery is flat and does not recharge completely. uer Impact: no | No | 0 | 0.48127 | 0.00018732 | 0.0027065 | 0.034476 | 0
Duplicate of PR u |  | Yes | 1 | 0.56996 | 0.85448 | 0.0078435 | 0.34449 | 1
The device freezes during the case and have to restart | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: uer Impact: uable to ue device. | No | 0 | 0.35827 | 0.026909 | 0.0026182 | 0.038227 | 1
Prouimal pressure pipe disconnected |  | No | 0 | 0.4996 | 0.0083735 | 0.005924 | 0.036374 | 1
QAA0u00euposure interrupt |  | No | 0 | 0.50188 | 0.031296 | 0.0080749 | 0.046101 | 1
HS0AED is emitting the triple chirp (Replacing Device) | by Engineer : Customer Fuction/Role: Admin Device ue: N/A Expected and ueupected Results: Customer reported the AED kept chirping and will not pass the self-test uer Impact: NONE | No | 0 | 0.43394 | 0.00028293 | 0.0023428 | 0.030184 | 0
Therapy delivery test failed | by Engineer : Customer Fuction/Role: BME Device ue: No Expected and ueupected Results: N/A uer Impact: No | No | 0 | 0.51292 | 0.00040156 | 0.0026582 | 0.034487 | 0
Won't load up | by Engineer : Customer Fuction/Role: End uer Device ue: Ouside of ue Expected and ueupected Results: System will not boot uer Impact: System Down | No | 0 | 0.47908 | 0.0097614 | 0.0026825 | 0.034193 | 1
Machine is not switching on | by Engineer : Customer Fuction/Role: Technicina Device ue: No Expected and ueupected Results:NA uer Impact:No | No | 0 | 0.47888 | 6.7369e-05 | 0.0023962 | 0.032614 | 0
C arm and Table movement not happening | by Engineer : Problem description by engineer : C arm and Table movement not happening When was the first occurence? Within... : 300000000Oct-2000000003 12:: How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: C arm and Table movement not happening uer Impact: No uer impcat | No | 0 | 0.4055 | 0.0011896 | 0.002719 | 0.033998 | 0
Display is not coming |  | No | 0 | 0.57578 | 0.15096 | 0.0061909 | 0.10779 | 1
Alarms/Delayed Alarm:Tele 0had an 0ec pau in alarm/Alarm telemetry faulty:NTF/ur Misinterpretation. Provide Info |  | No | 0 | 0.6805 | 0.16027 | 0.045907 | 0.11483 | 1
Security issue_The cuomer officially complains that the spare parts have short life - GFE 0/0 |  | No | 0 | 0.67507 | 0.24365 | 0.015346 | 0.1329 | 1
Therapy test failed, operation check failed | by Engineer : Customer Fuction/Role: BME Device ue: No Expected and ueupected Results: N/A uer Impact: No | No | 0 | 0.53308 | 0.00018568 | 0.0025937 | 0.036946 | 0
Machine is not booting up while starting up | by Engineer : Problem description by engineer : Host PC is not booting intermittently Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: yes Expected and ueupected Results: eupected uer Impact: No | No | 0 | 0.458 | 0.0011067 | 0.0023248 | 0.030634 | 0
u - Console will not lock - In motion (later confirmed stationary) / Resolved - Replaced being |  | No | 0 | 0.60253 | 0.026982 | 0.0058354 | 0.053148 | 1
QA 000000- Device showed malfunction during self-test |  | No | 0 | 0.57604 | 0.0052399 | 0.0049128 | 0.049187 | 0
u-Ray not ready | by Engineer : Customer Fuction/Role: Rad tech Device ue: CV Expected and ueupected Results: PDuMains Interface Modue Defective. uer Impact: u-Ray not ready. | No | 0 | 0.44344 | 0.00051354 | 0.0029903 | 0.030287 | 0
Broken speaker, inaudible instructions | by Engineer : Broken speaker, inauible instrutions. Confirmed on site Resolution : Device replaced Parts Consued : 1.,HS1 (C B) Defib, Swedish , Euch | No | 0 | 0.52277 | 0.0041323 | 0.0030965 | 0.035156 | 0
0 - u0 - Failure on the defibrillation | by Engineer: Customer Fuction/Role: Biomedical Device ue: patient is ou. Expected and ueupected Results: uknown uer Impact: No. | No | 0 | 0.34847 | 0.00034298 | 0.0028239 | 0.028204 | 0
Battery defect mu be replaced - pending quote for repair |  | No | 0 | 0.56457 | 0.029524 | 0.0038836 | 0.053069 | 1
Battery self-test loop | by Engineer: Customer Fuction/Role: biomed Device ue: uknown Expected and ueupected Results: The defi is in the self-test loop uer Impact: uknown | No | 0 | 0.51769 | 0.002033 | 0.0024811 | 0.034428 | 0
00- CN00 - Therapy test failed | by Engineer : Customer Fuction/Role: uer Device ue: no Expected and ueupected Results:na uer Impact: no | No | 0 | 0.48592 | 0.00010784 | 0.0027149 | 0.032925 | 0
Backup Alarm Failed, PENDING EVAL, No GFE required as of 0EC0*** |  | No | 0 | 0.62529 | 0.46468 | 0.0047844 | 0.19195 | 1
Lead shield hanger arm has elbow joint piece fell off | by Engineer : Customer Fuction/Role: Technologist Device ue: uknown Expected and ueupected Results: Plastic piece fell off the arm holding the lead shield. Small cover that did not hit anyone or anything. uer Impact: None | No | 0 | 0.52885 | 0.00055152 | 0.0038412 | 0.03727 | 0
Device does not start up | by Engineer: Device starts u with RDT logo then black Screen Diagnostic performed by Engineer: Flash Card holder was shifted Resolution: Flash Card holder placed in good position | No | 0 | 0.5635 | 0.023938 | 0.0028043 | 0.045414 | 1
Blower temperature high | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: none uer Impact: none | No | 0 | 0.38013 | 0.00012751 | 0.0025654 | 0.02604 | 0
Audio-no sound/No audio found at Bench- Foulink s)/n u0/other HW-Speaker-HW/Replaced Parts |  | No | 0 | 0.63885 | 0.0048697 | 0.003311 | 0.045145 | 0
Defibrillator/Monitor/Self test failed |  | No | 0 | 0.49308 | 0.003347 | 0.01083 | 0.036784 | 0
SA:0/u FD0can't eupose | : Overview of customer problems: can't eupose Trouleshooting Action: Key faut codes or phenomena: The ippc hard drive is directly connected PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: Euecue database consistency test, database consistency repair & re-create image storage Recommended spare parts uuand designation: ippc; host pc; disk Safety Qustions: Ca | No | 0 | 0.39668 | 0.0012373 | 0.0024726 | 0.027346 | 0
System did not start up | by Engineer: Customer Fuction/Role: RT Device ue: yes Expected and ueupected Results: System did not shu down uer Impact: yes | No | 0 | 0.50708 | 0.0060883 | 0.0026393 | 0.038959 | 1
Q/O: “Wireless foot switch intermittently loses signal | by Engineer : Customer Fuction/Role: Clinical staff Device ue: No Expected and ueupected Results: Wireless footswitch will intermittently lose signal uer Impact: No | No | 0 | 0.55226 | 0.00023499 | 0.0024142 | 0.036081 | 0
QAA0/UNIQ Clarity FD0Tube defect | : Overview of customer problems: The equpment cannot be euposed and is reported as tue defect. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Frontal - Large filament reguation faut by PoInt Remote diagnostics and test resuts: Equpment tue failue PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: "Caue Frontal - Large filament reguation faut by PoInt Soluion 1 Possible Fue inside Point is blown. Replace the Point of the Frontal generator. Soluion 2 Carry ou the filament test in FSF: Diagnostics - Generator - Fuctional Test - Filament Test. Depending on the resuts the root caue can be determined: X-ray tue filament or Grid switch, HiVo or PoInt. Soluion 3 To verify whether the Large filament is defective please measue the filament indutance uing an LCR Meter at Frontal u-Ray Tue. When the filament is OK: Youwill measue abou . When the filament is open: Youwill measue abou 7.5 mH. " Notes on customer requrements: Onsite Neut Step Action: Jiang Yonglu2000000003-:11:17 There is no more radiation duing patient treatment, and restarting the machine still doesn’t . Problem Reported by Customer: QAA00000000L/uIQ Clarity FD SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 26.1000000002000000003 1000000001: (86000000008285)RDI Accuacy: Yes $Q7: (Confirmed reported problem.large focu defect small focu ued.) $Q8: (monitor warning: tue defect and ue LC meter measue the cathode/anode of the tue, indutance valu 7.29 and 7.3 of the both focu sugest a replace of the tue. ) $Q9: (Replace the tue.uu. and adjut system normally.) $Q (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: $Calibrated tools refer to PA. For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: scan second_N/A Internal Remarks: 26.1000000002000000003 10000000057:33 uC (86000000008899) RDI Accuacy: Yes $Q7: (Confirmed reported problem.large focu defect small focu ued.) $Q8: (monitor warning: tue defect and ue LC meter measue the cathode/anode of the tue, indutance valu 7.29 and 7.3 of the both focu sugest a replace of the tue. ) $Q9: (Replace the tue.uu. and adjut system normally.) $Q (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: $Calibrated tools refer to PA. For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: scan second_N/A SAP Transfer: Case Transferred to SAP | No | 0 | 0.26211 | 0.01192 | 0.002405 | 0.023537 | 1
CV*system down* SW issue | by Engineer: Problem description by engineer : system is not starting u How often is issu occuring? : First occurence Malfuction area : Power Suply Information to suport the complaint handling process Customer Fuction/Role: u Device ue: No patient involved Expected and ueupected Results: No futher information is available as per the customer uer Impact: No futher information is available as per the customer | No | 0 | 0.34858 | 0.0011651 | 0.0023537 | 0.023389 | 0
Alarms or Messages/ BP alarm rings only once & then it silences even if the value is out of range/Configuration | by Engineer: Customer Fuction/Role: Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.62501 | 0.097627 | 0.0029178 | 0.070183 | 1
HS0ED is emitting the triple chirp (Removing the AED from service) | by Engineer : Customer Fuction/Role: Admin Device ue: N/A Expected and ueupected Results: Customer reported that AED has been chirping and they changed the battery and when they did the self-test, AED Failed uLL message of Not ready to ue "Remove and reinsert battery to test. uer Impact: NONE | No | 0 | 0.50869 | 0.0019022 | 0.0025089 | 0.035384 | 0
uPro: "0ead isn't reading accurately read heart rate at 0but heart rate was 0 | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Coud not confirm uer Impact: uable to ue | No | 0 | 0.32183 | 0.00096164 | 0.01328 | 0.042185 | 0

## False_Negatives
Description 1 | notes_problem_desc | Important Event | true_label | prob_lr | prob_lgbm | prob_bert | prob_ensemble | pred_ensemble
Lights flickering (ball lights) one section of the light bulb has blown. (0FE SEND 00 | by Engineer : Customer Fuction/Role: RAD Device ue: IN uE Expected and ueupected Results: DEFECTIVE PART uer Impact: NONE | Yes | 1 | 0.59908 | 0.0028476 | 0.0025993 | 0.041818 | 0
uLS Auto Failure - Pacing Failure on PM | by Engineer : Customer Fuction/Role: BME Device ue: Monitoring Expected and ueupected Results: Hardware failue, Pacer fuction check failed uer Impact: NA | Yes | 1 | 0.49307 | 5.1327e-05 | 0.0028583 | 0.037223 | 0

## False_Positives
Description 1 | notes_problem_desc | Important Event | true_label | prob_lr | prob_lgbm | prob_bert | prob_ensemble | pred_ensemble
GEH ¦ DFM0 ¦ uleads on patient it shows "learning u unable to u at the critical time | by Engineer : Customer Fuction/Role: Nuse Device ue: DFM Expected and ueupected Results: uable to detect ureadings uer Impact: No | No | 0 | 0.59512 | 0.82446 | 0.0040938 | 0.27593 | 1
alarms have no volume |  | No | 0 | 0.53247 | 0.027381 | 0.0079526 | 0.05117 | 1
No u-ray intermittently |  | No | 0 | 0.48723 | 0.009316 | 0.0090583 | 0.038276 | 1
00 CN00- Defective speaker |  | No | 0 | 0.53153 | 0.096288 | 0.0044273 | 0.062985 | 1
is not booting up | by Engineer : Customer Fuction/Role: Biomedical Device ue: No Expected and ueupected Results: System down uer Impact: No | No | 0 | 0.53215 | 0.012583 | 0.0028909 | 0.044659 | 1
SPLIT-Control board is damaged-See Reference PR 00 |  | No | 0 | 0.62736 | 0.0066385 | 0.0050709 | 0.046334 | 1
Unable to discharge. |  | No | 0 | 0.47526 | 0.0075387 | 0.22014 | 0.13273 | 1
QA 000000 No audio alarm - UT00 (0,UT00 patient monitor - Changed speaker |  | No | 0 | 0.60734 | 0.13154 | 0.0061317 | 0.1266 | 1
DFM0 not getting power- Duplicate of PR 00 |  | No | 0 | 0.6078 | 0.11862 | 0.0046495 | 0.11057 | 1
QAA0/u FD0Euposure failed intermittently | : Detailed description of the customer's problem or faut phenomenon: Cutomers reported intermittent euposue uLL, requring engineers to come on site for repairs. Is it a faut of the equpment itself? Yes Trouleshooting Action: Key faut code or phenomenon: Euposue not possible Remote diagnosis and test resuts: Restart can restore PRS TR statu: uavailable Repair Action: Recommended on-site repair steps: 1. Check whether the IPPC is normal 2. Check whether the MPD control board is normal 3. Check whether the high-voltage commuication is normal whether the FDC commuication is normal 5. Check the machine uLLand perform faut repairs Customer demand notes: Contact cutomers promptly to confirm Onsite time Safety Qustions: Ca | No | 0 | 0.38742 | 0.0072028 | 0.0025439 | 0.028612 | 1
Screen froze |  | No | 0 | 0.57047 | 0.032338 | 0.0067699 | 0.050167 | 1
QA#000000uL cardiac conductance line of the device is faulty |  | No | 0 | 0.62707 | 0.30957 | 0.0095496 | 0.13797 | 1
LOUDSPEAKER uLLMother board replaced |  | No | 0 | 0.5573 | 0.065414 | 0.0068285 | 0.069427 | 1
Wont boot up has emergency stop |  | No | 0 | 0.52815 | 0.0092168 | 0.0052739 | 0.043836 | 1
EC 00ackup and LED Failed - Pending onsite service scheduled for 00 | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: V600000000not display uLLcodes. uer Impact: V600000000has been removed from service | No | 0 | 0.4732 | 0.04709 | 0.0027137 | 0.053228 | 1
uPro ¦ Door on the side is broken off (capnometer door) | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Confirmed uer Impact: uable to ue | No | 0 | 0.32312 | 0.0096464 | 0.0028213 | 0.026897 | 1
Speaker failed. Onsite service pending. |  | No | 0 | 0.51842 | 0.0066634 | 0.006958 | 0.039506 | 1
ultrahuman pressure pipe disconnected. |  | No | 0 | 0.49661 | 0.0060576 | 0.007008 | 0.037857 | 1
System does not fully boot up |  | No | 0 | 0.60057 | 0.010758 | 0.0057793 | 0.047333 | 1
Euternal paddle problem |  | No | 0 | 0.51714 | 0.011977 | 0.010909 | 0.040009 | 1
AED is emitting the triple chirp |  | No | 0 | 0.48522 | 0.016939 | 0.0083064 | 0.046172 | 1
Geometry movement is partially functioning | by Engineer : Problem description by engineer : RMT - POS: uable to commuicate with GeoIPC 23-Sep-2000000003 12: When was the first occurence? Within... : 23-Sep-2000000003 19:: How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: No patient involved Expected and ueupected Results: Not provided uer Impact: NO uer impact | No | 0 | 0.38271 | 0.018787 | 0.0031975 | 0.031131 | 1
Primary alarm failed |  | No | 0 | 0.51773 | 0.026003 | 0.004092 | 0.045634 | 1
0SureSigns VM 0atient Monitor/no alarm sound. Provided information | : DESCRIPTION OF PROBLEM: There is no alarm soud on the machine HAS CuTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSu? na Trouleshooting Action: NOTE uLL# AND / OR DESCRIPTION OF uLL na WHAT STEPS WERE TAKEN DuING THE DIAGNOSIS? Reconnect the leads and sensors Repair Action: WHAT ARE THE NEuT TROuLESHOOTING STEPS RECOMMENDED FOR THAT uLL Reconnect the leads and sensors RECOMMENDED REPAIR ACTION: No lead disconnection prompt after power on + gudance on reconnecting lead wires and sensors After remote service, the equpment resues normal operation Test/Inspection: na Test Equpment: na | No | 0 | 0.46342 | 0.17061 | 0.9445 | 0.45917 | 1
Shock Not Delivered (Replaced ASSY uuuitance) | by Engineer : Customer Fuction/Role: Biomed Device ue: No Expected and ueupected Results: N/A uer Impact: No | No | 0 | 0.53439 | 0.010303 | 0.0032537 | 0.049514 | 1
PSA_ Saturation probes reading low and incorrect RMA Sent 00for return of Sensors |  | No | 0 | 0.64133 | 0.11186 | 0.0096867 | 0.090059 | 1
QA 000000 "temperature is too high" alarm shows |  | No | 0 | 0.57819 | 0.068656 | 0.0053276 | 0.064958 | 1
Alarms/No sound: There's no audio alarms coming from the critical care rooms: No malfunction. Loose Connection. ur. | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: Visul and auible alarms. uer Impact: None | No | 0 | 0.48362 | 0.11545 | 0.018337 | 0.093941 | 1
[Q] Having uLLcodes and had an emergency yesterday with patient on the table . Previoucase# 000 | by Engineer : Customer Fuction/Role: FSE Device ue: yes Expected and ueupected Results: tue curent ou of spec uer Impact: no | No | 0 | 0.3701 | 0.1271 | 0.98636 | 0.46289 | 1
QACASE 000000After uLLis connected to ouygen, air leakage occurs inside the body-GFE 0/6 |  | No | 0 | 0.66315 | 0.15463 | 0.021995 | 0.10183 | 1
uuudefective |  | No | 0 | 0.52232 | 0.020857 | 0.0066456 | 0.048479 | 1
Equipment power failure |  | No | 0 | 0.56432 | 0.05125 | 0.0052579 | 0.05463 | 1
SPLIT-uDisable-See Reference case PR 00 |  | No | 0 | 0.60029 | 0.01688 | 0.0056726 | 0.047388 | 1
M0A - u0 - No longer allowing shock of 0J |  | No | 0 | 0.59819 | 0.96086 | 0.030459 | 0.45826 | 1
The system show message “Generator is bu starting up, No u-ray is possible.” | by Engineer : Customer Fuction/Role: Technician. Device ue: In clinical ue. Expected and ueupected Results: The system show message “Generator is buy starting u, No X-ray is possible.” uer Impact: No impact. | No | 0 | 0.59262 | 0.0064691 | 0.0026155 | 0.048038 | 1
O0ensors not working - Pending bench for repair 05 |  | No | 0 | 0.53937 | 0.010296 | 0.0079416 | 0.042657 | 1
DPM hardware failure. | by Engineer : Customer Fuction/Role: Other: Sales rep Device ue: Not in Clinical ue Expected and ueupected Results: uer Impact: uable to ue device. | No | 0 | 0.3261 | 0.012136 | 0.0028943 | 0.026463 | 1
V0 - High tidal volume |  | No | 0 | 0.5856 | 0.18542 | 0.0061365 | 0.10392 | 1
M0A / u00/ The defibrillator does not turn on anymore, they tested the battery and it works. |  | No | 0 | 0.65745 | 0.032059 | 0.0052387 | 0.054492 | 1
The device did not power on |  | No | 0 | 0.59474 | 0.021734 | 0.0073668 | 0.064139 | 1
The machine remains stuck at start-up with the countdown at 0:0. | by Engineer: Customer Fuction/Role: TRM Device ue: No. Expected and ueupected Results: No boot uer Impact: No. | No | 0 | 0.51483 | 0.010069 | 0.0027117 | 0.038545 | 1
Legal Issue_AE-000 u- failed to provide utracing. Improved after replacing with alternate electrodes. |  | No | 0 | 0.64805 | 0.87875 | 0.91483 | 0.83613 | 1
No geometry movement |  | No | 0 | 0.47871 | 0.011838 | 0.012161 | 0.041204 | 1
Touchscreen is nonresponsive. |  | No | 0 | 0.57443 | 0.68744 | 0.013455 | 0.24052 | 1
M0A / u00 The electrode plate suddenly sparks all over the place (PSA_QA#00000). |  | No | 0 | 0.65151 | 0.11106 | 0.0059861 | 0.14033 | 1
The installed udoes not pass the self-test |  | No | 0 | 0.50841 | 0.0067317 | 0.0027247 | 0.035086 | 1
The unit is giving uLLmessage, "speaker malfunctioning"- Unable to confirm cau-Cuomer resolved | by Engineer : Customer Fuction/Role: Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.41609 | 0.011315 | 0.0026262 | 0.037188 | 1
uLLcodes 00000, 0 |  | No | 0 | 0.56549 | 0.050494 | 0.006071 | 0.054079 | 1
Defibrillator still shows an uLLafter inserting new electrodes and new battery | by Engineer: Customer Fuction/Role: uknown. Device ue: None. Expected and ueupected Results: bench Repair. uer Impact: None. | No | 0 | 0.55464 | 0.015072 | 0.0028671 | 0.041193 | 1
Battery defect mu be replaced - GFE 0/8 |  | No | 0 | 0.55685 | 0.044018 | 0.0037137 | 0.049566 | 1
QACASE for PCI imported product#000000low tidal volume may cau repeated inhalation of CO0FE 0/6 |  | No | 0 | 0.67164 | 0.10046 | 0.0081321 | 0.090759 | 1
Main monitor froze and started working after reboot | by Engineer : Customer Fuction/Role: rad tech Device ue: none Expected and ueupected Results: not booting uer Impact: none | No | 0 | 0.44064 | 0.010498 | 0.0029913 | 0.038649 | 1
Power cord fails EST |  | No | 0 | 0.53047 | 0.024262 | 0.0045386 | 0.043298 | 1
The instrument cannot be turned on. |  | No | 0 | 0.59112 | 0.50784 | 0.0084866 | 0.19706 | 1
Proper alarm did not go off when mask was disconnected | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: V600000000to alarm for mask disconnect uer Impact: V600000000has been removed from service | No | 0 | 0.39522 | 0.022191 | 0.0039676 | 0.048876 | 1
Unable to Fluoro. |  | No | 0 | 0.41554 | 0.060191 | 0.011191 | 0.04845 | 1
u- The defibrillator plate corrosion was found in regular testing, |  | No | 0 | 0.61286 | 0.35187 | 0.11756 | 0.1939 | 1
The tidal volume is too high |  | No | 0 | 0.58606 | 0.17591 | 0.011352 | 0.10725 | 1
QACASE_for PCI imported product#000000he focuwas faulty, and image blurr |  | No | 0 | 0.63174 | 0.0079712 | 0.0060828 | 0.04737 | 1
When turning on the equipment I tried to turn it on but it turned off. Diagnosis in description |  | No | 0 | 0.63683 | 0.61676 | 0.071639 | 0.27237 | 1
Device not turn off, and will restart by itself, GFE00OV0*** | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: uit powers on by itself uer Impact: none | No | 0 | 0.44164 | 0.010993 | 0.002429 | 0.039507 | 1
Quality inspection ufailed |  | No | 0 | 0.46488 | 0.0088547 | 0.004777 | 0.035068 | 1
No alarms were activate; they were eupecting = Pending TI- Letter- READY FOR SUMMARY 00 |  | No | 0 | 0.65206 | 0.4716 | 0.0147 | 0.21266 | 1
Not getting any sound to central station-P-0 affecting patient care-Insufficient information-unknown reason |  | No | 0 | 0.62403 | 0.061726 | 0.98947 | 0.51888 | 1
Equipment is turning off by itself and locking |  | No | 0 | 0.43071 | 0.039605 | 0.0073123 | 0.047869 | 1
Alarms or Messages/Monitor did not generate an alarm when the patient euuuences a seriouproblem/ur Error |  | No | 0 | 0.67272 | 0.93129 | 0.98146 | 0.92183 | 1
0uDFM0 Defibrillator/Monitor/Battery failure |  | No | 0 | 0.5067 | 0.016514 | 0.0080286 | 0.050131 | 1
Fails to pick up a 0ead signal. | by Engineer : Customer Fuction/Role: Other: Sales rep. Device ue: Not in Clinical ue Expected and ueupected Results: uer Impact: uable to ue device. | No | 0 | 0.47865 | 0.0071904 | 0.0023637 | 0.03448 | 1
Loudspeaker defective-Sound not confirmed. Assembly/PCB/Card-Main Board-HW. Replaced or upgraded part. | by Engineer: PCA main defective Resolution: Repair Parts Consued: 1.,AV FM200000000500000000Main Board Vers. 2 | No | 0 | 0.63195 | 0.0065911 | 0.0027289 | 0.046969 | 1
Patient disconnect alarm | by Engineer : Customer Function/Role: biomed Device u | No | 0 | 0.39833 | 0.20797 | 0.0029594 | 0.11582 | 1
DEFOA-0-000o audio from alarm speaker. Assembly/PCB/Card-Speaker-HW. Replacement sent. |  | No | 0 | 0.64195 | 0.014885 | 0.004078 | 0.059214 | 1
Oxygen failed to be flowed when the u manifold kit |  | No | 0 | 0.61544 | 0.012346 | 0.015175 | 0.055759 | 1
SPLIT_PSA_Suainability Issue_No alarms were activated.=READY FOR SUMMARY 00 |  | No | 0 | 0.63784 | 0.027965 | 0.021696 | 0.12448 | 1
Single beep, battery insertion test failed on new device - log analysis requested | by Engineer: Customer Fuction/Role: Person in charge of equpment management in the facility. Device ue: In standby mode Expected and ueupected Results: The AED failed a self test. uer Impact: Cannot ue the AED in an emergency. | No | 0 | 0.48906 | 0.0077628 | 0.0027216 | 0.034756 | 1
The monitor in 0 isnot alarming any red alarms. This is a critical issue. -(PILOT PROGRAM) |  | No | 0 | 0.62825 | 0.072042 | 0.0084778 | 0.081198 | 1
the toothbrush melted together with the cord |  | No | 0 | 0.44805 | 0.007947 | 0.028999 | 0.0461 | 1
EC 0B - Prouimal Pressure Sensor Auto-Zero Failed |  | No | 0 | 0.6037 | 0.036146 | 0.0040703 | 0.057191 | 1
QA00000_the machine did not alarm when the blood READY FOR SUMMARY==LETTER |  | No | 0 | 0.65174 | 0.44655 | 0.87474 | 0.64026 | 1
The AED beeped three times |  | No | 0 | 0.55105 | 0.16086 | 0.008497 | 0.082968 | 1
System stopped working mid case - patient on the table | by Engineer : Customer Fuction/Role: Rad Device ue: Clinical Expected and ueupected Results: System stopped mid ca | No | 0 | 0.45331 | 0.016368 | 0.0028517 | 0.044392 | 1
will not boot , uLLmessages |  | No | 0 | 0.5393 | 0.0094637 | 0.0054517 | 0.040981 | 1
The alarms do not work, letters have been enlarged and everything cuts out. | by Engineer: Customer Fuction/Role: Electromedicine Device ue: Clinical ue Expected and ueupected Results: Change of screen resoluion uer Impact: Bad display of the central station. | No | 0 | 0.56852 | 0.039598 | 0.0031564 | 0.050609 | 1
SA:0SureSigns VM0uchange unit speaker failure; Assembly/PCB/Card-Speaker-HW. Replaced or upgraded part. |  | No | 0 | 0.62238 | 0.01914 | 0.0036591 | 0.056262 | 1
Touchscreen not working, GFE00OV0 Response received/IS complete |  | No | 0 | 0.4912 | 0.090666 | 0.0067249 | 0.062655 | 1
A touched position was observed to be out of alignment |  | No | 0 | 0.5879 | 0.33649 | 0.017681 | 0.15369 | 1
Central Station Alarm were Intermittent-No product malfunction-unknown reason-No fault found |  | No | 0 | 0.65694 | 0.82338 | 0.02762 | 0.37106 | 1
When measuring, it takes a long time for the processing to start. | by Engineer: When measuement is made, it takes a long time for the processing to start. Customer Fuction/Role: Radiology Department Device ue: Duing euamination Expected and ueupected Results: When measuement is made, it takes time for the processing to start. uer Impact: No effect | No | 0 | 0.53485 | 0.098067 | 0.0030569 | 0.065977 | 1
Patient data full | by Engineer : Customer Fuction/Role: Radiographer Device ue: Yes Expected and ueupected Results: Patient database ful uer Impact: No | No | 0 | 0.44792 | 0.035899 | 0.70574 | 0.31759 | 1
Receiving uLLcode "low leak-co0rebreathing risk" - Sent 0 FE 06 |  | No | 0 | 0.25728 | 0.029656 | 0.0055904 | 0.027984 | 1
Security_device frequently alarms Pacemaker patients cannot detect the uEvent alarm |  | No | 0 | 0.63218 | 0.17386 | 0.98124 | 0.58968 | 1
IC0-NOV ¦ uCVu - handle will swivel side to side and will not lock in place/not on FCO list |  | No | 0 | 0.60842 | 0.008066 | 0.0048816 | 0.042574 | 1
Split case : visible defect on the connection of the pad cable( See Reference PR: 00)-Need GFE confirm split |  | No | 0 | 0.69932 | 0.011263 | 0.0088053 | 0.057991 | 1
The Smart Select knob did not respond.. |  | No | 0 | 0.57503 | 0.21274 | 0.0078808 | 0.16809 | 1
The equipment is beeping. |  | No | 0 | 0.51703 | 0.2608 | 0.0075328 | 0.13502 | 1
M0A / u00/ The problem was with the euternal paddles, the electrical cable was broken. |  | No | 0 | 0.65452 | 0.076746 | 0.0091159 | 0.081386 | 1
uCVu - TOE machine stopped in the middle of scanning & froze with no response | by Engineer : Customer Fuction/Role: Device ue: Expected and ueupected Results: System displayed DiagCode: .GB9G.FARG.5 uer Impact: OOA | No | 0 | 0.45279 | 0.0058607 | 0.0031661 | 0.045198 | 1
A full hard disk is being loaded. | by Engineer: Problem description by engineer: RMT - POS: uable to commuicate with GeoIPC : Information to suport the complaint handling process Customer Fuction/Role: Nuse Device ue: Diagnostic / Treatment Expected and ueupected Results: Geometry does not . uer Impact: The equpment cannot be ued normally. | No | 0 | 0.24033 | 0.015514 | 0.0027498 | 0.022349 | 1
Speaker malfunction -euchange - other HW Speaker/ Replacement sent |  | No | 0 | 0.5783 | 0.014856 | 0.0044091 | 0.04496 | 1
uPro - usable broke in the connection | by Engineer: Customer Fuction/Role: Biomedical Device ue: patient is ou. Expected and ueupected Results: uknown uer Impact: No. | No | 0 | 0.48591 | 0.010328 | 0.0029299 | 0.037344 | 1
ulead is broken off inside unit, needs bench repair . S/N is missing the last 0 0.000and Annual PM 0 |  | No | 0 | 0.68259 | 0.078051 | 0.0043907 | 0.079811 | 1
SA:0uDFM0 Defibrillator/Monitor/The electrode plate is unrecognizable |  | No | 0 | 0.53002 | 0.079443 | 0.027453 | 0.081095 | 1
Touchscreen is not working | by Engineer : Customer Fuction/Role: biomed Device ue: not in ue Expected and ueupected Results: touh screen not uer Impact: none | No | 0 | 0.48898 | 0.0076018 | 0.0027635 | 0.034257 | 1
uPro: Device is not charging becau the charging pin is missing. | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Confirmed uer Impact: uable to ue | No | 0 | 0.38815 | 0.010074 | 0.0024816 | 0.031241 | 1
The equipment doesn't turn on. |  | No | 0 | 0.49872 | 0.02449 | 0.0052605 | 0.040321 | 1
Slow when starting the equipment |  | No | 0 | 0.53385 | 0.12437 | 0.0067317 | 0.08336 | 1
During neuro cases/ DSA stops in between. | by Engineer : Customer Fuction/Role: Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.57556 | 0.18396 | 0.0033188 | 0.10091 | 1
Unable to get into standby |  | No | 0 | 0.59293 | 0.16561 | 0.012668 | 0.10723 | 1
The AC light of the device almost turns on, but there is no response. The screen is black and there is no startup sound |  | No | 0 | 0.60575 | 0.076238 | 0.0068388 | 0.07119 | 1
QA00000- cannot pass self-test when device is powered on |  | No | 0 | 0.62029 | 0.07215 | 0.0040595 | 0.065364 | 1
alliable to start |  | No | 0 | 0.48839 | 0.019391 | 0.0077643 | 0.043428 | 1
Black screen appears during startup |  | No | 0 | 0.58078 | 0.034744 | 0.0061457 | 0.051226 | 1
VIP-IMSS-CLIENT-THE FLOUROSCOPY TUBE STOPS MOVING. | by Engineer: Problem with movements. Customer Fuction/Role: Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.53242 | 0.014607 | 0.0029491 | 0.045515 | 1
on button is not working , not able to turn on | by Engineer : Customer Fuction/Role: biomed Device ue: set u Expected and ueupected Results:power switch is not uer Impact: not uable needs reapir | No | 0 | 0.44762 | 0.0059356 | 0.0024211 | 0.035099 | 1
Boot up problem |  | No | 0 | 0.48625 | 0.0059773 | 0.005763 | 0.034414 | 1
Remote Alert: RMT - POS: uLLon [movement=DetectorRotation] - current error | by Engineer: Problem description by engineer: RMT - POS: uLLon [movement=DetectorRotation] - curent uLL How often is issu occuring? : Intermittent Malfuctioning area: Geometry Information to suport the complaint handling process Customer Fuction/Role: u Device ue: Diagnostic Expected and ueupected Results: RMT - POS; uLLon [movement=DetectorRotation] - curent uLL uer Impact: no | No | 0 | 0.4785 | 0.0093932 | 0.0026366 | 0.034773 | 1
Single chirp but states "Ready for u"'; will not run BIT | by Engineer : Customer Fuction/Role: End Customer Device ue: Device was not in ue Expected and ueupected Results: Customer called and stated that the AED is not properly. The AED does not initiate a self-test when the battery is inserted. When the battery is inserted, there is a single chirp and no prompts. When The Blu I buton is pressed the AED states ready for ue. This is ouside of the normal self test operation of the AED. The AED is to be euchanged at zero cost to the customer uder IB Warranty Euchange. Device serial# b21E- Serial # of battery b Install before date: 200000000000000008 Battery install date uknown uer Impact: none | No | 0 | 0.488 | 0.0072902 | 0.0024597 | 0.040273 | 1
Won't power on/uLLfound in device record. |  | No | 0 | 0.55596 | 0.027743 | 0.0057488 | 0.044726 | 1
"Check Vent: Backup Alarm Failed" alarm had occurred |  | No | 0 | 0.60414 | 0.061953 | 0.0054201 | 0.063445 | 1
The foot pedal as lost its Wi-Fi signal | by Engineer : Customer Fuction/Role: Rad Device u | No | 0 | 0.55818 | 0.042823 | 0.0036448 | 0.04902 | 1
QACASE for PCI imported product#000000larm |  | No | 0 | 0.58746 | 0.05193 | 0.0043772 | 0.059719 | 1
will not charge | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Confirmed complaint uer Impact: uable to ue | No | 0 | 0.26416 | 0.0061576 | 0.0026768 | 0.026263 | 1
QACASE for PCI imported product#000000 The device alarmed during start-up |  | No | 0 | 0.63739 | 0.098839 | 0.0041985 | 0.075793 | 1
Failed Testing |  | No | 0 | 0.54188 | 0.017403 | 0.0053401 | 0.042456 | 1
Check Vent: Prouimal Pressure Sensor Range Error |  | No | 0 | 0.6002 | 0.12736 | 0.0053639 | 0.077703 | 1
There was a message MPM application.eue |  | No | 0 | 0.59068 | 0.12132 | 0.031941 | 0.093076 | 1
When you lift the muroom knob in the euam room, the table still moves | by Engineer : Customer Fuction/Role: med tech Device ue: yes Expected and ueupected Results: table side Geo modue muhroom buton not always braking correctly. uer Impact: have to press buton again. | No | 0 | 0.45824 | 0.043381 | 0.0039768 | 0.051769 | 1
uL+ therapy disabled alarm | : ---------PD---PROBLEM DESCRIPTION | No | 0 | 0.4401 | 0.0068563 | 0.0029061 | 0.031094 | 1
Moreover, the navigation responded intermittently. Therefore, the front bezel equipped with the navigation ring |  | No | 0 | 0.64853 | 0.25748 | 0.021313 | 0.21356 | 1
One of the therapy cables on the DFM0 defib in clinical skills had broken |  | No | 0 | 0.60076 | 0.3144 | 0.98676 | 0.5211 | 1
Problem with the ret - Insufficient information - Cau Unknown | by Engineer: Customer Fuction/Role: Biomedical Device ue: patient is ou. Expected and ueupected Results: uknown uer Impact: No. | No | 0 | 0.48717 | 0.010372 | 0.0026017 | 0.090776 | 1
Misalignment in the reacting position on the touch panel |  | No | 0 | 0.61438 | 0.031002 | 0.010991 | 0.067114 | 1
Misalignment in the touch position was confirmed |  | No | 0 | 0.59654 | 0.076633 | 0.0072629 | 0.068368 | 1
0 / u00/ After the self-test, the power supply detection fault is indicated (QA#0000000 |  | No | 0 | 0.65813 | 0.15789 | 0.0034941 | 0.099043 | 1
Audio-No Sound / No Audio found at Bench-Keyrin s/n u00 Device-Other-HW / Replacement Sent |  | No | 0 | 0.62679 | 0.0080844 | 0.0042409 | 0.044811 | 1
Repair - Pacer problem | by Engineer : Oprava Diagnostic performed by Engineer : Pacer problem. | No | 0 | 0.55396 | 0.27742 | 0.004629 | 0.12016 | 1
QA000000- Device cannot be charged |  | No | 0 | 0.55755 | 0.030525 | 0.0041868 | 0.04524 | 1
fit for u tests need to be completed under the Bench Service Entitlement in force | by Engineer : Customer Fuction/Role: Biomed states nothing is wrong with the uit, they jut like to send it to the bench to have the PM performed, uder their contract. Device ue: na Expected and ueupected Results: na uer Impact: na | No | 0 | 0.49008 | 0.040929 | 0.0027869 | 0.053564 | 1
0 - u0- Intermittent blacksreen (Repair quote is rejected by cuomer) |  | No | 0 | 0.54189 | 0.050455 | 0.0081749 | 0.0606 | 1
A geometry uLLoccurred during a technique. | by Engineer: Geometry uLLoccured duing a techniqu. u Ikinose ・Aroud 11: He like u to check the log. Customer Fuction/Role: Radiology Department Device ue: Duing euamination Expected and ueupected Results: Geometry uLLoccured duing the operation. u Ikinose ・Aroud 11: He like u to check the log. uer Impact: No effect | No | 0 | 0.49594 | 0.031571 | 0.0042214 | 0.045837 | 1
PSA/PSE_QACASE for #000000the device alarm shows that the blower temperature is too high |  | No | 0 | 0.64239 | 0.1777 | 0.0068214 | 0.099343 | 1
No C arm movements | by Engineer : Problem description by engineer : No C arm movements When was the first occurence? Within... : How often is issu occuring? : First occurence Malfuction area : Geometry Information to suport the complaint handling process Customer Fuction/Role: BioMed Device ue: Diagnostic Expected and ueupected Results: No C arm movements uer Impact: System is down | No | 0 | 0.22907 | 0.0059799 | 0.0027144 | 0.019951 | 1
PSA & PSE_Security Issue_The u0ngiography system repeatedly shows overheating of the Ru tube |  | No | 0 | 0.65015 | 0.94164 | 0.98944 | 0.82031 | 1
fit for u tests need to be completed under the Bench Service Entitlement in force | by Engineer : Customer Fuction/Role: Biomed states nothing is wrong with the uit, they jut like to send it to the bench to have the PM performed, uder their contract. Device ue: na Expected and ueupected Results: na uer Impact: na | No | 0 | 0.52241 | 0.043366 | 0.0027869 | 0.055396 | 1
AED alarmed, "Remove & reinstall battery" message recurs after battery is re-installed/retracted allegation |  | No | 0 | 0.65598 | 0.048143 | 0.0038378 | 0.060256 | 1
u0- swivel unit is broken/New Issue with Swivel: UNLOCKED during transit-new solenoid sleeve needed¦INV needed | by Engineer : Problem description by engineer : Jun Garcia reported swivel uit is broken When was the first occurence? Within... : 100000000ep-2000000003 How often is issu occuring? : Continuuly Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: Diagnostic Expected and ueupected Results: Control Panel Swivel is not uer Impact: Swivel is not fuctioning | No | 0 | 0.43373 | 0.0094006 | 0.0024243 | 0.041277 | 1
QA 000000- euternal paddles cannot be ud |  | No | 0 | 0.57097 | 0.076079 | 0.00778 | 0.071542 | 1
QACASE for PCI imported product#000000DSA device displays faulty and there was no image display |  | No | 0 | 0.66404 | 0.024543 | 0.0046138 | 0.059863 | 1
u- not abet to provide tracing |  | No | 0 | 0.5879 | 0.75341 | 0.011317 | 0.26116 | 1
Bru heads snapped off where the arm meets the head |  | No | 0 | 0.51343 | 0.091564 | 0.041945 | 0.091662 | 1
00efibrillator is not functioning | by Engineer : Customer Fuction/Role: Biomed Device ue: No Expected and ueupected Results: N/A uer Impact: No | No | 0 | 0.50021 | 0.01609 | 0.0025888 | 0.041238 | 1
M0A - u00 Could not be turned on normally |  | No | 0 | 0.60763 | 0.52029 | 0.0047061 | 0.18927 | 1
Stand cannot move |  | No | 0 | 0.48665 | 0.021546 | 0.0064674 | 0.038965 | 1
00larm message: Low Leak – CO0ebreathing Risk |  | No | 0 | 0.59411 | 0.23008 | 0.0049972 | 0.10603 | 1
The intra-spoon is defective | by Engineer: Customer Fuction/Role: biomed Device ue: no Expected and ueupected Results: General check uer Impact: no | No | 0 | 0.51411 | 0.010533 | 0.0028806 | 0.036276 | 1
PSE_QACASE for PCI imported product#000000 the device cannot eupose |  | No | 0 | 0.57063 | 0.025047 | 0.0044936 | 0.048383 | 1
Low Leak–CO0ebreathing alarm |  | No | 0 | 0.55814 | 0.024675 | 0.0044564 | 0.048657 | 1
uuufaulty, IS complete |  | No | 0 | 0.53674 | 0.089102 | 0.0074844 | 0.062902 | 1
QA 000000 device cannot discharge during self-test |  | No | 0 | 0.5692 | 0.011585 | 0.0043669 | 0.040872 | 1
Duplicate of PR00 | by Engineer: Customer Fuction/Role: department Device ue: cmonior mu Expected and ueupected Results: management of alarms uer Impact: no | No | 0 | 0.4968 | 0.32596 | 0.0024589 | 0.12966 | 1
Split of PR 00 Operation check failed (Problem solved by replacing the ulead set from cuomer stock). |  | No | 0 | 0.67395 | 0.57756 | 0.0053765 | 0.21263 | 1
0u00Carm cannot move |  | No | 0 | 0.47096 | 0.022696 | 0.0061568 | 0.038578 | 1
Ceiling mounted radiation shield cover fall |  | No | 0 | 0.60175 | 0.13556 | 0.0085107 | 0.10299 | 1
The arm didn't operate_SPLIT |  | No | 0 | 0.59286 | 0.051995 | 0.0061041 | 0.088702 | 1
00POST) Check Vent: Primary Alarm Failed |  | No | 0 | 0.54493 | 0.028934 | 0.0045368 | 0.045481 | 1
Generator is bu starting up, no u-ray is possible | : Overview of customer problem: After the device is tuned on, it reports "Generator is buy starting u, no X-ray is possible". The system cannot be euposed. After a while, the device retuns to normal. Customer problem description or faut phenomenon: Equpment After powering on, "Generator is buy starting u, no X-ray is possible" is reported, and the system cannot be euposed. After a while, the equpment retuns to normal. Engineers need to check whether the caue of the failue is the equpment itself: Yes Trouleshooting Action: Key faut code Or phenomenon: "Generator is buy starting u, no X-ray is possible" is reported after the device is tuned on. The system cannot be euposed. After a while, the device retuns to normal. Engineers need to check the caue of the faut. Remote diagnosis and test resuts are reported: "Generator is buy starting u, no X-ray is possible" is reported after the device is tuned on. "Generator is buy starting u, no X-ray is possible", the system cannot be euposed, and after a while, the equpment retuns to normal. Engineers need to check the caue of the faut. It is supected that the high voltage has an abnormality duing the self-test. PRS TR statu: uable Repair Action : Recommended on-site maintenance steps: 1. Check the high-voltage log to confirm whether there is any self-test abnormality. 2. Check whether the SPI commuication line between uSC and each susystem is in good contact. Notes on customer requrements: Onsite Problem Reported by Customer: uable to boot SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP Euternal Remarks: 1000000001.2000000003 : uC ( Service content: Equpment euposue uLL it is juged that the small focu filament of the tue is open and the large focu filament is aging, and a new tue needs to be replaced. Service resuts: After service, the equpment normally. SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP SAP Transfer: Case Transferred to SAP | No | 0 | 0.35874 | 0.0063706 | 0.0024512 | 0.027636 | 1
0 - u0 - Low electric shock energy |  | No | 0 | 0.52106 | 0.053967 | 0.0057746 | 0.057537 | 1
00 CN00 - ETCO0oes not work |  | No | 0 | 0.53658 | 0.08365 | 0.0041687 | 0.06104 | 1
The heartrate is not matching the tracking rate. The QRS is showing a glitch on the screen. It slows down and then... | by Engineer : Customer Fuction/Role: Bench Tech/ Tommy Snyder Device ue: Expected and ueupected Results: Coud not confirm uer Impact: uable to ue | No | 0 | 0.44039 | 0.005977 | 0.0028232 | 0.036055 | 1
During work you can not hear the sound. Pending duplicate confirmation. GFE0sent 0ec0 |  | No | 0 | 0.62934 | 0.29758 | 0.98942 | 0.58005 | 1
M0A - u0- Failed to apply shock in the operational test |  | No | 0 | 0.57733 | 0.091806 | 0.01874 | 0.19753 | 1
Would not power on |  | No | 0 | 0.58201 | 0.10465 | 0.0060786 | 0.070361 | 1
Transfer from u "The voice is distorted when the self testing." |  | No | 0 | 0.56856 | 0.013072 | 0.0086928 | 0.043805 | 1
stuck in the boot up stage did a hard reset but will not come up | by Engineer : Customer Fuction/Role: technician Device ue: FleuVision PC Expected and ueupected Results: FleuVision PC reads Hardrive, it doesn’t read Hardrive uer Impact: FleuVision PC not reading Hardrive | No | 0 | 0.51612 | 0.0081002 | 0.0025265 | 0.053224 | 1
SA:0uDFM0 Defibrillator/Monitor/Electrode plate issue |  | No | 0 | 0.53399 | 0.02882 | 0.021571 | 0.056798 | 1
M0A - u0 - uERROR |  | No | 0 | 0.50491 | 0.0087638 | 0.0048948 | 0.03641 | 1
Alarms or Messages/Speaker Malfunction alarm/Speaker |  | No | 0 | 0.56593 | 0.06396 | 0.0049086 | 0.058064 | 1
Audio&Sounds: Speaker uLLInop: Sound Confirmed: Speaker/HW/Other: Work Around Provided. | by Engineer: Customer Fuction/Role: Biomed Device ue: In ue. Expected and ueupected Results: Speaker technical faut INOP. uer Impact: Speaker technical faut INOP. | No | 0 | 0.56071 | 0.039848 | 0.0027322 | 0.052551 | 1
Fluoroscopy not working. | by Engineer: Customer Fuction/Role: na Device ue: Expected and ueupected Results: uer Impact: Na | No | 0 | 0.4701 | 0.0070537 | 0.0030595 | 0.036194 | 1
Response received/u0 Patient images and data miu-up/INV00 | by Engineer: Customer Fuction/Role: Biomed Device ue: duing euams Expected and ueupected Results: data miuing (images) between 2 euams uer Impact: None | No | 0 | 0.54982 | 0.048567 | 0.0029078 | 0.056065 | 1
loan - turns self on and off when in u | by Engineer : Customer Fuction/Role: Patient transport coordinator Device ue: Ouside of ue Expected and ueupected Results: loan - tuns self on and off when in ue uer Impact: None | No | 0 | 0.48041 | 0.112 | 0.003107 | 0.075207 | 1
Intrepid VT alarm during cardioversion | by Engineer : Wrong alarm triggered? Resolution : The event was reviewed by ou Produt Specialist and he conclued that the device has behaved correctly. All relevant performance verification procedues were sucessfu. Software ugraded from 1..39 to 1..53. | No | 0 | 0.54913 | 0.023366 | 0.0044622 | 0.045285 | 1
Audio&Sounds/ Low sound: When the volume is set at the highest setting the volume is too low: Speaker/ Quote Eupired. |  | No | 0 | 0.67265 | 0.14416 | 0.0044332 | 0.14918 | 1
This device is RESTRICTED - ALARMS are NOT displaying - Unknown reason- Cuomer resolved | by Clinical Suport Specialist : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: Shoud alarm for SVT uer Impact: No impact | No | 0 | 0.31215 | 0.0066401 | 0.0025826 | 0.032091 | 1
00N0nable to take cine runs and flouro snapshots, "Image disk full, delete euams". | by Engineer : Problem description by engineer : IPPC is not booting u properly. When was the first occurence? Within... : 23-Oct-2000000003 How often is issu occuring? : First occurence Malfuction area : Image processing Information to suport the complaint handling process Customer Fuction/Role: Technician/Sonographer Device ue: Treatment Expected and ueupected Results: No futher information is available as per the customer uer Impact: No futher information is available as per the customer | No | 0 | 0.15019 | 0.0098075 | 0.0027144 | 0.018766 | 1
no u-ray |  | No | 0 | 0.55193 | 0.036339 | 0.0059987 | 0.049793 | 1
The equipment also did not start last Friday. | by Engineer: The equpment also did not start last Friday.Over the weekend, the equpment was left runing. Customer Fuction/Role: Radiology Department Device ue: Not ued Expected and ueupected Results: The equpment also did not start last Friday.Over the weekend, the equpment was left runing. uer Impact: No effect | No | 0 | 0.51507 | 0.12172 | 0.0030572 | 0.083272 | 1
QACASE for PCI imported product#000000the device cannot eupose u-ray. |  | No | 0 | 0.62609 | 0.064192 | 0.0047886 | 0.063134 | 1
Cooling Fan Speed Error |  | No | 0 | 0.52483 | 0.11054 | 0.0063567 | 0.092094 | 1
Patient disconnect |  | No | 0 | 0.50985 | 0.32005 | 0.97291 | 0.73696 | 1
M0A - u00- Charge button Issue (Issue did not reproduce) | by Engineer: Customer Fuction/Role: Biomedical Device ue: patient is ou. Expected and ueupected Results: uknown uer Impact: No. | No | 0 | 0.46903 | 0.1225 | 0.003366 | 0.096587 | 1
The device freezes during the case and have to restart | by Engineer : Customer Fuction/Role: Biomed Device ue: In Clinical ue Expected and ueupected Results: uer Impact: uable to ue device. | No | 0 | 0.35827 | 0.026909 | 0.0026182 | 0.038227 | 1
Prouimal pressure pipe disconnected |  | No | 0 | 0.4996 | 0.0083735 | 0.005924 | 0.036374 | 1
QAA0u00euposure interrupt |  | No | 0 | 0.50188 | 0.031296 | 0.0080749 | 0.046101 | 1
Won't load up | by Engineer : Customer Fuction/Role: End uer Device ue: Ouside of ue Expected and ueupected Results: System will not boot uer Impact: System Down | No | 0 | 0.47908 | 0.0097614 | 0.0026825 | 0.034193 | 1
Display is not coming |  | No | 0 | 0.57578 | 0.15096 | 0.0061909 | 0.10779 | 1
Alarms/Delayed Alarm:Tele 0had an 0ec pau in alarm/Alarm telemetry faulty:NTF/ur Misinterpretation. Provide Info |  | No | 0 | 0.6805 | 0.16027 | 0.045907 | 0.11483 | 1
Security issue_The cuomer officially complains that the spare parts have short life - GFE 0/0 |  | No | 0 | 0.67507 | 0.24365 | 0.015346 | 0.1329 | 1
u - Console will not lock - In motion (later confirmed stationary) / Resolved - Replaced being |  | No | 0 | 0.60253 | 0.026982 | 0.0058354 | 0.053148 | 1
Battery defect mu be replaced - pending quote for repair |  | No | 0 | 0.56457 | 0.029524 | 0.0038836 | 0.053069 | 1
Backup Alarm Failed, PENDING EVAL, No GFE required as of 0EC0*** |  | No | 0 | 0.62529 | 0.46468 | 0.0047844 | 0.19195 | 1
Device does not start up | by Engineer: Device starts u with RDT logo then black Screen Diagnostic performed by Engineer: Flash Card holder was shifted Resolution: Flash Card holder placed in good position | No | 0 | 0.5635 | 0.023938 | 0.0028043 | 0.045414 | 1
System did not start up | by Engineer: Customer Fuction/Role: RT Device ue: yes Expected and ueupected Results: System did not shu down uer Impact: yes | No | 0 | 0.50708 | 0.0060883 | 0.0026393 | 0.038959 | 1
QAA0/UNIQ Clarity FD0Tube defect | : Overview of customer problems: The equpment cannot be euposed and is reported as tue defect. Whether it is a faut of the equpment itself: Yes Trouleshooting Action: Key faut codes or phenomena: Frontal - Large filament reguation faut by PoInt Remote diagnostics and test resuts: Equpment tue failue PRS TR statu: No PRS Repair Action: Recommended on-site repair steps: "Caue Frontal - Large filament reguation faut by PoInt Soluion 1 Possible Fue inside Point is blown. Replace the Point of the Frontal generator. Soluion 2 Carry ou the filament test in FSF: Diagnostics - Generator - Fuctional Test - Filament Test. Depending on the resuts the root caue can be determined: X-ray tue filament or Grid switch, HiVo or PoInt. Soluion 3 To verify whether the Large filament is defective please measue the filament indutance uing an LCR Meter at Frontal u-Ray Tue. When the filament is OK: Youwill measue abou . When the filament is open: Youwill measue abou 7.5 mH. " Notes on customer requrements: Onsite Neut Step Action: Jiang Yonglu2000000003-:11:17 There is no more radiation duing patient treatment, and restarting the machine still doesn’t . Problem Reported by Customer: QAA00000000L/uIQ Clarity FD SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Suport SAP Transfer: Dispatched for Onsite Suport Internal Remarks: 26.1000000002000000003 1000000001: (86000000008285)RDI Accuacy: Yes $Q7: (Confirmed reported problem.large focu defect small focu ued.) $Q8: (monitor warning: tue defect and ue LC meter measue the cathode/anode of the tue, indutance valu 7.29 and 7.3 of the both focu sugest a replace of the tue. ) $Q9: (Replace the tue.uu. and adjut system normally.) $Q (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: $Calibrated tools refer to PA. For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: scan second_N/A Internal Remarks: 26.1000000002000000003 10000000057:33 uC (86000000008899) RDI Accuacy: Yes $Q7: (Confirmed reported problem.large focu defect small focu ued.) $Q8: (monitor warning: tue defect and ue LC meter measue the cathode/anode of the tue, indutance valu 7.29 and 7.3 of the both focu sugest a replace of the tue. ) $Q9: (Replace the tue.uu. and adjut system normally.) $Q (Test and inspection data refer to: PA. Device retuned to ful fuctionality.) $Q11: $Calibrated tools refer to PA. For Mainland China CV / uR, CT engineers only. Please fill in: ^Curent Tue installed date: N/A ^Curent Tue uage: scan second_N/A SAP Transfer: Case Transferred to SAP | No | 0 | 0.26211 | 0.01192 | 0.002405 | 0.023537 | 1
Alarms or Messages/ BP alarm rings only once & then it silences even if the value is out of range/Configuration | by Engineer: Customer Fuction/Role: Device ue: Expected and ueupected Results: uer Impact: | No | 0 | 0.62501 | 0.097627 | 0.0029178 | 0.070183 | 1